## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2009-68

Agency: IOS, IOS, Ocean Sciences Division, Sidney, B.C.

Location: Broughton Archipelago	Project: Sea Lice Monitoring

Party Chief: Galbraith M.		Platform: Princeton

Date: 16 November 2009 –20 November 2009

Processed by: Germaine Gatien

Date of Processing: 22 January 2010 –  25 January 2010

Number of original CTD casts: 	20	Number of CTD casts processed: 20

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0334) was used.

# SUMMARY OF QUALITY AND CONCERNS

The pressure sensor is considered ±1db.

The configuration file used in acquisition contained the wrong conductivity and temperature sensor serial numbers. That information gets recorded in the header of the hex file, so even if the correct configuration file is used the data can’t be converted without fixing the headers.

The times in the log are in PDT; the times in the headers of the archived files are in UTC.

There was no salinity calibration sampling and no useful history for the conductivity sensor, so no comment can be made on the salinity accuracy. The CTD salinity precision is considered ±0.001psu when stopped. In motion the salinity is considered ±0.005psu. In areas of rapid temperature change the SBE25 does not perform this well, but for this data set the gradients were low.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data is in a single HEX file.

2. Preliminary Steps

The Daily Log was obtained. There was no note of any problems encountered.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

As was found for cruise 2009-69 there were was some confusion about what sensors were used. The con file used at sea and the headers in the hex file indicated that the T/C/P sensors were 2106/1729/0464 but converting with that assumption produced bad data. When it was assumed the CTD had sensors 5013/3531/0464 good results were achieved, but only after editing the hex file headers to match those numbers. The configuration parameters could not be independently confirmed, but they were obtained by Scott Rose directly from the factory, so are assumed to be correct. The T and C sensors have no entries in the history file, but the pressure sensor has been used on previous cruises.

The con file was adjusted to add a pressure offset of +0.7db since that has been found appropriate for other cruises (before and after this one) that used the same pressure sensor. The hex file was then converted to a CNV file. Salinity was included in the conversion to help assess if the sensors were correctly identified. The converted file contains CTD data from many casts. The section routine was used to separate the data into individual casts based on scan range for each cast that is recorded in the header. The casts names were chosen to match the event numbers in the log book. There are 20 files and there are 20 CTD entries in the log book.

As was found in 2009-69 the times in the headers differ from the CTD records by 1 hour. This is likely a computer clock problem. The log times look like PST, with work between 7:45 and 17:30, with the only activities after 16:00 being light traps. While those are reasonable times, it is not completely obvious that the computer times are not. However during 2009-69 the header times definitely looked less believable than the log times. The log will be assumed correct and times will be converted to UTC later in the processing.

Data from a few casts were plotted. The temperature is increasing with pressure, but the range is small.

The temperature and conductivity move to “in-water” values at pressures of ~0 so the pressure offset of +0.7db looks right for this data. The descent rate appears to have been kept reasonably high most of the time (between 0.4 and 0.6m/s) and while a little noisy, it is about as steady as possible for this type of deployment. The downcast data is much smoother than the upcast data.

The soak period was close to 1 minute, and seems to have been long enough to reach equilibrium.

The salinity channel was removed at this stage; it will be derived after filtering and/or CELLTM.

4. WILDEDIT

This step is usually skipped and since there is no evidence of spikes it was not run.

5. WFILTER

WFILTER was run using cosine filter, size 5 on the pressure, temperature and conductivity channels. Plots were made before and after this step. Reversals disappear from the pressure after this step but the trace is not overly smooth. When salinity was derived on filtered files it was much smoother than the unfiltered version, but temperature and salinity still have some reversals as expected.

6. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. For 2009-69 a variety of settings were tested, but all settings seemed to make the data worse. This was presumed to be because the routine is intended to correct for problems that arise when temperature changes fairly rapidly, and there was almost no variation for that data. Similar conditions were found in the 2009-68 data. One cast was tested and running CELLTM made the data slightly worse. So CELLTM was not run.

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

9. Checking Headers

HEADEDIT was used to add lines for latitude, longitude, station name and water depth.

A text editor was then used to enter those details to the headers of all the files.

The times in the headers are all set to the time of the first cast, but the list of scan ranges for each cast includes the start time for each cast as well. That information was used to correct the times in the headers of the HDR files. Those times are one hour later than the log book times which are assumed to be in PST, but this will be corrected in a separate step to avoid errors.

CLEAN was run to add event numbers which were missing from some files.

Header Check was run and a cross-reference listing was produced and details checked against the log records. Two errors were found and corrected in the HDR files. CLEAN was rerun.

ADD TIME CHANNEL was run to add 7 hours to the header times, so they are in UTC.

The surface check showed values between -0.6 and +1.8db with an average of -0.3db. The +1.8db readings was the first record in one file that was followed immediately by values of +0.1 and +0.2db, so that was just a spike. Most of the negative readings are from the upcasts. The corresponding conductivity values were ~0, so the CTD must have been very close to the surface, or quite possibly out of water. So the values are not considered a sign of poor pressure calibration.

##### 10. Test plots

Plots were made. The near-surface data looks unreliable, so before editing (but after DELETE) it would be efficient to clip the data as is often done with SBE25 data. Clipping to +0.8db would leave a little bad surface data, but clipping any deeper might lead to the loss of some good data.

11. Salinity bottles

Salinity

No salinity calibration samples were available.

12. SHIFT

Conductivity

Tests were run using values from -2 to +0.5. During 2007-13 and -14 when this sensor was used a setting of -0.6s looked best and for 2008-20 a setting of -0.25s looked best overall. For 2009-69 the salinity was pretty smooth anyway, so little change was noted, but the beset results varied from -0.6 to -1.5. The variation is at least partly due to varying descent rates. Individual settings for each cast might improve the data, but the differences in values are very small. A shift of -0.8 records (0.1seconds) was applied to all casts. For 2009-68 the conditions were similar, so one cast was tested and the results show that -0.8 records is a reasonable choice for these data as well.

All casts were put through SHIFT with a setting of -0.8 records for CONDUCTIVITY. Salinity was recalculated.

13. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

- DETAILED EDITING

CLIP was used to remove data above 0.8db.

Most files needed a little editing. CTDEDIT was used to clean salinity where unstable features looked likely to be caused by misalignment of T and C. Not all unstable features were removed. For 3 casts a single surface record was removed. The following casts required no editing: 11, 20, 33 and 34.

Notes of editing details were made in the headers.

- Inter-comparisons

Sensor History – This conductivity sensor was used for 2009-69 but there was no calibration sampling during that cruise. The pressure sensor had been used for 6 other cruises that have been processed since it was last calibrated at the factory. An offset of +0.7db was applied to all but one of them. For one cruise no offset was applied.

Comparison of repeat casts –There are no repeat casts.

Historic Ranges – The only climatology available for this area covered areas too large to reflect these inlet casts with the exception of casts #43 and 44 in Queen Charlotte Sound. The only excursion from the historic ranges for those 2 casts were a few salinity values that were very slightly low around 20db. That is more likely to reflect the limitations of the climatology than poor sensor calibration, because there are probably few data from this time of year and this close to shore.

- Bin Average and REMOVE

The files were bin averaged using 1db bins.

REORDER was run to remove Scan_Number, Conductivity:Primary, Descent Rate and Flag channels.

(Usually REMOVE is used, but that routine was not working at the time.)

- HEADER EDIT and final checks

Header Edit was used to fix formats and channel names and to add the following note to the headers:

The pressure sensor is considered ±1db

The CTD salinity precision is considered ±0.001psu when stopped. In motion

the salinity is considered ±0.005psu.

There were no salinity calibration samples and there is no calibration history for these

sensors since their most recent recalibration, so no comment can be made

on the accuracy.

The final files were named CTD.

It was discovered at this point that the bottom depths were in feet, so they were converted to metres.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors found.

The conductivity and pressure sensor history files were updated.

## Institute of Ocean Sciences

## CRUISE SUMMARY

Cruise ID#: 2009-68

Dates: Start: 16 November 2009 End: 20 November 2009

Location: Broughton Archipelago

Vessel: Princeton

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2009-68

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5013 | 30 Dec 08 | Factory |  |  |
| Conductivity | 3531 | 07 Jan 09 | Factory |  |  |
| Pressure | 0464 | 23 Mar 06 | Factory |  |  |