## Notes and Updates
SWL Salts 2010-5
Date | Name | Note
2011-01-31 | SZ | Set up spreadsheet for Doug Sieberg and Mary Steel who will be analysing salinity samples brought back from sea
Sampling protocol:
Please run a standard water at the start of each day (at least for the first 3 days) to confirm instrument stability.
Please run a deepwater reference between each cage of samples. Note there are two different batches of deep water in the salt lab.
Please see numbering/naming scheme in Autosal Data by Time used during at-sea analysis.
2011-02-23 | DG | Analysing Equipment
Guildline Autosal Model 8400B, serial # 69572
bath temperature of 24OC
accuracy of < +/- 0.0001 conductivity ratio
resolution <0.00001 conductivity ratio
The calibration checks were done using IAPSO Standard Seawater
K15 Value = 0.99997 Batch# P151 With "Use by" Date-May 20, 2012
The Salinometer remained stable and within the Guildline spec's of 0.002 PSU
Data Interpretation Equipment
Ocean Scientific International Salinometer Data Logger software was used to log the conductivity readings from the Guildline Autosal calculating salinity value.
Lab and Sample Conditions
Analysed in a temperature controlled lab with an ambient temperature changing between 22.2oc-23.7oc overall with a 1.0oc maximum change during analysing session.
Sample temperatures ranged from 21.1oc to 22.9oc overall with a 1.0oc maximum change during analysing session
The sample bottles were inverted and mixed prior to analysis.
all samples had minimum 3 flushes and minimum 2 reading
each reading is the average of 5 readings taken in a 5 second period
Analyst Observation
|  | 1 | The Autosal was stable making the analysing process very smooth
Autosal accuracy for electronics is < 0.002 PSU equivalents
the electronics of the Autosal 8400B drifted within the manufactures spec's during analysing
|  | 2 | Calibration checks were done at the start of each day before analysing started.
correction values were within manufactures spec's for all but one day.
|  | 3 | A Reference Samples was run every 24 samples(one cage) to check for drift in the electronics -all looks good -see DeepWaterCheck Records sheet
|  | 4 | All samples were stored in proper salinity bottles with separate inserts and caps which helps in minimizing evaporation and leakage
Analyst Suggestions
|  | 1 | continue to use proper salinity bottles with separate insert and cap mechanisms
|  | 2 | arctic waters are VERY cold -after sampling when the sample warms to ambient temp check that inserts have not popped up
|  | 3 | assure that transport of samples are protected from freezing
2011-02-25 | SZ | Just a few spreadsheet format edits made.
2011-03-03 | SZ | Made a few edits to sample numbers,
Created "Autosal-Data for Merging"
Added plot to "DeepWaterCheck Records" - check water looks good with avg value of 34.9565, consistant with expected value (w/in 0.001 psu)
Calculated Sp and made plots showing distrib of duplicate differences v salt. Sp = 0.0036 using n=47, after removing 1 outlier id'd using Chauvenet Criteria.
2015-01-21 | SZ | Changed Deep Water plot axis' date format

## flags
IOS QA/QC code from Joe - to be paired with final data
Flag = definition
a = Sample for this measurement was drawn from water bottle but not analyzed (not normally used)
b = Acceptable measurement (not normally used) | although data look odd they are real
c = Questionable measurement (no problem observed in sampling or analysis, but value is not trusted, nonetheless; includes outlyers) | keep - but questionable
d = Bad measurement (known problem with sampling or analysis, but not serious enough to completely discard the value)
e = Not reported (lost sample; unredeemably bad measurement) | don't use data point (left in working sheet but not transferred to final data)
f = Mean of replicate measurements (summary of replicates and table of replicate data should be included in the metadata file)
g = Manual chromatographic peak measurement
h = Irregular digital chromatographic peak integration
i = Sample not drawn for this measurement from this bottle (not normally used)
Other QA/QC
q = data point changed but requires another round of plotting (check again!)
qc = questionable measurment, checked by analyst & looks good from analysts view
j = missing sample

## Autosal-Data by Time
Salinity data for SWL 2010-05
Salinometer ID | Station | Sample ID (bottle label) | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 |  | K15 | 2011-02-15 09:26 | 24C | 0.99997 | 3.5403e-06 | 5.0664e-06 | 0.99997
SWL2010-05_1#1 | Slip 1 | 184 | 2011-02-15 09:29 | 24C | 0.92252 | 2.2391e-06 | 5.0664e-06 | 0.92253 | 31.9719 | 8.9577e-05
SWL2010-05_1#2 | Slip 1 | 185 | 2011-02-15 09:32 | 24C | 0.91638 | 2.2391e-06 | 5.0664e-06 | 0.91638 | 31.7336 | 8.5313e-05
SWL2010-05_1#3 | Slip 1 | 186 | 2011-02-15 09:34 | 24C | 0.91231 | 2.7433e-06 | 4.9472e-06 | 0.91231 | 31.5759 | 0.00010971
SWL2010-05_1#4 | Slip 1 | 187 | 2011-02-15 09:37 | 24C | 0.91007 | 1.1196e-06 | 5.126e-06 | 0.91007 | 31.4891 | 4.265e-05
SWL2010-05_1#5 | Slip 1 | 187 | 2011-02-15 09:39 | 24C | 0.90998 | 2.7425e-06 | 5.126e-06 | 0.90999 | 31.4859 | 0.00010447
SWL2010-05_1#6 | Slip 1 | 188 | 2011-02-15 09:42 | 24C | 0.90646 | 2.7425e-06 | 4.8876e-06 | 0.90646 | 31.3493 | 0.0001071
SWL2010-05_1#7 | Slip 1 | 189 | 2011-02-15 09:47 | 24C | 0.90355 | 2.2244e-06 | 5.0068e-06 | 0.90356 | 31.2369 | 8.5887e-05
|  | 190 | NO SAMPLE
SWL2010-05_1#8 | slip 2 | 191 | 2011-02-15 09:50 | 24C | 0.92419 | 3.8749e-06 | 5.0068e-06 | 0.9242 | 32.0367 | 0.00015052
SWL2010-05_1#9 | slip 2 | 192 | 2011-02-15 09:53 | 24C | 0.91821 | 2.4909e-06 | 5.0664e-06 | 0.91821 | 31.8045 | 9.6591e-05
SWL2010-05_1#10 | slip 2 | 193 | 2011-02-15 09:56 | 24C | 0.91379 | 3.4658e-06 | 4.9472e-06 | 0.91379 | 31.6332 | 0.0001334 | Bottle marked 194 but niskin# was 4 concluding sample must be #193
SWL2010-05_1#11 | slip 2 | 194 | 2011-02-15 09:59 | 24C | 0.90983 | 2.7261e-06 | 5.0068e-06 | 0.90983 | 31.4799 | 0.00010708
SWL2010-05_1#12 | slip 2 | 194 | 2011-02-15 10:07 | 24C | 0.90986 | 2.4916e-06 | 5.0664e-06 | 0.90986 | 31.4809 | 9.6523e-05
SWL2010-05_1#13 | slip 2 | 195 | 2011-02-15 10:10 | 24C | 0.90394 | 5.4475e-06 | 5.0068e-06 | 0.90394 | 31.2519 | 0.00021052
SWL2010-05_1#14 | slip 2 | 196 | 2011-02-15 10:13 | 24C | 0.89998 | 3.4658e-06 | 5.0068e-06 | 0.89999 | 31.0989 | 0.00013277
|  | 197 | NO SAMPLE
SWL2010-05_1#15 | slip 3 | 198 | 2011-02-15 11:14 | 24C | 0.92901 | 1.1063e-06 | 5.0664e-06 | 0.92901 | 32.2238 | 4.2703e-05
SWL2010-05_1#16 | slip 3 | 199 | 2011-02-15 11:16 | 24C | 0.92788 | 2.7425e-06 | 5.126e-06 | 0.92789 | 32.18 | 0.00010447
SWL2010-05_1#17 | slip 3 | 200 | 2011-02-15 11:18 | 24C | 0.91852 | 2.8661e-06 | 4.9472e-06 | 0.91852 | 31.8166 | 0.0001118
SWL2010-05_1#18 | slip 3 | 201 | 2011-02-15 11:20 | 24C | 0.91491 | 2.4744e-06 | 4.9472e-06 | 0.91492 | 31.6768 | 9.6496e-05
SWL2010-05_1#19 | slip 3 | 201 | 2011-02-15 11:22 | 24C | 0.91492 | 2.4777e-06 | 5.0664e-06 | 0.91493 | 31.6771 | 9.7078e-05 | Duplicate
SWL2010-05_1#20 | slip 3 | 202 | 2011-02-15 11:25 | 24C | 0.91323 | 3.6105e-06 | 5.0068e-06 | 0.91323 | 31.6115 | 0.00013926
SWL2010-05_1#21 | slip 3 | 203 | 2011-02-15 11:27 | 24C | 0.90874 | 2.7425e-06 | 5.0068e-06 | 0.90874 | 31.4377 | 0.00010709
SWL2010-05_1#22 | Reference Water | CB-6 | 2011-02-15 11:31 | 24C | 0.99886 | 2.7435e-06 | 5.0068e-06 | 0.99887 | 34.9553 | 0.00010447 | 1334/1335
SWL2010-05_1#23 | Loop 17 | Loop - 17 | 2011-02-15 11:34 | 24C | 0.9257 | 1.4046e-06 | 5.0068e-06 | 0.9257 | 32.0951 | 5.4351e-05 | 2011-07-19 05:25
SWL2010-05_1#24 | BRS 0 | 265 | 2011-02-15 11:37 | 24C | 0.90598 | 2.7425e-06 | 5.0068e-06 | 0.90598 | 31.3307 | 0.0001071
SWL2010-05_1#25 | BRS 0 | 266 | 2011-02-15 11:40 | 24C | 0.90589 | 2.2398e-06 | 5.0068e-06 | 0.9059 | 31.3275 | 8.5326e-05
SWL2010-05_1#26 | BRS 0 | 266 | 2011-02-15 11:42 | 24C | 0.90594 | 3.2141e-06 | 4.9472e-06 | 0.90595 | 31.3294 | 0.00012687 | Duplicate
SWL2010-05_1#27 | BRS 0 | 267 | 2011-02-15 11:44 | 24C | 0.90582 | 3.4659e-06 | 4.9472e-06 | 0.90583 | 31.3248 | 0.00013567
SWL2010-05_1#28 | BRS 0 | 268 | 2011-02-15 11:46 | 24C | 0.9058 | 2.2394e-06 | 5.0068e-06 | 0.90581 | 31.3241 | 8.7445e-05
SWL2010-05_1#29 | BRS 0 | 269 | 2011-02-15 11:49 | 24C | 0.89603 | 2.2725e-06 | 5.0068e-06 | 0.89604 | 30.9463 | 8.6796e-05
SWL2010-05_1#30 | UNT-1 | 270 | 2011-02-15 11:52 | 24C | 0.93132 | 2.4911e-06 | 4.9472e-06 | 0.93132 | 32.3136 | 9.87e-05
SWL2010-05_1#31 | UNT-1 | 270 | 2011-02-15 12:01 | 24C | 0.93121 | 2.4911e-06 | 5.0664e-06 | 0.93121 | 32.3092 | 9.3109e-05 | Duplicate
SWL2010-05_1#32 | UNT-1 | 271 | 2011-02-15 12:04 | 24C | 0.92956 | 1.4046e-06 | 5.0068e-06 | 0.92957 | 32.2454 | 5.2235e-05
SWL2010-05_1#33 | UNT-1 | 272 | 2011-02-15 12:09 | 24C | 0.92947 | 2.8901e-06 | 5.0068e-06 | 0.92948 | 32.2418 | 0.00011145
SWL2010-05_1#34 | UNT-1 | 273 | 2011-02-15 12:11 | 24C | 0.92551 | 2.8901e-06 | 5.0068e-06 | 0.92551 | 32.0878 | 0.00011603
SWL2010-05_1#35 | Unt-2 | 274 | 2011-02-15 12:13 | 24C | 0.94216 | 3.1414e-06 | 5.0664e-06 | 0.94216 | 32.7353 | 0.00012315
SWL2010-05_1#36 | Unt-2 | 275 | 2011-02-15 12:16 | 24C | 0.94278 | 3.8748e-06 | 5.0664e-06 | 0.94278 | 32.7594 | 0.00015052
SWL2010-05_1#37 | Unt-2 | 276 | 2011-02-15 12:18 | 24C | 0.94185 | 3.5406e-06 | 5.0068e-06 | 0.94186 | 32.7235 | 0.00014029
SWL2010-05_1#38 | Unt-2 | 276 | 2011-02-15 12:20 | 24C | 0.94184 | 2.4908e-06 | 5.0068e-06 | 0.94184 | 32.7229 | 9.6642e-05 | Duplicate
SWL2010-05_1#39 | Unt-2 | 277 | 2011-02-15 12:22 | 24C | 0.94198 | 3.3588e-06 | 4.9472e-06 | 0.94199 | 32.7285 | 0.00012945
SWL2010-05_1#40 | Unt-2 | 278 | 2011-02-15 12:25 | 24C | 0.9418 | 2.8897e-06 | 5.0068e-06 | 0.94181 | 32.7215 | 0.00011404
SWL2010-05_1#41 | Unt-3 | 279 | 2011-02-15 12:27 | 24C | 0.92777 | 2.0947e-06 | 5.0068e-06 | 0.92777 | 32.1755 | 8.1621e-05
SWL2010-05_1#42 | Unt-3 | 279 | 2011-02-15 12:30 | 24C | 0.92808 | 2.4909e-06 | 5.0664e-06 | 0.92809 | 32.1879 | 9.4938e-05 | Duplicate
SWL2010-05_1#43 | Unt-3 | 280 | 2011-02-15 12:33 | 24C | 0.9282 | 3.3594e-06 | 5.0068e-06 | 0.92821 | 32.1925 | 0.00013089
SWL2010-05_1#44 | Unt-3 | 281 | 2011-02-15 12:35 | 24C | 0.92763 | 6.664e-08 | 4.9472e-06 | 0.92764 | 32.1703 | 0
SWL2010-05_1#45 | Unt-3 | 282 | 2011-02-15 12:38 | 24C | 0.92789 | 2.2391e-06 | 4.9472e-06 | 0.92789 | 32.1803 | 8.8728e-05
SWL2010-05_1#46 | Unt-3 | 283 | 2011-02-15 12:40 | 24C | 0.91945 | 2.128e-06 | 5.126e-06 | 0.91946 | 31.8529 | 8.208e-05
SWL2010-05_1#47 | Reference Water | CB-6 | 2011-02-15 12:42 | 24C | 0.99891 | 2.89e-06 | 5.0068e-06 | 0.99892 | 34.9575 | 0.00011354 | 1334/1335
SWL2010-05_1#48 | BCL Deep | 142 | 2011-02-15 12:55 | 24C | 0.98494 | 2.4912e-06 | 5.0068e-06 | 0.98494 | 34.4082 | 9.6976e-05 | SZ- added missing sample number 142
SWL2010-05_1#49 | BCL Deep | 143 | 2011-02-15 12:58 | 24C | 0.98226 | 3.865e-06 | 5.0068e-06 | 0.98226 | 34.303 | 0.00014998
SWL2010-05_1#50 | BCL Deep | 143 | 2011-02-15 13:00 | 24C | 0.98233 | 7.0655e-06 | 4.9472e-06 | 0.98233 | 34.3057 | 0.0002782 | Duplicate
SWL2010-05_1#51 | BCL Deep | 144 | 2011-02-15 13:03 | 24C | 0.97908 | 2.4909e-06 | 4.9472e-06 | 0.97908 | 34.1782 | 9.4885e-05
SWL2010-05_1#52 | BCL Deep | 145 | 2011-02-15 13:05 | 24C | 0.97226 | 2.8907e-06 | 4.8876e-06 | 0.97227 | 33.9112 | 0.00011317
SWL2010-05_1#53 | BCL Deep | 146 | 2011-02-15 13:09 | 24C | 0.96591 | 4.2741e-06 | 4.9472e-06 | 0.96592 | 33.6625 | 0.00016611
SWL2010-05_1#54 | BCL Deep | 147 | 2011-02-15 13:11 | 24C | 0.96432 | 2.4911e-06 | 5.0068e-06 | 0.96433 | 33.6005 | 9.7495e-05
SWL2010-05_1#55 | BCL Deep | 148 | 2011-02-15 13:14 | 24C | 0.96226 | 2.0947e-06 | 4.9472e-06 | 0.96227 | 33.52 | 8.3896e-05
SWL2010-05_1#56 | BCL Deep | 149 | 2011-02-15 13:16 | 24C | 0.96088 | 2.4908e-06 | 5.0068e-06 | 0.96089 | 33.466 | 9.6609e-05
SWL2010-05_1#57 | BCL Deep | 149 | 2011-02-15 13:18 | 24C | 0.96061 | 2.7097e-06 | 4.9472e-06 | 0.96061 | 33.4552 | 0.00011074 | Duplicate
SWL2010-05_1#58 | BCL Deep | 150 | 2011-02-15 13:21 | 24C | 0.95872 | 3.4392e-06 | 5.0068e-06 | 0.95873 | 33.3815 | 0.00013386
SWL2010-05_1#59 | BCL Deep | 151 | 2011-02-15 13:23 | 24C | 0.95592 | 3.2641e-06 | 5.0068e-06 | 0.95592 | 33.272 | 0.00012596
SWL2010-05_1#60 | BCL Deep | 152 | 2011-02-15 13:26 | 24C | 0.95382 | 2.5369e-06 | 4.9472e-06 | 0.95382 | 33.1901 | 9.7279e-05
SWL2010-05_1#61 | BCL Deep | 153 | 2011-02-15 13:28 | 24C | 0.9505 | 2.7425e-06 | 5.0068e-06 | 0.9505 | 33.0603 | 0.00010659
SWL2010-05_1#62 | BCL Deep | 154 | 2011-02-15 13:31 | 24C | 0.94726 | 3.4659e-06 | 5.0664e-06 | 0.94726 | 32.9341 | 0.00013596
SWL2010-05_1#63 | BCL Deep | 155 | 2011-02-15 13:34 | 24C | 0.94555 | 1.1196e-06 | 5.0068e-06 | 0.94555 | 32.8674 | 4.265e-05
SWL2010-05_1#64 | BCL Deep | 156 | 2011-02-15 13:36 | 24C | 0.94556 | 1.3713e-06 | 5.126e-06 | 0.94557 | 32.8681 | 5.4351e-05
SWL2010-05_1#65 | BCL Deep | 157 | 2011-02-15 13:38 | 24C | 0.94537 | 1.3713e-06 | 5.0068e-06 | 0.94538 | 32.8606 | 5.2235e-05
SWL2010-05_1#66 | BCL Deep | 158 | 2011-02-15 13:41 | 24C | 0.94539 | 2.4909e-06 | 4.9472e-06 | 0.94539 | 32.8612 | 9.4938e-05
SWL2010-05_1#67 | BCL-3 | 159 | 2011-02-15 13:43 | 24C | 0.94184 | 3.1068e-06 | 5.0068e-06 | 0.94185 | 32.7231 | 0.00012131
SWL2010-05_1#68 | BCL-3 | 160 | 2011-02-15 13:46 | 24C | 0.94185 | 4.1138e-06 | 5.0068e-06 | 0.94186 | 32.7234 | 0.000163
SWL2010-05_1#69 | BCL-3 | 161 | 2011-02-15 13:48 | 24C | 0.9395 | 2.4909e-06 | 4.9472e-06 | 0.93951 | 32.6319 | 9.4885e-05
SWL2010-05_1#70 | BCL-3 | 162 | 2011-02-15 13:50 | 24C | 0.9316 | 4.3361e-06 | 5.0068e-06 | 0.93161 | 32.3247 | 0.00016849
SWL2010-05_1#71 | Reference Water | CB-6 | 2011-02-15 13:53 | 24C | 0.99888 | 3.0374e-06 | 5.126e-06 | 0.99889 | 34.9562 | 0.00011904
SWL2010-05_1#72 | UN-1 | 103 | 2011-02-15 14:01 | 24C | 0.92794 | 4.3844e-06 | 4.9472e-06 | 0.92794 | 32.1822 | 0.00017382
SWL2010-05_1#73 | UN-1 | 104 | 2011-02-15 14:04 | 24C | 0.92807 | 2.4909e-06 | 5.0068e-06 | 0.92808 | 32.1874 | 9.4885e-05
SWL2010-05_1#74 | UN-1 | 105 | 2011-02-15 14:06 | 24C | 0.92784 | 1.6605e-06 | 5.0068e-06 | 0.92785 | 32.1785 | 6.4968e-05
SWL2010-05_1#75 | UN-1 | 105 | 2011-02-15 14:10 | 24C | 0.92786 | 2.2394e-06 | 5.0068e-06 | 0.92787 | 32.1793 | 8.7067e-05 | Duplicate
SWL2010-05_1#76 | UN-1 | 106 | 2011-02-15 14:12 | 24C | 0.92778 | 3.5942e-06 | 5.0068e-06 | 0.92778 | 32.1759 | 0.00013907
SWL2010-05_1#77 | UN-1 | 107 | 2011-02-15 14:17 | 24C | 0.92778 | 2.4777e-06 | 5.0068e-06 | 0.92779 | 32.1761 | 9.7169e-05
SWL2010-05_1#78 | UN-1 | 108 | 2011-02-15 14:19 | 24C | 0.92445 | 3.9606e-06 | 5.0068e-06 | 0.92445 | 32.0466 | 0.00015496
SWL2010-05_1#79 | LOOP - 08 | LOOP - 08 | 2011-02-15 14:21 | 24C | 0.92405 | 2.4776e-06 | 4.9472e-06 | 0.92405 | 32.0312 | 9.7029e-05 | 2011-07-13 13:56
SWL2010-05_1#80 | BCL-1 | 109 | 2011-02-15 14:24 | 24C | 0.97395 | 3.4659e-06 | 5.0068e-06 | 0.97395 | 33.9771 | 0.00013341
SWL2010-05_1#81 | BCL-1 | 110 | 2011-02-15 14:27 | 24C | 0.97332 | 2.4909e-06 | 5.0068e-06 | 0.97333 | 33.9528 | 9.6599e-05
SWL2010-05_1#82 | BCL-1 | 111 | 2011-02-15 14:29 | 24C | 0.96425 | 2.4775e-06 | 4.9472e-06 | 0.96425 | 33.5976 | 9.4034e-05
SWL2010-05_1#83 | BCL-1 | 112 | 2011-02-15 14:31 | 24C | 0.95742 | 3.8647e-06 | 5.0068e-06 | 0.95742 | 33.3306 | 0.00014911
SWL2010-05_1#84 | BCL-1 | 113 | 2011-02-15 14:34 | 24C | 0.95723 | 2.8799e-06 | 5.0664e-06 | 0.95724 | 33.3233 | 0.00011318
SWL2010-05_1#85 | BCL-1 | 114 | 2011-02-15 14:36 | 24C | 0.95619 | 2.7425e-06 | 5.0068e-06 | 0.95619 | 33.2826 | 0.00010656
SWL2010-05_1#86 | BCL-1 | 115 | 2011-02-15 14:38 | 24C | 0.95378 | 2.4908e-06 | 5.0068e-06 | 0.95378 | 33.1884 | 9.6599e-05
SWL2010-05_1#87 | BCL-1 | 116 | 2011-02-15 14:40 | 24C | 0.95367 | 2.2391e-06 | 4.9472e-06 | 0.95367 | 33.1842 | 8.5326e-05
SWL2010-05_1#88 | BCL-1 | 117 | 2011-02-15 14:43 | 24C | 0.95169 | 2.4908e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 0.0001006
SWL2010-05_1#89 | BCL-1 | 118 | 2011-02-15 14:45 | 24C | 0.94813 | 2.4911e-06 | 5.0068e-06 | 0.94814 | 32.9682 | 9.4885e-05
SWL2010-05_1#90 | BCL-1 | 119 | 2011-02-15 14:48 | 24C | 0.94592 | 2.7425e-06 | 5.0068e-06 | 0.94592 | 32.8818 | 0.00010656
SWL2010-05_1#91 | BCL-1 | 120 | 2011-02-15 14:50 | 24C | 0.94397 | 2.2391e-06 | 4.9472e-06 | 0.94397 | 32.8058 | 8.8728e-05
SWL2010-05_1#92 | BCL-1 | 121 | 2011-02-15 14:52 | 24C | 0.94074 | 3.8626e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00014926
SWL2010-05_1#93 | BCL-1 | 122 | 2011-02-15 14:55 | 24C | 0.93871 | 3.3437e-06 | 4.8876e-06 | 0.93871 | 32.6011 | 0.00013184
SWL2010-05_1#94 | BCL-1 | 122 | 2011-02-15 14:57 | 24C | 0.93877 | 1.3712e-06 | 5.0068e-06 | 0.93877 | 32.6033 | 5.4326e-05 | Duplicate
SWL2010-05_1#95 | Reference Water | CB-6 | 2011-02-15 15:01 | 24C | 0.99888 | 2.4615e-06 | 5.0068e-06 | 0.99889 | 34.9563 | 9.4938e-05
SWL2010-05_1#96 | NP-1 | 01 | 2011-02-15 15:32 | 24C | 0.98465 | 2.2397e-06 | 5.0068e-06 | 0.98465 | 34.3968 | 8.6195e-05
SWL2010-05_1#97 | NP-1 | 02 | 2011-02-15 15:34 | 24C | 0.98162 | 4.583e-06 | 4.9472e-06 | 0.98162 | 34.2778 | 0.00017845
SWL2010-05_1#98 | NP-1 | 03 | 2011-02-15 15:36 | 24C | 0.97782 | 5.5594e-06 | 4.9472e-06 | 0.97783 | 34.129 | 0.00021878
SWL2010-05_1#99 | NP-1 | 04 | 2011-02-15 15:40 | 24C | 0.97493 | 2.4909e-06 | 5.0068e-06 | 0.97493 | 34.0155 | 9.869e-05
SWL2010-05_1#100 | NP-1 | 05 | 2011-02-15 15:42 | 24C | 0.97314 | 2.4909e-06 | 5.0068e-06 | 0.97314 | 33.9455 | 9.4938e-05
SWL2010-05_1#101 | NP-1 | 06 | 2011-02-15 15:44 | 24C | 0.97282 | 6.2115e-06 | 5.0664e-06 | 0.97282 | 33.9329 | 0.00024548
SWL2010-05_1#102 | NP-1 | 07 | 2011-02-15 15:47 | 24C | 0.97244 | 3.1414e-06 | 5.0664e-06 | 0.97245 | 33.9182 | 0.00012304
SWL2010-05_1#103 | NP-1 | 06 | 2011-02-15 15:49 | 24C | 0.97324 | 2.4754e-06 | 5.0068e-06 | 0.97325 | 33.9496 | 9.4938e-05 | Duplicate
SWL2010-05_1#104 | NP-1 | 08 | 2011-02-15 15:52 | 24C | 0.97156 | 2.4915e-06 | 5.0068e-06 | 0.97157 | 33.8838 | 9.6599e-05
SWL2010-05_1#105 | NP-1 | 09 | 2011-02-15 15:55 | 24C | 0.96875 | 2.8901e-06 | 4.9472e-06 | 0.96875 | 33.7735 | 0.00011151
SWL2010-05_1#106 | NP-1 | 10 | 2011-02-15 15:57 | 24C | 0.96497 | 2.4918e-06 | 4.828e-06 | 0.96498 | 33.6259 | 9.7495e-05
SWL2010-05_1#107 | NP-1 | 11 | 2011-02-15 15:59 | 24C | 0.95664 | 2.2394e-06 | 5.0068e-06 | 0.95665 | 33.3002 | 8.7077e-05
SWL2010-05_1#108 | NP-1 | 12 | 2011-02-15 16:01 | 24C | 0.94522 | 3.5403e-06 | 4.9472e-06 | 0.94522 | 32.8547 | 0.00014162
SWL2010-05_1#109 | NP-1 | 13 | 2011-02-15 16:03 | 24C | 0.93499 | 2.4909e-06 | 5.0068e-06 | 0.93499 | 32.4563 | 9.6599e-05
SWL2010-05_1#110 | NP-1 | 14 | 2011-02-15 16:05 | 24C | 0.92543 | 1.4046e-06 | 5.0068e-06 | 0.92544 | 32.0849 | 5.5371e-05 | 14
SWL2010-05_1#111 | NP-1 | 14 | 2011-02-15 16:08 | 24C | 0.92547 | 1.3712e-06 | 4.9472e-06 | 0.92547 | 32.0862 | 5.2235e-05 | Duplicate
SWL2010-05_1#112 | NP-1 | 15 | 2011-02-15 16:10 | 24C | 0.91924 | 2.4909e-06 | 5.0068e-06 | 0.91925 | 31.8448 | 9.4895e-05
SWL2010-05_1#113 | NP-1 | 16 | 2011-02-15 16:14 | 24C | 0.91899 | 1.3712e-06 | 4.9472e-06 | 0.919 | 31.835 | 5.1714e-05
SWL2010-05_1#114 | NP-1 | 17 | 2011-02-15 16:16 | 24C | 0.91891 | 2.475e-06 | 4.9472e-06 | 0.91891 | 31.8317 | 9.4364e-05
SWL2010-05_1#115 | NP-3 | 18 | 2011-02-15 16:18 | 24C | 0.98397 | 2.2391e-06 | 5.0664e-06 | 0.98398 | 34.3703 | 8.5352e-05
SWL2010-05_1#116 | NP-3 | 19 | 2011-02-15 16:20 | 24C | 0.9816 | 3.1418e-06 | 5.0068e-06 | 0.9816 | 34.277 | 0.00012103
SWL2010-05_1#117 | NP-3 | 20 | 2011-02-15 16:22 | 24C | 0.9772 | 2.4909e-06 | 4.9472e-06 | 0.97721 | 34.1047 | 9.8725e-05
SWL2010-05_1#118 | NP-3 | 20 | 2011-02-15 16:25 | 24C | 0.97724 | 1.1529e-06 | 4.9472e-06 | 0.97724 | 34.106 | 4.4374e-05 | Duplicate
SWL2010-05_1#119 | Reference Water | CB-6 | 2011-02-15 16:27 | 24C | 0.99887 | 2.5749e-06 | 5.0068e-06 | 0.99887 | 34.9556 | 0.00010121
SWL2010-05_1#120 | BCL-3 | 162 | 2011-02-15 16:35 | 24C | 0.93154 | 3.1338e-06 | 5.0068e-06 | 0.93155 | 32.3223 | 0.00012017
SWL2010-05_1#121 | BCL-3 | 163 | 2011-02-15 16:38 | 24C | 0.92728 | 2.2391e-06 | 5.0068e-06 | 0.92728 | 32.1566 | 8.7048e-05
SWL2010-05_1#122 | BCL-3 | 164 | 2011-02-15 16:40 | 24C | 0.92173 | 2.4909e-06 | 5.0068e-06 | 0.92174 | 31.9412 | 9.6496e-05
SWL2010-05_1#123 | BCL-3 | 165 | 2011-02-15 16:42 | 24C | 0.91926 | 2.4069e-06 | 5.0068e-06 | 0.91927 | 31.8455 | 9.4514e-05
SWL2010-05_1#124 | BCL-3 | 166 | 2011-02-15 16:45 | 24C | 0.91407 | 2.4925e-06 | 4.9472e-06 | 0.91408 | 31.6443 | 9.9652e-05
|  | 167 | NO SAMPLE
|  | 168 | NO SAMPLE
SWL2010-05_1#125 | BCL-4 | 169 | 2011-02-15 16:47 | 24C | 0.93716 | 1.7705e-06 | 5.0068e-06 | 0.93717 | 32.5409 | 6.8797e-05
SWL2010-05_1#126 | BCL-4 | 170 | 2011-02-15 16:50 | 24C | 0.93725 | 2.7432e-06 | 5.0068e-06 | 0.93725 | 32.5441 | 0.00010659
SWL2010-05_1#127 | BCL-4 | 171 | 2011-02-15 16:52 | 24C | 0.93344 | 2.7425e-06 | 4.8876e-06 | 0.93344 | 32.3959 | 0.00010656
SWL2010-05_1#128 | BCL-4 | 172 | 2011-02-15 16:54 | 24C | 0.9268 | 2.7425e-06 | 5.0068e-06 | 0.9268 | 32.1378 | 0.00010656
SWL2010-05_1#129 | BCL-4 | 172 | 2011-02-15 16:57 | 24C | 0.92694 | 2.7262e-06 | 4.8876e-06 | 0.92695 | 32.1435 | 0.00010656 | Duplicate
SWL2010-05_1#130 | BCL-4 | 173 | 2011-02-15 17:00 | 24C | 0.91616 | 2.7432e-06 | 5.126e-06 | 0.91616 | 31.725 | 0.00010395
SWL2010-05_1#131 | BCL-4 | 174 | 2011-02-15 17:02 | 24C | 0.90983 | 1.3712e-06 | 5.0664e-06 | 0.90984 | 31.48 | 5.2246e-05
SWL2010-05_1#132 | BCL-4 | 175 | 2011-02-15 17:05 | 24C | 0.9073 | 2.4911e-06 | 4.9472e-06 | 0.9073 | 31.3818 | 9.7041e-05
SWL2010-05_1#133 | LOOP-12 | LOOP - 12 | 2011-02-15 17:07 | 24C | 0.90477 | 1.7702e-06 | 5.0068e-06 | 0.90477 | 31.2839 | 6.8797e-05 | July 15/10 22:23
SWL2010-05_1#134 | BCL-5 | 176 | 2011-02-15 17:10 | 24C | 0.92218 | 2.2397e-06 | 4.9472e-06 | 0.92218 | 31.9586 | 8.5299e-05
SWL2010-05_1#135 | BCL-5 | 177 | 2011-02-15 17:12 | 24C | 0.92203 | 1.3551e-06 | 5.0664e-06 | 0.92203 | 31.9528 | 5.1714e-05
SWL2010-05_1#136 | BCL-5 | 177 | 2011-02-15 17:15 | 24C | 0.92202 | 1.1199e-06 | 4.9472e-06 | 0.92202 | 31.9524 | 4.4782e-05 | Duplicate
SWL2010-05_1#137 | BCL-5 | 178 | 2011-02-15 17:17 | 24C | 0.914 | 2.4912e-06 | 5.0068e-06 | 0.91401 | 31.6416 | 9.7041e-05
SWL2010-05_1#138 | BCL-5 | 179 | 2011-02-15 17:19 | 24C | 0.90887 | 2.8792e-06 | 5.0068e-06 | 0.90887 | 31.4427 | 0.00011137
SWL2010-05_1#139 | BCL-5 | 180 | 2011-02-15 17:21 | 24C | 0.90567 | 2.4908e-06 | 5.0068e-06 | 0.90567 | 31.3188 | 9.3946e-05
SWL2010-05_1#140 | BCL-5 | 181 | 2011-02-15 17:24 | 24C | 0.89056 | 3.9607e-06 | 4.9472e-06 | 0.89057 | 30.7352 | 0.00015076
SWL2010-05_1#141 | BCL-5 | 182 | 2011-02-15 17:26 | 24C | 0.88826 | 2.7432e-06 | 5.0068e-06 | 0.88827 | 30.6464 | 0.00010969
SWL2010-05_1#142 | Reference Water | CB-6 | 2011-02-15 17:29 | 24C | 0.99889 | 3.9082e-06 | 5.0068e-06 | 0.9989 | 34.9567 | 0.00015494
SWL2010-05_1#143 | NP-7 | 62 | 2011-02-15 17:34 | 24C | 0.95169 | 1.3713e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 5.2278e-05
SWL2010-05_1#144 | NP-7 | 63 | 2011-02-15 17:37 | 24C | 0.94074 | 4.1859e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00016101
SWL2010-05_1#145 | NP-7 | 64 | 2011-02-15 17:39 | 24C | 0.93822 | 2.8904e-06 | 5.0068e-06 | 0.93822 | 32.5819 | 0.00011603
SWL2010-05_1#146 | NP-7 | 65 | 2011-02-15 17:41 | 24C | 0.93775 | 2.8904e-06 | 4.9472e-06 | 0.93775 | 32.5637 | 0.00011403
SWL2010-05_1#147 | NP-7 | 66 | 2011-02-15 17:44 | 24C | 0.93771 | 2.7426e-06 | 5.0664e-06 | 0.93771 | 32.5621 | 0.00010451
SWL2010-05_1#148 | NP-7 | 67 | 2011-02-15 17:46 | 24C | 0.93776 | 2.4925e-06 | 5.0664e-06 | 0.93777 | 32.5642 | 9.9734e-05
SWL2010-05_1#149 | NP-7 | 68 | 2011-02-15 17:48 | 24C | 0.93776 | 3.4657e-06 | 5.0068e-06 | 0.93777 | 32.5643 | 0.00013619
SWL2010-05_1#150 | LOOP-4 | LOOP 4 | 2011-02-15 17:50 | 24C | 0.93724 | 2.7434e-06 | 5.0664e-06 | 0.93725 | 32.544 | 0.00011074 | July 10/10 07:30
SWL2010-05_1#151 | Reference Water | CB-6 | 2011-02-15 17:54 | 24C | 0.9989 | 2.1649e-06 | 5.0664e-06 | 0.9989 | 34.9569 | 8.2858e-05
Calibration#2 |  | K15 | 2011-02-16 09:24 | 24C | 0.99995 | 2.7425e-06 | 1.6987e-05 | 0.99997
Calibration#3 |  | K15 | 2011-02-16 09:38 | 24C | 0.99996 | 2.2391e-06 | 9.0599e-06 | 0.99997
SWL2010-05_1#152 | NP-9 | 69 | 2011-02-16 09:41 | 24C | 0.98509 | 1.3884e-06 | 8.9407e-06 | 0.9851 | 34.4144 | 5.4326e-05
SWL2010-05_1#153 | NP-9 | 70 | 2011-02-16 09:44 | 24C | 0.98327 | 2.4909e-06 | 9.0003e-06 | 0.98328 | 34.3427 | 9.4928e-05
SWL2010-05_1#154 | NP-9 | 71 | 2011-02-16 09:46 | 24C | 0.98045 | 3.0387e-06 | 9.0003e-06 | 0.98046 | 34.2323 | 0.00012157
SWL2010-05_1#155 | NP-9 | 71 | 2011-02-16 09:49 | 24C | 0.98043 | 3.1415e-06 | 9.0599e-06 | 0.98044 | 34.2313 | 0.00012514 | Duplicate
SWL2010-05_1#156 | NP-9 | 72 | 2011-02-16 09:51 | 24C | 0.97769 | 2.4777e-06 | 8.9407e-06 | 0.97769 | 34.1238 | 9.9543e-05
SWL2010-05_1#157 | NP-9 | 73 | 2011-02-16 09:53 | 24C | 0.9758 | 2.4744e-06 | 9.0599e-06 | 0.97581 | 34.0497 | 9.7029e-05
SWL2010-05_1#158 | NP-9 | 74 | 2011-02-16 09:55 | 24C | 0.97525 | 1.8038e-06 | 8.9407e-06 | 0.97526 | 34.0284 | 6.8797e-05 | 074 "
SWL2010-05_1#159 | NP-9 | 75 | 2011-02-16 09:57 | 24C | 0.97448 | 2.4909e-06 | 8.9407e-06 | 0.97449 | 33.9984 | 9.9198e-05
SWL2010-05_1#160 | NP-9 | 76 | 2011-02-16 09:59 | 24C | 0.97388 | 2.4911e-06 | 8.9407e-06 | 0.97389 | 33.9745 | 9.6976e-05
SWL2010-05_1#161 | NP-9 | 77 | 2011-02-16 10:01 | 24C | 0.97281 | 2.4911e-06 | 9.0599e-06 | 0.97282 | 33.9328 | 0.0001006
SWL2010-05_1#162 | NP-9 | 78 | 2011-02-16 10:03 | 24C | 0.97154 | 2.0804e-06 | 9.0003e-06 | 0.97154 | 33.8829 | 7.889e-05
SWL2010-05_1#163 | NP-9 | 79 | 2011-02-16 10:05 | 24C | 0.96575 | 2.4915e-06 | 8.9407e-06 | 0.96576 | 33.6566 | 9.3856e-05
SWL2010-05_1#164 | NP-9 | 80 | 2011-02-16 10:07 | 24C | 0.95408 | 1.1199e-06 | 9.0599e-06 | 0.95408 | 33.2002 | 4.5227e-05
SWL2010-05_1#165 | NP-9 | 81 | 2011-02-16 10:10 | 24C | 0.9495 | 1.372e-06 | 9.0003e-06 | 0.94951 | 33.0216 | 5.4326e-05
SWL2010-05_1#166 | NP-9 | 82 | 2011-02-16 10:12 | 24C | 0.94642 | 2.7426e-06 | 9.0599e-06 | 0.94643 | 32.9015 | 0.00010761
SWL2010-05_1#167 | NP-9 | 82 | 2011-02-16 10:17 | 24C | 0.94649 | 2.4909e-06 | 9.0003e-06 | 0.94649 | 32.9042 | 9.4885e-05 | Duplicate
SWL2010-05_1#168 | Reference Water | CB-6 | 2011-02-16 10:19 | 24C | 0.99889 | 3.8587e-06 | 9.0599e-06 | 0.9989 | 34.9565 | 0.00014953
SWL2010-05_1#169 | NP-9 | 83 | 2011-02-16 10:38 | 24C | 0.94606 | 2.8877e-06 | 8.9407e-06 | 0.94607 | 32.8875 | 0.00011008
SWL2010-05_1#170 | NP-9 | 84 | 2011-02-16 10:40 | 24C | 0.94603 | 2.7434e-06 | 9.0003e-06 | 0.94604 | 32.8865 | 0.00011074
SWL2010-05_1#171 | NP-9 | 85 | 2011-02-16 10:42 | 24C | 0.94597 | 3.143e-06 | 8.9407e-06 | 0.94598 | 32.884 | 0.00012347
SWL2010-05_1#172 | NP-11 | 86 | 2011-02-16 10:45 | 24C | 0.98548 | 3.4658e-06 | 8.9407e-06 | 0.98549 | 34.4296 | 0.00013656
SWL2010-05_1#173 | NP-11 | 87 | 2011-02-16 10:48 | 24C | 0.98357 | 4.2256e-06 | 9.0599e-06 | 0.98358 | 34.3548 | 0.00016833
SWL2010-05_1#174 | NP-11 | 88 | 2011-02-16 10:50 | 24C | 0.98046 | 2.7426e-06 | 9.0003e-06 | 0.98047 | 34.2327 | 0.00010865
SWL2010-05_1#175 | NP-11 | 88 | 2011-02-16 10:52 | 24C | 0.98051 | 3.2144e-06 | 9.0599e-06 | 0.98052 | 34.2346 | 0.00012553 | Duplicate
SWL2010-05_1#176 | NP-11 | 89 | 2011-02-16 10:55 | 24C | 0.9779 | 1.3712e-06 | 9.0599e-06 | 0.97791 | 34.1323 | 5.5371e-05
SWL2010-05_1#177 | NP-11 | 90 | 2011-02-16 10:57 | 24C | 0.97537 | 1.1196e-06 | 8.9407e-06 | 0.97538 | 34.0332 | 4.2703e-05
SWL2010-05_1#178 | NP-11 | 91 | 2011-02-16 10:59 | 24C | 0.97502 | 2.4908e-06 | 9.0599e-06 | 0.97503 | 34.0193 | 9.9553e-05
SWL2010-05_1#179 | NP-11 | 92 | 2011-02-16 11:02 | 24C | 0.97402 | 2.8901e-06 | 9.0003e-06 | 0.97403 | 33.9802 | 0.00011145
SWL2010-05_1#180 | NP-11 | 93 | 2011-02-16 11:04 | 24C | 0.97333 | 2.6582e-06 | 9.0599e-06 | 0.97334 | 33.9532 | 0.00010466
SWL2010-05_1#181 | NP-11 | 94 | 2011-02-16 11:07 | 24C | 0.97236 | 1.1203e-06 | 9.0003e-06 | 0.97237 | 33.9152 | 4.2703e-05
SWL2010-05_1#182 | NP-11 | 95 | 2011-02-16 11:10 | 24C | 0.96972 | 2.2401e-06 | 9.0003e-06 | 0.96973 | 33.8116 | 9.0453e-05
SWL2010-05_1#183 | NP-11 | 96 | 2011-02-16 11:12 | 24C | 0.96532 | 2.4909e-06 | 8.9407e-06 | 0.96533 | 33.6395 | 9.87e-05
SWL2010-05_1#184 | NP-11 | 96 | 2011-02-16 11:14 | 24C | 0.9652 | 2.8901e-06 | 9.0003e-06 | 0.96521 | 33.6349 | 0.00011213 | Duplicate
SWL2010-05_1#185 | NP-11 | 97 | 2011-02-16 11:17 | 24C | 0.95597 | 2.4911e-06 | 9.0599e-06 | 0.95598 | 33.2743 | 9.87e-05
SWL2010-05_1#186 | NP-11 | 98 | 2011-02-16 11:19 | 24C | 0.9488 | 2.2268e-06 | 9.0599e-06 | 0.94881 | 32.9944 | 8.7015e-05
SWL2010-05_1#187 | NP-11 | 99 | 2011-02-16 11:22 | 24C | 0.94613 | 2.2725e-06 | 8.9407e-06 | 0.94614 | 32.8902 | 8.9676e-05
SWL2010-05_1#188 | NP-11 | 100 | 2011-02-16 11:24 | 24C | 0.94712 | 4.6354e-06 | 8.9407e-06 | 0.94712 | 32.9287 | 0.00017833
SWL2010-05_1#189 | NP-11 | 101 | 2011-02-16 11:28 | 24C | 0.94016 | 3.1145e-06 | 9.0599e-06 | 0.94017 | 32.6577 | 0.00012001
SWL2010-05_1#190 | NP-11 | 102 | 2011-02-16 11:31 | 24C | 0.93472 | 2.2391e-06 | 9.0003e-06 | 0.93473 | 32.446 | 8.7024e-05
SWL2010-05_1#191 | Reference Water | CB-6 | 2011-02-16 11:33 | 24C | 0.99888 | 3.1867e-06 | 8.9407e-06 | 0.99889 | 34.9563 | 0.00012382
SWL2010-05_1#192 | UTBS-5 | 224 | 2011-02-16 11:41 | 24C | 0.93768 | 2.8901e-06 | 8.9407e-06 | 0.93769 | 32.5614 | 0.0001118
SWL2010-05_1#193 | UTBS-5 | 224 | 2011-02-16 11:44 | 24C | 0.93764 | 3.2141e-06 | 9.0003e-06 | 0.93765 | 32.5597 | 0.00012601 | Duplicate
SWL2010-05_1#194 | UTBS-5 | 225 | 2011-02-16 11:46 | 24C | 0.93599 | 2.1137e-06 | 9.0003e-06 | 0.93599 | 32.4952 | 8.1621e-05
SWL2010-05_1#195 | UTBS-5 | 226 | 2011-02-16 11:48 | 24C | 0.93425 | 3.1414e-06 | 9.0003e-06 | 0.93426 | 32.4276 | 0.00012104
SWL2010-05_1#196 | UTBS-5 | 227 | 2011-02-16 11:50 | 24C | 0.93391 | 1.7702e-06 | 9.0599e-06 | 0.93392 | 32.4145 | 6.8817e-05
SWL2010-05_1#197 | UTBS-5 | 228 | 2011-02-16 11:52 | 24C | 0.93373 | 2.2397e-06 | 9.0003e-06 | 0.93374 | 32.4075 | 8.7013e-05
|  | 229 | NO SAMPLE
SWL2010-05_1#198 | UTBS-2 | 230 | 2011-02-16 11:56 | 24C | 0.93698 | 4.0094e-06 | 9.0599e-06 | 0.93699 | 32.5339 | 0.00015582 | no 229
SWL2010-05_1#199 | UTBS-2 | 231 | 2011-02-16 11:58 | 24C | 0.93695 | 2.2391e-06 | 9.0003e-06 | 0.93695 | 32.5326 | 8.7048e-05
SWL2010-05_1#200 | UTBS-2 | 231 | 2011-02-16 12:00 | 24C | 0.93699 | 3.1414e-06 | 8.9407e-06 | 0.93699 | 32.5342 | 0.00012313 | Duplicate
SWL2010-05_1#201 | UTBS-2 | 232 | 2011-02-16 12:02 | 24C | 0.93719 | 1.1532e-06 | 9.0599e-06 | 0.9372 | 32.5422 | 4.4364e-05
SWL2010-05_1#202 | UTBS-2 | 233 | 2011-02-16 12:04 | 24C | 0.93146 | 3.3594e-06 | 9.0599e-06 | 0.93147 | 32.3192 | 0.0001316
SWL2010-05_1#203 | UTBS-2 | 234 | 2011-02-16 12:06 | 24C | 0.93041 | 2.2397e-06 | 8.9407e-06 | 0.93042 | 32.2784 | 8.8728e-05
SWL2010-05_1#204 | UTBS-4 | 235 | 2011-02-16 12:08 | 24C | 0.9439 | 4.6245e-06 | 8.9407e-06 | 0.94391 | 32.8034 | 0.00018283
SWL2010-05_1#205 | UTBS-4 | 236 | 2011-02-16 12:11 | 24C | 0.9434 | 1.3713e-06 | 9.0599e-06 | 0.94341 | 32.784 | 5.2235e-05
SWL2010-05_1#206 | UTBS-4 | 237 | 2011-02-16 12:13 | 24C | 0.9396 | 3.5403e-06 | 8.9407e-06 | 0.93961 | 32.6359 | 0.00013759
SWL2010-05_1#207 | UTBS-4 | 238 | 2011-02-16 12:15 | 24C | 0.93587 | 3.332e-08 | 8.9407e-06 | 0.93588 | 32.4907 | 0
SWL2010-05_1#208 | UTBS-4 | 238 | 2011-02-16 12:17 | 24C | 0.93587 | 4.2257e-06 | 9.0003e-06 | 0.93588 | 32.4908 | 0.00016353 | Duplicate
SWL2010-05_1#209 | UTBS-4 | 239 | 2011-02-16 12:19 | 24C | 0.93538 | 3.1417e-06 | 9.0599e-06 | 0.93539 | 32.4717 | 0.00012103 | SZ- changed sample number fro 238 to 239
SWL2010-05_1#210 | UTBS-1 | 240 | 2011-02-16 12:22 | 24C | 0.93791 | 5.0346e-06 | 9.0599e-06 | 0.93792 | 32.5701 | 0.00019692
SWL2010-05_1#211 | UTBS-1 | 240 | 2011-02-16 12:24 | 24C | 0.93774 | 2.4909e-06 | 9.0599e-06 | 0.93775 | 32.5634 | 9.87e-05 | Duplicate
SWL2010-05_1#212 | UTBS-1 | 241 | 2011-02-16 12:27 | 24C | 0.93771 | 3.466e-06 | 8.9407e-06 | 0.93772 | 32.5625 | 0.00013595
SWL2010-05_1#213 | UTBS-1 | 242 | 2011-02-16 12:29 | 24C | 0.93744 | 2.2396e-06 | 8.9407e-06 | 0.93745 | 32.552 | 8.5335e-05
SWL2010-05_1#214 | UTBS-1 | 243 | 2011-02-16 12:33 | 24C | 0.93498 | 2.7425e-06 | 9.0599e-06 | 0.93499 | 32.4563 | 0.00010973
SWL2010-05_1#215 | UTBS-1 | 244 | 2011-02-16 12:35 | 24C | 0.93342 | 4.2256e-06 | 9.0599e-06 | 0.93343 | 32.3956 | 0.00016286
SWL2010-05_1#216 | Reference Water | CB-6 | 2011-02-16 12:37 | 24C | 0.99887 | 3.4658e-06 | 9.0003e-06 | 0.99888 | 34.9559 | 0.00013927
SWL2010-05_1#217 | UTN-4 | 284 | 2011-02-16 12:42 | 24C | 0.93597 | 1.7702e-06 | 9.0003e-06 | 0.93598 | 32.4946 | 6.8797e-05
SWL2010-05_1#218 | UTN-4 | 285 | 2011-02-16 12:44 | 24C | 0.92955 | 3.865e-06 | 8.9407e-06 | 0.92956 | 32.2452 | 0.00014997
SWL2010-05_1#219 | UTN-4 | 286 | 2011-02-16 12:46 | 24C | 0.92897 | 2.4775e-06 | 9.0599e-06 | 0.92897 | 32.2223 | 9.4034e-05
SWL2010-05_1#220 | UTN-4 | 286 | 2011-02-16 12:49 | 24C | 0.929 | 2.8901e-06 | 9.0599e-06 | 0.92901 | 32.2237 | 0.00011008 | Duplicate
SWL2010-05_1#221 | UTN-4 | 287 | 2011-02-16 12:52 | 24C | 0.9282 | 2.128e-06 | 8.9407e-06 | 0.92821 | 32.1924 | 8.1169e-05
SWL2010-05_1#222 | UTN-4 | 288 | 2011-02-16 12:54 | 24C | 0.92467 | 5.3436e-06 | 8.9407e-06 | 0.92467 | 32.0553 | 0.0002072
SWL2010-05_1#223 | UTN-5 | 289 | 2011-02-16 12:56 | 24C | 0.93348 | 2.2391e-06 | 8.9407e-06 | 0.93348 | 32.3976 | 8.7239e-05
SWL2010-05_1#224 | UTN-5 | 290 | 2011-02-16 13:00 | 24C | 0.93383 | 2.4908e-06 | 8.9407e-06 | 0.93384 | 32.4114 | 9.7001e-05
SWL2010-05_1#225 | UTN-5 | 291 | 2011-02-16 13:02 | 24C | 0.93168 | 3.4517e-06 | 9.0003e-06 | 0.93169 | 32.3277 | 0.00013386
SWL2010-05_1#226 | UTN-5 | 292 | 2011-02-16 13:04 | 24C | 0.92743 | 1.7705e-06 | 8.9407e-06 | 0.92744 | 32.1627 | 6.8817e-05
SWL2010-05_1#227 | UTN-5 | 292 | 2011-02-16 13:07 | 24C | 0.92727 | 4.1893e-06 | 8.9407e-06 | 0.92728 | 32.1566 | 0.00016234 | Duplicate
SWL2010-05_1#228 | UTN-5 | 293 | 2011-02-16 13:09 | 24C | 0.92435 | 3.054e-06 | 8.9407e-06 | 0.92436 | 32.0429 | 0.00011839
SWL2010-05_1#229 | UTN-6 | 294 | 2011-02-16 13:12 | 24C | 0.92537 | 3.4657e-06 | 8.9407e-06 | 0.92538 | 32.0826 | 0.00013633
SWL2010-05_1#230 | UTN-6 | 295 | 2011-02-16 13:17 | 24C | 0.92523 | 3.865e-06 | 8.9407e-06 | 0.92523 | 32.077 | 0.00015135
SWL2010-05_1#231 | UTN-6 | 296 | 2011-02-16 13:19 | 24C | 0.9245 | 2.7425e-06 | 9.0003e-06 | 0.92451 | 32.049 | 0.00010865
SWL2010-05_1#232 | UTN-6 | 297 | 2011-02-16 13:23 | 24C | 0.92127 | 2.0873e-06 | 9.0599e-06 | 0.92128 | 31.9237 | 7.9556e-05
SWL2010-05_1#233 | UTN-6 | 297 | 2011-02-16 13:25 | 24C | 0.92125 | 3.4659e-06 | 9.0003e-06 | 0.92126 | 31.9227 | 0.00013374 | Duplicate
SWL2010-05_1#234 | UTN-6 | 298 | 2011-02-16 13:28 | 24C | 0.91896 | 3.204e-06 | 9.0003e-06 | 0.91897 | 31.8339 | 0.00012417
SWL2010-05_1#235 | UTN-7 | 299 | 2011-02-16 13:31 | 24C | 0.92586 | 2.0945e-06 | 9.0003e-06 | 0.92586 | 32.1015 | 8.1621e-05
SWL2010-05_1#236 | UTN-7 | 300 | 2011-02-16 13:33 | 24C | 0.92569 | 3.2141e-06 | 8.9407e-06 | 0.9257 | 32.0952 | 0.00012708
SWL2010-05_1#237 | UTN-7 | 300 | 2011-02-16 13:35 | 24C | 0.92568 | 4.1785e-06 | 8.9407e-06 | 0.92569 | 32.0946 | 0.00016445 | Duplicate
SWL2010-05_1#238 | UTN-7 | 301 | 2011-02-16 13:37 | 24C | 0.92383 | 3.53e-06 | 8.9407e-06 | 0.92384 | 32.0229 | 0.00013487
SWL2010-05_1#239 | UTN-7 | 302 | 2011-02-16 13:39 | 24C | 0.92089 | 2.2574e-06 | 9.0003e-06 | 0.92089 | 31.9086 | 8.7711e-05
SWL2010-05_1#240 | UTN-7 | 303 | 2011-02-16 13:41 | 24C | 0.9201 | 5.7875e-06 | 9.0003e-06 | 0.92011 | 31.8781 | 0.00022469
SWL2010-05_1#241 | Reference Water | CB-6 | 2011-02-16 13:47 | 24C | 0.99888 | 4.3195e-06 | 9.0003e-06 | 0.99889 | 34.9562 | 0.00017085
SWL2010-05_1#242 | PAG-4 | 344 | 2011-02-16 14:09 | 24C | 0.94193 | 4.31e-06 | 8.9407e-06 | 0.94194 | 32.7266 | 0.00016895
SWL2010-05_1#243 | PAG-4 | 345 | 2011-02-16 14:13 | 24C | 0.94076 | 3.2148e-06 | 9.0003e-06 | 0.94077 | 32.6811 | 0.00012554
SWL2010-05_1#244 | PAG-4 | 346 | 2011-02-16 14:15 | 24C | 0.93897 | 3.2145e-06 | 9.0599e-06 | 0.93898 | 32.6115 | 0.00012554
SWL2010-05_1#245 | PAG-4 | 347 | 2011-02-16 14:18 | 24C | 0.93838 | 2.8904e-06 | 9.0003e-06 | 0.93839 | 32.5884 | 0.00011213
SWL2010-05_1#246 | PAG-4 | 347 | 2011-02-16 14:20 | 24C | 0.93835 | 3.4551e-06 | 8.9407e-06 | 0.93836 | 32.5874 | 0.00013203 | Duplicate
SWL2010-05_1#247 | PAG-4 | 348 | 2011-02-16 14:22 | 24C | 0.93384 | 2.4911e-06 | 9.0003e-06 | 0.93385 | 32.4117 | 9.4885e-05
SWL2010-05_1#248 | PAG-4 | 349 | 2011-02-16 14:24 | 24C | 0.92737 | 3.4658e-06 | 8.9407e-06 | 0.92738 | 32.1603 | 0.00013726
SWL2010-05_1#249 | PAG-4 | 350 | 2011-02-16 14:26 | 24C | 0.92322 | 2.2262e-06 | 9.0599e-06 | 0.92323 | 31.999 | 8.7432e-05
SWL2010-05_1#250 | 1 Coke Cap 35 |  | 2011-02-16 14:30 | 24C | 1.0029 | 3.4671e-06 | 8.9407e-06 | 1.0029 | 35.1155 | 0.00013732 | NO INSERT!
SWL2010-05_1#251 | 8 Coke Cap 35 |  | 2011-02-16 14:33 | 24C | 1.0011 | 2.408e-06 | 8.9407e-06 | 1.0011 | 35.0425 | 9.4233e-05 | NO INSERT!
SWL2010-05_1#252 | Reference Water | CB-6 | 2011-02-16 14:39 | 24C | 0.99878 | 2.4909e-06 | 8.9407e-06 | 0.99879 | 34.9524 | 9.869e-05
SWL2010-05_1#253 | BCL-1 | 123 | 2011-02-16 14:45 | 24C | 0.9386 | 2.2391e-06 | 8.9407e-06 | 0.93861 | 32.5971 | 8.5406e-05
SWL2010-05_1#254 | BCL-1 | 124 | 2011-02-16 14:47 | 24C | 0.93911 | 2.4775e-06 | 8.9407e-06 | 0.93912 | 32.6168 | 9.4938e-05
SWL2010-05_1#255 | BCL-2 | 125 | 2011-02-16 14:51 | 24C | 0.98301 | 3.3587e-06 | 9.0599e-06 | 0.98302 | 34.3326 | 0.00013246
SWL2010-05_1#256 | BCL-2 | 126 | 2011-02-16 14:53 | 24C | 0.97924 | 3.6104e-06 | 8.9407e-06 | 0.97925 | 34.1848 | 0.00014403
SWL2010-05_1#257 | BCL-2 | 127 | 2011-02-16 14:55 | 24C | 0.97282 | 2.7426e-06 | 9.0003e-06 | 0.97283 | 33.9331 | 0.00010447
SWL2010-05_1#258 | BCL-2 | 127 | 2011-02-16 14:57 | 24C | 0.97275 | 2.2394e-06 | 8.9407e-06 | 0.97276 | 33.9304 | 8.7025e-05 | Duplicate
SWL2010-05_1#259 | BCL-2 | 128 | 2011-02-16 14:59 | 24C | 0.96596 | 2.8792e-06 | 8.9407e-06 | 0.96597 | 33.6646 | 0.00011145
SWL2010-05_1#260 | BCL-2 | 129 | 2011-02-16 15:02 | 24C | 0.96067 | 4.334e-06 | 8.9407e-06 | 0.96068 | 33.4578 | 0.00017005
SWL2010-05_1#261 | BCL-2 | 130 | 2011-02-16 15:04 | 24C | 0.95949 | 3.4657e-06 | 8.9407e-06 | 0.9595 | 33.4118 | 0.00013801
SWL2010-05_1#262 | BCL-2 | 131 | 2011-02-16 15:06 | 24C | 0.95829 | 3.466e-06 | 8.9407e-06 | 0.9583 | 33.365 | 0.0001334
SWL2010-05_1#263 | BCL-2 | 132 | 2011-02-16 15:08 | 24C | 0.95698 | 2.4908e-06 | 9.0599e-06 | 0.95699 | 33.3136 | 9.6599e-05
SWL2010-05_1#264 | BCL-2 | 133 | 2011-02-16 15:10 | 24C | 0.95578 | 3.4354e-06 | 9.0003e-06 | 0.95579 | 33.2668 | 0.0001334
SWL2010-05_1#265 | BCL-2 | 134 | 2011-02-16 15:12 | 24C | 0.95378 | 2.4909e-06 | 9.0003e-06 | 0.95379 | 33.1886 | 9.7462e-05
SWL2010-05_1#266 | BCL-2 | 135 | 2011-02-16 15:14 | 24C | 0.95283 | 3.5942e-06 | 9.0599e-06 | 0.95284 | 33.1518 | 0.00014204
SWL2010-05_1#267 | BCL-2 | 136 | 2011-02-16 15:17 | 24C | 0.95199 | 2.2394e-06 | 9.0599e-06 | 0.952 | 33.1187 | 9.0433e-05
SWL2010-05_1#268 | BCL-2 | 137 | 2011-02-16 15:35 | 24C | 0.9489 | 2.4909e-06 | 9.0003e-06 | 0.94891 | 32.9984 | 9.869e-05
SWL2010-05_1#269 | BCL-2 | 137 | 2011-02-16 15:37 | 24C | 0.94888 | 3.3455e-06 | 9.0003e-06 | 0.94889 | 32.9977 | 0.00012999 | Duplicate
SWL2010-05_1#270 | BCL-2 | 138 | 2011-02-16 15:40 | 24C | 0.94699 | 2.8907e-06 | 8.9407e-06 | 0.947 | 32.9239 | 0.00011317
SWL2010-05_1#271 | BCL-2 | 139 | 2011-02-16 15:42 | 24C | 0.94524 | 3.1415e-06 | 8.9407e-06 | 0.94525 | 32.8557 | 0.00011967
SWL2010-05_1#272 | BCL-2 | 140 | 2011-02-16 15:44 | 24C | 0.94027 | 2.7425e-06 | 8.9407e-06 | 0.94028 | 32.662 | 0.00010659
SWL2010-05_1#273 | BCL-2 | 141 | 2011-02-16 15:46 | 24C | 0.93864 | 4.9165e-06 | 9.0599e-06 | 0.93864 | 32.5984 | 0.00019281
SWL2010-05_1#274 | LOOP 9 |  | 2011-02-16 15:48 | 24C | 0.9386 | 3.865e-06 | 9.0599e-06 | 0.93861 | 32.597 | 0.00015042 | July 14/10 13:44, SZ-Change this from Loop 8 to Loop 9 as it fits timewise, missing 9 and already have an 8.
SWL2010-05_1#275 | Reference Water | CB-6 | 2011-02-16 15:54 | 24C | 0.99887 | 2.89e-06 | 9.0003e-06 | 0.99888 | 34.9561 | 0.00011604
SWL2010-05_1#276 | NP-5 | 42 | 2011-02-16 15:57 | 24C | 0.96704 | 2.4909e-06 | 9.0003e-06 | 0.96705 | 33.7068 | 9.6599e-05
SWL2010-05_1#277 | NP-5 | 43 | 2011-02-16 16:00 | 24C | 0.96298 | 1.6739e-06 | 9.0599e-06 | 0.96299 | 33.548 | 6.8483e-05
SWL2010-05_1#278 | NP-5 | 44 | 2011-02-16 16:02 | 24C | 0.95624 | 2.4911e-06 | 8.9407e-06 | 0.95625 | 33.2847 | 9.869e-05
SWL2010-05_1#279 | NP-5 | 44 | 2011-02-16 16:05 | 24C | 0.95628 | 1.1536e-06 | 9.0599e-06 | 0.95629 | 33.2864 | 4.265e-05 | Duplicate
SWL2010-05_1#280 | NP-5 | 45 | 2011-02-16 16:07 | 24C | 0.9494 | 2.4908e-06 | 9.0599e-06 | 0.94941 | 33.018 | 9.6976e-05
SWL2010-05_1#281 | NP-5 | 46 | 2011-02-16 16:09 | 24C | 0.93657 | 2.4908e-06 | 9.0003e-06 | 0.93658 | 32.5178 | 9.3005e-05
SWL2010-05_1#282 | NP-5 | 47 | 2011-02-16 16:11 | 24C | 0.93527 | 2.4908e-06 | 8.9407e-06 | 0.93528 | 32.4674 | 9.6976e-05
|  | 48 | NO SAMPLE
SWL2010-05_1#283 | NP-5 | 49 | 2011-02-16 16:14 | 24C | 0.93339 | 2.2405e-06 | 9.0599e-06 | 0.9334 | 32.3944 | 8.5406e-05
SWL2010-05_1#284 | NP-5 | 50 | 2011-02-16 16:16 | 24C | 0.93231 | 3.6109e-06 | 9.0003e-06 | 0.93232 | 32.3524 | 0.00013904
SWL2010-05_1#285 | NP-5 | 51 | 2011-02-16 16:18 | 24C | 0.93251 | 1.3712e-06 | 9.0599e-06 | 0.93252 | 32.36 | 5.4351e-05
SWL2010-05_1#286 | LOOP P2 | LOOP P2 | 2011-02-16 16:21 | 24C | 0.93207 | 2.0841e-06 | 9.0003e-06 | 0.93208 | 32.3431 | 8.1174e-05 | July 9/2010
SWL2010-05_1#287 | NP-7 | 52 | 2011-02-16 16:23 | 24C | 0.98453 | 1.7597e-06 | 9.0599e-06 | 0.98453 | 34.3921 | 6.6764e-05
SWL2010-05_1#288 | NP-7 | 53 | 2011-02-16 16:25 | 24C | 0.982 | 2.4911e-06 | 9.0003e-06 | 0.98201 | 34.2929 | 9.4885e-05
SWL2010-05_1#289 | NP-7 | 54 | 2011-02-16 16:28 | 24C | 0.97858 | 1.3712e-06 | 9.0003e-06 | 0.97859 | 34.1589 | 5.5395e-05
SWL2010-05_1#290 | NP-7 | 54 | 2011-02-16 16:30 | 24C | 0.97827 | 4.189e-06 | 9.0599e-06 | 0.97828 | 34.1468 | 0.00016484 | Duplicate ! Not close at all
SWL2010-05_1#291 | NP-7 | 55 | 2011-02-16 16:33 | 24C | 0.97517 | 0 | 9.0599e-06 | 0.97518 | 34.0252 | 0
SWL2010-05_1#292 | NP-7 | 56 | 2011-02-16 16:35 | 24C | 0.97322 | 2.7425e-06 | 9.0599e-06 | 0.97323 | 33.9487 | 0.00010865
SWL2010-05_1#293 | NP-7 | 57 | 2011-02-16 16:37 | 24C | 0.97231 | 6.4306e-06 | 9.0599e-06 | 0.97232 | 33.9131 | 0.00025094
SWL2010-05_1#294 | NP-7 | 58 | 2011-02-16 16:39 | 24C | 0.97246 | 3.6084e-06 | 9.0003e-06 | 0.97247 | 33.9191 | 0.00014251
SWL2010-05_1#295 | NP-7 | 59 | 2011-02-16 16:41 | 24C | 0.97153 | 2.4912e-06 | 8.9407e-06 | 0.97153 | 33.8825 | 9.9745e-05
SWL2010-05_1#296 | NP-7 | 60 | 2011-02-16 16:43 | 24C | 0.97031 | 2.89e-06 | 8.9407e-06 | 0.97032 | 33.835 | 0.00011008 | SZ- changed sample number from 660 to 60
SWL2010-05_1#297 | NP-7 | 61 | 2011-02-16 16:45 | 24C | 0.96418 | 1.1532e-06 | 9.0599e-06 | 0.96419 | 33.5951 | 4.5217e-05
SWL2010-05_1#298 | Reference Water | CB-6 | 2011-02-16 16:49 | 24C | 0.99893 | 2.89e-06 | 8.9407e-06 | 0.99894 | 34.9584 | 0.00011213
SWL2010-05_1#299 | BRS-4 | 245 | 2011-02-16 16:55 | 24C | 0.9257 | 2.7426e-06 | 9.0599e-06 | 0.92571 | 32.0955 | 0.00010865
SWL2010-05_1#300 | BRS-4 | 246 | 2011-02-16 16:58 | 24C | 0.9247 | 2.4915e-06 | 9.0599e-06 | 0.92471 | 32.0565 | 9.1806e-05
SWL2010-05_1#301 | BRS-4 | 246 | 2011-02-16 17:01 | 24C | 0.9247 | 1.1196e-06 | 8.9407e-06 | 0.92471 | 32.0567 | 4.2703e-05 | Duplicate
SWL2010-05_1#302 | BRS-4 | 247 | 2011-02-16 17:03 | 24C | 0.92459 | 3.2142e-06 | 9.0599e-06 | 0.9246 | 32.0524 | 0.00012519
SWL2010-05_1#303 | BRS-4 | 248 | 2011-02-16 17:06 | 24C | 0.92447 | 2.8897e-06 | 8.9407e-06 | 0.92447 | 32.0475 | 0.00011317
SWL2010-05_1#304 | BRS-4 | 249 | 2011-02-16 17:08 | 24C | 0.91952 | 2.2405e-06 | 9.0003e-06 | 0.91953 | 31.8555 | 8.5313e-05
SWL2010-05_1#305 | BRS-5 | 250 | 2011-02-16 17:10 | 24C | 0.94025 | 2.2391e-06 | 9.0003e-06 | 0.94026 | 32.6611 | 8.5406e-05
SWL2010-05_1#306 | BRS-5 | 250 | 2011-02-16 17:12 | 24C | 0.94027 | 4.189e-06 | 9.0003e-06 | 0.94028 | 32.662 | 0.00016283 | Duplicate
SWL2010-05_1#307 | BRS-5 | 251 | 2011-02-16 17:14 | 24C | 0.94044 | 2.7425e-06 | 8.9407e-06 | 0.94044 | 32.6684 | 0.00010761
SWL2010-05_1#308 | BRS-5 | 252 | 2011-02-16 17:16 | 24C | 0.94067 | 2.4912e-06 | 9.0599e-06 | 0.94068 | 32.6775 | 0.0001006
SWL2010-05_1#309 | BRS-5 | 253 | 2011-02-16 17:19 | 24C | 0.93773 | 2.4908e-06 | 8.9407e-06 | 0.93774 | 32.5632 | 9.6609e-05
SWL2010-05_1#310 | BRS-5 | 254 | 2011-02-16 17:21 | 24C | 0.93082 | 3.4662e-06 | 8.9407e-06 | 0.93083 | 32.2945 | 0.00013655
SWL2010-05_1#311 | BRS-3 | 255 | 2011-02-16 17:24 | 24C | 0.9239 | 3.36e-06 | 8.9407e-06 | 0.92391 | 32.0254 | 0.0001298
SWL2010-05_1#312 | BRS-3 | 256 | 2011-02-16 17:26 | 24C | 0.92428 | 2.7425e-06 | 8.9407e-06 | 0.92428 | 32.0401 | 0.00010451
SWL2010-05_1#313 | BRS-3 | 257 | 2011-02-16 17:28 | 24C | 0.92355 | 2.7432e-06 | 9.0599e-06 | 0.92356 | 32.012 | 0.00010659
SWL2010-05_1#314 | BRS-3 | 258 | 2011-02-16 17:30 | 24C | 0.91687 | 3.3591e-06 | 8.9407e-06 | 0.91688 | 31.753 | 0.00012967
SWL2010-05_1#315 | BRS-3 | 258 | 2011-02-16 17:32 | 24C | 0.91673 | 2.2394e-06 | 8.9407e-06 | 0.91673 | 31.7472 | 8.7006e-05 | Duplicate
SWL2010-05_1#316 | BRS-3 | 259 | 2011-02-16 17:34 | 24C | 0.91666 | 3.8747e-06 | 9.0003e-06 | 0.91667 | 31.7446 | 0.00015
SWL2010-05_1#317 | BRS-2 | 260 | 2011-02-16 17:37 | 24C | 0.91566 | 2.575e-06 | 9.0003e-06 | 0.91567 | 31.706 | 9.9156e-05
SWL2010-05_1#318 | BRS-2 | 260 | 2011-02-16 17:40 | 24C | 0.91558 | 2.7264e-06 | 8.9407e-06 | 0.91559 | 31.7029 | 0.00010396 | Duplicate
SWL2010-05_1#319 | BRS-2 | 261 | 2011-02-16 17:42 | 24C | 0.91489 | 2.4908e-06 | 8.9407e-06 | 0.9149 | 31.6761 | 9.6991e-05
SWL2010-05_1#320 | BRS-2 | 262 | 2011-02-16 17:45 | 24C | 0.91447 | 3.6082e-06 | 9.0599e-06 | 0.91448 | 31.66 | 0.00014255
SWL2010-05_1#321 | BRS-2 | 263 | 2011-02-16 17:48 | 24C | 0.91443 | 2.7425e-06 | 8.9407e-06 | 0.91443 | 31.6581 | 0.00010708
SWL2010-05_1#322 | BRS-2 | 264 | 2011-02-16 17:50 | 24C | 0.91417 | 2.4612e-06 | 9.0599e-06 | 0.91418 | 31.6483 | 9.4909e-05
SWL2010-05_1#323 | Reference Water | CB-6 | 2011-02-16 17:52 | 24C | 0.99889 | 3.1418e-06 | 9.0003e-06 | 0.9989 | 34.9566 | 0.00012312
Calibration#4 |  | K15 | 2011-02-17 09:11 | 24C | 0.99996 | 2.7425e-06 | 8.1062e-06 | 0.99997
SWL2010-05_1#324 | BC-2 | 323 | 2011-02-17 09:18 | 24C | 0.96411 | 3.4658e-06 | 8.0466e-06 | 0.96412 | 33.5923 | 0.0001334
SWL2010-05_1#325 | BC-2 | 324 | 2011-02-17 09:21 | 24C | 0.96216 | 2.7432e-06 | 7.987e-06 | 0.96216 | 33.5158 | 0.00010761
SWL2010-05_1#326 | BC-2 | 325 | 2011-02-17 09:23 | 24C | 0.95522 | 2.7425e-06 | 7.9274e-06 | 0.95523 | 33.245 | 0.00010451
SWL2010-05_1#327 | BC-2 | 326 | 2011-02-17 09:25 | 24C | 0.94689 | 2.89e-06 | 8.1062e-06 | 0.9469 | 32.9198 | 0.00011317
SWL2010-05_1#328 | BC-2 | 327 | 2011-02-17 09:27 | 24C | 0.94209 | 4.4536e-06 | 7.987e-06 | 0.9421 | 32.7328 | 0.0001688
SWL2010-05_1#329 | BC-2 | 327 | 2011-02-17 09:30 | 24C | 0.9421 | 3.2141e-06 | 8.1062e-06 | 0.94211 | 32.7334 | 0.00012361 | Duplicate
SWL2010-05_1#330 | BC-2 | 328 | 2011-02-17 09:32 | 24C | 0.9279 | 1.1199e-06 | 8.1062e-06 | 0.92791 | 32.1809 | 4.265e-05
SWL2010-05_1#331 | BC-2 | 329 | 2011-02-17 09:34 | 24C | 0.91093 | 3.8647e-06 | 8.0466e-06 | 0.91094 | 31.5227 | 0.00015041
SWL2010-05_1#332 | BC-2 | 330 | 2011-02-17 09:36 | 24C | 0.88721 | 2.0947e-06 | 8.1062e-06 | 0.88721 | 30.6059 | 8.1765e-05
SWL2010-05_1#333 | PAG-5 | 331 | 2011-02-17 09:39 | 24C | 0.95705 | 1.3713e-06 | 8.0466e-06 | 0.95706 | 33.3165 | 5.1197e-05
SWL2010-05_1#334 | PAG-5 | 332 | 2011-02-17 09:41 | 24C | 0.9466 | 2.4908e-06 | 8.1062e-06 | 0.94661 | 32.9087 | 9.4938e-05
SWL2010-05_1#335 | PAG-5 | 332 | 2011-02-17 09:44 | 24C | 0.94651 | 3.3587e-06 | 8.1062e-06 | 0.94652 | 32.905 | 0.00013477 | Duplicate
SWL2010-05_1#336 | PAG-5 | 333 | 2011-02-17 09:46 | 24C | 0.92577 | 1.7702e-06 | 8.0466e-06 | 0.92578 | 32.0983 | 6.9481e-05
|  | 334 | NO SAMPLE
SWL2010-05_1#337 | PAG-5 | 335 | 2011-02-17 09:49 | 24C | 0.91206 | 3.8748e-06 | 7.987e-06 | 0.91206 | 31.5662 | 0.00015214 | No Bottle #334
SWL2010-05_1#338 | PAG-5 | 336 | 2011-02-17 09:51 | 24C | 0.89843 | 2.4921e-06 | 7.987e-06 | 0.89843 | 31.0389 | 9.751e-05
SWL2010-05_1#339 | PAG-5 | 337 | 2011-02-17 09:54 | 24C | 0.8833 | 2.4909e-06 | 8.0466e-06 | 0.8833 | 30.4551 | 9.7017e-05
SWL2010-05_1#340 | PAG-6 | 338 | 2011-02-17 09:56 | 24C | 0.9294 | 3.6085e-06 | 8.0466e-06 | 0.92941 | 32.2391 | 0.00014094
SWL2010-05_1#341 | PAG-6 | 339 | 2011-02-17 09:59 | 24C | 0.92864 | 1.1199e-06 | 8.1062e-06 | 0.92865 | 32.2096 | 4.1875e-05
SWL2010-05_1#342 | PAG-6 | 339 | 2011-02-17 10:01 | 24C | 0.92861 | 2.7097e-06 | 8.1062e-06 | 0.92862 | 32.2086 | 0.00010447 | Duplicate
SWL2010-05_1#343 | PAG-6 | 340 | 2011-02-17 10:04 | 24C | 0.91602 | 2.7425e-06 | 8.0466e-06 | 0.91603 | 31.7199 | 0.00010347
SWL2010-05_1#344 | PAG-6 | 341 | 2011-02-17 10:06 | 24C | 0.89761 | 3.8746e-06 | 8.1062e-06 | 0.89762 | 31.0074 | 0.0001526
SWL2010-05_1#345 | PAG-6 | 342 | 2011-02-17 10:08 | 24C | 0.88673 | 3.1417e-06 | 8.1062e-06 | 0.88673 | 30.5874 | 0.00012138
SWL2010-05_1#346 | PAG-6 | 343 | 2011-02-17 10:11 | 24C | 0.84755 | 3.8434e-06 | 7.987e-06 | 0.84756 | 29.0824 | 0.00014761
SWL2010-05_1#347 | Reference Water | CB-6 | 2011-02-17 10:13 | 24C | 0.99875 | 2.8807e-06 | 7.9274e-06 | 0.99876 | 34.9513 | 0.00011216
SWL2010-05_1#348 | Slip-5 | 205 | 2011-02-17 10:43 | 24C | 0.93206 | 2.2398e-06 | 7.987e-06 | 0.93207 | 32.3425 | 8.801e-05
SWL2010-05_1#349 | Slip-5 | 205 | 2011-02-17 10:46 | 24C | 0.93205 | 2.7425e-06 | 7.987e-06 | 0.93206 | 32.3423 | 0.00010972 | Duplicate
SWL2010-05_1#350 | Slip-5 | 206 | 2011-02-17 10:49 | 24C | 0.93202 | 3.1414e-06 | 8.1062e-06 | 0.93203 | 32.3409 | 0.00012618
SWL2010-05_1#351 | Slip-5 | 207 | 2011-02-17 10:51 | 24C | 0.92459 | 2.4912e-06 | 7.9274e-06 | 0.9246 | 32.0523 | 9.4938e-05
SWL2010-05_1#352 | Slip-5 | 208 | 2011-02-17 10:53 | 24C | 0.91867 | 2.4909e-06 | 8.0466e-06 | 0.91868 | 31.8227 | 9.4895e-05
SWL2010-05_1#353 | Slip-5 | 209 | 2011-02-17 10:56 | 24C | 0.91403 | 1.7702e-06 | 8.0466e-06 | 0.91403 | 31.6425 | 6.9868e-05
SWL2010-05_1#354 | Slip-5 | 210 | 2011-02-17 10:58 | 24C | 0.9139 | 3.334e-06 | 7.987e-06 | 0.91391 | 31.6379 | 0.00012956
|  | 211 | NO SAMPLE
SWL2010-05_1#355 | Slip-4 | 212 | 2011-02-17 11:01 | 24C | 0.9341 | 4.2633e-06 | 7.987e-06 | 0.93411 | 32.422 | 0.00016473 | No Bottle #211
SWL2010-05_1#356 | Slip-4 | 213 | 2011-02-17 11:04 | 24C | 0.93341 | 3.6104e-06 | 7.987e-06 | 0.93342 | 32.3952 | 0.00013965
SWL2010-05_1#357 | Slip-4 | 214 | 2011-02-17 11:06 | 24C | 0.92376 | 2.7425e-06 | 7.987e-06 | 0.92377 | 32.0201 | 0.00010657
SWL2010-05_1#358 | Slip-4 | 214 | 2011-02-17 11:07 | 24C | 0.92382 | 2.7425e-06 | 8.0466e-06 | 0.92383 | 32.0223 | 0.00011077 | Duplicate
SWL2010-05_1#359 | Slip-4 | 215 | 2011-02-17 11:10 | 24C | 0.92084 | 3.4497e-06 | 7.987e-06 | 0.92085 | 31.9067 | 0.00013584
SWL2010-05_1#360 | Slip-4 | 216 | 2011-02-17 11:13 | 24C | 0.91378 | 3.5197e-06 | 7.987e-06 | 0.91379 | 31.633 | 0.00013827
SWL2010-05_1#361 | Slip-4 | 217 | 2011-02-17 11:15 | 24C | 0.91361 | 3.1417e-06 | 7.9274e-06 | 0.91362 | 31.6266 | 0.00011936
SWL2010-05_1#362 | BCL-6a | 218 | 2011-02-17 11:17 | 24C | 0.93275 | 2.4909e-06 | 7.987e-06 | 0.93276 | 32.3693 | 9.4885e-05
SWL2010-05_1#363 | BCL-6a | 219 | 2011-02-17 11:20 | 24C | 0.93259 | 2.2398e-06 | 8.1062e-06 | 0.9326 | 32.3632 | 8.5299e-05
SWL2010-05_1#364 | BCL-6a | 219 | 2011-02-17 11:22 | 24C | 0.93293 | 3.3588e-06 | 7.987e-06 | 0.93293 | 32.3762 | 0.00013277 | Duplicate
SWL2010-05_1#365 | BCL-6a | 220 | 2011-02-17 11:25 | 24C | 0.93203 | 3.1312e-06 | 8.1062e-06 | 0.93204 | 32.3415 | 0.00012108
SWL2010-05_1#366 | BCL-6a | 221 | 2011-02-17 11:27 | 24C | 0.91993 | 3.1313e-06 | 8.1062e-06 | 0.91993 | 31.8713 | 0.00012365
SWL2010-05_1#367 | BCL-6a | 222 | 2011-02-17 11:30 | 24C | 0.89369 | 5.3482e-06 | 8.1062e-06 | 0.8937 | 30.8561 | 0.00020704
SWL2010-05_1#368 | LOOP 14 | LOOP 14 | 2011-02-17 11:33 | 24C | 0.89535 | 3.8647e-06 | 7.987e-06 | 0.89536 | 30.9202 | 0.0001502 | July 17/10 15:16
SWL2010-05_1#369 | Reference Water | CB-6 | 2011-02-17 11:36 | 24C | 0.99891 | 4.6466e-06 | 8.0466e-06 | 0.99892 | 34.9574 | 0.00018033
SWL2010-05_1#370 | PAG-1 | 304 | 2011-02-17 11:42 | 24C | 0.92629 | 2.8794e-06 | 7.987e-06 | 0.9263 | 32.1184 | 0.00011151
SWL2010-05_1#371 | PAG-1 | 305 | 2011-02-17 11:44 | 24C | 0.92629 | 1.4046e-06 | 8.0466e-06 | 0.9263 | 32.1184 | 5.2235e-05
SWL2010-05_1#372 | PAG-1 | 306 | 2011-02-17 11:47 | 24C | 0.92272 | 3.2141e-06 | 8.1062e-06 | 0.92273 | 31.9798 | 0.00012654
SWL2010-05_1#373 | PAG-1 | 307 | 2011-02-17 11:49 | 24C | 0.92229 | 3.3456e-06 | 8.1062e-06 | 0.9223 | 31.963 | 0.00013
SWL2010-05_1#374 | PAG-1 | 308 | 2011-02-17 11:51 | 24C | 0.92226 | 1.3713e-06 | 8.0466e-06 | 0.92227 | 31.962 | 5.2246e-05
SWL2010-05_1#375 | PAG-1 | 308 | 2011-02-17 11:53 | 24C | 0.92216 | 1.1196e-06 | 8.1062e-06 | 0.92217 | 31.9579 | 4.265e-05 | Duplicate
SWL2010-05_1#376 | PAG-2 | 309 | 2011-02-17 11:55 | 24C | 0.92762 | 3.2141e-06 | 7.9274e-06 | 0.92763 | 32.1699 | 0.00012244
SWL2010-05_1#377 | PAG-2 | 309 | 2011-02-17 11:58 | 24C | 0.92751 | 4.7275e-06 | 8.1062e-06 | 0.92752 | 32.1657 | 0.00018259 | Duplicate
SWL2010-05_1#378 | PAG-2 | 310 | 2011-02-17 12:00 | 24C | 0.92741 | 2.7425e-06 | 8.1062e-06 | 0.92742 | 32.1619 | 0.00010865
SWL2010-05_1#379 | PAG-2 | 311 | 2011-02-17 12:03 | 24C | 0.92685 | 2.4908e-06 | 8.1062e-06 | 0.92686 | 32.14 | 9.6599e-05
SWL2010-05_1#380 | PAG-2 | 312 | 2011-02-17 12:05 | 24C | 0.92425 | 2.7425e-06 | 8.1062e-06 | 0.92426 | 32.0391 | 0.00010656
SWL2010-05_1#381 | PAG-2 | 313 | 2011-02-17 12:07 | 24C | 0.91 | 3.3588e-06 | 8.1062e-06 | 0.91 | 31.4864 | 0.00013116
SWL2010-05_1#382 | PAG-3 | 314 | 2011-02-17 12:10 | 24C | 0.9252 | 3.8654e-06 | 7.987e-06 | 0.92521 | 32.0759 | 0.00014998
SWL2010-05_1#383 | PAG-3 | 315 | 2011-02-17 12:12 | 24C | 0.92526 | 1.4046e-06 | 8.1062e-06 | 0.92527 | 32.0783 | 5.2235e-05
SWL2010-05_1#384 | PAG-3 | 315 | 2011-02-17 12:14 | 24C | 0.92524 | 2.2391e-06 | 7.987e-06 | 0.92525 | 32.0776 | 8.5299e-05 | Duplicate
SWL2010-05_1#385 | PAG-3 | 316 | 2011-02-17 12:17 | 24C | 0.92513 | 2.7435e-06 | 8.0466e-06 | 0.92513 | 32.0731 | 0.00010656
SWL2010-05_1#386 | PAG-3 | 317 | 2011-02-17 12:19 | 24C | 0.90869 | 2.7425e-06 | 8.1062e-06 | 0.9087 | 31.436 | 0.00010657
SWL2010-05_1#387 | PAG-3 | 318 | 2011-02-17 12:22 | 24C | 0.90765 | 3.6104e-06 | 7.987e-06 | 0.90766 | 31.3957 | 0.00013711
SWL2010-05_1#388 | BCL-7 | 319 | 2011-02-17 12:24 | 24C | 0.94684 | 4.3337e-06 | 7.987e-06 | 0.94685 | 32.918 | 0.00016797
SWL2010-05_1#389 | BCL-7 | 320 | 2011-02-17 12:26 | 24C | 0.94915 | 2.4909e-06 | 8.1062e-06 | 0.94916 | 33.008 | 9.5561e-05
SWL2010-05_1#390 | BCL-7 | 320 | 2011-02-17 12:29 | 24C | 0.94905 | 2.4909e-06 | 8.0466e-06 | 0.94906 | 33.0041 | 9.6609e-05 | Duplicate
SWL2010-05_1#391 | BCL-7 | 321 | 2011-02-17 12:32 | 24C | 0.93003 | 2.8343e-06 | 8.1062e-06 | 0.93004 | 32.2637 | 0.00011027
SWL2010-05_1#392 | BCL-7 | 322 | 2011-02-17 12:34 | 24C | 0.91856 | 3.4553e-06 | 8.1062e-06 | 0.91857 | 31.8185 | 0.00013659
SWL2010-05_1#393 | LOOP 18 | LOOP 18 | 2011-02-17 12:37 | 24C | 0.92059 | 2.7425e-06 | 7.987e-06 | 0.9206 | 31.8972 | 0.00010448
SWL2010-05_1#394 | Reference Water | CB-6 | 2011-02-17 12:39 | 24C | 0.99897 | 1.1396e-06 | 7.987e-06 | 0.99897 | 34.9596 | 4.5227e-05
SWL2010-05_1#395 | NP-3 | 21 | 2011-02-17 12:48 | 24C | 0.97218 | 2.4912e-06 | 8.1062e-06 | 0.97219 | 33.9081 | 9.6976e-05
SWL2010-05_1#396 | NP-3 | 22 | 2011-02-17 12:50 | 24C | 0.9696 | 3.1253e-06 | 8.1062e-06 | 0.96961 | 33.8072 | 0.00012381
SWL2010-05_1#397 | NP-3 | 23 | 2011-02-17 12:52 | 24C | 0.96901 | 2.2401e-06 | 8.0466e-06 | 0.96901 | 33.7838 | 8.7077e-05
SWL2010-05_1#398 | NP-3 | 24 | 2011-02-17 12:55 | 24C | 0.96876 | 2.2391e-06 | 8.1062e-06 | 0.96877 | 33.7741 | 8.8738e-05
SWL2010-05_1#399 | NP-3 | 25 | 2011-02-17 12:57 | 24C | 0.96655 | 1.1199e-06 | 7.987e-06 | 0.96656 | 33.6876 | 4.265e-05
SWL2010-05_1#400 | NP-3 | 26 | 2011-02-17 12:59 | 24C | 0.96249 | 2.226e-06 | 8.1062e-06 | 0.9625 | 33.5289 | 8.5352e-05
SWL2010-05_1#401 | NP-3 | 27 | 2011-02-17 13:02 | 24C | 0.95512 | 1.1196e-06 | 8.1062e-06 | 0.95513 | 33.2408 | 4.4374e-05
SWL2010-05_1#402 | NP-3 | 28 | 2011-02-17 13:04 | 24C | 0.94269 | 2.8897e-06 | 8.1062e-06 | 0.9427 | 32.7563 | 0.00011317
SWL2010-05_1#403 | NP-3 | 29 | 2011-02-17 13:06 | 24C | 0.93555 | 3.4551e-06 | 7.987e-06 | 0.93555 | 32.4781 | 0.00013412
SWL2010-05_1#404 | NP-3 | 30 | 2011-02-17 13:08 | 24C | 0.93235 | 3.1311e-06 | 8.1062e-06 | 0.93235 | 32.3537 | 0.00012313
SWL2010-05_1#405 | NP-3 | 31 | 2011-02-17 13:10 | 24C | 0.92932 | 3.3595e-06 | 7.987e-06 | 0.92933 | 32.2362 | 0.00012795
SWL2010-05_1#406 | NP-3 | 32 | 2011-02-17 13:12 | 24C | 0.92544 | 2.2276e-06 | 8.1062e-06 | 0.92545 | 32.0855 | 8.7239e-05
SWL2010-05_1#407 | NP-3 | 33 | 2011-02-17 13:15 | 24C | 0.92276 | 3.466e-06 | 8.0466e-06 | 0.92277 | 31.9814 | 0.00013375
SWL2010-05_1#408 | NP-3 | 34 | 2011-02-17 13:17 | 24C | 0.9228 | 3.2144e-06 | 8.1062e-06 | 0.92281 | 31.9827 | 0.00012686
SWL2010-05_1#409 | LOOP - 1 | LOOP - 1 | 2011-02-17 13:20 | 24C | 0.92841 | 2.4909e-06 | 8.0466e-06 | 0.92842 | 32.2007 | 9.4928e-05 | July 10/10 02:03
SWL2010-05_1#410 | NP-5 | 35 | 2011-02-17 13:22 | 24C | 0.98441 | 3.623e-06 | 8.0466e-06 | 0.98442 | 34.3877 | 0.0001426
SWL2010-05_1#411 | NP-5 | 36 | 2011-02-17 13:25 | 24C | 0.9818 | 3.2141e-06 | 8.0466e-06 | 0.98181 | 34.2852 | 0.0001289
SWL2010-05_1#412 | NP-5 | 37 | 2011-02-17 13:27 | 24C | 0.97763 | 2.4915e-06 | 8.1062e-06 | 0.97764 | 34.1217 | 9.4938e-05
SWL2010-05_1#413 | NP-5 | 38 | 2011-02-17 13:32 | 24C | 0.97533 | 6.1455e-06 | 8.1062e-06 | 0.97534 | 34.0315 | 0.00024173
SWL2010-05_1#414 | NP-5 | 38 | 2011-02-17 13:37 | 24C | 0.97444 | 2.9907e-06 | 8.0466e-06 | 0.97444 | 33.9964 | 0.0001146 | Duplicate??? Salinity values very different….
SWL2010-05_1#415 | NP-5 | 39 | 2011-02-17 13:40 | 24C | 0.97134 | 2.8904e-06 | 7.9274e-06 | 0.97135 | 33.8752 | 0.00011317
SWL2010-05_1#416 | NP-5 | 40 | 2011-02-17 13:42 | 24C | 0.97014 | 3.9746e-06 | 7.987e-06 | 0.97015 | 33.8282 | 0.00015494
SWL2010-05_1#417 | NP-5 | 41 | 2011-02-17 13:46 | 24C | 0.96862 | 2.2394e-06 | 7.987e-06 | 0.96863 | 33.7686 | 8.7876e-05
SWL2010-05_1#418 | Reference Water | CB-6 | 2011-02-17 13:48 | 24C | 0.99889 | 1.7705e-06 | 7.987e-06 | 0.9989 | 34.9567 | 7.0829e-05

## Autosal-Data for Merging
Salinity data for SWL 2010-05
|  |  | Primary | [9 empty] | Duplicates
orig order | Salinometer ID | Station | Sample ID (bottle label) | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments | Sample ID (bottle label) | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
99 | SWL2010-05_1#96 | NP-1 | 1 | 2011-02-15 15:32 | 24C | 0.98465 | 2.2397e-06 | 5.0068e-06 | 0.98465 | 34.3968 | 8.6195e-05
100 | SWL2010-05_1#97 | NP-1 | 2 | 2011-02-15 15:34 | 24C | 0.98162 | 4.583e-06 | 4.9472e-06 | 0.98162 | 34.2778 | 0.00017845
101 | SWL2010-05_1#98 | NP-1 | 3 | 2011-02-15 15:36 | 24C | 0.97782 | 5.5594e-06 | 4.9472e-06 | 0.97783 | 34.129 | 0.00021878
102 | SWL2010-05_1#99 | NP-1 | 4 | 2011-02-15 15:40 | 24C | 0.97493 | 2.4909e-06 | 5.0068e-06 | 0.97493 | 34.0155 | 9.869e-05
103 | SWL2010-05_1#100 | NP-1 | 5 | 2011-02-15 15:42 | 24C | 0.97314 | 2.4909e-06 | 5.0068e-06 | 0.97314 | 33.9455 | 9.4938e-05
104 | SWL2010-05_1#101 | NP-1 | 6 | 2011-02-15 15:44 | 24C | 0.97282 | 6.2115e-06 | 5.0664e-06 | 0.97282 | 33.9329 | 0.00024548 |  | 6 | 40589.6593 | 24C | 0.97324 | 2.4754e-06 | 5.0068e-06 | 0.97325 | 33.9496 | 9.4938e-05 | Duplicate
105 | SWL2010-05_1#102 | NP-1 | 7 | 2011-02-15 15:47 | 24C | 0.97244 | 3.1414e-06 | 5.0664e-06 | 0.97245 | 33.9182 | 0.00012304
107 | SWL2010-05_1#104 | NP-1 | 8 | 2011-02-15 15:52 | 24C | 0.97156 | 2.4915e-06 | 5.0068e-06 | 0.97157 | 33.8838 | 9.6599e-05
108 | SWL2010-05_1#105 | NP-1 | 9 | 2011-02-15 15:55 | 24C | 0.96875 | 2.8901e-06 | 4.9472e-06 | 0.96875 | 33.7735 | 0.00011151
109 | SWL2010-05_1#106 | NP-1 | 10 | 2011-02-15 15:57 | 24C | 0.96497 | 2.4918e-06 | 4.828e-06 | 0.96498 | 33.6259 | 9.7495e-05
110 | SWL2010-05_1#107 | NP-1 | 11 | 2011-02-15 15:59 | 24C | 0.95664 | 2.2394e-06 | 5.0068e-06 | 0.95665 | 33.3002 | 8.7077e-05
111 | SWL2010-05_1#108 | NP-1 | 12 | 2011-02-15 16:01 | 24C | 0.94522 | 3.5403e-06 | 4.9472e-06 | 0.94522 | 32.8547 | 0.00014162
112 | SWL2010-05_1#109 | NP-1 | 13 | 2011-02-15 16:03 | 24C | 0.93499 | 2.4909e-06 | 5.0068e-06 | 0.93499 | 32.4563 | 9.6599e-05
113 | SWL2010-05_1#110 | NP-1 | 14 | 2011-02-15 16:05 | 24C | 0.92543 | 1.4046e-06 | 5.0068e-06 | 0.92544 | 32.0849 | 5.5371e-05 |  | 14 | 40589.672 | 24C | 0.92547 | 1.3712e-06 | 4.9472e-06 | 0.92547 | 32.0862 | 5.2235e-05 | Duplicate
115 | SWL2010-05_1#112 | NP-1 | 15 | 2011-02-15 16:10 | 24C | 0.91924 | 2.4909e-06 | 5.0068e-06 | 0.91925 | 31.8448 | 9.4895e-05
116 | SWL2010-05_1#113 | NP-1 | 16 | 2011-02-15 16:14 | 24C | 0.91899 | 1.3712e-06 | 4.9472e-06 | 0.919 | 31.835 | 5.1714e-05
117 | SWL2010-05_1#114 | NP-1 | 17 | 2011-02-15 16:16 | 24C | 0.91891 | 2.475e-06 | 4.9472e-06 | 0.91891 | 31.8317 | 9.4364e-05
118 | SWL2010-05_1#115 | NP-3 | 18 | 2011-02-15 16:18 | 24C | 0.98397 | 2.2391e-06 | 5.0664e-06 | 0.98398 | 34.3703 | 8.5352e-05
119 | SWL2010-05_1#116 | NP-3 | 19 | 2011-02-15 16:20 | 24C | 0.9816 | 3.1418e-06 | 5.0068e-06 | 0.9816 | 34.277 | 0.00012103
120 | SWL2010-05_1#117 | NP-3 | 20 | 2011-02-15 16:22 | 24C | 0.9772 | 2.4909e-06 | 4.9472e-06 | 0.97721 | 34.1047 | 9.8725e-05 |  | 20 | 40589.6839 | 24C | 0.97724 | 1.1529e-06 | 4.9472e-06 | 0.97724 | 34.106 | 4.4374e-05 | Duplicate
407 | SWL2010-05_1#395 | NP-3 | 21 | 2011-02-17 12:48 | 24C | 0.97218 | 2.4912e-06 | 8.1062e-06 | 0.97219 | 33.9081 | 9.6976e-05
408 | SWL2010-05_1#396 | NP-3 | 22 | 2011-02-17 12:50 | 24C | 0.9696 | 3.1253e-06 | 8.1062e-06 | 0.96961 | 33.8072 | 0.00012381
409 | SWL2010-05_1#397 | NP-3 | 23 | 2011-02-17 12:52 | 24C | 0.96901 | 2.2401e-06 | 8.0466e-06 | 0.96901 | 33.7838 | 8.7077e-05
410 | SWL2010-05_1#398 | NP-3 | 24 | 2011-02-17 12:55 | 24C | 0.96876 | 2.2391e-06 | 8.1062e-06 | 0.96877 | 33.7741 | 8.8738e-05
411 | SWL2010-05_1#399 | NP-3 | 25 | 2011-02-17 12:57 | 24C | 0.96655 | 1.1199e-06 | 7.987e-06 | 0.96656 | 33.6876 | 4.265e-05
412 | SWL2010-05_1#400 | NP-3 | 26 | 2011-02-17 12:59 | 24C | 0.96249 | 2.226e-06 | 8.1062e-06 | 0.9625 | 33.5289 | 8.5352e-05
413 | SWL2010-05_1#401 | NP-3 | 27 | 2011-02-17 13:02 | 24C | 0.95512 | 1.1196e-06 | 8.1062e-06 | 0.95513 | 33.2408 | 4.4374e-05
414 | SWL2010-05_1#402 | NP-3 | 28 | 2011-02-17 13:04 | 24C | 0.94269 | 2.8897e-06 | 8.1062e-06 | 0.9427 | 32.7563 | 0.00011317
415 | SWL2010-05_1#403 | NP-3 | 29 | 2011-02-17 13:06 | 24C | 0.93555 | 3.4551e-06 | 7.987e-06 | 0.93555 | 32.4781 | 0.00013412
416 | SWL2010-05_1#404 | NP-3 | 30 | 2011-02-17 13:08 | 24C | 0.93235 | 3.1311e-06 | 8.1062e-06 | 0.93235 | 32.3537 | 0.00012313
417 | SWL2010-05_1#405 | NP-3 | 31 | 2011-02-17 13:10 | 24C | 0.92932 | 3.3595e-06 | 7.987e-06 | 0.92933 | 32.2362 | 0.00012795
418 | SWL2010-05_1#406 | NP-3 | 32 | 2011-02-17 13:12 | 24C | 0.92544 | 2.2276e-06 | 8.1062e-06 | 0.92545 | 32.0855 | 8.7239e-05
419 | SWL2010-05_1#407 | NP-3 | 33 | 2011-02-17 13:15 | 24C | 0.92276 | 3.466e-06 | 8.0466e-06 | 0.92277 | 31.9814 | 0.00013375
420 | SWL2010-05_1#408 | NP-3 | 34 | 2011-02-17 13:17 | 24C | 0.9228 | 3.2144e-06 | 8.1062e-06 | 0.92281 | 31.9827 | 0.00012686
422 | SWL2010-05_1#410 | NP-5 | 35 | 2011-02-17 13:22 | 24C | 0.98441 | 3.623e-06 | 8.0466e-06 | 0.98442 | 34.3877 | 0.0001426
423 | SWL2010-05_1#411 | NP-5 | 36 | 2011-02-17 13:25 | 24C | 0.9818 | 3.2141e-06 | 8.0466e-06 | 0.98181 | 34.2852 | 0.0001289
424 | SWL2010-05_1#412 | NP-5 | 37 | 2011-02-17 13:27 | 24C | 0.97763 | 2.4915e-06 | 8.1062e-06 | 0.97764 | 34.1217 | 9.4938e-05
425 | SWL2010-05_1#413 | NP-5 | 38 | 2011-02-17 13:32 | 24C | 0.97533 | 6.1455e-06 | 8.1062e-06 | 0.97534 | 34.0315 | 0.00024173 |  | 38 | 40591.567 | 24C | 0.97444 | 2.9907e-06 | 8.0466e-06 | 0.97444 | 33.9964 | 0.0001146 | Duplicate??? Salinity values very different….
427 | SWL2010-05_1#415 | NP-5 | 39 | 2011-02-17 13:40 | 24C | 0.97134 | 2.8904e-06 | 7.9274e-06 | 0.97135 | 33.8752 | 0.00011317
428 | SWL2010-05_1#416 | NP-5 | 40 | 2011-02-17 13:42 | 24C | 0.97014 | 3.9746e-06 | 7.987e-06 | 0.97015 | 33.8282 | 0.00015494
429 | SWL2010-05_1#417 | NP-5 | 41 | 2011-02-17 13:46 | 24C | 0.96862 | 2.2394e-06 | 7.987e-06 | 0.96863 | 33.7686 | 8.7876e-05
284 | SWL2010-05_1#276 | NP-5 | 42 | 2011-02-16 15:57 | 24C | 0.96704 | 2.4909e-06 | 9.0003e-06 | 0.96705 | 33.7068 | 9.6599e-05
285 | SWL2010-05_1#277 | NP-5 | 43 | 2011-02-16 16:00 | 24C | 0.96298 | 1.6739e-06 | 9.0599e-06 | 0.96299 | 33.548 | 6.8483e-05
286 | SWL2010-05_1#278 | NP-5 | 44 | 2011-02-16 16:02 | 24C | 0.95624 | 2.4911e-06 | 8.9407e-06 | 0.95625 | 33.2847 | 9.869e-05 |  | 44 | 40590.6698 | 24C | 0.95628 | 1.1536e-06 | 9.0599e-06 | 0.95629 | 33.2864 | 4.265e-05 | Duplicate
288 | SWL2010-05_1#280 | NP-5 | 45 | 2011-02-16 16:07 | 24C | 0.9494 | 2.4908e-06 | 9.0599e-06 | 0.94941 | 33.018 | 9.6976e-05
289 | SWL2010-05_1#281 | NP-5 | 46 | 2011-02-16 16:09 | 24C | 0.93657 | 2.4908e-06 | 9.0003e-06 | 0.93658 | 32.5178 | 9.3005e-05
290 | SWL2010-05_1#282 | NP-5 | 47 | 2011-02-16 16:11 | 24C | 0.93527 | 2.4908e-06 | 8.9407e-06 | 0.93528 | 32.4674 | 9.6976e-05
291 |  |  | 48 | NO SAMPLE
292 | SWL2010-05_1#283 | NP-5 | 49 | 2011-02-16 16:14 | 24C | 0.93339 | 2.2405e-06 | 9.0599e-06 | 0.9334 | 32.3944 | 8.5406e-05
293 | SWL2010-05_1#284 | NP-5 | 50 | 2011-02-16 16:16 | 24C | 0.93231 | 3.6109e-06 | 9.0003e-06 | 0.93232 | 32.3524 | 0.00013904
294 | SWL2010-05_1#285 | NP-5 | 51 | 2011-02-16 16:18 | 24C | 0.93251 | 1.3712e-06 | 9.0599e-06 | 0.93252 | 32.36 | 5.4351e-05
296 | SWL2010-05_1#287 | NP-7 | 52 | 2011-02-16 16:23 | 24C | 0.98453 | 1.7597e-06 | 9.0599e-06 | 0.98453 | 34.3921 | 6.6764e-05
297 | SWL2010-05_1#288 | NP-7 | 53 | 2011-02-16 16:25 | 24C | 0.982 | 2.4911e-06 | 9.0003e-06 | 0.98201 | 34.2929 | 9.4885e-05
298 | SWL2010-05_1#289 | NP-7 | 54 | 2011-02-16 16:28 | 24C | 0.97858 | 1.3712e-06 | 9.0003e-06 | 0.97859 | 34.1589 | 5.5395e-05 |  | 54 | 40590.6876 | 24C | 0.97827 | 4.189e-06 | 9.0599e-06 | 0.97828 | 34.1468 | 0.00016484 | Duplicate ! Not close at all
300 | SWL2010-05_1#291 | NP-7 | 55 | 2011-02-16 16:33 | 24C | 0.97517 | 0 | 9.0599e-06 | 0.97518 | 34.0252 | 0
301 | SWL2010-05_1#292 | NP-7 | 56 | 2011-02-16 16:35 | 24C | 0.97322 | 2.7425e-06 | 9.0599e-06 | 0.97323 | 33.9487 | 0.00010865
302 | SWL2010-05_1#293 | NP-7 | 57 | 2011-02-16 16:37 | 24C | 0.97231 | 6.4306e-06 | 9.0599e-06 | 0.97232 | 33.9131 | 0.00025094
303 | SWL2010-05_1#294 | NP-7 | 58 | 2011-02-16 16:39 | 24C | 0.97246 | 3.6084e-06 | 9.0003e-06 | 0.97247 | 33.9191 | 0.00014251
304 | SWL2010-05_1#295 | NP-7 | 59 | 2011-02-16 16:41 | 24C | 0.97153 | 2.4912e-06 | 8.9407e-06 | 0.97153 | 33.8825 | 9.9745e-05
306 | SWL2010-05_1#297 | NP-7 | 61 | 2011-02-16 16:45 | 24C | 0.96418 | 1.1532e-06 | 9.0599e-06 | 0.96419 | 33.5951 | 4.5217e-05
148 | SWL2010-05_1#143 | NP-7 | 62 | 2011-02-15 17:34 | 24C | 0.95169 | 1.3713e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 5.2278e-05
149 | SWL2010-05_1#144 | NP-7 | 63 | 2011-02-15 17:37 | 24C | 0.94074 | 4.1859e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00016101
150 | SWL2010-05_1#145 | NP-7 | 64 | 2011-02-15 17:39 | 24C | 0.93822 | 2.8904e-06 | 5.0068e-06 | 0.93822 | 32.5819 | 0.00011603
151 | SWL2010-05_1#146 | NP-7 | 65 | 2011-02-15 17:41 | 24C | 0.93775 | 2.8904e-06 | 4.9472e-06 | 0.93775 | 32.5637 | 0.00011403
152 | SWL2010-05_1#147 | NP-7 | 66 | 2011-02-15 17:44 | 24C | 0.93771 | 2.7426e-06 | 5.0664e-06 | 0.93771 | 32.5621 | 0.00010451
153 | SWL2010-05_1#148 | NP-7 | 67 | 2011-02-15 17:46 | 24C | 0.93776 | 2.4925e-06 | 5.0664e-06 | 0.93777 | 32.5642 | 9.9734e-05
154 | SWL2010-05_1#149 | NP-7 | 68 | 2011-02-15 17:48 | 24C | 0.93776 | 3.4657e-06 | 5.0068e-06 | 0.93777 | 32.5643 | 0.00013619
159 | SWL2010-05_1#152 | NP-9 | 69 | 2011-02-16 09:41 | 24C | 0.98509 | 1.3884e-06 | 8.9407e-06 | 0.9851 | 34.4144 | 5.4326e-05
160 | SWL2010-05_1#153 | NP-9 | 70 | 2011-02-16 09:44 | 24C | 0.98327 | 2.4909e-06 | 9.0003e-06 | 0.98328 | 34.3427 | 9.4928e-05
161 | SWL2010-05_1#154 | NP-9 | 71 | 2011-02-16 09:46 | 24C | 0.98045 | 3.0387e-06 | 9.0003e-06 | 0.98046 | 34.2323 | 0.00012157 |  | 71 | 40590.4087 | 24C | 0.98043 | 3.1415e-06 | 9.0599e-06 | 0.98044 | 34.2313 | 0.00012514 | Duplicate
163 | SWL2010-05_1#156 | NP-9 | 72 | 2011-02-16 09:51 | 24C | 0.97769 | 2.4777e-06 | 8.9407e-06 | 0.97769 | 34.1238 | 9.9543e-05
164 | SWL2010-05_1#157 | NP-9 | 73 | 2011-02-16 09:53 | 24C | 0.9758 | 2.4744e-06 | 9.0599e-06 | 0.97581 | 34.0497 | 9.7029e-05
165 | SWL2010-05_1#158 | NP-9 | 74 | 2011-02-16 09:55 | 24C | 0.97525 | 1.8038e-06 | 8.9407e-06 | 0.97526 | 34.0284 | 6.8797e-05
166 | SWL2010-05_1#159 | NP-9 | 75 | 2011-02-16 09:57 | 24C | 0.97448 | 2.4909e-06 | 8.9407e-06 | 0.97449 | 33.9984 | 9.9198e-05
167 | SWL2010-05_1#160 | NP-9 | 76 | 2011-02-16 09:59 | 24C | 0.97388 | 2.4911e-06 | 8.9407e-06 | 0.97389 | 33.9745 | 9.6976e-05
168 | SWL2010-05_1#161 | NP-9 | 77 | 2011-02-16 10:01 | 24C | 0.97281 | 2.4911e-06 | 9.0599e-06 | 0.97282 | 33.9328 | 0.0001006
169 | SWL2010-05_1#162 | NP-9 | 78 | 2011-02-16 10:03 | 24C | 0.97154 | 2.0804e-06 | 9.0003e-06 | 0.97154 | 33.8829 | 7.889e-05
170 | SWL2010-05_1#163 | NP-9 | 79 | 2011-02-16 10:05 | 24C | 0.96575 | 2.4915e-06 | 8.9407e-06 | 0.96576 | 33.6566 | 9.3856e-05
171 | SWL2010-05_1#164 | NP-9 | 80 | 2011-02-16 10:07 | 24C | 0.95408 | 1.1199e-06 | 9.0599e-06 | 0.95408 | 33.2002 | 4.5227e-05
172 | SWL2010-05_1#165 | NP-9 | 81 | 2011-02-16 10:10 | 24C | 0.9495 | 1.372e-06 | 9.0003e-06 | 0.94951 | 33.0216 | 5.4326e-05
173 | SWL2010-05_1#166 | NP-9 | 82 | 2011-02-16 10:12 | 24C | 0.94642 | 2.7426e-06 | 9.0599e-06 | 0.94643 | 32.9015 | 0.00010761 |  | 82 | 40590.4284 | 24C | 0.94649 | 2.4909e-06 | 9.0003e-06 | 0.94649 | 32.9042 | 9.4885e-05 | Duplicate
176 | SWL2010-05_1#169 | NP-9 | 83 | 2011-02-16 10:38 | 24C | 0.94606 | 2.8877e-06 | 8.9407e-06 | 0.94607 | 32.8875 | 0.00011008
177 | SWL2010-05_1#170 | NP-9 | 84 | 2011-02-16 10:40 | 24C | 0.94603 | 2.7434e-06 | 9.0003e-06 | 0.94604 | 32.8865 | 0.00011074
178 | SWL2010-05_1#171 | NP-9 | 85 | 2011-02-16 10:42 | 24C | 0.94597 | 3.143e-06 | 8.9407e-06 | 0.94598 | 32.884 | 0.00012347
179 | SWL2010-05_1#172 | NP-11 | 86 | 2011-02-16 10:45 | 24C | 0.98548 | 3.4658e-06 | 8.9407e-06 | 0.98549 | 34.4296 | 0.00013656
180 | SWL2010-05_1#173 | NP-11 | 87 | 2011-02-16 10:48 | 24C | 0.98357 | 4.2256e-06 | 9.0599e-06 | 0.98358 | 34.3548 | 0.00016833
181 | SWL2010-05_1#174 | NP-11 | 88 | 2011-02-16 10:50 | 24C | 0.98046 | 2.7426e-06 | 9.0003e-06 | 0.98047 | 34.2327 | 0.00010865 |  | 88 | 40590.453 | 24C | 0.98051 | 3.2144e-06 | 9.0599e-06 | 0.98052 | 34.2346 | 0.00012553 | Duplicate
183 | SWL2010-05_1#176 | NP-11 | 89 | 2011-02-16 10:55 | 24C | 0.9779 | 1.3712e-06 | 9.0599e-06 | 0.97791 | 34.1323 | 5.5371e-05
184 | SWL2010-05_1#177 | NP-11 | 90 | 2011-02-16 10:57 | 24C | 0.97537 | 1.1196e-06 | 8.9407e-06 | 0.97538 | 34.0332 | 4.2703e-05
185 | SWL2010-05_1#178 | NP-11 | 91 | 2011-02-16 10:59 | 24C | 0.97502 | 2.4908e-06 | 9.0599e-06 | 0.97503 | 34.0193 | 9.9553e-05
186 | SWL2010-05_1#179 | NP-11 | 92 | 2011-02-16 11:02 | 24C | 0.97402 | 2.8901e-06 | 9.0003e-06 | 0.97403 | 33.9802 | 0.00011145
187 | SWL2010-05_1#180 | NP-11 | 93 | 2011-02-16 11:04 | 24C | 0.97333 | 2.6582e-06 | 9.0599e-06 | 0.97334 | 33.9532 | 0.00010466
188 | SWL2010-05_1#181 | NP-11 | 94 | 2011-02-16 11:07 | 24C | 0.97236 | 1.1203e-06 | 9.0003e-06 | 0.97237 | 33.9152 | 4.2703e-05
189 | SWL2010-05_1#182 | NP-11 | 95 | 2011-02-16 11:10 | 24C | 0.96972 | 2.2401e-06 | 9.0003e-06 | 0.96973 | 33.8116 | 9.0453e-05
190 | SWL2010-05_1#183 | NP-11 | 96 | 2011-02-16 11:12 | 24C | 0.96532 | 2.4909e-06 | 8.9407e-06 | 0.96533 | 33.6395 | 9.87e-05 |  | 96 | 40590.4683 | 24C | 0.9652 | 2.8901e-06 | 9.0003e-06 | 0.96521 | 33.6349 | 0.00011213 | Duplicate
192 | SWL2010-05_1#185 | NP-11 | 97 | 2011-02-16 11:17 | 24C | 0.95597 | 2.4911e-06 | 9.0599e-06 | 0.95598 | 33.2743 | 9.87e-05
193 | SWL2010-05_1#186 | NP-11 | 98 | 2011-02-16 11:19 | 24C | 0.9488 | 2.2268e-06 | 9.0599e-06 | 0.94881 | 32.9944 | 8.7015e-05
194 | SWL2010-05_1#187 | NP-11 | 99 | 2011-02-16 11:22 | 24C | 0.94613 | 2.2725e-06 | 8.9407e-06 | 0.94614 | 32.8902 | 8.9676e-05
195 | SWL2010-05_1#188 | NP-11 | 100 | 2011-02-16 11:24 | 24C | 0.94712 | 4.6354e-06 | 8.9407e-06 | 0.94712 | 32.9287 | 0.00017833
196 | SWL2010-05_1#189 | NP-11 | 101 | 2011-02-16 11:28 | 24C | 0.94016 | 3.1145e-06 | 9.0599e-06 | 0.94017 | 32.6577 | 0.00012001
197 | SWL2010-05_1#190 | NP-11 | 102 | 2011-02-16 11:31 | 24C | 0.93472 | 2.2391e-06 | 9.0003e-06 | 0.93473 | 32.446 | 8.7024e-05
75 | SWL2010-05_1#72 | UN-1 | 103 | 2011-02-15 14:01 | 24C | 0.92794 | 4.3844e-06 | 4.9472e-06 | 0.92794 | 32.1822 | 0.00017382
76 | SWL2010-05_1#73 | UN-1 | 104 | 2011-02-15 14:04 | 24C | 0.92807 | 2.4909e-06 | 5.0068e-06 | 0.92808 | 32.1874 | 9.4885e-05
77 | SWL2010-05_1#74 | UN-1 | 105 | 2011-02-15 14:06 | 24C | 0.92784 | 1.6605e-06 | 5.0068e-06 | 0.92785 | 32.1785 | 6.4968e-05 |  | 105 | 40589.5901 | 24C | 0.92786 | 2.2394e-06 | 5.0068e-06 | 0.92787 | 32.1793 | 8.7067e-05 | Duplicate
79 | SWL2010-05_1#76 | UN-1 | 106 | 2011-02-15 14:12 | 24C | 0.92778 | 3.5942e-06 | 5.0068e-06 | 0.92778 | 32.1759 | 0.00013907
80 | SWL2010-05_1#77 | UN-1 | 107 | 2011-02-15 14:17 | 24C | 0.92778 | 2.4777e-06 | 5.0068e-06 | 0.92779 | 32.1761 | 9.7169e-05
81 | SWL2010-05_1#78 | UN-1 | 108 | 2011-02-15 14:19 | 24C | 0.92445 | 3.9606e-06 | 5.0068e-06 | 0.92445 | 32.0466 | 0.00015496
83 | SWL2010-05_1#80 | BCL-1 | 109 | 2011-02-15 14:24 | 24C | 0.97395 | 3.4659e-06 | 5.0068e-06 | 0.97395 | 33.9771 | 0.00013341
84 | SWL2010-05_1#81 | BCL-1 | 110 | 2011-02-15 14:27 | 24C | 0.97332 | 2.4909e-06 | 5.0068e-06 | 0.97333 | 33.9528 | 9.6599e-05
85 | SWL2010-05_1#82 | BCL-1 | 111 | 2011-02-15 14:29 | 24C | 0.96425 | 2.4775e-06 | 4.9472e-06 | 0.96425 | 33.5976 | 9.4034e-05
86 | SWL2010-05_1#83 | BCL-1 | 112 | 2011-02-15 14:31 | 24C | 0.95742 | 3.8647e-06 | 5.0068e-06 | 0.95742 | 33.3306 | 0.00014911
87 | SWL2010-05_1#84 | BCL-1 | 113 | 2011-02-15 14:34 | 24C | 0.95723 | 2.8799e-06 | 5.0664e-06 | 0.95724 | 33.3233 | 0.00011318
88 | SWL2010-05_1#85 | BCL-1 | 114 | 2011-02-15 14:36 | 24C | 0.95619 | 2.7425e-06 | 5.0068e-06 | 0.95619 | 33.2826 | 0.00010656
89 | SWL2010-05_1#86 | BCL-1 | 115 | 2011-02-15 14:38 | 24C | 0.95378 | 2.4908e-06 | 5.0068e-06 | 0.95378 | 33.1884 | 9.6599e-05
90 | SWL2010-05_1#87 | BCL-1 | 116 | 2011-02-15 14:40 | 24C | 0.95367 | 2.2391e-06 | 4.9472e-06 | 0.95367 | 33.1842 | 8.5326e-05
91 | SWL2010-05_1#88 | BCL-1 | 117 | 2011-02-15 14:43 | 24C | 0.95169 | 2.4908e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 0.0001006
92 | SWL2010-05_1#89 | BCL-1 | 118 | 2011-02-15 14:45 | 24C | 0.94813 | 2.4911e-06 | 5.0068e-06 | 0.94814 | 32.9682 | 9.4885e-05
93 | SWL2010-05_1#90 | BCL-1 | 119 | 2011-02-15 14:48 | 24C | 0.94592 | 2.7425e-06 | 5.0068e-06 | 0.94592 | 32.8818 | 0.00010656
94 | SWL2010-05_1#91 | BCL-1 | 120 | 2011-02-15 14:50 | 24C | 0.94397 | 2.2391e-06 | 4.9472e-06 | 0.94397 | 32.8058 | 8.8728e-05
95 | SWL2010-05_1#92 | BCL-1 | 121 | 2011-02-15 14:52 | 24C | 0.94074 | 3.8626e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00014926
96 | SWL2010-05_1#93 | BCL-1 | 122 | 2011-02-15 14:55 | 24C | 0.93871 | 3.3437e-06 | 4.8876e-06 | 0.93871 | 32.6011 | 0.00013184 |  | 122 | 40589.6231 | 24C | 0.93877 | 1.3712e-06 | 5.0068e-06 | 0.93877 | 32.6033 | 5.4326e-05 | Duplicate
261 | SWL2010-05_1#253 | BCL-1 | 123 | 2011-02-16 14:45 | 24C | 0.9386 | 2.2391e-06 | 8.9407e-06 | 0.93861 | 32.5971 | 8.5406e-05
262 | SWL2010-05_1#254 | BCL-1 | 124 | 2011-02-16 14:47 | 24C | 0.93911 | 2.4775e-06 | 8.9407e-06 | 0.93912 | 32.6168 | 9.4938e-05
263 | SWL2010-05_1#255 | BCL-2 | 125 | 2011-02-16 14:51 | 24C | 0.98301 | 3.3587e-06 | 9.0599e-06 | 0.98302 | 34.3326 | 0.00013246
264 | SWL2010-05_1#256 | BCL-2 | 126 | 2011-02-16 14:53 | 24C | 0.97924 | 3.6104e-06 | 8.9407e-06 | 0.97925 | 34.1848 | 0.00014403
265 | SWL2010-05_1#257 | BCL-2 | 127 | 2011-02-16 14:55 | 24C | 0.97282 | 2.7426e-06 | 9.0003e-06 | 0.97283 | 33.9331 | 0.00010447 |  | 127 | 40590.623 | 24C | 0.97275 | 2.2394e-06 | 8.9407e-06 | 0.97276 | 33.9304 | 8.7025e-05 | Duplicate
267 | SWL2010-05_1#259 | BCL-2 | 128 | 2011-02-16 14:59 | 24C | 0.96596 | 2.8792e-06 | 8.9407e-06 | 0.96597 | 33.6646 | 0.00011145
268 | SWL2010-05_1#260 | BCL-2 | 129 | 2011-02-16 15:02 | 24C | 0.96067 | 4.334e-06 | 8.9407e-06 | 0.96068 | 33.4578 | 0.00017005
269 | SWL2010-05_1#261 | BCL-2 | 130 | 2011-02-16 15:04 | 24C | 0.95949 | 3.4657e-06 | 8.9407e-06 | 0.9595 | 33.4118 | 0.00013801
270 | SWL2010-05_1#262 | BCL-2 | 131 | 2011-02-16 15:06 | 24C | 0.95829 | 3.466e-06 | 8.9407e-06 | 0.9583 | 33.365 | 0.0001334
271 | SWL2010-05_1#263 | BCL-2 | 132 | 2011-02-16 15:08 | 24C | 0.95698 | 2.4908e-06 | 9.0599e-06 | 0.95699 | 33.3136 | 9.6599e-05
272 | SWL2010-05_1#264 | BCL-2 | 133 | 2011-02-16 15:10 | 24C | 0.95578 | 3.4354e-06 | 9.0003e-06 | 0.95579 | 33.2668 | 0.0001334
273 | SWL2010-05_1#265 | BCL-2 | 134 | 2011-02-16 15:12 | 24C | 0.95378 | 2.4909e-06 | 9.0003e-06 | 0.95379 | 33.1886 | 9.7462e-05
274 | SWL2010-05_1#266 | BCL-2 | 135 | 2011-02-16 15:14 | 24C | 0.95283 | 3.5942e-06 | 9.0599e-06 | 0.95284 | 33.1518 | 0.00014204
275 | SWL2010-05_1#267 | BCL-2 | 136 | 2011-02-16 15:17 | 24C | 0.95199 | 2.2394e-06 | 9.0599e-06 | 0.952 | 33.1187 | 9.0433e-05
276 | SWL2010-05_1#268 | BCL-2 | 137 | 2011-02-16 15:35 | 24C | 0.9489 | 2.4909e-06 | 9.0003e-06 | 0.94891 | 32.9984 | 9.869e-05 |  | 137 | 40590.6509 | 24C | 0.94888 | 3.3455e-06 | 9.0003e-06 | 0.94889 | 32.9977 | 0.00012999 | Duplicate
278 | SWL2010-05_1#270 | BCL-2 | 138 | 2011-02-16 15:40 | 24C | 0.94699 | 2.8907e-06 | 8.9407e-06 | 0.947 | 32.9239 | 0.00011317
279 | SWL2010-05_1#271 | BCL-2 | 139 | 2011-02-16 15:42 | 24C | 0.94524 | 3.1415e-06 | 8.9407e-06 | 0.94525 | 32.8557 | 0.00011967
280 | SWL2010-05_1#272 | BCL-2 | 140 | 2011-02-16 15:44 | 24C | 0.94027 | 2.7425e-06 | 8.9407e-06 | 0.94028 | 32.662 | 0.00010659
281 | SWL2010-05_1#273 | BCL-2 | 141 | 2011-02-16 15:46 | 24C | 0.93864 | 4.9165e-06 | 9.0599e-06 | 0.93864 | 32.5984 | 0.00019281
51 | SWL2010-05_1#48 | BCL Deep | 142 | 2011-02-15 12:55 | 24C | 0.98494 | 2.4912e-06 | 5.0068e-06 | 0.98494 | 34.4082 | 9.6976e-05 | SZ- added missing sample number 142
52 | SWL2010-05_1#49 | BCL Deep | 143 | 2011-02-15 12:58 | 24C | 0.98226 | 3.865e-06 | 5.0068e-06 | 0.98226 | 34.303 | 0.00014998 |  | 143 | 40589.5417 | 24C | 0.98233 | 7.0655e-06 | 4.9472e-06 | 0.98233 | 34.3057 | 0.0002782 | Duplicate
54 | SWL2010-05_1#51 | BCL Deep | 144 | 2011-02-15 13:03 | 24C | 0.97908 | 2.4909e-06 | 4.9472e-06 | 0.97908 | 34.1782 | 9.4885e-05
55 | SWL2010-05_1#52 | BCL Deep | 145 | 2011-02-15 13:05 | 24C | 0.97226 | 2.8907e-06 | 4.8876e-06 | 0.97227 | 33.9112 | 0.00011317
56 | SWL2010-05_1#53 | BCL Deep | 146 | 2011-02-15 13:09 | 24C | 0.96591 | 4.2741e-06 | 4.9472e-06 | 0.96592 | 33.6625 | 0.00016611
57 | SWL2010-05_1#54 | BCL Deep | 147 | 2011-02-15 13:11 | 24C | 0.96432 | 2.4911e-06 | 5.0068e-06 | 0.96433 | 33.6005 | 9.7495e-05
58 | SWL2010-05_1#55 | BCL Deep | 148 | 2011-02-15 13:14 | 24C | 0.96226 | 2.0947e-06 | 4.9472e-06 | 0.96227 | 33.52 | 8.3896e-05
59 | SWL2010-05_1#56 | BCL Deep | 149 | 2011-02-15 13:16 | 24C | 0.96088 | 2.4908e-06 | 5.0068e-06 | 0.96089 | 33.466 | 9.6609e-05 |  | 149 | 40589.5541 | 24C | 0.96061 | 2.7097e-06 | 4.9472e-06 | 0.96061 | 33.4552 | 0.00011074 | Duplicate
61 | SWL2010-05_1#58 | BCL Deep | 150 | 2011-02-15 13:21 | 24C | 0.95872 | 3.4392e-06 | 5.0068e-06 | 0.95873 | 33.3815 | 0.00013386
62 | SWL2010-05_1#59 | BCL Deep | 151 | 2011-02-15 13:23 | 24C | 0.95592 | 3.2641e-06 | 5.0068e-06 | 0.95592 | 33.272 | 0.00012596
63 | SWL2010-05_1#60 | BCL Deep | 152 | 2011-02-15 13:26 | 24C | 0.95382 | 2.5369e-06 | 4.9472e-06 | 0.95382 | 33.1901 | 9.7279e-05
64 | SWL2010-05_1#61 | BCL Deep | 153 | 2011-02-15 13:28 | 24C | 0.9505 | 2.7425e-06 | 5.0068e-06 | 0.9505 | 33.0603 | 0.00010659
65 | SWL2010-05_1#62 | BCL Deep | 154 | 2011-02-15 13:31 | 24C | 0.94726 | 3.4659e-06 | 5.0664e-06 | 0.94726 | 32.9341 | 0.00013596
66 | SWL2010-05_1#63 | BCL Deep | 155 | 2011-02-15 13:34 | 24C | 0.94555 | 1.1196e-06 | 5.0068e-06 | 0.94555 | 32.8674 | 4.265e-05
67 | SWL2010-05_1#64 | BCL Deep | 156 | 2011-02-15 13:36 | 24C | 0.94556 | 1.3713e-06 | 5.126e-06 | 0.94557 | 32.8681 | 5.4351e-05
68 | SWL2010-05_1#65 | BCL Deep | 157 | 2011-02-15 13:38 | 24C | 0.94537 | 1.3713e-06 | 5.0068e-06 | 0.94538 | 32.8606 | 5.2235e-05
69 | SWL2010-05_1#66 | BCL Deep | 158 | 2011-02-15 13:41 | 24C | 0.94539 | 2.4909e-06 | 4.9472e-06 | 0.94539 | 32.8612 | 9.4938e-05
70 | SWL2010-05_1#67 | BCL-3 | 159 | 2011-02-15 13:43 | 24C | 0.94184 | 3.1068e-06 | 5.0068e-06 | 0.94185 | 32.7231 | 0.00012131
71 | SWL2010-05_1#68 | BCL-3 | 160 | 2011-02-15 13:46 | 24C | 0.94185 | 4.1138e-06 | 5.0068e-06 | 0.94186 | 32.7234 | 0.000163
72 | SWL2010-05_1#69 | BCL-3 | 161 | 2011-02-15 13:48 | 24C | 0.9395 | 2.4909e-06 | 4.9472e-06 | 0.93951 | 32.6319 | 9.4885e-05
73 | SWL2010-05_1#70 | BCL-3 | 162 | 2011-02-15 13:50 | 24C | 0.9316 | 4.3361e-06 | 5.0068e-06 | 0.93161 | 32.3247 | 0.00016849
123 | SWL2010-05_1#120 | BCL-3 | 162 | 2011-02-15 16:35 | 24C | 0.93154 | 3.1338e-06 | 5.0068e-06 | 0.93155 | 32.3223 | 0.00012017
124 | SWL2010-05_1#121 | BCL-3 | 163 | 2011-02-15 16:38 | 24C | 0.92728 | 2.2391e-06 | 5.0068e-06 | 0.92728 | 32.1566 | 8.7048e-05
125 | SWL2010-05_1#122 | BCL-3 | 164 | 2011-02-15 16:40 | 24C | 0.92173 | 2.4909e-06 | 5.0068e-06 | 0.92174 | 31.9412 | 9.6496e-05
126 | SWL2010-05_1#123 | BCL-3 | 165 | 2011-02-15 16:42 | 24C | 0.91926 | 2.4069e-06 | 5.0068e-06 | 0.91927 | 31.8455 | 9.4514e-05
127 | SWL2010-05_1#124 | BCL-3 | 166 | 2011-02-15 16:45 | 24C | 0.91407 | 2.4925e-06 | 4.9472e-06 | 0.91408 | 31.6443 | 9.9652e-05
128 |  |  | 167 | NO SAMPLE
129 |  |  | 168 | NO SAMPLE
130 | SWL2010-05_1#125 | BCL-4 | 169 | 2011-02-15 16:47 | 24C | 0.93716 | 1.7705e-06 | 5.0068e-06 | 0.93717 | 32.5409 | 6.8797e-05
131 | SWL2010-05_1#126 | BCL-4 | 170 | 2011-02-15 16:50 | 24C | 0.93725 | 2.7432e-06 | 5.0068e-06 | 0.93725 | 32.5441 | 0.00010659
132 | SWL2010-05_1#127 | BCL-4 | 171 | 2011-02-15 16:52 | 24C | 0.93344 | 2.7425e-06 | 4.8876e-06 | 0.93344 | 32.3959 | 0.00010656
133 | SWL2010-05_1#128 | BCL-4 | 172 | 2011-02-15 16:54 | 24C | 0.9268 | 2.7425e-06 | 5.0068e-06 | 0.9268 | 32.1378 | 0.00010656 |  | 172 | 40589.7063 | 24C | 0.92694 | 2.7262e-06 | 4.8876e-06 | 0.92695 | 32.1435 | 0.00010656 | Duplicate
135 | SWL2010-05_1#130 | BCL-4 | 173 | 2011-02-15 17:00 | 24C | 0.91616 | 2.7432e-06 | 5.126e-06 | 0.91616 | 31.725 | 0.00010395
136 | SWL2010-05_1#131 | BCL-4 | 174 | 2011-02-15 17:02 | 24C | 0.90983 | 1.3712e-06 | 5.0664e-06 | 0.90984 | 31.48 | 5.2246e-05
137 | SWL2010-05_1#132 | BCL-4 | 175 | 2011-02-15 17:05 | 24C | 0.9073 | 2.4911e-06 | 4.9472e-06 | 0.9073 | 31.3818 | 9.7041e-05
139 | SWL2010-05_1#134 | BCL-5 | 176 | 2011-02-15 17:10 | 24C | 0.92218 | 2.2397e-06 | 4.9472e-06 | 0.92218 | 31.9586 | 8.5299e-05
140 | SWL2010-05_1#135 | BCL-5 | 177 | 2011-02-15 17:12 | 24C | 0.92203 | 1.3551e-06 | 5.0664e-06 | 0.92203 | 31.9528 | 5.1714e-05 |  | 177 | 40589.7185 | 24C | 0.92202 | 1.1199e-06 | 4.9472e-06 | 0.92202 | 31.9524 | 4.4782e-05 | Duplicate
142 | SWL2010-05_1#137 | BCL-5 | 178 | 2011-02-15 17:17 | 24C | 0.914 | 2.4912e-06 | 5.0068e-06 | 0.91401 | 31.6416 | 9.7041e-05
143 | SWL2010-05_1#138 | BCL-5 | 179 | 2011-02-15 17:19 | 24C | 0.90887 | 2.8792e-06 | 5.0068e-06 | 0.90887 | 31.4427 | 0.00011137
144 | SWL2010-05_1#139 | BCL-5 | 180 | 2011-02-15 17:21 | 24C | 0.90567 | 2.4908e-06 | 5.0068e-06 | 0.90567 | 31.3188 | 9.3946e-05
145 | SWL2010-05_1#140 | BCL-5 | 181 | 2011-02-15 17:24 | 24C | 0.89056 | 3.9607e-06 | 4.9472e-06 | 0.89057 | 30.7352 | 0.00015076
146 | SWL2010-05_1#141 | BCL-5 | 182 | 2011-02-15 17:26 | 24C | 0.88826 | 2.7432e-06 | 5.0068e-06 | 0.88827 | 30.6464 | 0.00010969
2 | SWL2010-05_1#1 | Slip 1 | 184 | 2011-02-15 09:29 | 24C | 0.92252 | 2.2391e-06 | 5.0664e-06 | 0.92253 | 31.9719 | 8.9577e-05
3 | SWL2010-05_1#2 | Slip 1 | 185 | 2011-02-15 09:32 | 24C | 0.91638 | 2.2391e-06 | 5.0664e-06 | 0.91638 | 31.7336 | 8.5313e-05
4 | SWL2010-05_1#3 | Slip 1 | 186 | 2011-02-15 09:34 | 24C | 0.91231 | 2.7433e-06 | 4.9472e-06 | 0.91231 | 31.5759 | 0.00010971
5 | SWL2010-05_1#4 | Slip 1 | 187 | 2011-02-15 09:37 | 24C | 0.91007 | 1.1196e-06 | 5.126e-06 | 0.91007 | 31.4891 | 4.265e-05 |  | 187 | 40589.4022 | 24C | 0.90998 | 2.7425e-06 | 5.126e-06 | 0.90999 | 31.4859 | 0.00010447 | Duplicate
7 | SWL2010-05_1#6 | Slip 1 | 188 | 2011-02-15 09:42 | 24C | 0.90646 | 2.7425e-06 | 4.8876e-06 | 0.90646 | 31.3493 | 0.0001071
8 | SWL2010-05_1#7 | Slip 1 | 189 | 2011-02-15 09:47 | 24C | 0.90355 | 2.2244e-06 | 5.0068e-06 | 0.90356 | 31.2369 | 8.5887e-05
9 |  |  | 190 | NO SAMPLE
10 | SWL2010-05_1#8 | slip 2 | 191 | 2011-02-15 09:50 | 24C | 0.92419 | 3.8749e-06 | 5.0068e-06 | 0.9242 | 32.0367 | 0.00015052
11 | SWL2010-05_1#9 | slip 2 | 192 | 2011-02-15 09:53 | 24C | 0.91821 | 2.4909e-06 | 5.0664e-06 | 0.91821 | 31.8045 | 9.6591e-05
12 | SWL2010-05_1#10 | slip 2 | 193 | 2011-02-15 09:56 | 24C | 0.91379 | 3.4658e-06 | 4.9472e-06 | 0.91379 | 31.6332 | 0.0001334 | Bottle marked 194 but niskin# was 4 concluding sample must be #193
13 | SWL2010-05_1#11 | slip 2 | 194 | 2011-02-15 09:59 | 24C | 0.90983 | 2.7261e-06 | 5.0068e-06 | 0.90983 | 31.4799 | 0.00010708 |  | 194 | 40589.4217 | 24C | 0.90986 | 2.4916e-06 | 5.0664e-06 | 0.90986 | 31.4809 | 9.6523e-05 | Duplicate
15 | SWL2010-05_1#13 | slip 2 | 195 | 2011-02-15 10:10 | 24C | 0.90394 | 5.4475e-06 | 5.0068e-06 | 0.90394 | 31.2519 | 0.00021052
16 | SWL2010-05_1#14 | slip 2 | 196 | 2011-02-15 10:13 | 24C | 0.89998 | 3.4658e-06 | 5.0068e-06 | 0.89999 | 31.0989 | 0.00013277
17 |  |  | 197 | NO SAMPLE
18 | SWL2010-05_1#15 | slip 3 | 198 | 2011-02-15 11:14 | 24C | 0.92901 | 1.1063e-06 | 5.0664e-06 | 0.92901 | 32.2238 | 4.2703e-05
19 | SWL2010-05_1#16 | slip 3 | 199 | 2011-02-15 11:16 | 24C | 0.92788 | 2.7425e-06 | 5.126e-06 | 0.92789 | 32.18 | 0.00010447
20 | SWL2010-05_1#17 | slip 3 | 200 | 2011-02-15 11:18 | 24C | 0.91852 | 2.8661e-06 | 4.9472e-06 | 0.91852 | 31.8166 | 0.0001118
21 | SWL2010-05_1#18 | slip 3 | 201 | 2011-02-15 11:20 | 24C | 0.91491 | 2.4744e-06 | 4.9472e-06 | 0.91492 | 31.6768 | 9.6496e-05 |  | 201 | 40589.4739 | 24C | 0.91492 | 2.4777e-06 | 5.0664e-06 | 0.91493 | 31.6771 | 9.7078e-05 | Duplicate
23 | SWL2010-05_1#20 | slip 3 | 202 | 2011-02-15 11:25 | 24C | 0.91323 | 3.6105e-06 | 5.0068e-06 | 0.91323 | 31.6115 | 0.00013926
24 | SWL2010-05_1#21 | slip 3 | 203 | 2011-02-15 11:27 | 24C | 0.90874 | 2.7425e-06 | 5.0068e-06 | 0.90874 | 31.4377 | 0.00010709
359 | SWL2010-05_1#348 | Slip-5 | 205 | 2011-02-17 10:43 | 24C | 0.93206 | 2.2398e-06 | 7.987e-06 | 0.93207 | 32.3425 | 8.801e-05 |  | 205 | 40591.4486 | 24C | 0.93205 | 2.7425e-06 | 7.987e-06 | 0.93206 | 32.3423 | 0.00010972 | Duplicate
361 | SWL2010-05_1#350 | Slip-5 | 206 | 2011-02-17 10:49 | 24C | 0.93202 | 3.1414e-06 | 8.1062e-06 | 0.93203 | 32.3409 | 0.00012618
362 | SWL2010-05_1#351 | Slip-5 | 207 | 2011-02-17 10:51 | 24C | 0.92459 | 2.4912e-06 | 7.9274e-06 | 0.9246 | 32.0523 | 9.4938e-05
363 | SWL2010-05_1#352 | Slip-5 | 208 | 2011-02-17 10:53 | 24C | 0.91867 | 2.4909e-06 | 8.0466e-06 | 0.91868 | 31.8227 | 9.4895e-05
364 | SWL2010-05_1#353 | Slip-5 | 209 | 2011-02-17 10:56 | 24C | 0.91403 | 1.7702e-06 | 8.0466e-06 | 0.91403 | 31.6425 | 6.9868e-05
365 | SWL2010-05_1#354 | Slip-5 | 210 | 2011-02-17 10:58 | 24C | 0.9139 | 3.334e-06 | 7.987e-06 | 0.91391 | 31.6379 | 0.00012956
366 |  |  | 211 | NO SAMPLE
367 | SWL2010-05_1#355 | Slip-4 | 212 | 2011-02-17 11:01 | 24C | 0.9341 | 4.2633e-06 | 7.987e-06 | 0.93411 | 32.422 | 0.00016473 | No Bottle #211
368 | SWL2010-05_1#356 | Slip-4 | 213 | 2011-02-17 11:04 | 24C | 0.93341 | 3.6104e-06 | 7.987e-06 | 0.93342 | 32.3952 | 0.00013965
369 | SWL2010-05_1#357 | Slip-4 | 214 | 2011-02-17 11:06 | 24C | 0.92376 | 2.7425e-06 | 7.987e-06 | 0.92377 | 32.0201 | 0.00010657 |  | 214 | 40591.4635 | 24C | 0.92382 | 2.7425e-06 | 8.0466e-06 | 0.92383 | 32.0223 | 0.00011077 | Duplicate
371 | SWL2010-05_1#359 | Slip-4 | 215 | 2011-02-17 11:10 | 24C | 0.92084 | 3.4497e-06 | 7.987e-06 | 0.92085 | 31.9067 | 0.00013584
372 | SWL2010-05_1#360 | Slip-4 | 216 | 2011-02-17 11:13 | 24C | 0.91378 | 3.5197e-06 | 7.987e-06 | 0.91379 | 31.633 | 0.00013827
373 | SWL2010-05_1#361 | Slip-4 | 217 | 2011-02-17 11:15 | 24C | 0.91361 | 3.1417e-06 | 7.9274e-06 | 0.91362 | 31.6266 | 0.00011936
374 | SWL2010-05_1#362 | BCL-6a | 218 | 2011-02-17 11:17 | 24C | 0.93275 | 2.4909e-06 | 7.987e-06 | 0.93276 | 32.3693 | 9.4885e-05
375 | SWL2010-05_1#363 | BCL-6a | 219 | 2011-02-17 11:20 | 24C | 0.93259 | 2.2398e-06 | 8.1062e-06 | 0.9326 | 32.3632 | 8.5299e-05 |  | 219 | 40591.4738 | 24C | 0.93293 | 3.3588e-06 | 7.987e-06 | 0.93293 | 32.3762 | 0.00013277 | Duplicate
377 | SWL2010-05_1#365 | BCL-6a | 220 | 2011-02-17 11:25 | 24C | 0.93203 | 3.1312e-06 | 8.1062e-06 | 0.93204 | 32.3415 | 0.00012108
378 | SWL2010-05_1#366 | BCL-6a | 221 | 2011-02-17 11:27 | 24C | 0.91993 | 3.1313e-06 | 8.1062e-06 | 0.91993 | 31.8713 | 0.00012365
379 | SWL2010-05_1#367 | BCL-6a | 222 | 2011-02-17 11:30 | 24C | 0.89369 | 5.3482e-06 | 8.1062e-06 | 0.8937 | 30.8561 | 0.00020704
199 | SWL2010-05_1#192 | UTBS-5 | 224 | 2011-02-16 11:41 | 24C | 0.93768 | 2.8901e-06 | 8.9407e-06 | 0.93769 | 32.5614 | 0.0001118 |  | 224 | 40590.4888 | 24C | 0.93764 | 3.2141e-06 | 9.0003e-06 | 0.93765 | 32.5597 | 0.00012601 | Duplicate
201 | SWL2010-05_1#194 | UTBS-5 | 225 | 2011-02-16 11:46 | 24C | 0.93599 | 2.1137e-06 | 9.0003e-06 | 0.93599 | 32.4952 | 8.1621e-05
202 | SWL2010-05_1#195 | UTBS-5 | 226 | 2011-02-16 11:48 | 24C | 0.93425 | 3.1414e-06 | 9.0003e-06 | 0.93426 | 32.4276 | 0.00012104
203 | SWL2010-05_1#196 | UTBS-5 | 227 | 2011-02-16 11:50 | 24C | 0.93391 | 1.7702e-06 | 9.0599e-06 | 0.93392 | 32.4145 | 6.8817e-05
204 | SWL2010-05_1#197 | UTBS-5 | 228 | 2011-02-16 11:52 | 24C | 0.93373 | 2.2397e-06 | 9.0003e-06 | 0.93374 | 32.4075 | 8.7013e-05
205 |  |  | 229 | NO SAMPLE
206 | SWL2010-05_1#198 | UTBS-2 | 230 | 2011-02-16 11:56 | 24C | 0.93698 | 4.0094e-06 | 9.0599e-06 | 0.93699 | 32.5339 | 0.00015582 | no 229
207 | SWL2010-05_1#199 | UTBS-2 | 231 | 2011-02-16 11:58 | 24C | 0.93695 | 2.2391e-06 | 9.0003e-06 | 0.93695 | 32.5326 | 8.7048e-05 |  | 231 | 40590.4999 | 24C | 0.93699 | 3.1414e-06 | 8.9407e-06 | 0.93699 | 32.5342 | 0.00012313 | Duplicate
209 | SWL2010-05_1#201 | UTBS-2 | 232 | 2011-02-16 12:02 | 24C | 0.93719 | 1.1532e-06 | 9.0599e-06 | 0.9372 | 32.5422 | 4.4364e-05
210 | SWL2010-05_1#202 | UTBS-2 | 233 | 2011-02-16 12:04 | 24C | 0.93146 | 3.3594e-06 | 9.0599e-06 | 0.93147 | 32.3192 | 0.0001316
211 | SWL2010-05_1#203 | UTBS-2 | 234 | 2011-02-16 12:06 | 24C | 0.93041 | 2.2397e-06 | 8.9407e-06 | 0.93042 | 32.2784 | 8.8728e-05
212 | SWL2010-05_1#204 | UTBS-4 | 235 | 2011-02-16 12:08 | 24C | 0.9439 | 4.6245e-06 | 8.9407e-06 | 0.94391 | 32.8034 | 0.00018283
213 | SWL2010-05_1#205 | UTBS-4 | 236 | 2011-02-16 12:11 | 24C | 0.9434 | 1.3713e-06 | 9.0599e-06 | 0.94341 | 32.784 | 5.2235e-05
214 | SWL2010-05_1#206 | UTBS-4 | 237 | 2011-02-16 12:13 | 24C | 0.9396 | 3.5403e-06 | 8.9407e-06 | 0.93961 | 32.6359 | 0.00013759
215 | SWL2010-05_1#207 | UTBS-4 | 238 | 2011-02-16 12:15 | 24C | 0.93587 | 3.332e-08 | 8.9407e-06 | 0.93588 | 32.4907 | 0 |  | 238 | 40590.5119 | 24C | 0.93587 | 4.2257e-06 | 9.0003e-06 | 0.93588 | 32.4908 | 0.00016353 | Duplicate
217 | SWL2010-05_1#209 | UTBS-4 | 239 | 2011-02-16 12:19 | 24C | 0.93538 | 3.1417e-06 | 9.0599e-06 | 0.93539 | 32.4717 | 0.00012103 | SZ- changed sample number fro 238 to 239
218 | SWL2010-05_1#210 | UTBS-1 | 240 | 2011-02-16 12:22 | 24C | 0.93791 | 5.0346e-06 | 9.0599e-06 | 0.93792 | 32.5701 | 0.00019692 |  | 240 | 40590.5168 | 24C | 0.93774 | 2.4909e-06 | 9.0599e-06 | 0.93775 | 32.5634 | 9.87e-05 | Duplicate
220 | SWL2010-05_1#212 | UTBS-1 | 241 | 2011-02-16 12:27 | 24C | 0.93771 | 3.466e-06 | 8.9407e-06 | 0.93772 | 32.5625 | 0.00013595
221 | SWL2010-05_1#213 | UTBS-1 | 242 | 2011-02-16 12:29 | 24C | 0.93744 | 2.2396e-06 | 8.9407e-06 | 0.93745 | 32.552 | 8.5335e-05
222 | SWL2010-05_1#214 | UTBS-1 | 243 | 2011-02-16 12:33 | 24C | 0.93498 | 2.7425e-06 | 9.0599e-06 | 0.93499 | 32.4563 | 0.00010973
223 | SWL2010-05_1#215 | UTBS-1 | 244 | 2011-02-16 12:35 | 24C | 0.93342 | 4.2256e-06 | 9.0599e-06 | 0.93343 | 32.3956 | 0.00016286
308 | SWL2010-05_1#299 | BRS-4 | 245 | 2011-02-16 16:55 | 24C | 0.9257 | 2.7426e-06 | 9.0599e-06 | 0.92571 | 32.0955 | 0.00010865
309 | SWL2010-05_1#300 | BRS-4 | 246 | 2011-02-16 16:58 | 24C | 0.9247 | 2.4915e-06 | 9.0599e-06 | 0.92471 | 32.0565 | 9.1806e-05 |  | 246 | 40590.709 | 24C | 0.9247 | 1.1196e-06 | 8.9407e-06 | 0.92471 | 32.0567 | 4.2703e-05 | Duplicate
311 | SWL2010-05_1#302 | BRS-4 | 247 | 2011-02-16 17:03 | 24C | 0.92459 | 3.2142e-06 | 9.0599e-06 | 0.9246 | 32.0524 | 0.00012519
312 | SWL2010-05_1#303 | BRS-4 | 248 | 2011-02-16 17:06 | 24C | 0.92447 | 2.8897e-06 | 8.9407e-06 | 0.92447 | 32.0475 | 0.00011317
313 | SWL2010-05_1#304 | BRS-4 | 249 | 2011-02-16 17:08 | 24C | 0.91952 | 2.2405e-06 | 9.0003e-06 | 0.91953 | 31.8555 | 8.5313e-05
314 | SWL2010-05_1#305 | BRS-5 | 250 | 2011-02-16 17:10 | 24C | 0.94025 | 2.2391e-06 | 9.0003e-06 | 0.94026 | 32.6611 | 8.5406e-05 |  | 250 | 40590.7164 | 24C | 0.94027 | 4.189e-06 | 9.0003e-06 | 0.94028 | 32.662 | 0.00016283 | Duplicate
316 | SWL2010-05_1#307 | BRS-5 | 251 | 2011-02-16 17:14 | 24C | 0.94044 | 2.7425e-06 | 8.9407e-06 | 0.94044 | 32.6684 | 0.00010761
317 | SWL2010-05_1#308 | BRS-5 | 252 | 2011-02-16 17:16 | 24C | 0.94067 | 2.4912e-06 | 9.0599e-06 | 0.94068 | 32.6775 | 0.0001006
318 | SWL2010-05_1#309 | BRS-5 | 253 | 2011-02-16 17:19 | 24C | 0.93773 | 2.4908e-06 | 8.9407e-06 | 0.93774 | 32.5632 | 9.6609e-05
319 | SWL2010-05_1#310 | BRS-5 | 254 | 2011-02-16 17:21 | 24C | 0.93082 | 3.4662e-06 | 8.9407e-06 | 0.93083 | 32.2945 | 0.00013655
320 | SWL2010-05_1#311 | BRS-3 | 255 | 2011-02-16 17:24 | 24C | 0.9239 | 3.36e-06 | 8.9407e-06 | 0.92391 | 32.0254 | 0.0001298
321 | SWL2010-05_1#312 | BRS-3 | 256 | 2011-02-16 17:26 | 24C | 0.92428 | 2.7425e-06 | 8.9407e-06 | 0.92428 | 32.0401 | 0.00010451
322 | SWL2010-05_1#313 | BRS-3 | 257 | 2011-02-16 17:28 | 24C | 0.92355 | 2.7432e-06 | 9.0599e-06 | 0.92356 | 32.012 | 0.00010659
323 | SWL2010-05_1#314 | BRS-3 | 258 | 2011-02-16 17:30 | 24C | 0.91687 | 3.3591e-06 | 8.9407e-06 | 0.91688 | 31.753 | 0.00012967 |  | 258 | 40590.7308 | 24C | 0.91673 | 2.2394e-06 | 8.9407e-06 | 0.91673 | 31.7472 | 8.7006e-05 | Duplicate
325 | SWL2010-05_1#316 | BRS-3 | 259 | 2011-02-16 17:34 | 24C | 0.91666 | 3.8747e-06 | 9.0003e-06 | 0.91667 | 31.7446 | 0.00015
326 | SWL2010-05_1#317 | BRS-2 | 260 | 2011-02-16 17:37 | 24C | 0.91566 | 2.575e-06 | 9.0003e-06 | 0.91567 | 31.706 | 9.9156e-05 |  | 260 | 40590.7358 | 24C | 0.91558 | 2.7264e-06 | 8.9407e-06 | 0.91559 | 31.7029 | 0.00010396 | Duplicate
328 | SWL2010-05_1#319 | BRS-2 | 261 | 2011-02-16 17:42 | 24C | 0.91489 | 2.4908e-06 | 8.9407e-06 | 0.9149 | 31.6761 | 9.6991e-05
329 | SWL2010-05_1#320 | BRS-2 | 262 | 2011-02-16 17:45 | 24C | 0.91447 | 3.6082e-06 | 9.0599e-06 | 0.91448 | 31.66 | 0.00014255
330 | SWL2010-05_1#321 | BRS-2 | 263 | 2011-02-16 17:48 | 24C | 0.91443 | 2.7425e-06 | 8.9407e-06 | 0.91443 | 31.6581 | 0.00010708
331 | SWL2010-05_1#322 | BRS-2 | 264 | 2011-02-16 17:50 | 24C | 0.91417 | 2.4612e-06 | 9.0599e-06 | 0.91418 | 31.6483 | 9.4909e-05
27 | SWL2010-05_1#24 | BRS 0 | 265 | 2011-02-15 11:37 | 24C | 0.90598 | 2.7425e-06 | 5.0068e-06 | 0.90598 | 31.3307 | 0.0001071
28 | SWL2010-05_1#25 | BRS 0 | 266 | 2011-02-15 11:40 | 24C | 0.90589 | 2.2398e-06 | 5.0068e-06 | 0.9059 | 31.3275 | 8.5326e-05 |  | 266 | 40589.4874 | 24C | 0.90594 | 3.2141e-06 | 4.9472e-06 | 0.90595 | 31.3294 | 0.00012687 | Duplicate
30 | SWL2010-05_1#27 | BRS 0 | 267 | 2011-02-15 11:44 | 24C | 0.90582 | 3.4659e-06 | 4.9472e-06 | 0.90583 | 31.3248 | 0.00013567
31 | SWL2010-05_1#28 | BRS 0 | 268 | 2011-02-15 11:46 | 24C | 0.9058 | 2.2394e-06 | 5.0068e-06 | 0.90581 | 31.3241 | 8.7445e-05
32 | SWL2010-05_1#29 | BRS 0 | 269 | 2011-02-15 11:49 | 24C | 0.89603 | 2.2725e-06 | 5.0068e-06 | 0.89604 | 30.9463 | 8.6796e-05
33 | SWL2010-05_1#30 | UNT-1 | 270 | 2011-02-15 11:52 | 24C | 0.93132 | 2.4911e-06 | 4.9472e-06 | 0.93132 | 32.3136 | 9.87e-05 |  | 270 | 40589.501 | 24C | 0.93121 | 2.4911e-06 | 5.0664e-06 | 0.93121 | 32.3092 | 9.3109e-05 | Duplicate
35 | SWL2010-05_1#32 | UNT-1 | 271 | 2011-02-15 12:04 | 24C | 0.92956 | 1.4046e-06 | 5.0068e-06 | 0.92957 | 32.2454 | 5.2235e-05
36 | SWL2010-05_1#33 | UNT-1 | 272 | 2011-02-15 12:09 | 24C | 0.92947 | 2.8901e-06 | 5.0068e-06 | 0.92948 | 32.2418 | 0.00011145
37 | SWL2010-05_1#34 | UNT-1 | 273 | 2011-02-15 12:11 | 24C | 0.92551 | 2.8901e-06 | 5.0068e-06 | 0.92551 | 32.0878 | 0.00011603
38 | SWL2010-05_1#35 | Unt-2 | 274 | 2011-02-15 12:13 | 24C | 0.94216 | 3.1414e-06 | 5.0664e-06 | 0.94216 | 32.7353 | 0.00012315
39 | SWL2010-05_1#36 | Unt-2 | 275 | 2011-02-15 12:16 | 24C | 0.94278 | 3.8748e-06 | 5.0664e-06 | 0.94278 | 32.7594 | 0.00015052
40 | SWL2010-05_1#37 | Unt-2 | 276 | 2011-02-15 12:18 | 24C | 0.94185 | 3.5406e-06 | 5.0068e-06 | 0.94186 | 32.7235 | 0.00014029 |  | 276 | 40589.5138 | 24C | 0.94184 | 2.4908e-06 | 5.0068e-06 | 0.94184 | 32.7229 | 9.6642e-05 | Duplicate
42 | SWL2010-05_1#39 | Unt-2 | 277 | 2011-02-15 12:22 | 24C | 0.94198 | 3.3588e-06 | 4.9472e-06 | 0.94199 | 32.7285 | 0.00012945
43 | SWL2010-05_1#40 | Unt-2 | 278 | 2011-02-15 12:25 | 24C | 0.9418 | 2.8897e-06 | 5.0068e-06 | 0.94181 | 32.7215 | 0.00011404
44 | SWL2010-05_1#41 | Unt-3 | 279 | 2011-02-15 12:27 | 24C | 0.92777 | 2.0947e-06 | 5.0068e-06 | 0.92777 | 32.1755 | 8.1621e-05 |  | 279 | 40589.5205 | 24C | 0.92808 | 2.4909e-06 | 5.0664e-06 | 0.92809 | 32.1879 | 9.4938e-05 | Duplicate
46 | SWL2010-05_1#43 | Unt-3 | 280 | 2011-02-15 12:33 | 24C | 0.9282 | 3.3594e-06 | 5.0068e-06 | 0.92821 | 32.1925 | 0.00013089
47 | SWL2010-05_1#44 | Unt-3 | 281 | 2011-02-15 12:35 | 24C | 0.92763 | 6.664e-08 | 4.9472e-06 | 0.92764 | 32.1703 | 0
48 | SWL2010-05_1#45 | Unt-3 | 282 | 2011-02-15 12:38 | 24C | 0.92789 | 2.2391e-06 | 4.9472e-06 | 0.92789 | 32.1803 | 8.8728e-05
49 | SWL2010-05_1#46 | Unt-3 | 283 | 2011-02-15 12:40 | 24C | 0.91945 | 2.128e-06 | 5.126e-06 | 0.91946 | 31.8529 | 8.208e-05
225 | SWL2010-05_1#217 | UTN-4 | 284 | 2011-02-16 12:42 | 24C | 0.93597 | 1.7702e-06 | 9.0003e-06 | 0.93598 | 32.4946 | 6.8797e-05
226 | SWL2010-05_1#218 | UTN-4 | 285 | 2011-02-16 12:44 | 24C | 0.92955 | 3.865e-06 | 8.9407e-06 | 0.92956 | 32.2452 | 0.00014997
227 | SWL2010-05_1#219 | UTN-4 | 286 | 2011-02-16 12:46 | 24C | 0.92897 | 2.4775e-06 | 9.0599e-06 | 0.92897 | 32.2223 | 9.4034e-05 |  | 286 | 40590.5339 | 24C | 0.929 | 2.8901e-06 | 9.0599e-06 | 0.92901 | 32.2237 | 0.00011008 | Duplicate
229 | SWL2010-05_1#221 | UTN-4 | 287 | 2011-02-16 12:52 | 24C | 0.9282 | 2.128e-06 | 8.9407e-06 | 0.92821 | 32.1924 | 8.1169e-05
230 | SWL2010-05_1#222 | UTN-4 | 288 | 2011-02-16 12:54 | 24C | 0.92467 | 5.3436e-06 | 8.9407e-06 | 0.92467 | 32.0553 | 0.0002072
231 | SWL2010-05_1#223 | UTN-5 | 289 | 2011-02-16 12:56 | 24C | 0.93348 | 2.2391e-06 | 8.9407e-06 | 0.93348 | 32.3976 | 8.7239e-05
232 | SWL2010-05_1#224 | UTN-5 | 290 | 2011-02-16 13:00 | 24C | 0.93383 | 2.4908e-06 | 8.9407e-06 | 0.93384 | 32.4114 | 9.7001e-05
233 | SWL2010-05_1#225 | UTN-5 | 291 | 2011-02-16 13:02 | 24C | 0.93168 | 3.4517e-06 | 9.0003e-06 | 0.93169 | 32.3277 | 0.00013386
234 | SWL2010-05_1#226 | UTN-5 | 292 | 2011-02-16 13:04 | 24C | 0.92743 | 1.7705e-06 | 8.9407e-06 | 0.92744 | 32.1627 | 6.8817e-05 |  | 292 | 40590.5463 | 24C | 0.92727 | 4.1893e-06 | 8.9407e-06 | 0.92728 | 32.1566 | 0.00016234 | Duplicate
236 | SWL2010-05_1#228 | UTN-5 | 293 | 2011-02-16 13:09 | 24C | 0.92435 | 3.054e-06 | 8.9407e-06 | 0.92436 | 32.0429 | 0.00011839
237 | SWL2010-05_1#229 | UTN-6 | 294 | 2011-02-16 13:12 | 24C | 0.92537 | 3.4657e-06 | 8.9407e-06 | 0.92538 | 32.0826 | 0.00013633
238 | SWL2010-05_1#230 | UTN-6 | 295 | 2011-02-16 13:17 | 24C | 0.92523 | 3.865e-06 | 8.9407e-06 | 0.92523 | 32.077 | 0.00015135
239 | SWL2010-05_1#231 | UTN-6 | 296 | 2011-02-16 13:19 | 24C | 0.9245 | 2.7425e-06 | 9.0003e-06 | 0.92451 | 32.049 | 0.00010865
240 | SWL2010-05_1#232 | UTN-6 | 297 | 2011-02-16 13:23 | 24C | 0.92127 | 2.0873e-06 | 9.0599e-06 | 0.92128 | 31.9237 | 7.9556e-05 |  | 297 | 40590.5593 | 24C | 0.92125 | 3.4659e-06 | 9.0003e-06 | 0.92126 | 31.9227 | 0.00013374 | Duplicate
242 | SWL2010-05_1#234 | UTN-6 | 298 | 2011-02-16 13:28 | 24C | 0.91896 | 3.204e-06 | 9.0003e-06 | 0.91897 | 31.8339 | 0.00012417
243 | SWL2010-05_1#235 | UTN-7 | 299 | 2011-02-16 13:31 | 24C | 0.92586 | 2.0945e-06 | 9.0003e-06 | 0.92586 | 32.1015 | 8.1621e-05
244 | SWL2010-05_1#236 | UTN-7 | 300 | 2011-02-16 13:33 | 24C | 0.92569 | 3.2141e-06 | 8.9407e-06 | 0.9257 | 32.0952 | 0.00012708 |  | 300 | 40590.5657 | 24C | 0.92568 | 4.1785e-06 | 8.9407e-06 | 0.92569 | 32.0946 | 0.00016445 | Duplicate
246 | SWL2010-05_1#238 | UTN-7 | 301 | 2011-02-16 13:37 | 24C | 0.92383 | 3.53e-06 | 8.9407e-06 | 0.92384 | 32.0229 | 0.00013487
247 | SWL2010-05_1#239 | UTN-7 | 302 | 2011-02-16 13:39 | 24C | 0.92089 | 2.2574e-06 | 9.0003e-06 | 0.92089 | 31.9086 | 8.7711e-05
248 | SWL2010-05_1#240 | UTN-7 | 303 | 2011-02-16 13:41 | 24C | 0.9201 | 5.7875e-06 | 9.0003e-06 | 0.92011 | 31.8781 | 0.00022469
382 | SWL2010-05_1#370 | PAG-1 | 304 | 2011-02-17 11:42 | 24C | 0.92629 | 2.8794e-06 | 7.987e-06 | 0.9263 | 32.1184 | 0.00011151
383 | SWL2010-05_1#371 | PAG-1 | 305 | 2011-02-17 11:44 | 24C | 0.92629 | 1.4046e-06 | 8.0466e-06 | 0.9263 | 32.1184 | 5.2235e-05
384 | SWL2010-05_1#372 | PAG-1 | 306 | 2011-02-17 11:47 | 24C | 0.92272 | 3.2141e-06 | 8.1062e-06 | 0.92273 | 31.9798 | 0.00012654
385 | SWL2010-05_1#373 | PAG-1 | 307 | 2011-02-17 11:49 | 24C | 0.92229 | 3.3456e-06 | 8.1062e-06 | 0.9223 | 31.963 | 0.00013
386 | SWL2010-05_1#374 | PAG-1 | 308 | 2011-02-17 11:51 | 24C | 0.92226 | 1.3713e-06 | 8.0466e-06 | 0.92227 | 31.962 | 5.2246e-05 |  | 308 | 40591.4949 | 24C | 0.92216 | 1.1196e-06 | 8.1062e-06 | 0.92217 | 31.9579 | 4.265e-05 | Duplicate
388 | SWL2010-05_1#376 | PAG-2 | 309 | 2011-02-17 11:55 | 24C | 0.92762 | 3.2141e-06 | 7.9274e-06 | 0.92763 | 32.1699 | 0.00012244 |  | 309 | 40591.4983 | 24C | 0.92751 | 4.7275e-06 | 8.1062e-06 | 0.92752 | 32.1657 | 0.00018259 | Duplicate
390 | SWL2010-05_1#378 | PAG-2 | 310 | 2011-02-17 12:00 | 24C | 0.92741 | 2.7425e-06 | 8.1062e-06 | 0.92742 | 32.1619 | 0.00010865
391 | SWL2010-05_1#379 | PAG-2 | 311 | 2011-02-17 12:03 | 24C | 0.92685 | 2.4908e-06 | 8.1062e-06 | 0.92686 | 32.14 | 9.6599e-05
392 | SWL2010-05_1#380 | PAG-2 | 312 | 2011-02-17 12:05 | 24C | 0.92425 | 2.7425e-06 | 8.1062e-06 | 0.92426 | 32.0391 | 0.00010656
393 | SWL2010-05_1#381 | PAG-2 | 313 | 2011-02-17 12:07 | 24C | 0.91 | 3.3588e-06 | 8.1062e-06 | 0.91 | 31.4864 | 0.00013116
394 | SWL2010-05_1#382 | PAG-3 | 314 | 2011-02-17 12:10 | 24C | 0.9252 | 3.8654e-06 | 7.987e-06 | 0.92521 | 32.0759 | 0.00014998
395 | SWL2010-05_1#383 | PAG-3 | 315 | 2011-02-17 12:12 | 24C | 0.92526 | 1.4046e-06 | 8.1062e-06 | 0.92527 | 32.0783 | 5.2235e-05 |  | 315 | 40591.5099 | 24C | 0.92524 | 2.2391e-06 | 7.987e-06 | 0.92525 | 32.0776 | 8.5299e-05 | Duplicate
397 | SWL2010-05_1#385 | PAG-3 | 316 | 2011-02-17 12:17 | 24C | 0.92513 | 2.7435e-06 | 8.0466e-06 | 0.92513 | 32.0731 | 0.00010656
398 | SWL2010-05_1#386 | PAG-3 | 317 | 2011-02-17 12:19 | 24C | 0.90869 | 2.7425e-06 | 8.1062e-06 | 0.9087 | 31.436 | 0.00010657
399 | SWL2010-05_1#387 | PAG-3 | 318 | 2011-02-17 12:22 | 24C | 0.90765 | 3.6104e-06 | 7.987e-06 | 0.90766 | 31.3957 | 0.00013711
400 | SWL2010-05_1#388 | BCL-7 | 319 | 2011-02-17 12:24 | 24C | 0.94684 | 4.3337e-06 | 7.987e-06 | 0.94685 | 32.918 | 0.00016797
401 | SWL2010-05_1#389 | BCL-7 | 320 | 2011-02-17 12:26 | 24C | 0.94915 | 2.4909e-06 | 8.1062e-06 | 0.94916 | 33.008 | 9.5561e-05 |  | 320 | 40591.5202 | 24C | 0.94905 | 2.4909e-06 | 8.0466e-06 | 0.94906 | 33.0041 | 9.6609e-05 | Duplicate
403 | SWL2010-05_1#391 | BCL-7 | 321 | 2011-02-17 12:32 | 24C | 0.93003 | 2.8343e-06 | 8.1062e-06 | 0.93004 | 32.2637 | 0.00011027
404 | SWL2010-05_1#392 | BCL-7 | 322 | 2011-02-17 12:34 | 24C | 0.91856 | 3.4553e-06 | 8.1062e-06 | 0.91857 | 31.8185 | 0.00013659
334 | SWL2010-05_1#324 | BC-2 | 323 | 2011-02-17 09:18 | 24C | 0.96411 | 3.4658e-06 | 8.0466e-06 | 0.96412 | 33.5923 | 0.0001334
335 | SWL2010-05_1#325 | BC-2 | 324 | 2011-02-17 09:21 | 24C | 0.96216 | 2.7432e-06 | 7.987e-06 | 0.96216 | 33.5158 | 0.00010761
336 | SWL2010-05_1#326 | BC-2 | 325 | 2011-02-17 09:23 | 24C | 0.95522 | 2.7425e-06 | 7.9274e-06 | 0.95523 | 33.245 | 0.00010451
337 | SWL2010-05_1#327 | BC-2 | 326 | 2011-02-17 09:25 | 24C | 0.94689 | 2.89e-06 | 8.1062e-06 | 0.9469 | 32.9198 | 0.00011317
338 | SWL2010-05_1#328 | BC-2 | 327 | 2011-02-17 09:27 | 24C | 0.94209 | 4.4536e-06 | 7.987e-06 | 0.9421 | 32.7328 | 0.0001688 |  | 327 | 40591.3955 | 24C | 0.9421 | 3.2141e-06 | 8.1062e-06 | 0.94211 | 32.7334 | 0.00012361 | Duplicate
340 | SWL2010-05_1#330 | BC-2 | 328 | 2011-02-17 09:32 | 24C | 0.9279 | 1.1199e-06 | 8.1062e-06 | 0.92791 | 32.1809 | 4.265e-05
341 | SWL2010-05_1#331 | BC-2 | 329 | 2011-02-17 09:34 | 24C | 0.91093 | 3.8647e-06 | 8.0466e-06 | 0.91094 | 31.5227 | 0.00015041
342 | SWL2010-05_1#332 | BC-2 | 330 | 2011-02-17 09:36 | 24C | 0.88721 | 2.0947e-06 | 8.1062e-06 | 0.88721 | 30.6059 | 8.1765e-05
343 | SWL2010-05_1#333 | PAG-5 | 331 | 2011-02-17 09:39 | 24C | 0.95705 | 1.3713e-06 | 8.0466e-06 | 0.95706 | 33.3165 | 5.1197e-05
344 | SWL2010-05_1#334 | PAG-5 | 332 | 2011-02-17 09:41 | 24C | 0.9466 | 2.4908e-06 | 8.1062e-06 | 0.94661 | 32.9087 | 9.4938e-05 |  | 332 | 40591.4055 | 24C | 0.94651 | 3.3587e-06 | 8.1062e-06 | 0.94652 | 32.905 | 0.00013477 | Duplicate
346 | SWL2010-05_1#336 | PAG-5 | 333 | 2011-02-17 09:46 | 24C | 0.92577 | 1.7702e-06 | 8.0466e-06 | 0.92578 | 32.0983 | 6.9481e-05
347 |  |  | 334 | NO SAMPLE
348 | SWL2010-05_1#337 | PAG-5 | 335 | 2011-02-17 09:49 | 24C | 0.91206 | 3.8748e-06 | 7.987e-06 | 0.91206 | 31.5662 | 0.00015214 | No Bottle #334
349 | SWL2010-05_1#338 | PAG-5 | 336 | 2011-02-17 09:51 | 24C | 0.89843 | 2.4921e-06 | 7.987e-06 | 0.89843 | 31.0389 | 9.751e-05
350 | SWL2010-05_1#339 | PAG-5 | 337 | 2011-02-17 09:54 | 24C | 0.8833 | 2.4909e-06 | 8.0466e-06 | 0.8833 | 30.4551 | 9.7017e-05
351 | SWL2010-05_1#340 | PAG-6 | 338 | 2011-02-17 09:56 | 24C | 0.9294 | 3.6085e-06 | 8.0466e-06 | 0.92941 | 32.2391 | 0.00014094
352 | SWL2010-05_1#341 | PAG-6 | 339 | 2011-02-17 09:59 | 24C | 0.92864 | 1.1199e-06 | 8.1062e-06 | 0.92865 | 32.2096 | 4.1875e-05 |  | 339 | 40591.4172 | 24C | 0.92861 | 2.7097e-06 | 8.1062e-06 | 0.92862 | 32.2086 | 0.00010447 | Duplicate
354 | SWL2010-05_1#343 | PAG-6 | 340 | 2011-02-17 10:04 | 24C | 0.91602 | 2.7425e-06 | 8.0466e-06 | 0.91603 | 31.7199 | 0.00010347
355 | SWL2010-05_1#344 | PAG-6 | 341 | 2011-02-17 10:06 | 24C | 0.89761 | 3.8746e-06 | 8.1062e-06 | 0.89762 | 31.0074 | 0.0001526
356 | SWL2010-05_1#345 | PAG-6 | 342 | 2011-02-17 10:08 | 24C | 0.88673 | 3.1417e-06 | 8.1062e-06 | 0.88673 | 30.5874 | 0.00012138
357 | SWL2010-05_1#346 | PAG-6 | 343 | 2011-02-17 10:11 | 24C | 0.84755 | 3.8434e-06 | 7.987e-06 | 0.84756 | 29.0824 | 0.00014761
250 | SWL2010-05_1#242 | PAG-4 | 344 | 2011-02-16 14:09 | 24C | 0.94193 | 4.31e-06 | 8.9407e-06 | 0.94194 | 32.7266 | 0.00016895
251 | SWL2010-05_1#243 | PAG-4 | 345 | 2011-02-16 14:13 | 24C | 0.94076 | 3.2148e-06 | 9.0003e-06 | 0.94077 | 32.6811 | 0.00012554
252 | SWL2010-05_1#244 | PAG-4 | 346 | 2011-02-16 14:15 | 24C | 0.93897 | 3.2145e-06 | 9.0599e-06 | 0.93898 | 32.6115 | 0.00012554
253 | SWL2010-05_1#245 | PAG-4 | 347 | 2011-02-16 14:18 | 24C | 0.93838 | 2.8904e-06 | 9.0003e-06 | 0.93839 | 32.5884 | 0.00011213 |  | 347 | 40590.597 | 24C | 0.93835 | 3.4551e-06 | 8.9407e-06 | 0.93836 | 32.5874 | 0.00013203 | Duplicate
255 | SWL2010-05_1#247 | PAG-4 | 348 | 2011-02-16 14:22 | 24C | 0.93384 | 2.4911e-06 | 9.0003e-06 | 0.93385 | 32.4117 | 9.4885e-05
256 | SWL2010-05_1#248 | PAG-4 | 349 | 2011-02-16 14:24 | 24C | 0.92737 | 3.4658e-06 | 8.9407e-06 | 0.92738 | 32.1603 | 0.00013726
257 | SWL2010-05_1#249 | PAG-4 | 350 | 2011-02-16 14:26 | 24C | 0.92322 | 2.2262e-06 | 9.0599e-06 | 0.92323 | 31.999 | 8.7432e-05
305 | SWL2010-05_1#296 | NP-7 | 60 | 2011-02-16 16:43 | 24C | 0.97031 | 2.89e-06 | 8.9407e-06 | 0.97032 | 33.835 | 0.00011008 | SZ- changed sample number from 660 to 60
421 | SWL2010-05_1#409 | LOOP - 1 | LOOP - 1 | 2011-02-17 13:20 | 24C | 0.92841 | 2.4909e-06 | 8.0466e-06 | 0.92842 | 32.2007 | 9.4928e-05 | July 10/10 02:03
295 | SWL2010-05_1#286 | LOOP P2 | LOOP P2 | 2011-02-16 16:21 | 24C | 0.93207 | 2.0841e-06 | 9.0003e-06 | 0.93208 | 32.3431 | 8.1174e-05 | July 9/2010
155 | SWL2010-05_1#150 | LOOP-4 | LOOP 4 | 2011-02-15 17:50 | 24C | 0.93724 | 2.7434e-06 | 5.0664e-06 | 0.93725 | 32.544 | 0.00011074 | July 10/10 07:30
82 | SWL2010-05_1#79 | LOOP - 08 | LOOP - 08 | 2011-02-15 14:21 | 24C | 0.92405 | 2.4776e-06 | 4.9472e-06 | 0.92405 | 32.0312 | 9.7029e-05 | 2011-07-13 13:56
282 | SWL2010-05_1#274 | LOOP 9 |  | 2011-02-16 15:48 | 24C | 0.9386 | 3.865e-06 | 9.0599e-06 | 0.93861 | 32.597 | 0.00015042 | July 14/10 13:44, SZ-Change this from Loop 8 to Loop 9 as it fits timewise, missing 9 and already have an 8.
138 | SWL2010-05_1#133 | LOOP-12 | LOOP - 12 | 2011-02-15 17:07 | 24C | 0.90477 | 1.7702e-06 | 5.0068e-06 | 0.90477 | 31.2839 | 6.8797e-05 | July 15/10 22:23
380 | SWL2010-05_1#368 | LOOP 14 | LOOP 14 | 2011-02-17 11:33 | 24C | 0.89535 | 3.8647e-06 | 7.987e-06 | 0.89536 | 30.9202 | 0.0001502 | July 17/10 15:16
26 | SWL2010-05_1#23 | Loop 17 | Loop - 17 | 2011-02-15 11:34 | 24C | 0.9257 | 1.4046e-06 | 5.0068e-06 | 0.9257 | 32.0951 | 5.4351e-05 | 2011-07-19 05:25
405 | SWL2010-05_1#393 | LOOP 18 | LOOP 18 | 2011-02-17 12:37 | 24C | 0.92059 | 2.7425e-06 | 7.987e-06 | 0.9206 | 31.8972 | 0.00010448
258 | SWL2010-05_1#250 | 1 Coke Cap 35 |  | 2011-02-16 14:30 | 24C | 1.0029 | 3.4671e-06 | 8.9407e-06 | 1.0029 | 35.1155 | 0.00013732 | NO INSERT!
259 | SWL2010-05_1#251 | 8 Coke Cap 35 |  | 2011-02-16 14:33 | 24C | 1.0011 | 2.408e-06 | 8.9407e-06 | 1.0011 | 35.0425 | 9.4233e-05 | NO INSERT!

## Timeline Summary
Summary of Analysis - One row per Cage
SWL 2010-05
|  |  |  | Starting Standby # |  | Temperature
CASE | Date (d-mon-yr) | Station | Samples Analysed | Starting | Ending | Room | Sample | Comments
1 | 2011-02-15 | Slip 1 | 184-189 | 24+5987 |  | 22.8 | --
|  | slip 2 | 191-196
|  | slip 3 | 198-203 |  | ---
2 | 2011-02-15 | Loop 17 | LOOP | --- |  | 23 | 22.4
|  | BRS 0 | 265-269
|  | UNT-1 | 270-273
|  | Unt-2 | 274-278
|  | Unt-3 | 279-283 |  | 24+5989
3 | 2011-02-15 | BCL Deep | 142-158 | 24-5988 |  | 22.7 | 22.3
|  | BCL -3 | 159-261 |  | 24+5988
4 | 2011-02-15 | UN-1 | 103-108 | 24+5988 |  | 23.6 | 22.1
|  | LOOP-8 | LOOP
|  | BCL-1 | 109-122 |  | 24-5988
5 | 2011-02-15 | NP-1 | 2011-01-17 | 24+5988 |  | 23 | 22.9
|  | NP-3 | 18-20 |  | ---
6 | 2011-02-15 | BCL-3 | 162-166 | 24-5988 |  | 23.7 | --
|  | BCL-4 | 169-175
|  | LOOP-12 | LOOP
|  | BCL-5 | 176-182 |  | 24-5988
7 | 2011-02-15 | NP-7 | 62-68 | 24+5988 |  | 23.6 | --
|  | LOOP-4 | LOOP |  | 24+5988
| 2011-02-16 | NP-9 | 69-82 | 24+5988 | 24+5989 | 22.2 | --
8 | 2011-02-16 | NP-9 | 83-85 | 24+5989 |  | 23 | 22
|  | NP-11 | 86-102 |  | 24+5989
9 | 2011-02-16 | UTBS-5 | 224-228 | 24+5989 |  | 23.3 | 21.9
|  | UTBS-2 | 230-234
|  | UTBS-4 | 235-239
|  | UTBS-1 | 240-244 |  | 24+5990
10 | 2011-02-16 | UTN-4 | 284-288 | 24+5989 |  | 23.3 | 22.2
|  | UTN-5 | 289-293
|  | UTN-6 | 294-298
|  | UTN-7 | 299-303 |  | 24+5990
11 | 2011-02-16 | PAG-4 | 344-350 | 24+5990 |  | 23 | 22.1
|  | COKE CAP | 1
|  | COKE CAP | 8 |  | 24+5990
12 | 2011-02-16 | BCL-1 | 123, 124 | 24+5990 |  | 23.2 | 22
|  | BCL-2 | 125-141
|  | LOOP-8 | LOOP |  | 24+5990
13 | 2011-02-16 | NP-5 | 42-51 | 24+5990 |  | 23.2 | 22.1
|  | LOOP P2 | LOOP
|  | NP-7 | 52-61 |  | 24+5989
14 | 2011-02-16 | BRS-4 | 245-249 | 24+5989 |  | 23.2 | 21.6
|  | BRS-5 | 250-254
|  | BRS-3 | 255-259
|  | BRS-2 | 260-264 |  | 24+5989
15 | 2011-02-17 | BC-2 | 323-330 | 24+5989 |  | 23.3 | 22.8
|  | PAG-5 | 331-337
|  | PAG-6 | 338-343 |  | 24+5989
16 | 2011-02-17 | SLIP-5 | 205-210 | 24+5989 |  | 23.1 | 22.1
|  | SLIP-4 | 212-217
|  | BCL-6a | 218-222
|  | LOOP-14 | LOOP |  | ---
17 | 2011-02-17 | PAG-1 | 304-308 | 24+5989 |  | 23.3 | 22.9
|  | PAG-2 | 309-313
|  | PAG-3 | 314-318
|  | BCL-7 | 319-322
|  | LOOP-18 | LOOP |  | 24+5989
18 | 2011-02-17 | NP-3 | 21-34 | 24+5989 |  | 23.1 | 21.8
|  | LOOP-1 | LOOP
|  | NP-5 | 35-41 |  | 24+5989

## Cal Records
Example data in pale yellow
IAPSO STD Water |  | Actual Data forSWL 2010-05 Salinity in Black
|  | Standby # |  | Read Value | [5 empty] | IAPSO Standard Seawater Bottle Info.
Date/Time | Sample ID | Before | After | @ Start Calibration | @ End Calibration | Zero | Uncorrected Ratio | Correction | Adjusted Ratio | K15 | Batch # | Batch Date | Salinity | Comments
2010-09-18 13:07 | Calibration#1 | 24.6099 | 24.609 |  |  | 0 | 0.99996 | 8.1062e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2010-09-18 13:08 | Calibration#2 | 24.609 | 24.609 |  |  | 0 | 0.99997 | 1.9073e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2011-02-15 09:26 | Calibration#1 | 24+5987 | 24+5987 | --- | --- | 0 | 0.99997 | 5.0664e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2011-02-16 09:24 | Calibration#2 | 24+5988 | 24+5988 | --- | --- | 0 | 0.99995 | 1.6987e-05 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2011-02-16 09:38 | Calibration#3 | 24+5988 | 24+5988 | --- | --- | 0 | 0.99996 | 9.0599e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2011-02-17 09:11 | Calibration#4 | 24+5989 | 24+5989 | --- | --- | 0 | 0.99996 | 8.1062e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999

## DeepWaterCheck Records
SWL 2010-05 Salinity
Deep Water Check
Cast No. | Station | DWS ID | Sample mix | DateTime | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments | Average | STDEV |  | Calculated Salinity
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 11:43 | 0.99892 | 2.5747e-06 | -1.967e-06 | 0.99892 | 34.9576 | 0.00010077 |  |  |  |  | 34.9576
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 14:09 | 0.99891 | 3.6232e-06 | -2.0862e-06 | 0.99891 | 34.957 | 0.00014261 | DELETED READING 1 - LOWER THAN OTHER 2 |  |  |  | 34.957
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 15:12 | 0.9989 | 4.7529e-06 | -2.1458e-06 | 0.99889 | 34.9565 | 0.00018681 | DELETED READING 1 - LOW |  |  |  | 34.9565
SWL2010-05_1#22 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 11:31 | 0.99886 | 2.7435e-06 | 5.0068e-06 | 0.99887 | 34.9553 | 0.00010447 |  |  |  |  | 34.9553
SWL2010-05_1#47 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 12:42 | 0.99891 | 2.89e-06 | 5.0068e-06 | 0.99892 | 34.9575 | 0.00011354 |  |  |  |  | 34.9575
SWL2010-05_1#71 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 13:53 | 0.99888 | 3.0374e-06 | 5.126e-06 | 0.99889 | 34.9562 | 0.00011904 |  |  |  |  | 34.9562
SWL2010-05_1#95 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 15:01 | 0.99888 | 2.4615e-06 | 5.0068e-06 | 0.99889 | 34.9563 | 9.4938e-05 |  |  |  |  | 34.9563
SWL2010-05_1#119 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 16:27 | 0.99887 | 2.5749e-06 | 5.0068e-06 | 0.99887 | 34.9556 | 0.00010121 |  |  |  |  | 34.9556
SWL2010-05_1#142 | CB-6 | Reference Water | 1334/1335 | 2011-02-15 17:29 | 0.99889 | 3.9082e-06 | 5.0068e-06 | 0.9989 | 34.9567 | 0.00015494 |  |  |  |  | 34.9567
| CB-6 | Reference Water | 1334-1335 | 2011-02-15 17:54 | 0.9989 | 2.1649e-06 | 5.0664e-06 | 0.9989 | 34.9569 | 8.2858e-05 |  |  |  |  | 34.9569
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 10:19 | 0.99889 | 3.8587e-06 | 9.0599e-06 | 0.9989 | 34.9565 | 0.00014953 |  |  |  |  | 34.9565
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 11:33 | 0.99888 | 3.1867e-06 | 8.9407e-06 | 0.99889 | 34.9563 | 0.00012382 |  |  |  |  | 34.9563
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 12:37 | 0.99887 | 3.4658e-06 | 9.0003e-06 | 0.99888 | 34.9559 | 0.00013927 |  |  |  |  | 34.9559
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 13:47 | 0.99888 | 4.3195e-06 | 9.0003e-06 | 0.99889 | 34.9562 | 0.00017085 |  |  |  |  | 34.9562
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 14:39 | 0.99878 | 2.4909e-06 | 8.9407e-06 | 0.99879 | 34.9524 | 9.869e-05 |  |  |  |  | 34.9524
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 15:54 | 0.99887 | 2.89e-06 | 9.0003e-06 | 0.99888 | 34.9561 | 0.00011604 |  |  |  |  | 34.9561
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 16:49 | 0.99893 | 2.89e-06 | 8.9407e-06 | 0.99894 | 34.9584 | 0.00011213 |  |  |  |  | 34.9584
| CB-6 | Reference Water | 1334-1335 | 2011-02-16 17:52 | 0.99889 | 3.1418e-06 | 9.0003e-06 | 0.9989 | 34.9566 | 0.00012312 |  |  |  |  | 34.9566
| CB-6 | Reference Water | 1334-1335 | 2011-02-17 10:13 | 0.99875 | 2.8807e-06 | 7.9274e-06 | 0.99876 | 34.9513 | 0.00011216 |  |  |  |  | 34.9513
| CB-6 | Reference Water | 1334-1335 | 2011-02-17 11:36 | 0.99891 | 4.6466e-06 | 8.0466e-06 | 0.99892 | 34.9574 | 0.00018033 |  |  |  |  | 34.9574
| CB-6 | Reference Water | 1334-1335 | 2011-02-17 12:39 | 0.99897 | 1.1396e-06 | 7.987e-06 | 0.99897 | 34.9596 | 4.5227e-05 |  |  |  |  | 34.9596
| CB-6 | Reference Water | 1334-1335 | 2011-02-17 13:48 | 0.99889 | 1.7705e-06 | 7.987e-06 | 0.9989 | 34.9567 | 7.0829e-05 |  |  |  |  | 34.9567
[14 empty] | Avg | 34.9562
[14 empty] | std | 0.0018336

## RAW_SWL2010-05_1
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2011-02-15 09:26 | 24C | 0.99997 | 3.5403e-06 | 5.0664e-06 | 0.99997
SWL2010-05_1#1 | 184 Slip-1 | 2011-02-15 09:29 | 24C | 0.92252 | 2.2391e-06 | 5.0664e-06 | 0.92253 | 31.9719 | 8.9577e-05
SWL2010-05_1#2 | 185 " | 2011-02-15 09:32 | 24C | 0.91638 | 2.2391e-06 | 5.0664e-06 | 0.91638 | 31.7336 | 8.5313e-05
SWL2010-05_1#3 | 186 " | 2011-02-15 09:34 | 24C | 0.91231 | 2.7433e-06 | 4.9472e-06 | 0.91231 | 31.5759 | 0.00010971
SWL2010-05_1#4 | 187 " | 2011-02-15 09:37 | 24C | 0.91007 | 1.1196e-06 | 5.126e-06 | 0.91007 | 31.4891 | 4.265e-05
SWL2010-05_1#5 | 187 Duplicate | 2011-02-15 09:39 | 24C | 0.90998 | 2.7425e-06 | 5.126e-06 | 0.90999 | 31.4859 | 0.00010447
SWL2010-05_1#6 | 188 " | 2011-02-15 09:42 | 24C | 0.90646 | 2.7425e-06 | 4.8876e-06 | 0.90646 | 31.3493 | 0.0001071
SWL2010-05_1#7 | 189 " | 2011-02-15 09:47 | 24C | 0.90355 | 2.2244e-06 | 5.0068e-06 | 0.90356 | 31.2369 | 8.5887e-05
SWL2010-05_1#8 | 191 Slip-2 | 2011-02-15 09:50 | 24C | 0.92419 | 3.8749e-06 | 5.0068e-06 | 0.9242 | 32.0367 | 0.00015052
SWL2010-05_1#9 | 192 " | 2011-02-15 09:53 | 24C | 0.91821 | 2.4909e-06 | 5.0664e-06 | 0.91821 | 31.8045 | 9.6591e-05
SWL2010-05_1#10 | 194 " | 2011-02-15 09:56 | 24C | 0.91379 | 3.4658e-06 | 4.9472e-06 | 0.91379 | 31.6332 | 0.0001334 | Where is 193?
SWL2010-05_1#11 | 194 " | 2011-02-15 09:59 | 24C | 0.90983 | 2.7261e-06 | 5.0068e-06 | 0.90983 | 31.4799 | 0.00010708
SWL2010-05_1#12 | 191 Duplicate | 2011-02-15 10:07 | 24C | 0.90986 | 2.4916e-06 | 5.0664e-06 | 0.90986 | 31.4809 | 9.6523e-05
SWL2010-05_1#13 | 195 " | 2011-02-15 10:10 | 24C | 0.90394 | 5.4475e-06 | 5.0068e-06 | 0.90394 | 31.2519 | 0.00021052
SWL2010-05_1#14 | 196 " | 2011-02-15 10:13 | 24C | 0.89998 | 3.4658e-06 | 5.0068e-06 | 0.89999 | 31.0989 | 0.00013277
SWL2010-05_1#15 | 198 Slip-3 | 2011-02-15 11:14 | 24C | 0.92901 | 1.1063e-06 | 5.0664e-06 | 0.92901 | 32.2238 | 4.2703e-05
SWL2010-05_1#16 | 199 " | 2011-02-15 11:16 | 24C | 0.92788 | 2.7425e-06 | 5.126e-06 | 0.92789 | 32.18 | 0.00010447
SWL2010-05_1#17 | 200 " | 2011-02-15 11:18 | 24C | 0.91852 | 2.8661e-06 | 4.9472e-06 | 0.91852 | 31.8166 | 0.0001118
SWL2010-05_1#18 | 201 " | 2011-02-15 11:20 | 24C | 0.91491 | 2.4744e-06 | 4.9472e-06 | 0.91492 | 31.6768 | 9.6496e-05
SWL2010-05_1#19 | 201 " Duplicate | 2011-02-15 11:22 | 24C | 0.91492 | 2.4777e-06 | 5.0664e-06 | 0.91493 | 31.6771 | 9.7078e-05 | Duplicate
SWL2010-05_1#20 | 202 " | 2011-02-15 11:25 | 24C | 0.91323 | 3.6105e-06 | 5.0068e-06 | 0.91323 | 31.6115 | 0.00013926
SWL2010-05_1#21 | 203 " | 2011-02-15 11:27 | 24C | 0.90874 | 2.7425e-06 | 5.0068e-06 | 0.90874 | 31.4377 | 0.00010709
SWL2010-05_1#22 | Ref CB-06 | 2011-02-15 11:31 | 24C | 0.99886 | 2.7435e-06 | 5.0068e-06 | 0.99887 | 34.9553 | 0.00010447
SWL2010-05_1#23 | Loop - 17 July 19 5:25 | 2011-02-15 11:34 | 24C | 0.9257 | 1.4046e-06 | 5.0068e-06 | 0.9257 | 32.0951 | 5.4351e-05 | 2011-07-19 05:25
SWL2010-05_1#24 | 265 BRS-0 | 2011-02-15 11:37 | 24C | 0.90598 | 2.7425e-06 | 5.0068e-06 | 0.90598 | 31.3307 | 0.0001071
SWL2010-05_1#25 | 266 " | 2011-02-15 11:40 | 24C | 0.90589 | 2.2398e-06 | 5.0068e-06 | 0.9059 | 31.3275 | 8.5326e-05
SWL2010-05_1#26 | 266 Duplicate | 2011-02-15 11:42 | 24C | 0.90594 | 3.2141e-06 | 4.9472e-06 | 0.90595 | 31.3294 | 0.00012687 | Duplicate
SWL2010-05_1#27 | 267 " | 2011-02-15 11:44 | 24C | 0.90582 | 3.4659e-06 | 4.9472e-06 | 0.90583 | 31.3248 | 0.00013567
SWL2010-05_1#28 | 268 " | 2011-02-15 11:46 | 24C | 0.9058 | 2.2394e-06 | 5.0068e-06 | 0.90581 | 31.3241 | 8.7445e-05
SWL2010-05_1#29 | 269 " | 2011-02-15 11:49 | 24C | 0.89603 | 2.2725e-06 | 5.0068e-06 | 0.89604 | 30.9463 | 8.6796e-05
SWL2010-05_1#30 | 270 UTN - 1 | 2011-02-15 11:52 | 24C | 0.93132 | 2.4911e-06 | 4.9472e-06 | 0.93132 | 32.3136 | 9.87e-05
SWL2010-05_1#31 | 270 Duplicate | 2011-02-15 12:01 | 24C | 0.93121 | 2.4911e-06 | 5.0664e-06 | 0.93121 | 32.3092 | 9.3109e-05 | Duplicate
SWL2010-05_1#32 | 271 UTN - 1 | 2011-02-15 12:04 | 24C | 0.92956 | 1.4046e-06 | 5.0068e-06 | 0.92957 | 32.2454 | 5.2235e-05
SWL2010-05_1#33 | 272 " | 2011-02-15 12:09 | 24C | 0.92947 | 2.8901e-06 | 5.0068e-06 | 0.92948 | 32.2418 | 0.00011145
SWL2010-05_1#34 | 273 " | 2011-02-15 12:11 | 24C | 0.92551 | 2.8901e-06 | 5.0068e-06 | 0.92551 | 32.0878 | 0.00011603
SWL2010-05_1#35 | 274 UTN - 2 | 2011-02-15 12:13 | 24C | 0.94216 | 3.1414e-06 | 5.0664e-06 | 0.94216 | 32.7353 | 0.00012315
SWL2010-05_1#36 | 275 " | 2011-02-15 12:16 | 24C | 0.94278 | 3.8748e-06 | 5.0664e-06 | 0.94278 | 32.7594 | 0.00015052
SWL2010-05_1#37 | 276 " | 2011-02-15 12:18 | 24C | 0.94185 | 3.5406e-06 | 5.0068e-06 | 0.94186 | 32.7235 | 0.00014029
SWL2010-05_1#38 | 276 Duplicate | 2011-02-15 12:20 | 24C | 0.94184 | 2.4908e-06 | 5.0068e-06 | 0.94184 | 32.7229 | 9.6642e-05 | Duplicate
SWL2010-05_1#39 | 277 UTN -2 | 2011-02-15 12:22 | 24C | 0.94198 | 3.3588e-06 | 4.9472e-06 | 0.94199 | 32.7285 | 0.00012945
SWL2010-05_1#40 | 278 " | 2011-02-15 12:25 | 24C | 0.9418 | 2.8897e-06 | 5.0068e-06 | 0.94181 | 32.7215 | 0.00011404
SWL2010-05_1#41 | 279 UTN - 3 | 2011-02-15 12:27 | 24C | 0.92777 | 2.0947e-06 | 5.0068e-06 | 0.92777 | 32.1755 | 8.1621e-05
SWL2010-05_1#42 | 279 Duplicate | 2011-02-15 12:30 | 24C | 0.92808 | 2.4909e-06 | 5.0664e-06 | 0.92809 | 32.1879 | 9.4938e-05 | Duplicate
SWL2010-05_1#43 | 280 UTN - 3 | 2011-02-15 12:33 | 24C | 0.9282 | 3.3594e-06 | 5.0068e-06 | 0.92821 | 32.1925 | 0.00013089
SWL2010-05_1#44 | 281 " | 2011-02-15 12:35 | 24C | 0.92763 | 6.664e-08 | 4.9472e-06 | 0.92764 | 32.1703 | 0
SWL2010-05_1#45 | 282 " | 2011-02-15 12:38 | 24C | 0.92789 | 2.2391e-06 | 4.9472e-06 | 0.92789 | 32.1803 | 8.8728e-05
SWL2010-05_1#46 | 283 " | 2011-02-15 12:40 | 24C | 0.91945 | 2.128e-06 | 5.126e-06 | 0.91946 | 31.8529 | 8.208e-05
SWL2010-05_1#47 | Ref. CB-06 | 2011-02-15 12:42 | 24C | 0.99891 | 2.89e-06 | 5.0068e-06 | 0.99892 | 34.9575 | 0.00011354
SWL2010-05_1#48 | BCL Deep | 2011-02-15 12:55 | 24C | 0.98494 | 2.4912e-06 | 5.0068e-06 | 0.98494 | 34.4082 | 9.6976e-05
SWL2010-05_1#49 | 143 " | 2011-02-15 12:58 | 24C | 0.98226 | 3.865e-06 | 5.0068e-06 | 0.98226 | 34.303 | 0.00014998
SWL2010-05_1#50 | 143 Duplicate | 2011-02-15 13:00 | 24C | 0.98233 | 7.0655e-06 | 4.9472e-06 | 0.98233 | 34.3057 | 0.0002782 | Duplicate
SWL2010-05_1#51 | 144 BCL - Deep | 2011-02-15 13:03 | 24C | 0.97908 | 2.4909e-06 | 4.9472e-06 | 0.97908 | 34.1782 | 9.4885e-05
SWL2010-05_1#52 | 145 BCL - Deep | 2011-02-15 13:05 | 24C | 0.97226 | 2.8907e-06 | 4.8876e-06 | 0.97227 | 33.9112 | 0.00011317
SWL2010-05_1#53 | 146 " | 2011-02-15 13:09 | 24C | 0.96591 | 4.2741e-06 | 4.9472e-06 | 0.96592 | 33.6625 | 0.00016611
SWL2010-05_1#54 | 147 " | 2011-02-15 13:11 | 24C | 0.96432 | 2.4911e-06 | 5.0068e-06 | 0.96433 | 33.6005 | 9.7495e-05
SWL2010-05_1#55 | 148 " | 2011-02-15 13:14 | 24C | 0.96226 | 2.0947e-06 | 4.9472e-06 | 0.96227 | 33.52 | 8.3896e-05
SWL2010-05_1#56 | 149 " | 2011-02-15 13:16 | 24C | 0.96088 | 2.4908e-06 | 5.0068e-06 | 0.96089 | 33.466 | 9.6609e-05
SWL2010-05_1#57 | 149 Duplicate | 2011-02-15 13:18 | 24C | 0.96061 | 2.7097e-06 | 4.9472e-06 | 0.96061 | 33.4552 | 0.00011074 | Duplicate
SWL2010-05_1#58 | 150 BCL - Deep | 2011-02-15 13:21 | 24C | 0.95872 | 3.4392e-06 | 5.0068e-06 | 0.95873 | 33.3815 | 0.00013386
SWL2010-05_1#59 | 151 " | 2011-02-15 13:23 | 24C | 0.95592 | 3.2641e-06 | 5.0068e-06 | 0.95592 | 33.272 | 0.00012596
SWL2010-05_1#60 | 152 " | 2011-02-15 13:26 | 24C | 0.95382 | 2.5369e-06 | 4.9472e-06 | 0.95382 | 33.1901 | 9.7279e-05
SWL2010-05_1#61 | 153 " | 2011-02-15 13:28 | 24C | 0.9505 | 2.7425e-06 | 5.0068e-06 | 0.9505 | 33.0603 | 0.00010659
SWL2010-05_1#62 | 154 " | 2011-02-15 13:31 | 24C | 0.94726 | 3.4659e-06 | 5.0664e-06 | 0.94726 | 32.9341 | 0.00013596
SWL2010-05_1#63 | 155 " | 2011-02-15 13:34 | 24C | 0.94555 | 1.1196e-06 | 5.0068e-06 | 0.94555 | 32.8674 | 4.265e-05
SWL2010-05_1#64 | 156 " | 2011-02-15 13:36 | 24C | 0.94556 | 1.3713e-06 | 5.126e-06 | 0.94557 | 32.8681 | 5.4351e-05
SWL2010-05_1#65 | 157 " | 2011-02-15 13:38 | 24C | 0.94537 | 1.3713e-06 | 5.0068e-06 | 0.94538 | 32.8606 | 5.2235e-05
SWL2010-05_1#66 | 158 " | 2011-02-15 13:41 | 24C | 0.94539 | 2.4909e-06 | 4.9472e-06 | 0.94539 | 32.8612 | 9.4938e-05
SWL2010-05_1#67 | 159 BCL - 3 | 2011-02-15 13:43 | 24C | 0.94184 | 3.1068e-06 | 5.0068e-06 | 0.94185 | 32.7231 | 0.00012131
SWL2010-05_1#68 | 160 " | 2011-02-15 13:46 | 24C | 0.94185 | 4.1138e-06 | 5.0068e-06 | 0.94186 | 32.7234 | 0.000163
SWL2010-05_1#69 | 161 " | 2011-02-15 13:48 | 24C | 0.9395 | 2.4909e-06 | 4.9472e-06 | 0.93951 | 32.6319 | 9.4885e-05
SWL2010-05_1#70 | 162 " | 2011-02-15 13:50 | 24C | 0.9316 | 4.3361e-06 | 5.0068e-06 | 0.93161 | 32.3247 | 0.00016849
SWL2010-05_1#71 | Ref. CB-06 | 2011-02-15 13:53 | 24C | 0.99888 | 3.0374e-06 | 5.126e-06 | 0.99889 | 34.9562 | 0.00011904
SWL2010-05_1#72 | 103 UN - 1 | 2011-02-15 14:01 | 24C | 0.92794 | 4.3844e-06 | 4.9472e-06 | 0.92794 | 32.1822 | 0.00017382
SWL2010-05_1#73 | 104 " | 2011-02-15 14:04 | 24C | 0.92807 | 2.4909e-06 | 5.0068e-06 | 0.92808 | 32.1874 | 9.4885e-05
SWL2010-05_1#74 | 105 " | 2011-02-15 14:06 | 24C | 0.92784 | 1.6605e-06 | 5.0068e-06 | 0.92785 | 32.1785 | 6.4968e-05
SWL2010-05_1#75 | 105 Duplicate | 2011-02-15 14:10 | 24C | 0.92786 | 2.2394e-06 | 5.0068e-06 | 0.92787 | 32.1793 | 8.7067e-05 | Duplicate
SWL2010-05_1#76 | 106 UN - 1 | 2011-02-15 14:12 | 24C | 0.92778 | 3.5942e-06 | 5.0068e-06 | 0.92778 | 32.1759 | 0.00013907
SWL2010-05_1#77 | 107 " | 2011-02-15 14:17 | 24C | 0.92778 | 2.4777e-06 | 5.0068e-06 | 0.92779 | 32.1761 | 9.7169e-05
SWL2010-05_1#78 | 108 " | 2011-02-15 14:19 | 24C | 0.92445 | 3.9606e-06 | 5.0068e-06 | 0.92445 | 32.0466 | 0.00015496
SWL2010-05_1#79 | LOOP - 08 | 2011-02-15 14:21 | 24C | 0.92405 | 2.4776e-06 | 4.9472e-06 | 0.92405 | 32.0312 | 9.7029e-05 | 2011-07-13 13:56
SWL2010-05_1#80 | 109 BCL -1 | 2011-02-15 14:24 | 24C | 0.97395 | 3.4659e-06 | 5.0068e-06 | 0.97395 | 33.9771 | 0.00013341
SWL2010-05_1#81 | 110 " | 2011-02-15 14:27 | 24C | 0.97332 | 2.4909e-06 | 5.0068e-06 | 0.97333 | 33.9528 | 9.6599e-05
SWL2010-05_1#82 | 111 " | 2011-02-15 14:29 | 24C | 0.96425 | 2.4775e-06 | 4.9472e-06 | 0.96425 | 33.5976 | 9.4034e-05
SWL2010-05_1#83 | 112 " | 2011-02-15 14:31 | 24C | 0.95742 | 3.8647e-06 | 5.0068e-06 | 0.95742 | 33.3306 | 0.00014911
SWL2010-05_1#84 | 113 " | 2011-02-15 14:34 | 24C | 0.95723 | 2.8799e-06 | 5.0664e-06 | 0.95724 | 33.3233 | 0.00011318
SWL2010-05_1#85 | 114 " | 2011-02-15 14:36 | 24C | 0.95619 | 2.7425e-06 | 5.0068e-06 | 0.95619 | 33.2826 | 0.00010656
SWL2010-05_1#86 | 115 " | 2011-02-15 14:38 | 24C | 0.95378 | 2.4908e-06 | 5.0068e-06 | 0.95378 | 33.1884 | 9.6599e-05
SWL2010-05_1#87 | 116 " | 2011-02-15 14:40 | 24C | 0.95367 | 2.2391e-06 | 4.9472e-06 | 0.95367 | 33.1842 | 8.5326e-05
SWL2010-05_1#88 | 117 " | 2011-02-15 14:43 | 24C | 0.95169 | 2.4908e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 0.0001006
SWL2010-05_1#89 | 118 " | 2011-02-15 14:45 | 24C | 0.94813 | 2.4911e-06 | 5.0068e-06 | 0.94814 | 32.9682 | 9.4885e-05
SWL2010-05_1#90 | 119 " | 2011-02-15 14:48 | 24C | 0.94592 | 2.7425e-06 | 5.0068e-06 | 0.94592 | 32.8818 | 0.00010656
SWL2010-05_1#91 | 120 " | 2011-02-15 14:50 | 24C | 0.94397 | 2.2391e-06 | 4.9472e-06 | 0.94397 | 32.8058 | 8.8728e-05
SWL2010-05_1#92 | 121 " | 2011-02-15 14:52 | 24C | 0.94074 | 3.8626e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00014926
SWL2010-05_1#93 | 122 " | 2011-02-15 14:55 | 24C | 0.93871 | 3.3437e-06 | 4.8876e-06 | 0.93871 | 32.6011 | 0.00013184
SWL2010-05_1#94 | 122 Duplicate | 2011-02-15 14:57 | 24C | 0.93877 | 1.3712e-06 | 5.0068e-06 | 0.93877 | 32.6033 | 5.4326e-05 | Duplicate
SWL2010-05_1#95 | Ref. CB-06 | 2011-02-15 15:01 | 24C | 0.99888 | 2.4615e-06 | 5.0068e-06 | 0.99889 | 34.9563 | 9.4938e-05
SWL2010-05_1#96 | 001 NP - 1 | 2011-02-15 15:32 | 24C | 0.98465 | 2.2397e-06 | 5.0068e-06 | 0.98465 | 34.3968 | 8.6195e-05
SWL2010-05_1#97 | 002 " | 2011-02-15 15:34 | 24C | 0.98162 | 4.583e-06 | 4.9472e-06 | 0.98162 | 34.2778 | 0.00017845
SWL2010-05_1#98 | 003 " | 2011-02-15 15:36 | 24C | 0.97782 | 5.5594e-06 | 4.9472e-06 | 0.97783 | 34.129 | 0.00021878
SWL2010-05_1#99 | 004 " | 2011-02-15 15:40 | 24C | 0.97493 | 2.4909e-06 | 5.0068e-06 | 0.97493 | 34.0155 | 9.869e-05
SWL2010-05_1#100 | 005 " | 2011-02-15 15:42 | 24C | 0.97314 | 2.4909e-06 | 5.0068e-06 | 0.97314 | 33.9455 | 9.4938e-05
SWL2010-05_1#101 | 006 " | 2011-02-15 15:44 | 24C | 0.97282 | 6.2115e-06 | 5.0664e-06 | 0.97282 | 33.9329 | 0.00024548
SWL2010-05_1#102 | 007 " | 2011-02-15 15:47 | 24C | 0.97244 | 3.1414e-06 | 5.0664e-06 | 0.97245 | 33.9182 | 0.00012304
SWL2010-05_1#103 | 006 Duplicate sample 101 | 2011-02-15 15:49 | 24C | 0.97324 | 2.4754e-06 | 5.0068e-06 | 0.97325 | 33.9496 | 9.4938e-05 | Duplicate
SWL2010-05_1#104 | 008 NP - 1 | 2011-02-15 15:52 | 24C | 0.97156 | 2.4915e-06 | 5.0068e-06 | 0.97157 | 33.8838 | 9.6599e-05
SWL2010-05_1#105 | 009 " | 2011-02-15 15:55 | 24C | 0.96875 | 2.8901e-06 | 4.9472e-06 | 0.96875 | 33.7735 | 0.00011151
SWL2010-05_1#106 | 010 " | 2011-02-15 15:57 | 24C | 0.96497 | 2.4918e-06 | 4.828e-06 | 0.96498 | 33.6259 | 9.7495e-05
SWL2010-05_1#107 | 011 " | 2011-02-15 15:59 | 24C | 0.95664 | 2.2394e-06 | 5.0068e-06 | 0.95665 | 33.3002 | 8.7077e-05
SWL2010-05_1#108 | 012 " | 2011-02-15 16:01 | 24C | 0.94522 | 3.5403e-06 | 4.9472e-06 | 0.94522 | 32.8547 | 0.00014162
SWL2010-05_1#109 | 013 " | 2011-02-15 16:03 | 24C | 0.93499 | 2.4909e-06 | 5.0068e-06 | 0.93499 | 32.4563 | 9.6599e-05
SWL2010-05_1#110 | 14 | 2011-02-15 16:05 | 24C | 0.92543 | 1.4046e-06 | 5.0068e-06 | 0.92544 | 32.0849 | 5.5371e-05 | 14
SWL2010-05_1#111 | 014 Duplicate | 2011-02-15 16:08 | 24C | 0.92547 | 1.3712e-06 | 4.9472e-06 | 0.92547 | 32.0862 | 5.2235e-05 | Duplicate
SWL2010-05_1#112 | 015 NP - 1 | 2011-02-15 16:10 | 24C | 0.91924 | 2.4909e-06 | 5.0068e-06 | 0.91925 | 31.8448 | 9.4895e-05
SWL2010-05_1#113 | 016 " | 2011-02-15 16:14 | 24C | 0.91899 | 1.3712e-06 | 4.9472e-06 | 0.919 | 31.835 | 5.1714e-05
SWL2010-05_1#114 | 017 " | 2011-02-15 16:16 | 24C | 0.91891 | 2.475e-06 | 4.9472e-06 | 0.91891 | 31.8317 | 9.4364e-05
SWL2010-05_1#115 | 018 NP - 3 | 2011-02-15 16:18 | 24C | 0.98397 | 2.2391e-06 | 5.0664e-06 | 0.98398 | 34.3703 | 8.5352e-05
SWL2010-05_1#116 | 019 " | 2011-02-15 16:20 | 24C | 0.9816 | 3.1418e-06 | 5.0068e-06 | 0.9816 | 34.277 | 0.00012103
SWL2010-05_1#117 | 020 " | 2011-02-15 16:22 | 24C | 0.9772 | 2.4909e-06 | 4.9472e-06 | 0.97721 | 34.1047 | 9.8725e-05
SWL2010-05_1#118 | 020 Duplicate | 2011-02-15 16:25 | 24C | 0.97724 | 1.1529e-06 | 4.9472e-06 | 0.97724 | 34.106 | 4.4374e-05 | Duplicate
SWL2010-05_1#119 | Ref. CB - 06 | 2011-02-15 16:27 | 24C | 0.99887 | 2.5749e-06 | 5.0068e-06 | 0.99887 | 34.9556 | 0.00010121
SWL2010-05_1#120 | 162 BCL - 3 | 2011-02-15 16:35 | 24C | 0.93154 | 3.1338e-06 | 5.0068e-06 | 0.93155 | 32.3223 | 0.00012017
SWL2010-05_1#121 | 163 " | 2011-02-15 16:38 | 24C | 0.92728 | 2.2391e-06 | 5.0068e-06 | 0.92728 | 32.1566 | 8.7048e-05
SWL2010-05_1#122 | 164 " | 2011-02-15 16:40 | 24C | 0.92173 | 2.4909e-06 | 5.0068e-06 | 0.92174 | 31.9412 | 9.6496e-05
SWL2010-05_1#123 | 165 " | 2011-02-15 16:42 | 24C | 0.91926 | 2.4069e-06 | 5.0068e-06 | 0.91927 | 31.8455 | 9.4514e-05
SWL2010-05_1#124 | 166 " | 2011-02-15 16:45 | 24C | 0.91407 | 2.4925e-06 | 4.9472e-06 | 0.91408 | 31.6443 | 9.9652e-05
SWL2010-05_1#125 | 169 BCL - 4 | 2011-02-15 16:47 | 24C | 0.93716 | 1.7705e-06 | 5.0068e-06 | 0.93717 | 32.5409 | 6.8797e-05
SWL2010-05_1#126 | 170 " | 2011-02-15 16:50 | 24C | 0.93725 | 2.7432e-06 | 5.0068e-06 | 0.93725 | 32.5441 | 0.00010659
SWL2010-05_1#127 | 171 " | 2011-02-15 16:52 | 24C | 0.93344 | 2.7425e-06 | 4.8876e-06 | 0.93344 | 32.3959 | 0.00010656
SWL2010-05_1#128 | 172 " | 2011-02-15 16:54 | 24C | 0.9268 | 2.7425e-06 | 5.0068e-06 | 0.9268 | 32.1378 | 0.00010656
SWL2010-05_1#129 | 172 Duplicate | 2011-02-15 16:57 | 24C | 0.92694 | 2.7262e-06 | 4.8876e-06 | 0.92695 | 32.1435 | 0.00010656 | Duplicate
SWL2010-05_1#130 | 173 BCL - 4 | 2011-02-15 17:00 | 24C | 0.91616 | 2.7432e-06 | 5.126e-06 | 0.91616 | 31.725 | 0.00010395
SWL2010-05_1#131 | 174 " | 2011-02-15 17:02 | 24C | 0.90983 | 1.3712e-06 | 5.0664e-06 | 0.90984 | 31.48 | 5.2246e-05
SWL2010-05_1#132 | 175 " | 2011-02-15 17:05 | 24C | 0.9073 | 2.4911e-06 | 4.9472e-06 | 0.9073 | 31.3818 | 9.7041e-05
SWL2010-05_1#133 | LOOP - 12 | 2011-02-15 17:07 | 24C | 0.90477 | 1.7702e-06 | 5.0068e-06 | 0.90477 | 31.2839 | 6.8797e-05 | July 15/10 22:23
SWL2010-05_1#134 | 176 BCL - 5 | 2011-02-15 17:10 | 24C | 0.92218 | 2.2397e-06 | 4.9472e-06 | 0.92218 | 31.9586 | 8.5299e-05
SWL2010-05_1#135 | 177 " | 2011-02-15 17:12 | 24C | 0.92203 | 1.3551e-06 | 5.0664e-06 | 0.92203 | 31.9528 | 5.1714e-05
SWL2010-05_1#136 | 177 Duplicate | 2011-02-15 17:15 | 24C | 0.92202 | 1.1199e-06 | 4.9472e-06 | 0.92202 | 31.9524 | 4.4782e-05 | Duplicate
SWL2010-05_1#137 | 178 BCL - 5 | 2011-02-15 17:17 | 24C | 0.914 | 2.4912e-06 | 5.0068e-06 | 0.91401 | 31.6416 | 9.7041e-05
SWL2010-05_1#138 | 179 " | 2011-02-15 17:19 | 24C | 0.90887 | 2.8792e-06 | 5.0068e-06 | 0.90887 | 31.4427 | 0.00011137
SWL2010-05_1#139 | 180 " | 2011-02-15 17:21 | 24C | 0.90567 | 2.4908e-06 | 5.0068e-06 | 0.90567 | 31.3188 | 9.3946e-05
SWL2010-05_1#140 | 181 " | 2011-02-15 17:24 | 24C | 0.89056 | 3.9607e-06 | 4.9472e-06 | 0.89057 | 30.7352 | 0.00015076
SWL2010-05_1#141 | 182 " | 2011-02-15 17:26 | 24C | 0.88826 | 2.7432e-06 | 5.0068e-06 | 0.88827 | 30.6464 | 0.00010969
SWL2010-05_1#142 | Ref. CB - 06 | 2011-02-15 17:29 | 24C | 0.99889 | 3.9082e-06 | 5.0068e-06 | 0.9989 | 34.9567 | 0.00015494
SWL2010-05_1#143 | 062 NP - 7 | 2011-02-15 17:34 | 24C | 0.95169 | 1.3713e-06 | 5.0068e-06 | 0.95169 | 33.1068 | 5.2278e-05
SWL2010-05_1#144 | 063 " | 2011-02-15 17:37 | 24C | 0.94074 | 4.1859e-06 | 5.0068e-06 | 0.94075 | 32.6803 | 0.00016101
SWL2010-05_1#145 | 064 " | 2011-02-15 17:39 | 24C | 0.93822 | 2.8904e-06 | 5.0068e-06 | 0.93822 | 32.5819 | 0.00011603
SWL2010-05_1#146 | 065 " | 2011-02-15 17:41 | 24C | 0.93775 | 2.8904e-06 | 4.9472e-06 | 0.93775 | 32.5637 | 0.00011403
SWL2010-05_1#147 | 066 " | 2011-02-15 17:44 | 24C | 0.93771 | 2.7426e-06 | 5.0664e-06 | 0.93771 | 32.5621 | 0.00010451
SWL2010-05_1#148 | 067 " | 2011-02-15 17:46 | 24C | 0.93776 | 2.4925e-06 | 5.0664e-06 | 0.93777 | 32.5642 | 9.9734e-05
SWL2010-05_1#149 | 068" | 2011-02-15 17:48 | 24C | 0.93776 | 3.4657e-06 | 5.0068e-06 | 0.93777 | 32.5643 | 0.00013619
SWL2010-05_1#150 | LOOP 4 | 2011-02-15 17:50 | 24C | 0.93724 | 2.7434e-06 | 5.0664e-06 | 0.93725 | 32.544 | 0.00011074 | July 10/10 07:30
SWL2010-05_1#151 | Ref. CB-06 | 2011-02-15 17:54 | 24C | 0.9989 | 2.1649e-06 | 5.0664e-06 | 0.9989 | 34.9569 | 8.2858e-05
Calibration#2 | K15 | 2011-02-16 09:24 | 24C | 0.99995 | 2.7425e-06 | 1.6987e-05 | 0.99997
Calibration#3 | K15 | 2011-02-16 09:38 | 24C | 0.99996 | 2.2391e-06 | 9.0599e-06 | 0.99997
SWL2010-05_1#152 | 069 NP - 9 | 2011-02-16 09:41 | 24C | 0.98509 | 1.3884e-06 | 8.9407e-06 | 0.9851 | 34.4144 | 5.4326e-05
SWL2010-05_1#153 | 070 " | 2011-02-16 09:44 | 24C | 0.98327 | 2.4909e-06 | 9.0003e-06 | 0.98328 | 34.3427 | 9.4928e-05
SWL2010-05_1#154 | 071 " | 2011-02-16 09:46 | 24C | 0.98045 | 3.0387e-06 | 9.0003e-06 | 0.98046 | 34.2323 | 0.00012157
SWL2010-05_1#155 | 071 Duplicate | 2011-02-16 09:49 | 24C | 0.98043 | 3.1415e-06 | 9.0599e-06 | 0.98044 | 34.2313 | 0.00012514 | Duplicate
SWL2010-05_1#156 | 072 NP - 9 | 2011-02-16 09:51 | 24C | 0.97769 | 2.4777e-06 | 8.9407e-06 | 0.97769 | 34.1238 | 9.9543e-05
SWL2010-05_1#157 | 073 " | 2011-02-16 09:53 | 24C | 0.9758 | 2.4744e-06 | 9.0599e-06 | 0.97581 | 34.0497 | 9.7029e-05
SWL2010-05_1#158 | 074 " | 2011-02-16 09:55 | 24C | 0.97525 | 1.8038e-06 | 8.9407e-06 | 0.97526 | 34.0284 | 6.8797e-05 | 074 "
SWL2010-05_1#159 | 076 " | 2011-02-16 09:57 | 24C | 0.97448 | 2.4909e-06 | 8.9407e-06 | 0.97449 | 33.9984 | 9.9198e-05
SWL2010-05_1#160 | 076 " | 2011-02-16 09:59 | 24C | 0.97388 | 2.4911e-06 | 8.9407e-06 | 0.97389 | 33.9745 | 9.6976e-05
SWL2010-05_1#161 | 077 " | 2011-02-16 10:01 | 24C | 0.97281 | 2.4911e-06 | 9.0599e-06 | 0.97282 | 33.9328 | 0.0001006
SWL2010-05_1#162 | 078 " | 2011-02-16 10:03 | 24C | 0.97154 | 2.0804e-06 | 9.0003e-06 | 0.97154 | 33.8829 | 7.889e-05
SWL2010-05_1#163 | 079 " | 2011-02-16 10:05 | 24C | 0.96575 | 2.4915e-06 | 8.9407e-06 | 0.96576 | 33.6566 | 9.3856e-05
SWL2010-05_1#164 | 080 " | 2011-02-16 10:07 | 24C | 0.95408 | 1.1199e-06 | 9.0599e-06 | 0.95408 | 33.2002 | 4.5227e-05
SWL2010-05_1#165 | 081 " | 2011-02-16 10:10 | 24C | 0.9495 | 1.372e-06 | 9.0003e-06 | 0.94951 | 33.0216 | 5.4326e-05
SWL2010-05_1#166 | 082 " | 2011-02-16 10:12 | 24C | 0.94642 | 2.7426e-06 | 9.0599e-06 | 0.94643 | 32.9015 | 0.00010761
SWL2010-05_1#167 | 082 Duplicate | 2011-02-16 10:17 | 24C | 0.94649 | 2.4909e-06 | 9.0003e-06 | 0.94649 | 32.9042 | 9.4885e-05 | Duplicate
SWL2010-05_1#168 | Ref. CB - 06 | 2011-02-16 10:19 | 24C | 0.99889 | 3.8587e-06 | 9.0599e-06 | 0.9989 | 34.9565 | 0.00014953
SWL2010-05_1#169 | 083 NP - 9 | 2011-02-16 10:38 | 24C | 0.94606 | 2.8877e-06 | 8.9407e-06 | 0.94607 | 32.8875 | 0.00011008
SWL2010-05_1#170 | 084 " | 2011-02-16 10:40 | 24C | 0.94603 | 2.7434e-06 | 9.0003e-06 | 0.94604 | 32.8865 | 0.00011074
SWL2010-05_1#171 | 085 " | 2011-02-16 10:42 | 24C | 0.94597 | 3.143e-06 | 8.9407e-06 | 0.94598 | 32.884 | 0.00012347
SWL2010-05_1#172 | 086 NP - 11 | 2011-02-16 10:45 | 24C | 0.98548 | 3.4658e-06 | 8.9407e-06 | 0.98549 | 34.4296 | 0.00013656
SWL2010-05_1#173 | 087 " | 2011-02-16 10:48 | 24C | 0.98357 | 4.2256e-06 | 9.0599e-06 | 0.98358 | 34.3548 | 0.00016833
SWL2010-05_1#174 | 088 " | 2011-02-16 10:50 | 24C | 0.98046 | 2.7426e-06 | 9.0003e-06 | 0.98047 | 34.2327 | 0.00010865
SWL2010-05_1#175 | 088 Duplicate | 2011-02-16 10:52 | 24C | 0.98051 | 3.2144e-06 | 9.0599e-06 | 0.98052 | 34.2346 | 0.00012553 | Duplicate
SWL2010-05_1#176 | 089 NP - 11 | 2011-02-16 10:55 | 24C | 0.9779 | 1.3712e-06 | 9.0599e-06 | 0.97791 | 34.1323 | 5.5371e-05
SWL2010-05_1#177 | 090 " | 2011-02-16 10:57 | 24C | 0.97537 | 1.1196e-06 | 8.9407e-06 | 0.97538 | 34.0332 | 4.2703e-05
SWL2010-05_1#178 | 091 " | 2011-02-16 10:59 | 24C | 0.97502 | 2.4908e-06 | 9.0599e-06 | 0.97503 | 34.0193 | 9.9553e-05
SWL2010-05_1#179 | 092 " | 2011-02-16 11:02 | 24C | 0.97402 | 2.8901e-06 | 9.0003e-06 | 0.97403 | 33.9802 | 0.00011145
SWL2010-05_1#180 | 093 " | 2011-02-16 11:04 | 24C | 0.97333 | 2.6582e-06 | 9.0599e-06 | 0.97334 | 33.9532 | 0.00010466
SWL2010-05_1#181 | 094 " | 2011-02-16 11:07 | 24C | 0.97236 | 1.1203e-06 | 9.0003e-06 | 0.97237 | 33.9152 | 4.2703e-05
SWL2010-05_1#182 | 095 " | 2011-02-16 11:10 | 24C | 0.96972 | 2.2401e-06 | 9.0003e-06 | 0.96973 | 33.8116 | 9.0453e-05
SWL2010-05_1#183 | 096 " | 2011-02-16 11:12 | 24C | 0.96532 | 2.4909e-06 | 8.9407e-06 | 0.96533 | 33.6395 | 9.87e-05
SWL2010-05_1#184 | 096 Duplicate | 2011-02-16 11:14 | 24C | 0.9652 | 2.8901e-06 | 9.0003e-06 | 0.96521 | 33.6349 | 0.00011213 | Duplicate
SWL2010-05_1#185 | 097 NP - 11 | 2011-02-16 11:17 | 24C | 0.95597 | 2.4911e-06 | 9.0599e-06 | 0.95598 | 33.2743 | 9.87e-05
SWL2010-05_1#186 | 098 " | 2011-02-16 11:19 | 24C | 0.9488 | 2.2268e-06 | 9.0599e-06 | 0.94881 | 32.9944 | 8.7015e-05
SWL2010-05_1#187 | 099 " | 2011-02-16 11:22 | 24C | 0.94613 | 2.2725e-06 | 8.9407e-06 | 0.94614 | 32.8902 | 8.9676e-05
SWL2010-05_1#188 | 100 " | 2011-02-16 11:24 | 24C | 0.94712 | 4.6354e-06 | 8.9407e-06 | 0.94712 | 32.9287 | 0.00017833
SWL2010-05_1#189 | 101 " | 2011-02-16 11:28 | 24C | 0.94016 | 3.1145e-06 | 9.0599e-06 | 0.94017 | 32.6577 | 0.00012001
SWL2010-05_1#190 | 102 " | 2011-02-16 11:31 | 24C | 0.93472 | 2.2391e-06 | 9.0003e-06 | 0.93473 | 32.446 | 8.7024e-05
SWL2010-05_1#191 | Ref. CB - 06 | 2011-02-16 11:33 | 24C | 0.99888 | 3.1867e-06 | 8.9407e-06 | 0.99889 | 34.9563 | 0.00012382
SWL2010-05_1#192 | 224 UTBS-5 | 2011-02-16 11:41 | 24C | 0.93768 | 2.8901e-06 | 8.9407e-06 | 0.93769 | 32.5614 | 0.0001118
SWL2010-05_1#193 | 224 Duplicate | 2011-02-16 11:44 | 24C | 0.93764 | 3.2141e-06 | 9.0003e-06 | 0.93765 | 32.5597 | 0.00012601 | Duplicate
SWL2010-05_1#194 | 225 UTBS-5 | 2011-02-16 11:46 | 24C | 0.93599 | 2.1137e-06 | 9.0003e-06 | 0.93599 | 32.4952 | 8.1621e-05
SWL2010-05_1#195 | 226 " | 2011-02-16 11:48 | 24C | 0.93425 | 3.1414e-06 | 9.0003e-06 | 0.93426 | 32.4276 | 0.00012104
SWL2010-05_1#196 | 227 " | 2011-02-16 11:50 | 24C | 0.93391 | 1.7702e-06 | 9.0599e-06 | 0.93392 | 32.4145 | 6.8817e-05
SWL2010-05_1#197 | 228 " | 2011-02-16 11:52 | 24C | 0.93373 | 2.2397e-06 | 9.0003e-06 | 0.93374 | 32.4075 | 8.7013e-05
SWL2010-05_1#198 | 230 UTBS - 2 | 2011-02-16 11:56 | 24C | 0.93698 | 4.0094e-06 | 9.0599e-06 | 0.93699 | 32.5339 | 0.00015582 | no 229
SWL2010-05_1#199 | 231 " | 2011-02-16 11:58 | 24C | 0.93695 | 2.2391e-06 | 9.0003e-06 | 0.93695 | 32.5326 | 8.7048e-05
SWL2010-05_1#200 | 231 Duplicate | 2011-02-16 12:00 | 24C | 0.93699 | 3.1414e-06 | 8.9407e-06 | 0.93699 | 32.5342 | 0.00012313 | Duplicate
SWL2010-05_1#201 | 232 UTBS - 2 | 2011-02-16 12:02 | 24C | 0.93719 | 1.1532e-06 | 9.0599e-06 | 0.9372 | 32.5422 | 4.4364e-05
SWL2010-05_1#202 | 233 " | 2011-02-16 12:04 | 24C | 0.93146 | 3.3594e-06 | 9.0599e-06 | 0.93147 | 32.3192 | 0.0001316
SWL2010-05_1#203 | 234 " | 2011-02-16 12:06 | 24C | 0.93041 | 2.2397e-06 | 8.9407e-06 | 0.93042 | 32.2784 | 8.8728e-05
SWL2010-05_1#204 | 235 UTBS - 4 | 2011-02-16 12:08 | 24C | 0.9439 | 4.6245e-06 | 8.9407e-06 | 0.94391 | 32.8034 | 0.00018283
SWL2010-05_1#205 | 236 " | 2011-02-16 12:11 | 24C | 0.9434 | 1.3713e-06 | 9.0599e-06 | 0.94341 | 32.784 | 5.2235e-05
SWL2010-05_1#206 | 237 " | 2011-02-16 12:13 | 24C | 0.9396 | 3.5403e-06 | 8.9407e-06 | 0.93961 | 32.6359 | 0.00013759
SWL2010-05_1#207 | 238 " | 2011-02-16 12:15 | 24C | 0.93587 | 3.332e-08 | 8.9407e-06 | 0.93588 | 32.4907 | 0
SWL2010-05_1#208 | 238 Duplicate | 2011-02-16 12:17 | 24C | 0.93587 | 4.2257e-06 | 9.0003e-06 | 0.93588 | 32.4908 | 0.00016353 | Duplicate
SWL2010-05_1#209 | 239 UTBS - 4 | 2011-02-16 12:19 | 24C | 0.93538 | 3.1417e-06 | 9.0599e-06 | 0.93539 | 32.4717 | 0.00012103
SWL2010-05_1#210 | 240 UTBS - 1 | 2011-02-16 12:22 | 24C | 0.93791 | 5.0346e-06 | 9.0599e-06 | 0.93792 | 32.5701 | 0.00019692
SWL2010-05_1#211 | 240 Duplicate | 2011-02-16 12:24 | 24C | 0.93774 | 2.4909e-06 | 9.0599e-06 | 0.93775 | 32.5634 | 9.87e-05 | Duplicate
SWL2010-05_1#212 | 241 UTBS - 1 | 2011-02-16 12:27 | 24C | 0.93771 | 3.466e-06 | 8.9407e-06 | 0.93772 | 32.5625 | 0.00013595
SWL2010-05_1#213 | 242 " | 2011-02-16 12:29 | 24C | 0.93744 | 2.2396e-06 | 8.9407e-06 | 0.93745 | 32.552 | 8.5335e-05
SWL2010-05_1#214 | 243 " | 2011-02-16 12:33 | 24C | 0.93498 | 2.7425e-06 | 9.0599e-06 | 0.93499 | 32.4563 | 0.00010973
SWL2010-05_1#215 | 244 " | 2011-02-16 12:35 | 24C | 0.93342 | 4.2256e-06 | 9.0599e-06 | 0.93343 | 32.3956 | 0.00016286
SWL2010-05_1#216 | Ref. CB - 06 | 2011-02-16 12:37 | 24C | 0.99887 | 3.4658e-06 | 9.0003e-06 | 0.99888 | 34.9559 | 0.00013927
SWL2010-05_1#217 | 284 UTN - 4 | 2011-02-16 12:42 | 24C | 0.93597 | 1.7702e-06 | 9.0003e-06 | 0.93598 | 32.4946 | 6.8797e-05
SWL2010-05_1#218 | 285 " | 2011-02-16 12:44 | 24C | 0.92955 | 3.865e-06 | 8.9407e-06 | 0.92956 | 32.2452 | 0.00014997
SWL2010-05_1#219 | 286 " | 2011-02-16 12:46 | 24C | 0.92897 | 2.4775e-06 | 9.0599e-06 | 0.92897 | 32.2223 | 9.4034e-05
SWL2010-05_1#220 | 286 Duplicate | 2011-02-16 12:49 | 24C | 0.929 | 2.8901e-06 | 9.0599e-06 | 0.92901 | 32.2237 | 0.00011008 | Duplicate
SWL2010-05_1#221 | 287 UTN - 4 | 2011-02-16 12:52 | 24C | 0.9282 | 2.128e-06 | 8.9407e-06 | 0.92821 | 32.1924 | 8.1169e-05
SWL2010-05_1#222 | 288 " | 2011-02-16 12:54 | 24C | 0.92467 | 5.3436e-06 | 8.9407e-06 | 0.92467 | 32.0553 | 0.0002072
SWL2010-05_1#223 | 289 UTN - 5 | 2011-02-16 12:56 | 24C | 0.93348 | 2.2391e-06 | 8.9407e-06 | 0.93348 | 32.3976 | 8.7239e-05
SWL2010-05_1#224 | 290 " | 2011-02-16 13:00 | 24C | 0.93383 | 2.4908e-06 | 8.9407e-06 | 0.93384 | 32.4114 | 9.7001e-05
SWL2010-05_1#225 | 291 " | 2011-02-16 13:02 | 24C | 0.93168 | 3.4517e-06 | 9.0003e-06 | 0.93169 | 32.3277 | 0.00013386
SWL2010-05_1#226 | 292 " | 2011-02-16 13:04 | 24C | 0.92743 | 1.7705e-06 | 8.9407e-06 | 0.92744 | 32.1627 | 6.8817e-05
SWL2010-05_1#227 | 292 Duplicate | 2011-02-16 13:07 | 24C | 0.92727 | 4.1893e-06 | 8.9407e-06 | 0.92728 | 32.1566 | 0.00016234 | Duplicate
SWL2010-05_1#228 | 293 " | 2011-02-16 13:09 | 24C | 0.92435 | 3.054e-06 | 8.9407e-06 | 0.92436 | 32.0429 | 0.00011839
SWL2010-05_1#229 | 294 UTN - 6 | 2011-02-16 13:12 | 24C | 0.92537 | 3.4657e-06 | 8.9407e-06 | 0.92538 | 32.0826 | 0.00013633
SWL2010-05_1#230 | 295 " | 2011-02-16 13:17 | 24C | 0.92523 | 3.865e-06 | 8.9407e-06 | 0.92523 | 32.077 | 0.00015135
SWL2010-05_1#231 | 296 " | 2011-02-16 13:19 | 24C | 0.9245 | 2.7425e-06 | 9.0003e-06 | 0.92451 | 32.049 | 0.00010865
SWL2010-05_1#232 | 297 " | 2011-02-16 13:23 | 24C | 0.92127 | 2.0873e-06 | 9.0599e-06 | 0.92128 | 31.9237 | 7.9556e-05
SWL2010-05_1#233 | 297 Duplicate | 2011-02-16 13:25 | 24C | 0.92125 | 3.4659e-06 | 9.0003e-06 | 0.92126 | 31.9227 | 0.00013374 | Duplicate
SWL2010-05_1#234 | 298 UTN - 6 | 2011-02-16 13:28 | 24C | 0.91896 | 3.204e-06 | 9.0003e-06 | 0.91897 | 31.8339 | 0.00012417
SWL2010-05_1#235 | 299 UTN - 7 | 2011-02-16 13:31 | 24C | 0.92586 | 2.0945e-06 | 9.0003e-06 | 0.92586 | 32.1015 | 8.1621e-05
SWL2010-05_1#236 | 300 " | 2011-02-16 13:33 | 24C | 0.92569 | 3.2141e-06 | 8.9407e-06 | 0.9257 | 32.0952 | 0.00012708
SWL2010-05_1#237 | 300 Duplicate | 2011-02-16 13:35 | 24C | 0.92568 | 4.1785e-06 | 8.9407e-06 | 0.92569 | 32.0946 | 0.00016445 | Duplicate
SWL2010-05_1#238 | 301 UTN - 7 | 2011-02-16 13:37 | 24C | 0.92383 | 3.53e-06 | 8.9407e-06 | 0.92384 | 32.0229 | 0.00013487
SWL2010-05_1#239 | 302 " | 2011-02-16 13:39 | 24C | 0.92089 | 2.2574e-06 | 9.0003e-06 | 0.92089 | 31.9086 | 8.7711e-05
SWL2010-05_1#240 | 303 " | 2011-02-16 13:41 | 24C | 0.9201 | 5.7875e-06 | 9.0003e-06 | 0.92011 | 31.8781 | 0.00022469
SWL2010-05_1#241 | Ref. CB - 06 | 2011-02-16 13:47 | 24C | 0.99888 | 4.3195e-06 | 9.0003e-06 | 0.99889 | 34.9562 | 0.00017085
SWL2010-05_1#242 | 344 PAG - 4 | 2011-02-16 14:09 | 24C | 0.94193 | 4.31e-06 | 8.9407e-06 | 0.94194 | 32.7266 | 0.00016895
SWL2010-05_1#243 | 345 " | 2011-02-16 14:13 | 24C | 0.94076 | 3.2148e-06 | 9.0003e-06 | 0.94077 | 32.6811 | 0.00012554
SWL2010-05_1#244 | 346 " | 2011-02-16 14:15 | 24C | 0.93897 | 3.2145e-06 | 9.0599e-06 | 0.93898 | 32.6115 | 0.00012554
SWL2010-05_1#245 | 347 " | 2011-02-16 14:18 | 24C | 0.93838 | 2.8904e-06 | 9.0003e-06 | 0.93839 | 32.5884 | 0.00011213
SWL2010-05_1#246 | 347 Duplicate | 2011-02-16 14:20 | 24C | 0.93835 | 3.4551e-06 | 8.9407e-06 | 0.93836 | 32.5874 | 0.00013203 | Duplicate
SWL2010-05_1#247 | 348 PAG - 4 | 2011-02-16 14:22 | 24C | 0.93384 | 2.4911e-06 | 9.0003e-06 | 0.93385 | 32.4117 | 9.4885e-05
SWL2010-05_1#248 | 349 " | 2011-02-16 14:24 | 24C | 0.92737 | 3.4658e-06 | 8.9407e-06 | 0.92738 | 32.1603 | 0.00013726
SWL2010-05_1#249 | 350 " | 2011-02-16 14:26 | 24C | 0.92322 | 2.2262e-06 | 9.0599e-06 | 0.92323 | 31.999 | 8.7432e-05
SWL2010-05_1#250 | 1 Coke Cap 35 | 2011-02-16 14:30 | 24C | 1.0029 | 3.4671e-06 | 8.9407e-06 | 1.0029 | 35.1155 | 0.00013732 | NO INSERT!
SWL2010-05_1#251 | 8 Coke Cap 35 | 2011-02-16 14:33 | 24C | 1.0011 | 2.408e-06 | 8.9407e-06 | 1.0011 | 35.0425 | 9.4233e-05 | NO INSERT!
SWL2010-05_1#252 | Ref. CB - 06 | 2011-02-16 14:39 | 24C | 0.99878 | 2.4909e-06 | 8.9407e-06 | 0.99879 | 34.9524 | 9.869e-05
SWL2010-05_1#253 | 123 BCL - 1 | 2011-02-16 14:45 | 24C | 0.9386 | 2.2391e-06 | 8.9407e-06 | 0.93861 | 32.5971 | 8.5406e-05
SWL2010-05_1#254 | 124 " | 2011-02-16 14:47 | 24C | 0.93911 | 2.4775e-06 | 8.9407e-06 | 0.93912 | 32.6168 | 9.4938e-05
SWL2010-05_1#255 | 125 BCL - 2 | 2011-02-16 14:51 | 24C | 0.98301 | 3.3587e-06 | 9.0599e-06 | 0.98302 | 34.3326 | 0.00013246
SWL2010-05_1#256 | 126 " | 2011-02-16 14:53 | 24C | 0.97924 | 3.6104e-06 | 8.9407e-06 | 0.97925 | 34.1848 | 0.00014403
SWL2010-05_1#257 | 127 " | 2011-02-16 14:55 | 24C | 0.97282 | 2.7426e-06 | 9.0003e-06 | 0.97283 | 33.9331 | 0.00010447
SWL2010-05_1#258 | 127 Duplicate | 2011-02-16 14:57 | 24C | 0.97275 | 2.2394e-06 | 8.9407e-06 | 0.97276 | 33.9304 | 8.7025e-05 | Duplicate
SWL2010-05_1#259 | 128 BCL - 2 | 2011-02-16 14:59 | 24C | 0.96596 | 2.8792e-06 | 8.9407e-06 | 0.96597 | 33.6646 | 0.00011145
SWL2010-05_1#260 | 129 " | 2011-02-16 15:02 | 24C | 0.96067 | 4.334e-06 | 8.9407e-06 | 0.96068 | 33.4578 | 0.00017005
SWL2010-05_1#261 | 130 " | 2011-02-16 15:04 | 24C | 0.95949 | 3.4657e-06 | 8.9407e-06 | 0.9595 | 33.4118 | 0.00013801
SWL2010-05_1#262 | 131 " | 2011-02-16 15:06 | 24C | 0.95829 | 3.466e-06 | 8.9407e-06 | 0.9583 | 33.365 | 0.0001334
SWL2010-05_1#263 | 132 " | 2011-02-16 15:08 | 24C | 0.95698 | 2.4908e-06 | 9.0599e-06 | 0.95699 | 33.3136 | 9.6599e-05
SWL2010-05_1#264 | 133 " | 2011-02-16 15:10 | 24C | 0.95578 | 3.4354e-06 | 9.0003e-06 | 0.95579 | 33.2668 | 0.0001334
SWL2010-05_1#265 | 134 " | 2011-02-16 15:12 | 24C | 0.95378 | 2.4909e-06 | 9.0003e-06 | 0.95379 | 33.1886 | 9.7462e-05
SWL2010-05_1#266 | 135 " | 2011-02-16 15:14 | 24C | 0.95283 | 3.5942e-06 | 9.0599e-06 | 0.95284 | 33.1518 | 0.00014204
SWL2010-05_1#267 | 136 " | 2011-02-16 15:17 | 24C | 0.95199 | 2.2394e-06 | 9.0599e-06 | 0.952 | 33.1187 | 9.0433e-05
SWL2010-05_1#268 | 137 " | 2011-02-16 15:35 | 24C | 0.9489 | 2.4909e-06 | 9.0003e-06 | 0.94891 | 32.9984 | 9.869e-05
SWL2010-05_1#269 | 137 Duplicate | 2011-02-16 15:37 | 24C | 0.94888 | 3.3455e-06 | 9.0003e-06 | 0.94889 | 32.9977 | 0.00012999 | Duplicate
SWL2010-05_1#270 | 138 BCL - 2 | 2011-02-16 15:40 | 24C | 0.94699 | 2.8907e-06 | 8.9407e-06 | 0.947 | 32.9239 | 0.00011317
SWL2010-05_1#271 | 139 " | 2011-02-16 15:42 | 24C | 0.94524 | 3.1415e-06 | 8.9407e-06 | 0.94525 | 32.8557 | 0.00011967
SWL2010-05_1#272 | 140 " | 2011-02-16 15:44 | 24C | 0.94027 | 2.7425e-06 | 8.9407e-06 | 0.94028 | 32.662 | 0.00010659
SWL2010-05_1#273 | 141 " | 2011-02-16 15:46 | 24C | 0.93864 | 4.9165e-06 | 9.0599e-06 | 0.93864 | 32.5984 | 0.00019281
SWL2010-05_1#274 | LOOP 8 | 2011-02-16 15:48 | 24C | 0.9386 | 3.865e-06 | 9.0599e-06 | 0.93861 | 32.597 | 0.00015042 | July 14/10 13:44
SWL2010-05_1#275 | Ref. CB - 06 | 2011-02-16 15:54 | 24C | 0.99887 | 2.89e-06 | 9.0003e-06 | 0.99888 | 34.9561 | 0.00011604
SWL2010-05_1#276 | 042 NP - 5 | 2011-02-16 15:57 | 24C | 0.96704 | 2.4909e-06 | 9.0003e-06 | 0.96705 | 33.7068 | 9.6599e-05
SWL2010-05_1#277 | 043 " | 2011-02-16 16:00 | 24C | 0.96298 | 1.6739e-06 | 9.0599e-06 | 0.96299 | 33.548 | 6.8483e-05
SWL2010-05_1#278 | 044 " | 2011-02-16 16:02 | 24C | 0.95624 | 2.4911e-06 | 8.9407e-06 | 0.95625 | 33.2847 | 9.869e-05
SWL2010-05_1#279 | 044 Duplicate | 2011-02-16 16:05 | 24C | 0.95628 | 1.1536e-06 | 9.0599e-06 | 0.95629 | 33.2864 | 4.265e-05 | Duplicate
SWL2010-05_1#280 | 045 NP - 5 | 2011-02-16 16:07 | 24C | 0.9494 | 2.4908e-06 | 9.0599e-06 | 0.94941 | 33.018 | 9.6976e-05
SWL2010-05_1#281 | 046 " | 2011-02-16 16:09 | 24C | 0.93657 | 2.4908e-06 | 9.0003e-06 | 0.93658 | 32.5178 | 9.3005e-05
SWL2010-05_1#282 | 047 " | 2011-02-16 16:11 | 24C | 0.93527 | 2.4908e-06 | 8.9407e-06 | 0.93528 | 32.4674 | 9.6976e-05
SWL2010-05_1#283 | 049 " | 2011-02-16 16:14 | 24C | 0.93339 | 2.2405e-06 | 9.0599e-06 | 0.9334 | 32.3944 | 8.5406e-05
SWL2010-05_1#284 | 050 " | 2011-02-16 16:16 | 24C | 0.93231 | 3.6109e-06 | 9.0003e-06 | 0.93232 | 32.3524 | 0.00013904
SWL2010-05_1#285 | 051 " | 2011-02-16 16:18 | 24C | 0.93251 | 1.3712e-06 | 9.0599e-06 | 0.93252 | 32.36 | 5.4351e-05
SWL2010-05_1#286 | LOOP P2 | 2011-02-16 16:21 | 24C | 0.93207 | 2.0841e-06 | 9.0003e-06 | 0.93208 | 32.3431 | 8.1174e-05 | July 9/2010
SWL2010-05_1#287 | 052 NP - 7 | 2011-02-16 16:23 | 24C | 0.98453 | 1.7597e-06 | 9.0599e-06 | 0.98453 | 34.3921 | 6.6764e-05
SWL2010-05_1#288 | 053 " | 2011-02-16 16:25 | 24C | 0.982 | 2.4911e-06 | 9.0003e-06 | 0.98201 | 34.2929 | 9.4885e-05
SWL2010-05_1#289 | 054 " | 2011-02-16 16:28 | 24C | 0.97858 | 1.3712e-06 | 9.0003e-06 | 0.97859 | 34.1589 | 5.5395e-05
SWL2010-05_1#290 | 054 Duplicate | 2011-02-16 16:30 | 24C | 0.97827 | 4.189e-06 | 9.0599e-06 | 0.97828 | 34.1468 | 0.00016484 | Duplicate ! Not close at all
SWL2010-05_1#291 | 055 NP - 7 | 2011-02-16 16:33 | 24C | 0.97517 | 0 | 9.0599e-06 | 0.97518 | 34.0252 | 0
SWL2010-05_1#292 | 056 " | 2011-02-16 16:35 | 24C | 0.97322 | 2.7425e-06 | 9.0599e-06 | 0.97323 | 33.9487 | 0.00010865
SWL2010-05_1#293 | 057 " | 2011-02-16 16:37 | 24C | 0.97231 | 6.4306e-06 | 9.0599e-06 | 0.97232 | 33.9131 | 0.00025094
SWL2010-05_1#294 | 058 " | 2011-02-16 16:39 | 24C | 0.97246 | 3.6084e-06 | 9.0003e-06 | 0.97247 | 33.9191 | 0.00014251
SWL2010-05_1#295 | 059 " | 2011-02-16 16:41 | 24C | 0.97153 | 2.4912e-06 | 8.9407e-06 | 0.97153 | 33.8825 | 9.9745e-05
SWL2010-05_1#296 | 060 " | 2011-02-16 16:43 | 24C | 0.97031 | 2.89e-06 | 8.9407e-06 | 0.97032 | 33.835 | 0.00011008
SWL2010-05_1#297 | 061 " | 2011-02-16 16:45 | 24C | 0.96418 | 1.1532e-06 | 9.0599e-06 | 0.96419 | 33.5951 | 4.5217e-05
SWL2010-05_1#298 | Ref. CB - 06 | 2011-02-16 16:49 | 24C | 0.99893 | 2.89e-06 | 8.9407e-06 | 0.99894 | 34.9584 | 0.00011213
SWL2010-05_1#299 | 245 BRS - 4 | 2011-02-16 16:55 | 24C | 0.9257 | 2.7426e-06 | 9.0599e-06 | 0.92571 | 32.0955 | 0.00010865
SWL2010-05_1#300 | 246 " | 2011-02-16 16:58 | 24C | 0.9247 | 2.4915e-06 | 9.0599e-06 | 0.92471 | 32.0565 | 9.1806e-05
SWL2010-05_1#301 | 246 Duplicate | 2011-02-16 17:01 | 24C | 0.9247 | 1.1196e-06 | 8.9407e-06 | 0.92471 | 32.0567 | 4.2703e-05 | Duplicate
SWL2010-05_1#302 | 247 BRS - 4 | 2011-02-16 17:03 | 24C | 0.92459 | 3.2142e-06 | 9.0599e-06 | 0.9246 | 32.0524 | 0.00012519
SWL2010-05_1#303 | 248 " | 2011-02-16 17:06 | 24C | 0.92447 | 2.8897e-06 | 8.9407e-06 | 0.92447 | 32.0475 | 0.00011317
SWL2010-05_1#304 | 249 " | 2011-02-16 17:08 | 24C | 0.91952 | 2.2405e-06 | 9.0003e-06 | 0.91953 | 31.8555 | 8.5313e-05
SWL2010-05_1#305 | 250 BRS - 5 | 2011-02-16 17:10 | 24C | 0.94025 | 2.2391e-06 | 9.0003e-06 | 0.94026 | 32.6611 | 8.5406e-05
SWL2010-05_1#306 | 250 Duplicate | 2011-02-16 17:12 | 24C | 0.94027 | 4.189e-06 | 9.0003e-06 | 0.94028 | 32.662 | 0.00016283 | Duplicate
SWL2010-05_1#307 | 251 BRS - 5 | 2011-02-16 17:14 | 24C | 0.94044 | 2.7425e-06 | 8.9407e-06 | 0.94044 | 32.6684 | 0.00010761
SWL2010-05_1#308 | 252 " | 2011-02-16 17:16 | 24C | 0.94067 | 2.4912e-06 | 9.0599e-06 | 0.94068 | 32.6775 | 0.0001006
SWL2010-05_1#309 | 253 " | 2011-02-16 17:19 | 24C | 0.93773 | 2.4908e-06 | 8.9407e-06 | 0.93774 | 32.5632 | 9.6609e-05
SWL2010-05_1#310 | 254 " | 2011-02-16 17:21 | 24C | 0.93082 | 3.4662e-06 | 8.9407e-06 | 0.93083 | 32.2945 | 0.00013655
SWL2010-05_1#311 | 255 BRS - 3 | 2011-02-16 17:24 | 24C | 0.9239 | 3.36e-06 | 8.9407e-06 | 0.92391 | 32.0254 | 0.0001298
SWL2010-05_1#312 | 256 " | 2011-02-16 17:26 | 24C | 0.92428 | 2.7425e-06 | 8.9407e-06 | 0.92428 | 32.0401 | 0.00010451
SWL2010-05_1#313 | 257 " | 2011-02-16 17:28 | 24C | 0.92355 | 2.7432e-06 | 9.0599e-06 | 0.92356 | 32.012 | 0.00010659
SWL2010-05_1#314 | 258 " | 2011-02-16 17:30 | 24C | 0.91687 | 3.3591e-06 | 8.9407e-06 | 0.91688 | 31.753 | 0.00012967
SWL2010-05_1#315 | 258 Duplicate | 2011-02-16 17:32 | 24C | 0.91673 | 2.2394e-06 | 8.9407e-06 | 0.91673 | 31.7472 | 8.7006e-05 | Duplicate
SWL2010-05_1#316 | 259 BRS - 3 | 2011-02-16 17:34 | 24C | 0.91666 | 3.8747e-06 | 9.0003e-06 | 0.91667 | 31.7446 | 0.00015
SWL2010-05_1#317 | 260 BRS - 2 | 2011-02-16 17:37 | 24C | 0.91566 | 2.575e-06 | 9.0003e-06 | 0.91567 | 31.706 | 9.9156e-05
SWL2010-05_1#318 | 260 Duplicate | 2011-02-16 17:40 | 24C | 0.91558 | 2.7264e-06 | 8.9407e-06 | 0.91559 | 31.7029 | 0.00010396 | Duplicate
SWL2010-05_1#319 | 261 BRS - 2 | 2011-02-16 17:42 | 24C | 0.91489 | 2.4908e-06 | 8.9407e-06 | 0.9149 | 31.6761 | 9.6991e-05
SWL2010-05_1#320 | 262 " | 2011-02-16 17:45 | 24C | 0.91447 | 3.6082e-06 | 9.0599e-06 | 0.91448 | 31.66 | 0.00014255
SWL2010-05_1#321 | 263 " | 2011-02-16 17:48 | 24C | 0.91443 | 2.7425e-06 | 8.9407e-06 | 0.91443 | 31.6581 | 0.00010708
SWL2010-05_1#322 | 264 " | 2011-02-16 17:50 | 24C | 0.91417 | 2.4612e-06 | 9.0599e-06 | 0.91418 | 31.6483 | 9.4909e-05
SWL2010-05_1#323 | Ref. CB - 06 | 2011-02-16 17:52 | 24C | 0.99889 | 3.1418e-06 | 9.0003e-06 | 0.9989 | 34.9566 | 0.00012312
Calibration#4 | K15 | 2011-02-17 09:11 | 24C | 0.99996 | 2.7425e-06 | 8.1062e-06 | 0.99997
SWL2010-05_1#324 | 323 BC - 2 | 2011-02-17 09:18 | 24C | 0.96411 | 3.4658e-06 | 8.0466e-06 | 0.96412 | 33.5923 | 0.0001334
SWL2010-05_1#325 | 324 " | 2011-02-17 09:21 | 24C | 0.96216 | 2.7432e-06 | 7.987e-06 | 0.96216 | 33.5158 | 0.00010761
SWL2010-05_1#326 | 325 " | 2011-02-17 09:23 | 24C | 0.95522 | 2.7425e-06 | 7.9274e-06 | 0.95523 | 33.245 | 0.00010451
SWL2010-05_1#327 | 326 " | 2011-02-17 09:25 | 24C | 0.94689 | 2.89e-06 | 8.1062e-06 | 0.9469 | 32.9198 | 0.00011317
SWL2010-05_1#328 | 327 " | 2011-02-17 09:27 | 24C | 0.94209 | 4.4536e-06 | 7.987e-06 | 0.9421 | 32.7328 | 0.0001688
SWL2010-05_1#329 | 327 Duplicate | 2011-02-17 09:30 | 24C | 0.9421 | 3.2141e-06 | 8.1062e-06 | 0.94211 | 32.7334 | 0.00012361 | Duplicate
SWL2010-05_1#330 | 328 BC - 2 | 2011-02-17 09:32 | 24C | 0.9279 | 1.1199e-06 | 8.1062e-06 | 0.92791 | 32.1809 | 4.265e-05
SWL2010-05_1#331 | 329 " | 2011-02-17 09:34 | 24C | 0.91093 | 3.8647e-06 | 8.0466e-06 | 0.91094 | 31.5227 | 0.00015041
SWL2010-05_1#332 | 330 " | 2011-02-17 09:36 | 24C | 0.88721 | 2.0947e-06 | 8.1062e-06 | 0.88721 | 30.6059 | 8.1765e-05
SWL2010-05_1#333 | 331 PAG - 5 | 2011-02-17 09:39 | 24C | 0.95705 | 1.3713e-06 | 8.0466e-06 | 0.95706 | 33.3165 | 5.1197e-05
SWL2010-05_1#334 | 332 " | 2011-02-17 09:41 | 24C | 0.9466 | 2.4908e-06 | 8.1062e-06 | 0.94661 | 32.9087 | 9.4938e-05
SWL2010-05_1#335 | 332 Duplicate | 2011-02-17 09:44 | 24C | 0.94651 | 3.3587e-06 | 8.1062e-06 | 0.94652 | 32.905 | 0.00013477 | Duplicate
SWL2010-05_1#336 | 333 PAG - 5 | 2011-02-17 09:46 | 24C | 0.92577 | 1.7702e-06 | 8.0466e-06 | 0.92578 | 32.0983 | 6.9481e-05
SWL2010-05_1#337 | 335 " | 2011-02-17 09:49 | 24C | 0.91206 | 3.8748e-06 | 7.987e-06 | 0.91206 | 31.5662 | 0.00015214 | No Bottle #334
SWL2010-05_1#338 | 336 " | 2011-02-17 09:51 | 24C | 0.89843 | 2.4921e-06 | 7.987e-06 | 0.89843 | 31.0389 | 9.751e-05
SWL2010-05_1#339 | 337 " | 2011-02-17 09:54 | 24C | 0.8833 | 2.4909e-06 | 8.0466e-06 | 0.8833 | 30.4551 | 9.7017e-05
SWL2010-05_1#340 | 338 PAG - 6 | 2011-02-17 09:56 | 24C | 0.9294 | 3.6085e-06 | 8.0466e-06 | 0.92941 | 32.2391 | 0.00014094
SWL2010-05_1#341 | 339 " | 2011-02-17 09:59 | 24C | 0.92864 | 1.1199e-06 | 8.1062e-06 | 0.92865 | 32.2096 | 4.1875e-05
SWL2010-05_1#342 | 339 Duplicate | 2011-02-17 10:01 | 24C | 0.92861 | 2.7097e-06 | 8.1062e-06 | 0.92862 | 32.2086 | 0.00010447 | Duplicate
SWL2010-05_1#343 | 340 PAG - 6 | 2011-02-17 10:04 | 24C | 0.91602 | 2.7425e-06 | 8.0466e-06 | 0.91603 | 31.7199 | 0.00010347
SWL2010-05_1#344 | 341 " | 2011-02-17 10:06 | 24C | 0.89761 | 3.8746e-06 | 8.1062e-06 | 0.89762 | 31.0074 | 0.0001526
SWL2010-05_1#345 | 342 " | 2011-02-17 10:08 | 24C | 0.88673 | 3.1417e-06 | 8.1062e-06 | 0.88673 | 30.5874 | 0.00012138
SWL2010-05_1#346 | 343 " | 2011-02-17 10:11 | 24C | 0.84755 | 3.8434e-06 | 7.987e-06 | 0.84756 | 29.0824 | 0.00014761
SWL2010-05_1#347 | Ref. CB - 06 | 2011-02-17 10:13 | 24C | 0.99875 | 2.8807e-06 | 7.9274e-06 | 0.99876 | 34.9513 | 0.00011216
SWL2010-05_1#348 | 205 Slip - 5 | 2011-02-17 10:43 | 24C | 0.93206 | 2.2398e-06 | 7.987e-06 | 0.93207 | 32.3425 | 8.801e-05
SWL2010-05_1#349 | 205 Duplicate | 2011-02-17 10:46 | 24C | 0.93205 | 2.7425e-06 | 7.987e-06 | 0.93206 | 32.3423 | 0.00010972 | Duplicate
SWL2010-05_1#350 | 206 Slip - 5 | 2011-02-17 10:49 | 24C | 0.93202 | 3.1414e-06 | 8.1062e-06 | 0.93203 | 32.3409 | 0.00012618
SWL2010-05_1#351 | 207 " | 2011-02-17 10:51 | 24C | 0.92459 | 2.4912e-06 | 7.9274e-06 | 0.9246 | 32.0523 | 9.4938e-05
SWL2010-05_1#352 | 208 " | 2011-02-17 10:53 | 24C | 0.91867 | 2.4909e-06 | 8.0466e-06 | 0.91868 | 31.8227 | 9.4895e-05
SWL2010-05_1#353 | 209 " | 2011-02-17 10:56 | 24C | 0.91403 | 1.7702e-06 | 8.0466e-06 | 0.91403 | 31.6425 | 6.9868e-05
SWL2010-05_1#354 | 210 " | 2011-02-17 10:58 | 24C | 0.9139 | 3.334e-06 | 7.987e-06 | 0.91391 | 31.6379 | 0.00012956
SWL2010-05_1#355 | 212 Slip - 4 | 2011-02-17 11:01 | 24C | 0.9341 | 4.2633e-06 | 7.987e-06 | 0.93411 | 32.422 | 0.00016473 | No Bottle #211
SWL2010-05_1#356 | 213 " | 2011-02-17 11:04 | 24C | 0.93341 | 3.6104e-06 | 7.987e-06 | 0.93342 | 32.3952 | 0.00013965
SWL2010-05_1#357 | 214 " | 2011-02-17 11:06 | 24C | 0.92376 | 2.7425e-06 | 7.987e-06 | 0.92377 | 32.0201 | 0.00010657
SWL2010-05_1#358 | 214 Duplicate | 2011-02-17 11:07 | 24C | 0.92382 | 2.7425e-06 | 8.0466e-06 | 0.92383 | 32.0223 | 0.00011077 | Duplicate
SWL2010-05_1#359 | 215 Slip - 4 | 2011-02-17 11:10 | 24C | 0.92084 | 3.4497e-06 | 7.987e-06 | 0.92085 | 31.9067 | 0.00013584
SWL2010-05_1#360 | 216 " | 2011-02-17 11:13 | 24C | 0.91378 | 3.5197e-06 | 7.987e-06 | 0.91379 | 31.633 | 0.00013827
SWL2010-05_1#361 | 217 " | 2011-02-17 11:15 | 24C | 0.91361 | 3.1417e-06 | 7.9274e-06 | 0.91362 | 31.6266 | 0.00011936
SWL2010-05_1#362 | 218 BCL - 6a | 2011-02-17 11:17 | 24C | 0.93275 | 2.4909e-06 | 7.987e-06 | 0.93276 | 32.3693 | 9.4885e-05
SWL2010-05_1#363 | 219 " | 2011-02-17 11:20 | 24C | 0.93259 | 2.2398e-06 | 8.1062e-06 | 0.9326 | 32.3632 | 8.5299e-05
SWL2010-05_1#364 | 219 Duplicate | 2011-02-17 11:22 | 24C | 0.93293 | 3.3588e-06 | 7.987e-06 | 0.93293 | 32.3762 | 0.00013277 | Duplicate
SWL2010-05_1#365 | 220 BCL - 6a | 2011-02-17 11:25 | 24C | 0.93203 | 3.1312e-06 | 8.1062e-06 | 0.93204 | 32.3415 | 0.00012108
SWL2010-05_1#366 | 221 " | 2011-02-17 11:27 | 24C | 0.91993 | 3.1313e-06 | 8.1062e-06 | 0.91993 | 31.8713 | 0.00012365
SWL2010-05_1#367 | 222 " | 2011-02-17 11:30 | 24C | 0.89369 | 5.3482e-06 | 8.1062e-06 | 0.8937 | 30.8561 | 0.00020704
SWL2010-05_1#368 | LOOP 14 | 2011-02-17 11:33 | 24C | 0.89535 | 3.8647e-06 | 7.987e-06 | 0.89536 | 30.9202 | 0.0001502 | July 17/10 15:16
SWL2010-05_1#369 | Ref. CB - 06 | 2011-02-17 11:36 | 24C | 0.99891 | 4.6466e-06 | 8.0466e-06 | 0.99892 | 34.9574 | 0.00018033
SWL2010-05_1#370 | 304 PAG - 1 | 2011-02-17 11:42 | 24C | 0.92629 | 2.8794e-06 | 7.987e-06 | 0.9263 | 32.1184 | 0.00011151
SWL2010-05_1#371 | 305 " | 2011-02-17 11:44 | 24C | 0.92629 | 1.4046e-06 | 8.0466e-06 | 0.9263 | 32.1184 | 5.2235e-05
SWL2010-05_1#372 | 306 " | 2011-02-17 11:47 | 24C | 0.92272 | 3.2141e-06 | 8.1062e-06 | 0.92273 | 31.9798 | 0.00012654
SWL2010-05_1#373 | 307 " | 2011-02-17 11:49 | 24C | 0.92229 | 3.3456e-06 | 8.1062e-06 | 0.9223 | 31.963 | 0.00013
SWL2010-05_1#374 | 308 " | 2011-02-17 11:51 | 24C | 0.92226 | 1.3713e-06 | 8.0466e-06 | 0.92227 | 31.962 | 5.2246e-05
SWL2010-05_1#375 | 308 Duplicate | 2011-02-17 11:53 | 24C | 0.92216 | 1.1196e-06 | 8.1062e-06 | 0.92217 | 31.9579 | 4.265e-05 | Duplicate
SWL2010-05_1#376 | 309 PAG - 2 | 2011-02-17 11:55 | 24C | 0.92762 | 3.2141e-06 | 7.9274e-06 | 0.92763 | 32.1699 | 0.00012244
SWL2010-05_1#377 | 309 Duplicate | 2011-02-17 11:58 | 24C | 0.92751 | 4.7275e-06 | 8.1062e-06 | 0.92752 | 32.1657 | 0.00018259 | Duplicate
SWL2010-05_1#378 | 310 PAG - 2 | 2011-02-17 12:00 | 24C | 0.92741 | 2.7425e-06 | 8.1062e-06 | 0.92742 | 32.1619 | 0.00010865
SWL2010-05_1#379 | 311 " | 2011-02-17 12:03 | 24C | 0.92685 | 2.4908e-06 | 8.1062e-06 | 0.92686 | 32.14 | 9.6599e-05
SWL2010-05_1#380 | 312 " | 2011-02-17 12:05 | 24C | 0.92425 | 2.7425e-06 | 8.1062e-06 | 0.92426 | 32.0391 | 0.00010656
SWL2010-05_1#381 | 313 " | 2011-02-17 12:07 | 24C | 0.91 | 3.3588e-06 | 8.1062e-06 | 0.91 | 31.4864 | 0.00013116
SWL2010-05_1#382 | 314 PAG - 3 | 2011-02-17 12:10 | 24C | 0.9252 | 3.8654e-06 | 7.987e-06 | 0.92521 | 32.0759 | 0.00014998
SWL2010-05_1#383 | 315 " | 2011-02-17 12:12 | 24C | 0.92526 | 1.4046e-06 | 8.1062e-06 | 0.92527 | 32.0783 | 5.2235e-05
SWL2010-05_1#384 | 315 Duplicate | 2011-02-17 12:14 | 24C | 0.92524 | 2.2391e-06 | 7.987e-06 | 0.92525 | 32.0776 | 8.5299e-05 | Duplicate
SWL2010-05_1#385 | 316 PAG - 3 | 2011-02-17 12:17 | 24C | 0.92513 | 2.7435e-06 | 8.0466e-06 | 0.92513 | 32.0731 | 0.00010656
SWL2010-05_1#386 | 317 " | 2011-02-17 12:19 | 24C | 0.90869 | 2.7425e-06 | 8.1062e-06 | 0.9087 | 31.436 | 0.00010657
SWL2010-05_1#387 | 318 " | 2011-02-17 12:22 | 24C | 0.90765 | 3.6104e-06 | 7.987e-06 | 0.90766 | 31.3957 | 0.00013711
SWL2010-05_1#388 | 319 BCL - 7" | 2011-02-17 12:24 | 24C | 0.94684 | 4.3337e-06 | 7.987e-06 | 0.94685 | 32.918 | 0.00016797
SWL2010-05_1#389 | 320 " | 2011-02-17 12:26 | 24C | 0.94915 | 2.4909e-06 | 8.1062e-06 | 0.94916 | 33.008 | 9.5561e-05
SWL2010-05_1#390 | 320 Duplicate | 2011-02-17 12:29 | 24C | 0.94905 | 2.4909e-06 | 8.0466e-06 | 0.94906 | 33.0041 | 9.6609e-05 | Duplicate
SWL2010-05_1#391 | 321 BCL - 7 | 2011-02-17 12:32 | 24C | 0.93003 | 2.8343e-06 | 8.1062e-06 | 0.93004 | 32.2637 | 0.00011027
SWL2010-05_1#392 | 322 " | 2011-02-17 12:34 | 24C | 0.91856 | 3.4553e-06 | 8.1062e-06 | 0.91857 | 31.8185 | 0.00013659
SWL2010-05_1#393 | LOOP 18 | 2011-02-17 12:37 | 24C | 0.92059 | 2.7425e-06 | 7.987e-06 | 0.9206 | 31.8972 | 0.00010448
SWL2010-05_1#394 | Ref. CB - 06 | 2011-02-17 12:39 | 24C | 0.99897 | 1.1396e-06 | 7.987e-06 | 0.99897 | 34.9596 | 4.5227e-05
SWL2010-05_1#395 | 021 NP - 3 | 2011-02-17 12:48 | 24C | 0.97218 | 2.4912e-06 | 8.1062e-06 | 0.97219 | 33.9081 | 9.6976e-05
SWL2010-05_1#396 | 022 " | 2011-02-17 12:50 | 24C | 0.9696 | 3.1253e-06 | 8.1062e-06 | 0.96961 | 33.8072 | 0.00012381
SWL2010-05_1#397 | 023 " | 2011-02-17 12:52 | 24C | 0.96901 | 2.2401e-06 | 8.0466e-06 | 0.96901 | 33.7838 | 8.7077e-05
SWL2010-05_1#398 | 024 " | 2011-02-17 12:55 | 24C | 0.96876 | 2.2391e-06 | 8.1062e-06 | 0.96877 | 33.7741 | 8.8738e-05
SWL2010-05_1#399 | 025 " | 2011-02-17 12:57 | 24C | 0.96655 | 1.1199e-06 | 7.987e-06 | 0.96656 | 33.6876 | 4.265e-05
SWL2010-05_1#400 | 026 " | 2011-02-17 12:59 | 24C | 0.96249 | 2.226e-06 | 8.1062e-06 | 0.9625 | 33.5289 | 8.5352e-05
SWL2010-05_1#401 | 027 " | 2011-02-17 13:02 | 24C | 0.95512 | 1.1196e-06 | 8.1062e-06 | 0.95513 | 33.2408 | 4.4374e-05
SWL2010-05_1#402 | 028 " | 2011-02-17 13:04 | 24C | 0.94269 | 2.8897e-06 | 8.1062e-06 | 0.9427 | 32.7563 | 0.00011317
SWL2010-05_1#403 | 029 " | 2011-02-17 13:06 | 24C | 0.93555 | 3.4551e-06 | 7.987e-06 | 0.93555 | 32.4781 | 0.00013412
SWL2010-05_1#404 | 030 " | 2011-02-17 13:08 | 24C | 0.93235 | 3.1311e-06 | 8.1062e-06 | 0.93235 | 32.3537 | 0.00012313
SWL2010-05_1#405 | 031 " | 2011-02-17 13:10 | 24C | 0.92932 | 3.3595e-06 | 7.987e-06 | 0.92933 | 32.2362 | 0.00012795
SWL2010-05_1#406 | 032 " | 2011-02-17 13:12 | 24C | 0.92544 | 2.2276e-06 | 8.1062e-06 | 0.92545 | 32.0855 | 8.7239e-05
SWL2010-05_1#407 | 033 " | 2011-02-17 13:15 | 24C | 0.92276 | 3.466e-06 | 8.0466e-06 | 0.92277 | 31.9814 | 0.00013375
SWL2010-05_1#408 | 034 " | 2011-02-17 13:17 | 24C | 0.9228 | 3.2144e-06 | 8.1062e-06 | 0.92281 | 31.9827 | 0.00012686
SWL2010-05_1#409 | LOOP - 1 | 2011-02-17 13:20 | 24C | 0.92841 | 2.4909e-06 | 8.0466e-06 | 0.92842 | 32.2007 | 9.4928e-05 | July 10/10 02:03
SWL2010-05_1#410 | 035 NP - 5 | 2011-02-17 13:22 | 24C | 0.98441 | 3.623e-06 | 8.0466e-06 | 0.98442 | 34.3877 | 0.0001426
SWL2010-05_1#411 | 036 " | 2011-02-17 13:25 | 24C | 0.9818 | 3.2141e-06 | 8.0466e-06 | 0.98181 | 34.2852 | 0.0001289
SWL2010-05_1#412 | 037 " | 2011-02-17 13:27 | 24C | 0.97763 | 2.4915e-06 | 8.1062e-06 | 0.97764 | 34.1217 | 9.4938e-05
SWL2010-05_1#413 | 038 " | 2011-02-17 13:32 | 24C | 0.97533 | 6.1455e-06 | 8.1062e-06 | 0.97534 | 34.0315 | 0.00024173
SWL2010-05_1#414 | 038 Duplicate ?? | 2011-02-17 13:37 | 24C | 0.97444 | 2.9907e-06 | 8.0466e-06 | 0.97444 | 33.9964 | 0.0001146 | Duplicate???
SWL2010-05_1#415 | 039 NP - 5 | 2011-02-17 13:40 | 24C | 0.97134 | 2.8904e-06 | 7.9274e-06 | 0.97135 | 33.8752 | 0.00011317
SWL2010-05_1#416 | 040 " | 2011-02-17 13:42 | 24C | 0.97014 | 3.9746e-06 | 7.987e-06 | 0.97015 | 33.8282 | 0.00015494
SWL2010-05_1#417 | 041 " | 2011-02-17 13:46 | 24C | 0.96862 | 2.2394e-06 | 7.987e-06 | 0.96863 | 33.7686 | 8.7876e-05
SWL2010-05_1#418 | Ref. CB - 06 | 2011-02-17 13:48 | 24C | 0.99889 | 1.7705e-06 | 7.987e-06 | 0.9989 | 34.9567 | 7.0829e-05

## Sp
Example data in pale yellow
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded | [16 empty] | Should abs(diff) be used?
Cast No. | Station | Bottle | Pressure | Sample ID | Salt-1 | Salt-2 | diff | abs(diff) | diff^2 | Chauvenet | Accept or Reject |  | QA/QC - Analyst | QA/QC - Data | mean abs(diff) | sd diff | Chauvenet | Accept or Reject | mean diff | sd diff
|  |  |  | 6 | 34.7256 | 34.7255 |  | 8.0109e-05 | 6.4174e-09 | 0.65858 | ACCEPT |  |  |  | 0.0039961 | 0.005946 |  |  | 0.00097442 | 0.0071199
|  |  |  | 11 | 34.3057 | 34.3064 |  | 0.00060272 | 3.6327e-07 | 0.57069 | ACCEPT |  |  |  | 0.00097442 | 0.0071199
|  |  |  | 6 | 33.9329 | 33.9496 | -0.016663 | 0.016663 | 0.00027764 | 2.1303 | ACCEPT | [5 empty] | 2.4771 | ACCEPT
|  |  |  | 14 | 32.0849 | 32.0862 | -0.0012589 | 0.0012589 | 1.5847e-06 | 0.46034 | ACCEPT | [5 empty] | 0.31366 | ACCEPT
|  |  |  | 20 | 34.1047 | 34.106 | -0.0013123 | 0.0013123 | 1.722e-06 | 0.45136 | ACCEPT | [5 empty] | 0.32116 | ACCEPT
|  |  |  | 38 | 34.0315 | 33.9964 | 0.035122 | 0.035122 | 0.0012335 | 5.2348 | REJECT | [5 empty] | 4.796 | REJECT
|  |  |  | 44 | 33.2847 | 33.2864 | -0.0017128 | 0.0017128 | 2.9337e-06 | 0.384 | ACCEPT | [5 empty] | 0.37742 | ACCEPT
|  |  |  | 54 | 34.1589 | 34.1468 | 0.012154 | 0.012154 | 0.00014771 | 1.3719 | ACCEPT | [5 empty] | 1.5701 | ACCEPT
|  |  |  | 71 | 34.2323 | 34.2313 | 0.0010719 | 0.0010719 | 1.149e-06 | 0.49178 | ACCEPT | [5 empty] | 0.013696 | ACCEPT
|  |  |  | 82 | 32.9015 | 32.9042 | -0.0026894 | 0.0026894 | 7.2327e-06 | 0.21976 | ACCEPT | [5 empty] | 0.51458 | ACCEPT
|  |  |  | 88 | 34.2327 | 34.2346 | -0.0019264 | 0.0019264 | 3.7111e-06 | 0.34807 | ACCEPT | [5 empty] | 0.40742 | ACCEPT
|  |  |  | 96 | 33.6395 | 33.6349 | 0.0046005 | 0.0046005 | 2.1165e-05 | 0.10166 | ACCEPT | [5 empty] | 0.50929 | ACCEPT
|  |  |  | 105 | 32.1785 | 32.1793 | -0.00072098 | 0.00072098 | 5.1981e-07 | 0.5508 | ACCEPT | [5 empty] | 0.23812 | ACCEPT
|  |  |  | 122 | 32.6011 | 32.6033 | -0.0022507 | 0.0022507 | 5.0655e-06 | 0.29354 | ACCEPT | [5 empty] | 0.45297 | ACCEPT
|  |  |  | 127 | 33.9331 | 33.9304 | 0.0026474 | 0.0026474 | 7.0087e-06 | 0.22682 | ACCEPT | SAL ALL |  |  |  |  | 0.23497 | ACCEPT
|  |  |  | 137 | 32.9984 | 32.9977 | 0.00067902 | 0.00067902 | 4.6106e-07 | 0.55786 | ACCEPT | N tot |  | 48 |  |  | 0.041489 | ACCEPT
|  |  |  | 143 | 34.303 | 34.3057 | -0.0026932 | 0.0026932 | 7.2532e-06 | 0.21912 | ACCEPT | N used |  | 47 |  |  | 0.51512 | ACCEPT
|  |  |  | 149 | 33.466 | 33.4552 | 0.010784 | 0.010784 | 0.0001163 | 1.1416 | ACCEPT | sum of sqrs |  | 0.0011946 |  |  | 1.3778 | ACCEPT
|  |  |  | 172 | 32.1378 | 32.1435 | -0.0056534 | 0.0056534 | 3.1961e-05 | 0.27873 | ACCEPT |  |  | 1.2709e-05 |  |  | 0.93088 | ACCEPT
|  |  |  | 177 | 31.9528 | 31.9524 | 0.00036621 | 0.00036621 | 1.3411e-07 | 0.61047 | ACCEPT | Sp |  | 0.0035649 |  |  | 0.085423 | ACCEPT
|  |  |  | 187 | 31.4891 | 31.4859 | 0.0031757 | 0.0031757 | 1.0085e-05 | 0.13796 | ACCEPT | N discard |  | 1 |  |  | 0.30918 | ACCEPT
|  |  |  | 194 | 31.4799 | 31.4809 | -0.00098991 | 0.00098991 | 9.7993e-07 | 0.50557 | ACCEPT | STD for compare |  | 0.0050899 |  |  | 0.27589 | ACCEPT
|  |  |  | 201 | 31.6768 | 31.6771 | -0.00028992 | 0.00028992 | 8.4052e-08 | 0.6233 | ACCEPT | mean for compare |  | 0.00024787 |  |  | 0.17758 | ACCEPT
|  |  |  | 205 | 32.3425 | 32.3423 | 0.00016785 | 0.00016785 | 2.8173e-08 | 0.64383 | ACCEPT | Critical values for Chauvenet statistic |  |  |  |  | 0.11328 | ACCEPT
|  |  |  | 214 | 32.0201 | 32.0223 | -0.0021973 | 0.0021973 | 4.828e-06 | 0.30252 | ACCEPT | n | Chauvenet |  |  |  | 0.44547 | ACCEPT
|  |  |  | 219 | 32.3632 | 32.3762 | -0.012985 | 0.012985 | 0.00016862 | 1.5118 | ACCEPT | 3 | 1.37 |  |  |  | 1.9606 | ACCEPT
|  |  |  | 224 | 32.5614 | 32.5597 | 0.0016403 | 0.0016403 | 2.6906e-06 | 0.39619 | ACCEPT | 4 | 1.53 |  |  |  | 0.093527 | ACCEPT
|  |  |  | 231 | 32.5326 | 32.5342 | -0.0015945 | 0.0015945 | 2.5426e-06 | 0.40389 | ACCEPT | 5 | 1.64 |  |  |  | 0.36081 | ACCEPT
|  |  |  | 238 | 32.4907 | 32.4908 | -0.00019455 | 0.00019455 | 3.785e-08 | 0.63934 | ACCEPT | 6 | 1.73 |  |  |  | 0.16418 | ACCEPT
|  |  |  | 240 | 32.5701 | 32.5634 | 0.0066719 | 0.0066719 | 4.4514e-05 | 0.45003 | ACCEPT | 7 | 1.8 |  |  |  | 0.80022 | ACCEPT
|  |  |  | 246 | 32.0565 | 32.0567 | -0.00026703 | 0.00026703 | 7.1304e-08 | 0.62715 | ACCEPT | 8 | 1.87 |  |  |  | 0.17436 | ACCEPT
|  |  |  | 250 | 32.6611 | 32.662 | -0.00087738 | 0.00087738 | 7.698e-07 | 0.5245 | ACCEPT | 9 | 1.91 |  |  |  | 0.26009 | ACCEPT
|  |  |  | 258 | 31.753 | 31.7472 | 0.0057602 | 0.0057602 | 3.318e-05 | 0.29669 | ACCEPT | 10 | 1.97 |  |  |  | 0.67217 | ACCEPT
|  |  |  | 260 | 31.706 | 31.7029 | 0.0031376 | 0.0031376 | 9.8445e-06 | 0.14438 | ACCEPT | 20 | 2.24 |  |  |  | 0.30382 | ACCEPT
|  |  |  | 266 | 31.3275 | 31.3294 | -0.0019341 | 0.0019341 | 3.7406e-06 | 0.34679 | ACCEPT | 30 | 2.39 |  |  |  | 0.4085 | ACCEPT
|  |  |  | 270 | 32.3136 | 32.3092 | 0.0043182 | 0.0043182 | 1.8647e-05 | 0.054185 | ACCEPT | 50 | 2.57 |  |  |  | 0.46964 | ACCEPT
|  |  |  | 276 | 32.7235 | 32.7229 | 0.00053406 | 0.00053406 | 2.8522e-07 | 0.58224 | ACCEPT | 80 | 2.74 |  |  |  | 0.061849 | ACCEPT
|  |  |  | 279 | 32.1755 | 32.1879 | -0.012329 | 0.012329 | 0.00015201 | 1.4015 | ACCEPT | 100 | 2.81 |  |  |  | 1.8685 | ACCEPT
|  |  |  | 286 | 32.2223 | 32.2237 | -0.0014191 | 0.0014191 | 2.0138e-06 | 0.4334 | ACCEPT | 150 | 2.93 |  |  |  | 0.33617 | ACCEPT
|  |  |  | 292 | 32.1627 | 32.1566 | 0.0061569 | 0.0061569 | 3.7908e-05 | 0.36342 | ACCEPT | 200 | 3.02 |  |  |  | 0.72789 | ACCEPT
|  |  |  | 297 | 31.9237 | 31.9227 | 0.0010071 | 0.0010071 | 1.0142e-06 | 0.50269 | ACCEPT | [5 empty] | 0.0045876 | ACCEPT
|  |  |  | 300 | 32.0952 | 32.0946 | 0.00061798 | 0.00061798 | 3.819e-07 | 0.56812 | ACCEPT | [5 empty] | 0.050062 | ACCEPT
|  |  |  | 308 | 31.962 | 31.9579 | 0.0041523 | 0.0041523 | 1.7242e-05 | 0.026277 | ACCEPT | [5 empty] | 0.44634 | ACCEPT
|  |  |  | 309 | 32.1699 | 32.1657 | 0.0042381 | 0.0042381 | 1.7962e-05 | 0.040712 | ACCEPT | Pooled standard deviation of pairs of samples (Sp) was calculated by: |  |  |  |  | 0.45839 | ACCEPT
|  |  |  | 315 | 32.0783 | 32.0776 | 0.0006485 | 0.0006485 | 4.2055e-07 | 0.56299 | ACCEPT |  | Sp = SQRT{sum (d*d)/2k} |  |  |  | 0.045775 | ACCEPT
|  |  |  | 320 | 33.008 | 33.0041 | 0.0039444 | 0.0039444 | 1.5558e-05 | 0.0086877 | ACCEPT |  | where k = no. of pairs and d = difference between pairs |  |  |  | 0.41714 | ACCEPT
|  |  |  | 327 | 32.7328 | 32.7334 | -0.00056076 | 0.00056076 | 3.1445e-07 | 0.57775 | ACCEPT | [5 empty] | 0.21562 | ACCEPT
|  |  |  | 332 | 32.9087 | 32.905 | 0.0036278 | 0.0036278 | 1.3161e-05 | 0.061937 | ACCEPT | [5 empty] | 0.37267 | ACCEPT
|  |  |  | 339 | 32.2096 | 32.2086 | 0.0010757 | 0.0010757 | 1.1572e-06 | 0.49114 | ACCEPT | [5 empty] | 0.014232 | ACCEPT
|  |  |  | 347 | 32.5884 | 32.5874 | 0.00099182 | 0.00099182 | 9.8371e-07 | 0.50525 | ACCEPT | [5 empty] | 0.0024445 | ACCEPT

## Manual Salt Calc
Worksheet in case any sample needs to be recalcualted
Sample # | Bath temp | Pressure | M | Autosal Display of Cond. Ratio (Flush 1) | Autosal Display of Cond. Ratio (Flush 2) | Autosal Display of Cond. Ratio (Flush 3) | Autosal Display of Cond. Ratio (Flush 4) | Average Conductivity Ratio | Correction to Ratio | Salinity
223 | 24 | 0 | 2 | 1.9596 | 1.9597 | 1.9598 |  | 1.9597 | -9.5367e-07 | #VALUE!
1480 | 24 | 0 | ?? | 1.9957 | 1.9958 | 1.9959 |  | 1.9958 |  | 34.9168
Did not use this to calulate salinity
Mary Steel

## Autosal Data Examplesheet
2010-07 Sal Data | Example data in pale yellow
coloured data is to be entered on final paper in that format
Salinometer ID | Station | Sample ID (bottle label) | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#2 | P151 |  | 2010-09-18 13:08 | 24C | 0.99997 | 2.7426e-06 | 1.9073e-06 | 0.99997
2010-07#23 | CABOS | 23 | 2010-09-18 14:29 | 24C | 0.9972 | 4.6253e-06 | 1.967e-06 | 0.9972 | 34.8897 | 0.00017952
2010-07#24 | CABOS | 24 | 2010-09-18 14:33 | 24C | 0.99698 | 3.2141e-06 | 2.0266e-06 | 0.99699 | 34.8814 | 0.00012599
2010-07#25 | CABOS | 25 | 2010-09-18 14:35 | 24C | 0.99679 | 5.0935e-06 | 2.0266e-06 | 0.99679 | 34.8738 | 0.00020105
2010-07#26 | CABOS | 26 | 2010-09-18 16:01 | 24C | 0.99659 | 3.5943e-06 | 2.0266e-06 | 0.9966 | 34.866 | 0.00013924
2010-07#27 | CABOS | 27 | 2010-09-18 16:05 | 24C | 0.99637 | 3.8544e-06 | 1.967e-06 | 0.99637 | 34.8572 | 0.00015042
2010-07#28 | CABOS | 27 | 2010-09-18 16:07 | 24C | 0.99637 | 3.4659e-06 | 2.0266e-06 | 0.99637 | 34.8573 | 0.00013203
2010-07#46 | 2010-08 LOOP 15 |  | 2010-09-18 22:18 | 24C | 0.73328 | 4.8273e-06 | 2.0862e-06 | 0.73328 | 24.7563 | 0.00018047
2010-07#47 | 2010-08 LOOP 16 |  | 2010-09-18 22:22 | 24C | 0.74613 | 2.7425e-06 | 2.0266e-06 | 0.74614 | 25.2382 | 0.00010447
2010-07#48 | 2010-08 LOOP 17 |  | 2010-09-18 22:24 | 24C | 0.77025 | 2.2241e-06 | 1.967e-06 | 0.77026 | 26.1456 | 8.3074e-05
Calibration#3 | P151 |  | 2010-09-22 01:17 | 24C | 0.99995 | 3.5403e-06 | 1.508e-05 | 0.99997
Calibration#4 | P151 |  | 2010-09-22 01:19 | 24C | 0.99997 | 9.0741e-06 | -1.967e-06 | 0.99997
2010-07#75 | Deep Water Standard (2010-07) | CB-28b | 2010-09-22 03:45 | 24C | 0.99697 | 4.6245e-06 | -2.0266e-06 | 0.99697 | 34.8808 | 0.00018091
2010-07#78 | CB-31b | 45 | 2010-09-22 03:58 | 24C | 0.99878 | 4.1861e-06 | -1.9073e-06 | 0.99878 | 34.952 | 0.00016514
2010-07#79 | CB-31b | 46 | 2010-09-22 04:01 | 24C | 0.99807 | 3.8626e-06 | -1.967e-06 | 0.99807 | 34.9241 | 0.00015407
2010-07#80 | CB-31b | 47 | 2010-09-22 04:05 | 24C | 0.99751 | 6.4383e-06 | -2.0266e-06 | 0.99751 | 34.902 | 0.00025197
2010-07#81 | CB-31b | 48 | 2010-09-22 04:09 | 24C | 0.99711 | 4.3205e-06 | -1.9073e-06 | 0.99711 | 34.8864 | 0.00016818
EXAMPLE SHEET ONLY