## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18 January 2018 | Changed channel name “Fluorescence:Wetlabs:CDOM” to “Fluorescence_CDOM:URU:Wetlabs” in the *.CTD files. J.L. |

## PROCESSING NOTES

Cruise: 2010-05

Agency: OSD

Location: North Pacific	/ Bering Sea / Chukchi Sea

Project: C3O/IPY

Party Chief: Vagle S.

Platform: Sir Wilfrid Laurier

Date: July 6, 2010 - July 22, 2010

Processed by: Germaine Gatien

Date of Processing: 9 July 2012 – 5 February 2016

Number of original HEX files: 43

Number of CTD files:  43	Number of CHE files:	43

Number of original TSG files: 21 (To be processed later)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#993DR), an SBE 43 DO sensor (#1117), a SeaPoint Fluorometer (#2336), a Wet Labs CDOM Fluorometer (#1076), a Biospherical QSP-200L4S PAR sensor (#70123), an ISUS (#02) and an altimeter (#1252).

22 OceanTest Equipment 10L bottles were used mounted on the rosette.

A thermosalinograph (SeaBird 21 S/N 3274) was mounted with a fluorometer (SCF2859) with 30X cable and a remote temperature sensor #0271.

The data logging computer was W9B6IOS101V55.

The data acquisition program was Seasave 7.20c.

The deck unit was a Seabird model 11, serial # 11P53201-0800.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572. Bottles were analyzed in February 2011.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order. Information about bottle sampling was found in a spreadsheet.

Salinity samples were analyzed 7 months after the end of the cruise. While the method for sealing salinity bottles during this cruise appears to perform much better than that used in most Pacific cruises, there is still some evidence of evaporation in many samples. It is recommended that analysis be carried out within a few months. Poor flushing of Niskin bottles is also a source of error of the same sign. Fortunately, post-cruise calibrations indicate that the calibration drift in temperature and conductivity sensors was small so no recalibration was applied.

A SeaPoint fluorometer was mounted on the CTD for this cruise. It reads higher than the extracted chlorophyll in most cases. However, the 1.5m distance between the bottles and CTD would lead to fluorometer readings being higher above the CHL maxima, which is the case for the majority of the bottles. The deepest samples, which are below the maxima, have SeaPoint readings generally higher than the extracted chlorophyll by roughly a factor of ~1.5, but the extracted CHL values are low at that level. At high chlorophyll values the fluorometer is often lower than the extracted CHL.

There was no dissolved oxygen sampling, so the error due to slow response time cannot be assessed. A crude estimate of the error due to calibration drift was made by comparing results when casts were converted using pre-cruise and post-cruise calibrations. The results were extremely noisy especially near the surface where temperature and DO gradients were high, and the only conclusion to be found in the noise, is that using the pre-cruise calibration produced lower values than the post-cruise calibration, by up to 5%. Some of that drift would have occurred both during and after the cruise. DO surface saturation in the early part of the cruise in the open Pacific Ocean was about 105%, which does not suggests DO was reading low then. No recalibration was applied and the data should be considered nominal and results are reported with fewer significant figures than usual.

Processing data from the Laurier cruises would be much simpler if the final information for each variable was reported separately using templates with separate sheets for raw data, final data, precision study and notes and it gives the analyst a way to say what information should go into the metadata and what format should be used for each variable based on the precision study. The combined spreadsheets produced from the Laurier cruises do not always contain all the information needed, while the individual reports do not always make clear what the final values are. The template approach has been adopted by some other programs and has been found to improve processing efficiency and reduce errors. The variables can be merged in processing and a master spreadsheet can be produced with both CTD and sample values at the end of processing.

An error in the PAR configuration parameters was corrected by a linear recalibration.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Files had non-standard names, so those were fixed before conversion.

##### Preliminary Steps

The Daily Science Log Book was obtained and a cruise report was available. There were no rosette log sheets but file 2012-05_SWL_Chem_20120-06-06.xls was provided and contains a sampling log with details on bottles fired.

Extracted chlorophyll, nutrients, salinity, DIC and Alkalinity were provided in a single spreadsheet. (Late in processing O18 data were obtained so some steps had to be repeated.)

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. Most sensors had been recalibrated in late 2009 and there were post-cruise calibrations from January 2011 for the temperature and conductivity sensors and from Dec. 2010 for the dissolved oxygen sensor. There was an error in the PAR parameters that was mentioned in a note with the data. There were also some inconsistencies about which transmissometer and SeaPoint fluorometer were in use. According to the Arctic Group’s calibration spreadsheet fluorometer #2336 was written off in June 2010 before this cruise began. It is likely that sensor #2745 was used since that # is recorded in a set-up page. There was an error in the date of calibration of the SPAR sensor.

There was a change of configuration starting with cast #15 when the fluorometer cable was changed.  Corrections were made to the 2 different conversion files and these were saved as 2012-05-ctd1 and 2012-05-ctd2.xmlcon.

A second version of each of these files was prepared with the post-cruise DO calibration parameters to see how large a change there was between the calibrations.

A third version of the first of these files was prepared with parameters for a different transmissometer because there is some doubt about which was in use.

Tests were then run as follows:
1. Since there was no calibration sampling for dissolved oxygen, a quick study was made to get a rough idea about how much the calibration drifted between the pre-cruise and post-cruise factory checks. The first 14 files were converted using the two different CON files. The differences were plotted against pre-cruise DO, but there was too much noise to make a good judgment, especially near the surface where large temperature gradients likely account for the worst noise. A rough estimate would be that using the pre-cruise calibration produces values that are low by 2% to 5%, but the method of comparison was too inexact to use this to recalibrate. Moreover, some of this drift is likely to have occurred “on the shelf” before or after the cruise. (See 2010-05-DO-drift.xlsx.)

The pre-cruise calibrations were chosen for processing.

2. The first 8 files were converted using version 1 and 3, to test the transmissivity. Those files were then converted to IOS Header format to get %/m and the deep transmissivity values were compared with those from files in the same area and time of year from 2007. Typical values at around 1000db were 60%/m if the sensor is assumed to be #993 and 47%/m if it is sensor #662. During 2007 in the same area the transmissivity was about 58%/m and values of 47%/m seem most unlikely in this region, so the log and header entry of #993 will be assumed to be correct.

All files were converted using either 2010-05-ctd1.xmlcon or 2010-05-ctd2.xmlcon, as appropriate.

Voltage channels were converted for Transmissivity, SeaPoint Fluorescence, WetLabs Fluorescence:CDOM, altimeter and PAR channels. Plots were made to ensure that the voltage channels and the channels derived from them have the same shapes and they do.

All channels were plotted for 1 cast each using the 2 configurations. The temperature and conductivity channels track well on downcasts but are much noisier during upcasts. The primary channels look spikier than the secondary though there was at least one large spike in secondary conductivity. Altimetry looks useful. The transmissivity, PAR, SPAR, fluorescence and dissolved oxygen have the right shapes. Descent rates are high but noisy. The SeaPoint fluorescence offset between downcast and upcast is larger than that of the temperature traces by from about 15 to 50 records. The usual alignment correction is +24 records but that is possibly not enough. However, there is so much noise in the upcasts and many stops for bottles so this is hard to judge.

##### Initial Rosette File Conversion and DO Calibration Study

There was no calibration sampling, so this step was skipped.

##### Hysteresis Study

The nominal hysteresis factor was applied, but there was no sampling below 1000m, so no tests were run to fine-tune the factor.

##### BOTTLE FILE PREPARATION

The ROS files were converted in the same way as the full files except that bottle number and bottle position channels were included and oxygen concentration and salinity were derived.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files. There were a few spikes in the secondary salinity. CTDEDIT was used to remove such spikes from casts #2, 4, 21 and 27; a few near-surface records were also removed from cast #2. The edited files were copied to BOT.

A preliminary header check turned up no problems. The log indicates that the fluorescence went off-scale during cast #14, but that presumably happened at a point not captured in the bottle files.

The BOT files were averaged on bottle number to enable preparation of file ADDSAMP.csv. That file was edited to add sample numbers and was then used to add sample numbers to the BOT files, creating SAM files. REMOVE was used to remove the voltage channels. The files were then bin-averaged on bottle number to create SAMAVG files. A second set of bin-averaged files were created with voltage channels and standard deviations included. Those files were named SAMAVGSTD and will be stored.

Most bottle sample data were presented in a single spreadsheet, 2010-05-SWL_Chem.xls. This was simplified to include only the columns needed for addition to the BOT files and saved as 2010-05-samples.csv.  The spreadsheet indicates that sample “a” is to be archived even where duplicates are available; if the first sample was rejected the two samples were reversed in the spreadsheet. Later, O18 data was delivered, so that was added to the 2010-05-samples.csv file.

The analyst did a precision study using the Chauvenet criterion to identify 1 outlier among the 48 duplicates. With the 1 exclusion the pooled standard deviation was 0.0036. See file 2010-05 SWL Salt Data 2015-01-21.xls)

Some changes were made to the data provided:

- Sample 38 – SAL – The duplicates were rejected as a Chauvenet outlier in precision study; one of the samples is much closer to the bottle than the other, so it was selected and flagged 2.

- Sample #57 – ALK – No value was entered and flag was 4; changed to pad value and flag 5. Comment was available.

- Samples #167 and 168 – DIC – Flag 2 with no comment – A comment was found in another file.

- Sample #190 – Nutrients – Flag 5 was entered with original values – values replaced with pad values.

- Sample #251 – O18 - the entry “=1.12” which should have been “-1.12”.

- Sample #314 – Nutrients - NO3 has a slightly negative value – replaced with 0.

File 2010-05-samples.csv was then converted to individual MRG1 files. Those were put through CLEAN and SORT (on bottle #) and then merged with the SAMAVG files with output MRG. These files were put through CLEAN to remove SeaBird headers and comments and to add 0 flags where flag channels were blank and to remove channels with no data.

Next, text file 2010-05-bot-hdr.txt was prepared to add an explanation of quality flags and some general comments from analysts.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was a lot of scatter in the plot of (CTD SAL – BOTTLE SAL) versus pressure. Possible explanations are:

- In high-gradient waters the 1.5m vertical offset between bottle and CTD would typically lead to the differences being +ive. There are some, but not many, of those in this data and most are near the surface, as expected.

- At the bottom we sometimes see a +ive difference possibly because flushing is poor as the CTD slows suddenly, bringing water from above to fill the Niskin bottles. So the bottles would have lower salinity than the CTD.

- Poor flushing on the way up leads to –ive differences as shed wakes carry water from below, but in open waters like these there is generally sufficient ship motion to minimize that problem.

- A 7-month wait for analysis is likely to lead to some sample evaporation, which would make them saltier than the CTD, hence –ive differences. Moreover, even if there were no evaporation, adsorption has been reported to raise salinity in standard glass bottles by as much as 0.007 during the first 6 months.

To sort out what samples should be flagged, the spreadsheet from COMPARE was used to prepare a table with columns for Event #, Sample #, Pressure, CTD Salinity, CTD Salinity Std. Dev., CTD – Bottle Salinity, plus columns to indicate if the sample was from the bottom of a cast, and whether the CTD was higher than the bottle by >0.005 or >0.01 or lower than the bottle by >0.005 or >0.01. Comments were entered together with suggested flags. In most cases the CTD salinity was noisy and the differences <0.01 so that local gradients likely explain the differences; those cases were not investigated further. Where the information was not sufficient to make a judgement based on the information in the spreadsheet, plots were made. Looking for salinity to match the bottle values enabled decisions about flags. Comments were entered as to why a flag was or was not recommended. This information is in spreadsheet 2010-05-salinity-flag-study.xls.

After consultation with Sarah Zimmermann, the following samples were flagged:

- Sample #35 –Flag 3 – Bottom bottle, bottle value higher by >0.005 but wrong direction to be explained by flushing problem; CTD close to an STD value.

- Sample #39 – Flag 3 – CTD data quiet, bottle value higher by >0.01, corresponds to CTD 10m lower. No reason to expect poor flushing.

- Sample #54 – Flag 3 – Probably not an analysis problem; likely shed wake corruption of bottle contents.

- Sample #86 – Flag 4 - Outlier in comparison with CTD; bottom bottle and no local source of water that salty.

- Sample #87 - Flag 4 - Outlier in comparison with CTD; no source of water that salty within 10m.

- Sample #90 - Flag 3 - Outlier in comparison with CTD; possibly shed wake corruption of bottle contents.

- Sample #91 - Flag 3 - Outlier in comparison with CTD; possibly shed wake corruption of bottle contents.

- Sample #104 - Flag 3 - Outlier in comparison with CTD; no source of water that salty

- Sample #124 – Flag 4 – Outlier in comparison with CTD; near-surface well mixed with no salinity as high as bottle.

- Sample #125 – Flag 4 - Outlier in comparison with CTD; near-bottom – no salinity that high in profile.

- Sample #126 – Flag 3 - Outlier in comparison with CTD; possibly due to shed wake corruption of bottle contents but descent rate suggests good flushing.

- Sample #131 – Flag 3 - Outlier in comparison with CTD; bottle value corresponds to CTD 11m below.

- Sample #133 – Flag 3 – Outlier in comparison with CTD; bottle value corresponds to CTD about 15m below.

- Sample #135 – Flag 3 – Outlier in comparison with CTD; bottle value corresponds to CTD about 13m below.

- Sample #136 – Flag 3 – Outlier in comparison with CTD; could be due to incomplete flushing.

- Sample #144 – Flag 4 – Outlier in comparison with CTD; no source of water this salty within 20m.

- Sample #191 – Flag 3 - Outlier in comparison with CTD; bottom bottle; bottle value corresponds to CTD about 9m above so possibly poor flushing of bottle.

- Sample #214 – Flag 4 - Outlier in comparison with CTD; no water this salty found anywhere in profile.

- Sample #216 – Flag 4 - Outlier in comparison with CTD; well-mixed surface water, no water this salty within 20m.

- Sample #272 – Flag 3 - Outlier in comparison with CTD; bottle value corresponds to CTD about 10m below so possibly incomplete flushing of bottle.

- Sample #274 – Flag 4 - Outlier in comparison with CTD; bottom bottle; no water this salty found anywhere in profile.

- Sample #275 – Flag 5 – Extreme outlier in comparison with CTD; no salinity that high in whole profile. Pad value entered.

- Sample #319 – Flag 3 – Major outlier in comparison with CTD; bottom bottle; fairly noisy CTD data and bottle value found within 10m.

- Sample #331 – Flag 3 - Major outlier in comparison with CTD; bottom bottle; fairly noisy CTD data and bottle value found about 11m above.

- Sample #344 – Flag 3 - Outlier in comparison with CTD; bottom bottle; fairly noisy CTD data and bottle value found within 5m.

When only bottles that are within 0.005 of the CTD and for which the standard deviation in the CTD salinity is <0.0008 are included (92 bottles), the primary CTD salinity is low by an average of 0.0023 and the secondary is low by 0.001. The standard deviation in both fits is ~0.002. There was a post-cruise calibration of T and C sensors that showed that the primary salinity was low by ~0.0013 and the secondary was low by <0.001 in January 2011. So the fits based on the 92 bottles (out of 339) are within 0.001 of what is expected from the post-cruise calibrations.

Note that bottles that read lower than the CTD salinity but were not flagged may well have been affected by the long wait for analysis; there were just other likely explanations for high bottle values.

For full details for the COMPARE run see file 2010-05-sal-comp1.xls.

Fluorescence

COMPARE was run using the SeaPoint Fluorescence and the Extracted Chlorophyll from bottles. As usual the fit is poor for very low chlorophyll and there were many very low values in these data. The ratio of CTD fluorescence to extracted chlorophyll varied from 0.3 to 40, when CHL<0.1 are excluded. When points are further excluded based on the standard deviation in the CTD fluorescence the ratio comes down and is noisy but close to 1 for most cases when the standard deviation is <15% of the fluorescence value. There are some very high gradients in many of these casts, with maxima at 25 to 40m, so we would expect the fluorometer to read higher than the chlorophyll for the majority of bottles due to the 1.5m displacement between the bottle and CTD, but incomplete flushing of Niskin bottles would have the opposite effect.

For more detail see file 2010-05-fl-chl-comp1.xls.

The flag changes mentioned above were applied to file 2010-05-samples.csv which was reconverted to individual cast files with extension MRG1. Once again those files were put through CLEAN and SORT (on bottle #) and then merged with the SAMREM files with output MRG. These files were put through CLEAN to remove SeaBird headers and comments.

##### Conversion of Full Files from Raw Data

All files were converted using 2010-05-ctd1-con and 2010-05-ctd2.con, as appropriate.

A few casts were examined and all expected channels are present.

The descent rate is extremely noisy for many casts with many complete reversals of direction during the descent and obvious shed wake corruption.

The two temperature channels are fairly close during the downcasts though both channels have occasional spikes. During the upcasts, traces differ more and there are odd excursions that are often seen in upcasts. This is likely something to do with how the CTD is mounted. The conductivity channels are similar to temperature.

Altimetry looks useful when the CTD got near the bottom, dissolved oxygen, fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is noisy on the upcast and there were many stops for bottles. There is always a slight shift in values between upcast and downcast so that matching features works better than just looking at the vertical offset of the trace. An advance of 5s looks reasonable.

ALIGNCTD was used to advance the DO Voltage by 5s relative to the pressure.

##### CELLTM

The upcast data are extremely noisy and there were many stops for bottles, so the usual tests for CELLTM settings were a little hard to interpret. For the deeper casts in places where the upcast data was fairly smooth, all settings tested improved the data somewhat, with a choice of (α = 0.0245, β=9.5) looking as good as any other choice; these values are default values. For the shallower casts all settings appear to make things worse, but an examination of profiles shows that the default setting has only a minor effect on the downcasts, but affects the upcasts much more. So CELLTM will be applied in the usual way.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2010-05-0002 | 500 950 | ~0 +0.0004 | +0.0001 +0.0001 | +0.0008 +0.0008 | High, Noisy |
| 2010-05-0002 | 500 | ~0 | +0.0001 | +0.0013 | High, Noisy |
| 2010-05-0002 | 500 950 | +0.0001 +0.0004 | +0.0001 +0.0001 | +0.0009 +0.0008 | High, Noisy |

All the differences are small and there is no obvious pressure or temporal dependence, though none of the later casts are deep enough to be sure about the variation with time. The differences between salinity channels are consistent with the salinity comparison, though slightly smaller than the difference between post-cruise calibrations. That may mean there was further drift in the primary after this cruise.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are negative pressures and negative values for many channels.

Speeds look reasonable.

Surface check was run and shows an average surface pressure for the cruise was 0.19db but the associated salinity values were mostly close to 0 so this looks reasonable. Examination of a few casts suggests that the pressure is within 0.2db. During downcasts the CTD often sat at about 0.2db with salinity values showing it is in water. But the pumps were not on and the CTD could have been dipping in and out of the water. Near the beginning of cast #6 the pressure was about 0.2db with salinity values that look like the CTD was in water or right at the surface. Then pressure was reduced briefly to <0 and the salinity values fell sharply. This would indicate that the pressure setting is reliable. At the end of cast #6 the acquisition system was left running while the CTD was on deck and the pressure reads ~0.1db which looks reasonable. The factory calibrations for pressure look appropriate, as we would expect so close to recalibration.

The cross-reference check was compared with the log book and the only problems found were in station names for casts #6, 11, 12, 24, 36, 37, 38, 42 and 43. There were fixed in the CLN and SAMAVG files.

The altimeter readings and bottom depths from the headers of the CLN files were exported to a spreadsheet. A sample of casts was checked and no errors were found in the altimetry. There were many discrepancies between bottom depths and logs and no entry for some casts. Each case was checked by adding the maximum sampling depth (adjusted from pressure) plus altimetry. In most cases the log agreed with that, so the header entry was changed to match the log. These were fixed in the CLN files and SAM files. (SAMAVG, MERGE and CLEAN steps were rerun after these corrections.)

The MRGCLN2 files were recreated after these corrections.

The cruise track was plotted and added to the end of this report. No problems were found.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence trace look like that of the temperature trace. This task was complicated by:

- Noisy upcast temperature data so the vertical offset is confused.

- Many stops for bottles on upcasts so that fluorescence gets opportunities to “catch up”

- Noisy descent and ascent rates resulting in very confused traces.

- Very low fluorescence values so there is little variability to aid estimate of vertical offset.

The offset tests were not useful, but applying the advance of +24 records that has been found best for many data sets did lead to more convincing downcast fluorescence with sharp increases just below the base of large temperature gradients.

SHIFT was run on all casts to advance the SeaPoint fluorescence channel by +24 records.

Conductivity

Tests were run on the conductivity channels. The results were plotted for 3 casts using a variety of shifts and the best results were with an advance of -0.4 records for the primary and -0.6 records for the secondary.

SHIFT was run on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier.

CDOM

The CDOM signal is so noisy that any vertical offset cannot be detected. Since the sensor was not pumped it is not expected to need alignment.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in November 2009 and there were post-cruise calibrations in January 2011. The post-cruise calibration of temperature and conductivity implies drift of about -0.001 in the primary salinity and -0.0005 in the secondary salinity.

2. Dissolved Oxygen

The DO sensor was recalibrated in Nov. 2009 and there was a post-cruise calibration in December 2010. This cruise was the only one using this sensor between the two calibrations. Based on the study described in section 2 the drift would be <5%.

3. Pressure

The sensor was recalibrated in December 2009 and this cruise was the first time it was used. It was used for 4 cruises during 2011. The factory settings were used for 3 of them. For the last one in June 2011 a higher offset was used at sea, but those data have not yet been processed, so a final value has not yet been established.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Those were only available for the first 2 casts but larger-scale climatology was available for 2 other casts. There were no excursions from those ranges though salinity at the bottom of the mixed layer was at the minimum for station NP-1.

Repeat Casts –

There were no repeat casts.

##### DETAILED EDITING

The bottle comparison shows that the secondary salinity is closest to bottles, but it is much noisier than the primary. A revisit was made to the SHIFT stage to see if there was a better choice for the secondary conductivity, but -0.6 records does look best. There are patches where the primary T-S data look spiky, but overall the choice of primary temperature and salinity look best, so they were chosen for editing. Even the primary is very noisy, having both random spikes and corruption from shed wakes. The spikes may be associated with high decelerations of the CTD.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. Most of the deeper casts had extremely noisy descent rates so that despite a high average rate, there are signs of shed wakes corrupting data. Heavy editing was required for those casts.

The most difficult editing for the shallower casts was in the near-surface high-gradient zone, with temperature drops on the order of 7C° in 4m with little change in salinity. There are many unstable features in that zone where small reversals in temperature lead to larger reversals in salinity. These unstable features occur while the descent rate looks reasonable steady. Tests were run to see if CELLTM had caused these instabilities but skipping that step did not improve the data. Salinity was cleaned in those zones to remove spikes that look instrumental in nature; while some unstable features remained they are small and most will be removed by bin-averaging.

All EDU files were copied to EDT.

All files were put through Bin Average using 0.5db bins and producing standard deviations. Plots were made to see if further editing was required. Many unstable features remain but most are fine-scale. A few were re-examined in CTDEDIT to see if further editing would be useful but the features do not have obvious instrumental source, so were left in case they are real.

##### Initial Recalibration

The pressure looks ok.

We have insufficient information on which to base a recalibration of dissolved oxygen.

The primary salinity was found to be low by 0.0023 in COMPARE, but there were many outliers in that comparison probably due to either evaporation or adsorption of samples or poor flushing of bottles which both lead to high bottle values. A post-cruise calibration of temperature and conductivity sensors implies the primary salinity was low by 0.0013 when it reached the factory. Since the drift is small and some of it may have occurred on the shelf after the cruise, no recalibration was applied.

##### Fluorometer Processing

There were no off-scale fluorescence data.

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 0.5	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Some unstable features remain as described in section 16, chiefly in high-gradient zones in shallow water. Very heavy editing would be needed to remove them and often the unstable feature is due to a very small reversal in salinity. No further editing was applied.

Bin Average was run a second time to include standard deviations for all channels. Those files have extensions AVGSTD.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on the AVG files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate,Voltage:0, Voltage:1, Voltage:3, Voltage:4, Voltage:6, Uploy0 and Flag

REMOVE was run on the AVGSTD files to remove (Output: *.REMSTD):

Scan_Number, Scan_Number:STD, Status:Pump, Status:Pump:STD, Flag and Flag:STD

A second SBE DO channel (with umol/kg units) was added to both the REM and REMSTD.

REORDER was run to get the two DO channels together.

An error in the PAR sensor parameters was found and fixed using CALIBRATE and file 2010-05-recal1.ccf.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:SeaPoint, Fluorescence:WetLabs:CDOM, PAR and

PAR:Reference data are nominal and unedited except that some records were removed

in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, so values are nominal and

are reported with with 1 less significant figure than usual.

A comparison of pre-cruise and post-cruise calibrations suggests the reported DO could

be low by up to 5%, but the surface saturation of DO does look reasonable.

A post-cruise calibration showed drift in the Salinity:T0:S0 channel to be no more

than 0.0013;no recalibration was applied.

At the end of processing, an error was found in the PAR calibration parameters; this

was corrected.

For further processing details see the processing report 2010-05_Processing_Report.doc                                                                                                                                       .

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

A second set of files was prepared for the use of Sarah Zimmermann with standard deviations and original voltages included, as well all T, C and S channels. These have extension CTDSTD.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 97% to 124% with the higher values all near shore, and most values between 102% and 106%. The results for the early casts in the open Pacific Ocean are in the expected range; no judgment can be made on the other casts in the absence of DO samples as variability is likely to be high.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on the MRGSORT files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate, and Flag,

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together. There are no DO samples, so the mass units derivation will not be applied to bottle data.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and notes about the CTD data.

A header check was run on the final files and a few problems were found and corrected. Then bottle data were exported to file 2010-05-bottles-final.csv.

Standards check was run on all files and no problems were found except that the formats for dissolved oxygen and nutrients were not those normally used, but this was deliberate as the nutrient formats were changed years after this analysis was done to reflect improved results and the dissolved oxygen data were considered of lower quality than usual.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found. The plot was added to the end of this report.

# Particulars

# 4. Missed bottle at 35, 2 bottles at 25 (14 and 15)

# 6. Seasave left running after cast – long file

# 8. Pumps ran in air for 40s

# 11. Lots of jelly seen at surface

# 14. Fluorescence saturates @40m

# 15. Fluorometer gain cable changed from 10X to 3X

# 36. Whitecaps

# Institute of Ocean Sciences Cruise Summary

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 25Nov2009 | Factory | 6Jan 2011 |  |
| Conductivity | 3579 | 25Nov2009 | Factory | 8Jan2011 | Factory |
| Secondary Temp. | 5073 | 25Nov2009 | Factory | 6Jan 2011 |  |
| Secondary Cond. | 3581 | 25Nov2009 | Factory | 8Jan2011 | Factory |
| Transmissometer | CST662DR | 26May2008 |  |  |  |
| SBE 43 DO sensor | 1117 | 25Nov2009 | Factory | 29Dec2010 | Factory |
| SeaPoint Fluorometer | 2745 |  |  |  |  |
| PAR | 70123 | 30Aug2009 |  |  |  |
| Wetlabs CDOM | 1076 | 6Nov2006 |  |  |  |
| ISUS | 072 |  |  |  |  |
| Surface PAR | 20281 | 1Apr2007 |  |  |  |
| Pressure Sensor | 941 | 07Dec2009 | Factory |  |  |
| Altimeter | 1076 |  |  |  |  |