## Notes and Updates
Date | Name | Note
2011-01-31 | SZ | Set up spreadsheet for Doug Sieberg and Mary Steel who will be analysing salinity samples brought back from sea
Sampling protocol:
Please run a standard water at the start of each day (at least for the first 3 days) to confirm instrument stability.
Please run a deepwater reference between each cage of samples. Note there are two different batches of deep water in the salt lab.
Please see numbering/naming scheme in Autosal Data by Time used during at-sea analysis.
2011-02-15 | DG | Analysing Equipment
Guildline Autosal Model 8400B, serial # 69572
bath temperature of 24OC
accuracy of < +/- 0.0001 conductivity ratio
resolution <0.00001 conductivity ratio
The calibration checks were done using IAPSO Standard Seawater
K15 Value = 0.99997 Batch# P151 "Use by Date": May 20, 2012
The Salinometer remained stable and within the Guildline spec's of 0.002 PSU
Data Interpretation Equipment
Ocean Scientific International Salinometer Data Logger software was used to log the conductivity readings from the Guildline Autosal calculating salinity value.
Lab and Sample Conditions
Analysed in a temperature controlled lab with an ambient temperature changing between 20.4oc-23.6oc overall with a 1.0oc maximum change during analysing session.
Sample temperatures ranged from 20.3oc to 24.1oc overall with a 1.9oc maximum change during analysing session
The sample bottles were inverted and mixed prior to analysis.
all samples had minimum 3 flushes and minimum 2 reading
each reading is the average of 5 readings taken in a 5 second period
Analyst Observation
|  | 1 | The Autosal was stable making the analysing process very smooth
Autosal accuracy for electronics is < 0.002 PSU equivalents
the electronics of the Autosal 8400B drifted within the manufactures spec's during analysing
|  | 2 | Calibration checks were done at the start of each day before analysing started.
correction values were within manufactures spec's for all but one day.
Standardizing was necessary on February 9, 2010
the Rs was adjusted from 5.56 to 5.54 and the Standby from 24+5984 to 24+5986
|  | 3 | A Reference Samples was ran every 24 samples(one cage) to check for drift in the electronics -all looks good -see DeepWaterCheck Records sheet
|  | 4 | All samples were stored in proper salinity bottles with separate inserts and caps which helps in minimizing evaporation and leakage
Analyst Suggestions
|  | 1 | continue to use proper salinity bottles with separate insert and cap mechanisms
|  | 2 | arctic waters are VERY cold -after sampling when the sample warms to ambient temp check that inserts have not popped up
|  | 3 | assure that transport of samples are protected from freezing
2011-03-02 | SZ | Minor formatting changes
made Autosal-Data for Merging page:
sort by sample number and make column showing sample number diff and check sample #s are correct - found a couple of typos
all 'd's are correct and show those run out of time order, so would like to use these as dups
sort by 'd', then sort by sample number which puts 'd' as first occurance throughout
Check-Water: CB-3 averages to 34.956 and CB-6 averages to 34.957
Comparison of 23 samples where primary run at sea and duplicate run on shore shows at sea > IOS by 0.001
Average difference btw the primary and duplicates run on shore is also 0.001, so live with error?
Looks like secondary CTD salt drifts -0.001 based on bottom water atarting with cast 57. It will be good to compare repeat stations (start and end of cruise).
Primary CTD has clearly shifted..
Bottom water shows Casts 54 and 62 are -0.003 to -0.004 below the expected value. Casts 31, 34 57 and 66 are -0.001 below.
Transition to IOS analyses happened with cast 49.
2011-03-29 | SZ | It looks like there is a 0.001 offset between primaries and duplicates based on when the samples were run at IOS. All duplicates were run on the Feb 9th.
Primaries run before the 9th are 0.0016 higher…and those run after are the same. This is not consistant with the Deep Water Check data that show CB-3,
run before the 9th, is 0.001 lower than expected (34.956) and CB-25 run after the 9th is as expected (34.957). In other words, before the 9th,
the duplicates are reading high but the DWC is reading low.
Just need to accept +/- 0.001 error…

## flags
IOS QA/QC code from Joe - to be paired with final data | In other words…
Flag = definition
a = Sample for this measurement was drawn from water bottle but not analyzed (not normally used)
b = Acceptable measurement (not normally used) | although data look odd they are real
c = Questionable measurement (no problem observed in sampling or analysis, but value is not trusted, nonetheless; includes outlyers) | keep - but questionable
d = Bad measurement (known problem with sampling or analysis, but not serious enough to completely discard the value)
e = Not reported (lost sample; unredeemably bad measurement) | don't use data point (left in working sheet but not transferred to final data)
f = Mean of replicate measurements (summary of replicates and table of replicate data should be included in the metadata file)
g = Manual chromatographic peak measurement
h = Irregular digital chromatographic peak integration
i = Sample not drawn for this measurement from this bottle (not normally used)
Other QA/QC
q = data point changed but requires another round of plotting (check again!)
qc = questionable measurment, checked by analyst & looks good from analysts view
j = missing sample

## Autosal-Data by Time
Salinometer ID | Station | Sample ID (bottle label) | Duplicates | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
Sample ID | Station | Bottle Label |  | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 |  | K15 |  | 2011-02-02 10:19 | 24C | 0.99992 | 2.2397e-06 | 4.5896e-05 | 0.99997
LSSL2010#1 | CB-DW | 1262 |  | 2011-02-02 11:30 | 24C | 0.99858 | 3.1311e-06 | 4.6015e-05 | 0.99862 | 34.9458 | 0.00012618
LSSL2010#2 | CB-DW | 1263 |  | 2011-02-02 11:32 | 24C | 0.9983 | 3.8749e-06 | 4.6015e-05 | 0.99835 | 34.935 | 0.00014763
LSSL2010#3 | CB-DW | 1264 |  | 2011-02-02 11:35 | 24C | 0.99788 | 6.0358e-06 | 4.6015e-05 | 0.99793 | 34.9186 | 0.00023836
LSSL2010#4 | CB-DW | 1265 |  | 2011-02-02 11:38 | 24C | 0.99705 | 3.4659e-06 | 4.6074e-05 | 0.9971 | 34.8858 | 0.00013799
LSSL2010#5 | CB-DW | 1266 |  | 2011-02-02 11:40 | 24C | 0.99664 | 5.0211e-06 | 4.5896e-05 | 0.99668 | 34.8695 | 0.00019505
LSSL2010#6 | CB-DW | 1267 |  | 2011-02-02 11:43 | 24C | 0.9963 | 2.7423e-06 | 4.6015e-05 | 0.99634 | 34.856 | 0.00010656
LSSL2010#7 | CB-DW | 1268 |  | 2011-02-02 11:46 | 24C | 0.9961 | 2.4909e-06 | 4.6015e-05 | 0.99615 | 34.8485 | 9.9734e-05
LSSL2010#8 | CB-DW | 1269 |  | 2011-02-02 11:48 | 24C | 0.99553 | 6.2135e-06 | 4.6015e-05 | 0.99558 | 34.8259 | 0.00024024
LSSL2010#9 | CB-DW | 1270 |  | 2011-02-02 11:51 | 24C | 0.99564 | 7.1714e-06 | 4.5955e-05 | 0.99569 | 34.8302 | 0.00028287
LSSL2010#10 | CB-DW | 1271 |  | 2011-02-02 11:53 | 24C | 0.99532 | 5.5828e-06 | 4.6015e-05 | 0.99537 | 34.8178 | 0.00021731
LSSL2010#11 | CB-DW | 1272 |  | 2011-02-02 11:56 | 24C | 0.99442 | 4.4786e-06 | 4.6015e-05 | 0.99446 | 34.7822 | 0.00017744
LSSL2010#12 | CB-DW | 1273 |  | 2011-02-02 11:59 | 24C | 0.99201 | 2.7425e-06 | 4.6015e-05 | 0.99206 | 34.6876 | 0.00010447
LSSL2010#13 | CB-DW | 1274 |  | 2011-02-02 12:02 | 24C | 0.98578 | 3.3587e-06 | 4.6015e-05 | 0.98583 | 34.4429 | 0.00013108
LSSL2010#14 | CB-DW | 1275 |  | 2011-02-02 12:05 | 24C | 0.97479 | 4.9007e-06 | 4.6074e-05 | 0.97483 | 34.0117 | 0.00019277
LSSL2010#15 | CB-DW | 1276 |  | 2011-02-02 12:07 | 24C | 0.96762 | 3.6235e-06 | 4.6015e-05 | 0.96766 | 33.7309 | 0.00014164
LSSL2010#16 | CB-DW | 1277 |  | 2011-02-02 12:10 | 24C | 0.95007 | 4.8959e-06 | 4.5955e-05 | 0.95012 | 33.0455 | 0.0001943
LSSL2010#17 | CB-DW | 1278 |  | 2011-02-02 12:13 | 24C | 0.9466 | 4.7277e-06 | 4.6015e-05 | 0.94665 | 32.9101 | 0.00018258
LSSL2010#18 | CB-DW | 1279 |  | 2011-02-02 12:15 | 24C | 0.93913 | 3.4658e-06 | 4.6015e-05 | 0.93918 | 32.6191 | 0.00013386
LSSL2010#19 | CB-DW | 1280 |  | 2011-02-02 12:18 | 24C | 0.92996 | 4.3843e-06 | 4.6015e-05 | 0.93 | 32.2623 | 0.0001726
LSSL2010#20 | CB-DW | 1281 |  | 2011-02-02 12:21 | 24C | 0.91984 | 6.1978e-06 | 4.5955e-05 | 0.91989 | 31.8695 | 0.00024055
LSSL2010#21 | CB-DW | 1282 |  | 2011-02-02 12:24 | 24C | 0.90726 | 4.1856e-06 | 4.6015e-05 | 0.9073 | 31.3818 | 0.0001651
LSSL2010#22 | CB-DW | 1283 |  | 2011-02-02 12:26 | 24C | 0.83034 | 3.3834e-06 | 4.6015e-05 | 0.83039 | 28.4264 | 0.00012852
LSSL2010#23 | CB-DW | 1284 |  | 2011-02-02 12:29 | 24C | 0.77006 | 4.9815e-06 | 4.6015e-05 | 0.77011 | 26.1399 | 0.00018882
LSSL2010#24 | CB-DW | 1285 |  | 2011-02-02 12:32 | 24C | 0.76375 | 2.7301e-06 | 4.5896e-05 | 0.7638 | 25.9022 | 0.00010253
LSSL2010#25 | reference sample | CB-3 |  | 2011-02-02 12:35 | 24C | 0.99882 | 6.5389e-06 | 4.5955e-05 | 0.99886 | 34.9553 | 0.0002569 | 614-615
LSSL2010#26 | reference sample | CB-3 |  | 2011-02-02 12:37 | 24C | 0.99885 | 4.3332e-06 | 4.5955e-05 | 0.9989 | 34.9567 | 0.00017044 | 614-615
LSSL2010#27 | CB-6 | 1334 |  | 2011-02-02 14:03 | 24C | 0.99886 | 3.264e-06 | 4.5955e-05 | 0.99891 | 34.9571 | 0.00012912
LSSL2010#28 | CB-6 | 1335 |  | 2011-02-02 14:06 | 24C | 0.99886 | 5.6675e-06 | 4.5955e-05 | 0.9989 | 34.9568 | 0.00022572
LSSL2010#29 | CB-6 | 1336 |  | 2011-02-02 14:09 | 24C | 0.99927 | 2.7426e-06 | 4.6015e-05 | 0.99931 | 34.973 | 0.00010456
LSSL2010#30 | CB-6 | 1337 |  | 2011-02-02 14:14 | 24C | 0.99853 | 3.8402e-06 | 4.6015e-05 | 0.99857 | 34.9439 | 0.00014859
LSSL2010#31 | CB-6 | 1338 |  | 2011-02-02 14:17 | 24C | 0.99771 | 1.3713e-06 | 4.6015e-05 | 0.99776 | 34.9118 | 5.4351e-05
LSSL2010#32 | CB-6 | 1339 |  | 2011-02-02 14:20 | 24C | 0.99689 | 3.8748e-06 | 4.6015e-05 | 0.99694 | 34.8794 | 0.00015047
LSSL2010#33 | CB-6 | 1340 |  | 2011-02-02 14:22 | 24C | 0.99647 | 8.4334e-06 | 4.6015e-05 | 0.99652 | 34.8631 | 0.00033289
LSSL2010#34 | CB-6 | 1341 |  | 2011-02-02 14:24 | 24C | 0.99615 | 5.5979e-06 | 4.6015e-05 | 0.99619 | 34.8502 | 0.00022175
LSSL2010#35 | CB-6 | 1342 |  | 2011-02-02 14:27 | 24C | 0.99577 | 4.5027e-06 | 4.6015e-05 | 0.99581 | 34.8353 | 0.00017851
LSSL2010#36 | CB-6 | 1343 |  | 2011-02-02 14:30 | 24C | 0.99565 | 1.3713e-06 | 4.6015e-05 | 0.9957 | 34.8308 | 5.2235e-05
LSSL2010#37 | CB-6 | 1344 |  | 2011-02-02 14:33 | 24C | 0.99469 | 5.4847e-06 | 4.6074e-05 | 0.99473 | 34.7928 | 0.00021364
LSSL2010#38 | CB-6 | 1345 |  | 2011-02-02 14:35 | 24C | 0.99286 | 7.7817e-06 | 4.6015e-05 | 0.99291 | 34.7211 | 0.00030835
LSSL2010#39 | CB-6 | 1346 |  | 2011-02-02 14:38 | 24C | 0.98572 | 4.5125e-06 | 4.6015e-05 | 0.98577 | 34.4405 | 0.00017485
LSSL2010#40 | CB-6 | 1347 |  | 2011-02-02 14:40 | 24C | 0.96722 | 2.1278e-06 | 4.6015e-05 | 0.96726 | 33.7153 | 8.1169e-05
LSSL2010#41 | CB-6 | 1348 |  | 2011-02-02 14:42 | 24C | 0.96711 | 6.6064e-06 | 4.6074e-05 | 0.96716 | 33.7111 | 0.00025526
LSSL2010#42 | CB-6 | 1349 |  | 2011-02-02 14:45 | 24C | 0.95239 | 3.3587e-06 | 4.6015e-05 | 0.95243 | 33.1358 | 0.00013383
LSSL2010#43 | CB-6 | 1350 |  | 2011-02-02 14:48 | 24C | 0.94684 | 8.6176e-06 | 4.6015e-05 | 0.94689 | 32.9196 | 0.00033553
LSSL2010#44 | CB-6 | 1351 |  | 2011-02-02 14:52 | 24C | 0.9394 | 6.2131e-06 | 4.5896e-05 | 0.93944 | 32.6294 | 0.00024166
LSSL2010#45 | CB-6 | 1352 |  | 2011-02-02 14:55 | 24C | 0.93189 | 3.0662e-06 | 4.6015e-05 | 0.93194 | 32.3375 | 0.00011836
LSSL2010#46 | CB-6 | 1353 |  | 2011-02-02 14:57 | 24C | 0.92048 | 3.862e-06 | 4.5896e-05 | 0.92052 | 31.8941 | 0.00015135
LSSL2010#47 | CB-6 | 1357 |  | 2011-02-02 15:01 | 24C | 0.73233 | 2.8904e-06 | 4.5896e-05 | 0.73238 | 24.7226 | 0.00010975
LSSL2010#48 | reference sample | CB-3 | 1 | 2011-02-02 15:13 | 24C | 0.99894 | 2.7425e-06 | 4.6015e-05 | 0.99899 | 34.9601 | 0.00010659 | 614-615
LSSL2010#49 | CB-21 | 1358 |  | 2011-02-02 15:17 | 24C | 0.99705 | 3.3588e-06 | 4.5955e-05 | 0.9971 | 34.8858 | 0.00013233
LSSL2010#50 | CB-21 | 1359 |  | 2011-02-02 15:20 | 24C | 0.99701 | 4.7299e-06 | 4.6015e-05 | 0.99705 | 34.8841 | 0.00018569
LSSL2010#51 | CB-21 | 1360 |  | 2011-02-02 15:23 | 24C | 0.99694 | 6.1012e-06 | 4.6074e-05 | 0.99698 | 34.8813 | 0.00023971
LSSL2010#52 | CB-21 | 1361 |  | 2011-02-02 15:25 | 24C | 0.997 | 6.3002e-06 | 4.6074e-05 | 0.99705 | 34.8839 | 0.00024981
LSSL2010#53 | CB-21 | 1362 |  | 2011-02-02 15:27 | 24C | 0.99656 | 3.8648e-06 | 4.6015e-05 | 0.99661 | 34.8666 | 0.00015243
LSSL2010#54 | CB-21 | 1363 |  | 2011-02-02 15:30 | 24C | 0.99652 | 4.5125e-06 | 4.6015e-05 | 0.99657 | 34.8651 | 0.00017433
LSSL2010#55 | CB-21 | 1364 |  | 2011-02-02 15:32 | 24C | 0.99657 | 4.1136e-06 | 4.6015e-05 | 0.99662 | 34.8668 | 0.00016089
LSSL2010#56 | CB-21 | 1365 |  | 2011-02-02 15:35 | 24C | 0.99657 | 7.0607e-06 | 4.5955e-05 | 0.99662 | 34.8669 | 0.0002774
LSSL2010#57 | CB-21 | 1366 |  | 2011-02-02 15:37 | 24C | 0.99612 | 5.5575e-06 | 4.6015e-05 | 0.99617 | 34.8492 | 0.00022085
LSSL2010#58 | CB-21 | 1367 |  | 2011-02-02 15:40 | 24C | 0.99609 | 2.4908e-06 | 4.6074e-05 | 0.99614 | 34.8479 | 9.6976e-05
LSSL2010#59 | CB-21 | 1368 |  | 2011-02-02 15:44 | 24C | 0.99619 | 3.8619e-06 | 4.5955e-05 | 0.99623 | 34.8517 | 0.00015179
LSSL2010#60 | CB-21 | 1369 |  | 2011-02-02 15:46 | 24C | 0.99621 | 3.6085e-06 | 4.6015e-05 | 0.99625 | 34.8526 | 0.00014196
LSSL2010#61 | CB-21 | 1370 |  | 2011-02-02 15:49 | 24C | 0.99528 | 3.8323e-06 | 4.6015e-05 | 0.99533 | 34.8162 | 0.00014921
LSSL2010#62 | CB-21 | 1371 |  | 2011-02-02 15:51 | 24C | 0.99529 | 5.8623e-06 | 4.6015e-05 | 0.99534 | 34.8166 | 0.00023023
LSSL2010#63 | CB-21 | 1372 |  | 2011-02-02 15:53 | 24C | 0.99573 | 3.9989e-06 | 4.6015e-05 | 0.99577 | 34.8336 | 0.00015782
LSSL2010#64 | CB-21 | 1373 |  | 2011-02-02 15:57 | 24C | 0.99542 | 4.3587e-06 | 4.5955e-05 | 0.99546 | 34.8215 | 0.00016984
LSSL2010#65 | CB-21 | 1374 |  | 2011-02-02 16:01 | 24C | 0.98885 | 8.3446e-06 | 4.5896e-05 | 0.9889 | 34.5635 | 0.00032792
LSSL2010#66 | CB-21 | 1375 |  | 2011-02-02 16:05 | 24C | 0.98885 | 5.4849e-06 | 4.6134e-05 | 0.9889 | 34.5635 | 0.00021388
LSSL2010#67 | CB-21 | 1376 |  | 2011-02-02 16:09 | 24C | 0.98889 | 7.3776e-06 | 4.5955e-05 | 0.98894 | 34.565 | 0.00029054
LSSL2010#68 | CB-21 | 1377 |  | 2011-02-02 16:11 | 24C | 0.98891 | 6.8046e-06 | 4.5896e-05 | 0.98895 | 34.5656 | 0.00027202
LSSL2010#69 | CB-21 | 1378 |  | 2011-02-02 16:14 | 24C | 0.74173 | 2.4776e-06 | 4.6015e-05 | 0.74177 | 25.0744 | 9.2326e-05
LSSL2010#70 | CB-21 | 1379 |  | 2011-02-02 16:17 | 24C | 0.74127 | 6.7174e-06 | 4.6015e-05 | 0.74131 | 25.0573 | 0.00024994
LSSL2010#71 | CB-21 | 1380 |  | 2011-02-02 16:19 | 24C | 0.74123 | 3.8619e-06 | 4.6015e-05 | 0.74128 | 25.0559 | 0.00014712
LSSL2010#72 | CB-21 | 1381 |  | 2011-02-02 16:22 | 24C | 0.74119 | 4.9029e-06 | 4.6015e-05 | 0.74124 | 25.0545 | 0.00018162
LSSL2010#73 | reference sample | CB-3 | 2 | 2011-02-02 16:26 | 24C | 0.99884 | 5.084e-06 | 4.5955e-05 | 0.99889 | 34.9563 | 0.00020024 | 614-615
Calibration#2 |  | K15 |  | 2011-02-03 09:22 | 24C | 0.99995 | 2.2391e-06 | 2.408e-05 | 0.99997
LSSL2010#74 | CB-15 | 1118 |  | 2011-02-03 09:32 | 24C | 0.9988 | 5.871e-06 | 2.408e-05 | 0.99882 | 34.9536 | 0.00023096
LSSL2010#75 | CB-15 | 1119 |  | 2011-02-03 09:35 | 24C | 0.99863 | 3.4658e-06 | 2.4021e-05 | 0.99866 | 34.9471 | 0.00013341
LSSL2010#76 | CB-15 | 1120 |  | 2011-02-03 09:37 | 24C | 0.99875 | 4.3223e-06 | 2.3961e-05 | 0.99878 | 34.9519 | 0.00017063
LSSL2010#77 | CB-15 | 1121 |  | 2011-02-03 09:39 | 24C | 0.99847 | 5.9206e-06 | 2.408e-05 | 0.99849 | 34.9406 | 0.00023374
LSSL2010#78 | CB-15 | 1122 |  | 2011-02-03 09:42 | 24C | 0.9978 | 4.4373e-06 | 2.408e-05 | 0.99782 | 34.9143 | 0.00017504
LSSL2010#79 | CB-15 | 1123 |  | 2011-02-03 09:45 | 24C | 0.99695 | 3.5949e-06 | 2.3961e-05 | 0.99697 | 34.8808 | 0.00014346
LSSL2010#80 | CB-15 | 1124 |  | 2011-02-03 09:47 | 24C | 0.99664 | 3.6104e-06 | 2.3961e-05 | 0.99666 | 34.8686 | 0.00014379
LSSL2010#81 | CB-15 | 1125 |  | 2011-02-03 09:49 | 24C | 0.99629 | 6.5199e-06 | 2.3961e-05 | 0.99632 | 34.855 | 0.00025501
LSSL2010#82 | CB-15 | 1126 |  | 2011-02-03 09:51 | 24C | 0.99623 | 2.2396e-06 | 2.3961e-05 | 0.99625 | 34.8525 | 8.9681e-05
LSSL2010#83 | CB-15 | 1127 |  | 2011-02-03 09:54 | 24C | 0.99575 | 2.8901e-06 | 2.3901e-05 | 0.99577 | 34.8337 | 0.00011606
LSSL2010#84 | CB-15 | 1128 |  | 2011-02-03 09:56 | 24C | 0.99545 | 2.4908e-06 | 2.3961e-05 | 0.99547 | 34.822 | 9.869e-05
LSSL2010#85 | CB-15 | 1129 |  | 2011-02-03 09:58 | 24C | 0.9943 | 1.3713e-06 | 2.3961e-05 | 0.99433 | 34.7768 | 5.2235e-05
LSSL2010#86 | CB-15 | 1130 |  | 2011-02-03 10:00 | 24C | 0.98943 | 6.6227e-06 | 2.4021e-05 | 0.98946 | 34.5854 | 0.00025908
LSSL2010#87 | CB-15 | 1131 |  | 2011-02-03 10:03 | 24C | 0.9761 | 2.4744e-06 | 2.3961e-05 | 0.97613 | 34.0623 | 9.6652e-05
LSSL2010#88 | CB-15 | 1132 |  | 2011-02-03 10:05 | 24C | 0.97646 | 3.781e-06 | 2.4021e-05 | 0.97648 | 34.0763 | 0.00014929
LSSL2010#89 | CB-15 | 1133 |  | 2011-02-03 10:07 | 24C | 0.962 | 5.8084e-06 | 2.3901e-05 | 0.96203 | 33.5105 | 0.00022729
LSSL2010#90 | CB-15 | 1134 |  | 2011-02-03 10:10 | 24C | 0.95338 | 2.2392e-06 | 2.3961e-05 | 0.95341 | 33.1737 | 8.7015e-05
LSSL2010#91 | CB-15 | 1135 |  | 2011-02-03 10:12 | 24C | 0.94763 | 9.0252e-06 | 2.3961e-05 | 0.94765 | 32.9494 | 0.00035108
LSSL2010#92 | CB-15 | 1136 |  | 2011-02-03 10:15 | 24C | 0.9388 | 7.1241e-06 | 2.4021e-05 | 0.93882 | 32.6053 | 0.00027659
LSSL2010#93 | CB-15 | 1137 |  | 2011-02-03 10:17 | 24C | 0.93319 | 6.8045e-06 | 2.4021e-05 | 0.93321 | 32.3871 | 0.00026278
LSSL2010#94 | CB-15 | 1138 |  | 2011-02-03 10:19 | 24C | 0.92299 | 4.7007e-06 | 2.3901e-05 | 0.92302 | 31.9909 | 0.00018663
LSSL2010#95 | CB-15 | 1139 |  | 2011-02-03 10:22 | 24C | 0.86631 | 6.2687e-06 | 2.408e-05 | 0.86634 | 29.8025 | 0.00024276
LSSL2010#96 | CB-15 | 1140 |  | 2011-02-03 10:25 | 24C | 0.83407 | 3.8489e-06 | 2.3961e-05 | 0.83409 | 28.5676 | 0.00014713
LSSL2010#97 | CB-15 | 1141 |  | 2011-02-03 10:27 | 24C | 0.7848 | 8.9936e-06 | 2.3901e-05 | 0.78482 | 26.6956 | 0.00033962
LSSL2010#98 | reference sample | CB-3 | 3 | 2011-02-03 10:31 | 24C | 0.99884 | 8.9631e-06 | 2.3961e-05 | 0.99886 | 34.9553 | 0.00035389 | 614-615
LSSL2010#99 | CB-13 | 1046 |  | 2011-02-03 11:12 | 24C | 0.99891 | 2.8901e-06 | 2.3961e-05 | 0.99893 | 34.9579 | 0.00011213
LSSL2010#100 | CB-13 | 1047 |  | 2011-02-03 11:14 | 24C | 0.9989 | 5.3052e-06 | 2.4021e-05 | 0.99892 | 34.9575 | 0.00020517
LSSL2010#101 | CB-13 | 1048 |  | 2011-02-03 11:16 | 24C | 0.99891 | 3.8619e-06 | 2.4021e-05 | 0.99894 | 34.9581 | 0.00015532
LSSL2010#102 | CB-13 | 1049 |  | 2011-02-03 11:19 | 24C | 0.9985 | 3.1415e-06 | 2.3961e-05 | 0.99853 | 34.942 | 0.00012313
LSSL2010#103 | CB-13 | 1050 |  | 2011-02-03 11:21 | 24C | 0.99782 | 2.7425e-06 | 2.3961e-05 | 0.99784 | 34.9152 | 0.0001097
LSSL2010#104 | CB-13 | 1051 |  | 2011-02-03 11:23 | 24C | 0.99692 | 4.9037e-06 | 2.4021e-05 | 0.99694 | 34.8796 | 0.00019225
LSSL2010#105 | CB-13 | 1052 |  | 2011-02-03 11:25 | 24C | 0.99658 | 9.3855e-06 | 2.4021e-05 | 0.9966 | 34.8662 | 0.00036876
LSSL2010#106 | CB-13 | 1053 |  | 2011-02-03 11:28 | 24C | 0.99611 | 5.1551e-06 | 2.408e-05 | 0.99614 | 34.8481 | 0.00020366
LSSL2010#107 | CB-13 | 1054 |  | 2011-02-03 11:30 | 24C | 0.99602 | 6.8489e-06 | 2.3961e-05 | 0.99605 | 34.8444 | 0.00027
LSSL2010#108 | CB-13 | 1055 |  | 2011-02-03 11:32 | 24C | 0.99587 | 6.3301e-06 | 2.3961e-05 | 0.9959 | 34.8385 | 0.00024835
LSSL2010#109 | CB-13 | 1056 |  | 2011-02-03 11:35 | 24C | 0.99508 | 7.4722e-06 | 2.4021e-05 | 0.9951 | 34.8072 | 0.00029601
LSSL2010#110 | CB-13 | 1057 |  | 2011-02-03 11:37 | 24C | 0.99387 | 2.7301e-06 | 2.4021e-05 | 0.9939 | 34.7599 | 0.00010708
LSSL2010#111 | CB-13 | 1058 |  | 2011-02-03 11:39 | 24C | 0.98949 | 2.7426e-06 | 2.408e-05 | 0.98952 | 34.5877 | 0.00010241
LSSL2010#112 | CB-13 | 1059 |  | 2011-02-03 11:42 | 24C | 0.98177 | 5.4721e-06 | 2.3961e-05 | 0.98179 | 34.2844 | 0.0002165
LSSL2010#113 | CB-13 | 1060 |  | 2011-02-03 11:44 | 24C | 0.97318 | 3.6104e-06 | 2.3961e-05 | 0.9732 | 33.9478 | 0.00014182
LSSL2010#114 | CB-13 | 1061 |  | 2011-02-03 11:47 | 24C | 0.95881 | 4.6063e-06 | 2.3961e-05 | 0.95883 | 33.3857 | 0.00017979
LSSL2010#115 | CB-13 | 1062 |  | 2011-02-03 11:49 | 24C | 0.95393 | 3.8756e-06 | 2.3901e-05 | 0.95395 | 33.195 | 0.00014951
LSSL2010#116 | CB-13 | 1063 |  | 2011-02-03 11:52 | 24C | 0.94382 | 3.392e-06 | 2.4021e-05 | 0.94385 | 32.801 | 0.00013165
LSSL2010#117 | CB-13 | 1064 |  | 2011-02-03 11:54 | 24C | 0.93725 | 3.8437e-06 | 2.3901e-05 | 0.93728 | 32.5452 | 0.00014723
LSSL2010#118 | CB-13 | 1065 |  | 2011-02-03 11:57 | 24C | 0.93194 | 5.7638e-06 | 2.3961e-05 | 0.93197 | 32.3387 | 0.00022378
LSSL2010#119 | CB-13 | 1066 |  | 2011-02-03 11:59 | 24C | 0.91967 | 6.3446e-06 | 2.408e-05 | 0.91969 | 31.8618 | 0.00024698
LSSL2010#120 | CB-13 | 1067 |  | 2011-02-03 12:01 | 24C | 0.88833 | 7.1397e-06 | 2.3961e-05 | 0.88835 | 30.6498 | 0.00027429
LSSL2010#121 | CB-13 | 1068 |  | 2011-02-03 12:03 | 24C | 0.83546 | 8.0422e-06 | 2.3961e-05 | 0.83548 | 28.6207 | 0.00031011
LSSL2010#122 | CB-13 | 1069 |  | 2011-02-03 12:06 | 24C | 0.79683 | 1.3713e-06 | 2.4021e-05 | 0.79686 | 27.1512 | 5.2235e-05
LSSL2010#123 | reference sample | CB-3 | 4 | 2011-02-03 12:08 | 24C | 0.99886 | 7.2013e-06 | 2.4021e-05 | 0.99888 | 34.9559 | 0.00028346 | 614-615
LSSL2010#124 | PP-6 | 1214 |  | 2011-02-03 12:29 | 24C | 0.99707 | 2.5749e-06 | 2.3961e-05 | 0.9971 | 34.8857 | 9.9229e-05
LSSL2010#125 | PP-6 | 1215 |  | 2011-02-03 12:34 | 24C | 0.99709 | 5.1558e-06 | 2.3961e-05 | 0.99711 | 34.8863 | 0.00020151
LSSL2010#126 | PP-6 | 1216 |  | 2011-02-03 12:37 | 24C | 0.99706 | 5.168e-06 | 2.4021e-05 | 0.99708 | 34.8852 | 0.00019939
LSSL2010#127 | PP-6 | 1217 |  | 2011-02-03 12:39 | 24C | 0.99707 | 1.3712e-06 | 2.3961e-05 | 0.9971 | 34.8857 | 5.4326e-05
LSSL2010#128 | PP-6 | 1218 |  | 2011-02-03 12:42 | 24C | 0.99672 | 6.9617e-06 | 2.4021e-05 | 0.99674 | 34.8719 | 0.00027607
LSSL2010#129 | PP-6 | 1219 |  | 2011-02-03 12:44 | 24C | 0.99653 | 5.5668e-06 | 2.408e-05 | 0.99656 | 34.8646 | 0.00021841
LSSL2010#130 | PP-6 | 1220 |  | 2011-02-03 12:46 | 24C | 0.99679 | 1.0181e-05 | 2.4021e-05 | 0.99681 | 34.8746 | 0.00040166
LSSL2010#131 | PP-6 | 1221 |  | 2011-02-03 12:48 | 24C | 0.9967 | 3.3588e-06 | 2.4021e-05 | 0.99672 | 34.871 | 0.00012971
LSSL2010#132 | PP-6 | 1222 |  | 2011-02-03 12:51 | 24C | 0.99627 | 2.2394e-06 | 2.408e-05 | 0.99629 | 34.854 | 8.7024e-05
LSSL2010#133 | PP-6 | 1223 |  | 2011-02-03 12:53 | 24C | 0.99637 | 4.1735e-06 | 2.408e-05 | 0.9964 | 34.8582 | 0.00016273
LSSL2010#134 | PP-6 | 1224 |  | 2011-02-03 12:56 | 24C | 0.99637 | 5.2931e-06 | 2.3961e-05 | 0.99639 | 34.8582 | 0.00020805
LSSL2010#135 | PP-6 | 1225 |  | 2011-02-03 13:01 | 24C | 0.99637 | 5.2035e-06 | 2.3961e-05 | 0.9964 | 34.8582 | 0.0002015
LSSL2010#136 | PP-6 | 1226 |  | 2011-02-03 13:03 | 24C | 0.99565 | 5.5668e-06 | 2.408e-05 | 0.99567 | 34.8298 | 0.00021849
LSSL2010#137 | PP-6 | 1227 |  | 2011-02-03 13:05 | 24C | 0.99573 | 7.3523e-06 | 2.4021e-05 | 0.99575 | 34.8328 | 0.00028976
LSSL2010#138 | PP-6 | 1228 |  | 2011-02-03 13:08 | 24C | 0.9958 | 4.9029e-06 | 2.3961e-05 | 0.99583 | 34.8358 | 0.00019366
LSSL2010#139 | PP-6 | 1229 |  | 2011-02-03 13:10 | 24C | 0.99565 | 4.9162e-06 | 2.4021e-05 | 0.99567 | 34.8296 | 0.0001954
LSSL2010#140 | PP-6 | 1230 |  | 2011-02-03 13:12 | 24C | 0.99154 | 7.4722e-06 | 2.4021e-05 | 0.99157 | 34.6683 | 0.00028939
LSSL2010#141 | PP-6 | 1231 |  | 2011-02-03 13:15 | 24C | 0.99148 | 8.3443e-06 | 2.3961e-05 | 0.99151 | 34.666 | 0.00032832
LSSL2010#142 | PP-6 | 1232 |  | 2011-02-03 13:17 | 24C | 0.99141 | 4.6066e-06 | 2.3961e-05 | 0.99144 | 34.6632 | 0.00017824
LSSL2010#143 | PP-6 | 1233 |  | 2011-02-03 13:19 | 24C | 0.99125 | 3.854e-06 | 2.3961e-05 | 0.99127 | 34.6568 | 0.00015227
LSSL2010#144 | PP-6 | 1234 |  | 2011-02-03 13:22 | 24C | 0.78013 | 4.6354e-06 | 2.4021e-05 | 0.78015 | 26.5191 | 0.00017619
LSSL2010#145 | PP-6 | 1235 |  | 2011-02-03 13:25 | 24C | 0.78006 | 3.6105e-06 | 2.3961e-05 | 0.78009 | 26.5167 | 0.00013716
LSSL2010#146 | PP-6 | 1236 |  | 2011-02-03 13:28 | 24C | 0.78005 | 9.5635e-06 | 2.4021e-05 | 0.78008 | 26.5163 | 0.00035717
LSSL2010#147 | PP-6 | 1237 |  | 2011-02-03 13:30 | 24C | 0.78004 | 5.4694e-06 | 2.3901e-05 | 0.78007 | 26.516 | 0.00020784
LSSL2010#148 | reference sample | CB-3 | 5 | 2011-02-03 13:33 | 24C | 0.99885 | 4.9731e-06 | 2.408e-05 | 0.99887 | 34.9555 | 0.00019541 | 614-615
LSSL2010#149 | CB-17 | 1142 |  | 2011-02-03 13:45 | 24C | 0.99886 | 7.8373e-06 | 2.3961e-05 | 0.99888 | 34.9561 | 0.00030983
LSSL2010#150 | CB-17 | 1143 |  | 2011-02-03 13:47 | 24C | 0.99886 | 4.2256e-06 | 2.4021e-05 | 0.99888 | 34.956 | 0.00016938
LSSL2010#151 | CB-17 | 1144 |  | 2011-02-03 13:49 | 24C | 0.99876 | 3.8623e-06 | 2.3961e-05 | 0.99879 | 34.9523 | 0.00015198
LSSL2010#152 | CB-17 | 1145 |  | 2011-02-03 13:52 | 24C | 0.99848 | 7.277e-06 | 2.4021e-05 | 0.9985 | 34.941 | 0.00028506
LSSL2010#153 | CB-17 | 1146 |  | 2011-02-03 13:54 | 24C | 0.9977 | 4.1602e-06 | 2.3961e-05 | 0.99773 | 34.9105 | 0.00016507
LSSL2010#154 | CB-17 | 1147 |  | 2011-02-03 13:57 | 24C | 0.99696 | 6.6094e-06 | 2.408e-05 | 0.99698 | 34.8813 | 0.00025993
LSSL2010#155 | CB-17 | 1148 |  | 2011-02-03 13:59 | 24C | 0.99672 | 5.1679e-06 | 2.408e-05 | 0.99674 | 34.8719 | 0.000204
|  | 1149 |  | NO SAMPLE
LSSL2010#156 | CB-17 | 1150 |  | 2011-02-03 14:02 | 24C | 0.9959 | 5.346e-06 | 2.4021e-05 | 0.99592 | 34.8395 | 0.00020888
LSSL2010#157 | CB-17 | 1151 |  | 2011-02-03 14:04 | 24C | 0.99564 | 5.6518e-06 | 2.3961e-05 | 0.99566 | 34.8293 | 0.00022341
LSSL2010#158 | CB-17 | 1152 |  | 2011-02-03 15:15 | 24C | 0.99524 | 7.8752e-06 | 2.3961e-05 | 0.99527 | 34.8137 | 0.00030593
LSSL2010#159 | CB-17 | 1153 |  | 2011-02-03 15:18 | 24C | 0.99453 | 7.1644e-06 | 2.3901e-05 | 0.99455 | 34.7858 | 0.00028045
LSSL2010#160 | CB-17 | 1154 |  | 2011-02-03 15:20 | 24C | 0.99344 | 4.9004e-06 | 2.3961e-05 | 0.99346 | 34.7429 | 0.0001893
LSSL2010#161 | CB-17 | 1155 |  | 2011-02-03 15:23 | 24C | 0.98595 | 2.2398e-06 | 2.4021e-05 | 0.98597 | 34.4485 | 9.0443e-05
LSSL2010#162 | CB-17 | 1156 |  | 2011-02-03 15:25 | 24C | 0.97925 | 2.8897e-06 | 2.3961e-05 | 0.97928 | 34.1858 | 0.00011403
LSSL2010#163 | CB-17 | 1157 |  | 2011-02-03 15:28 | 24C | 0.96571 | 2.4909e-06 | 2.4021e-05 | 0.96574 | 33.6555 | 9.3081e-05
LSSL2010#164 | CB-17 | 1158 |  | 2011-02-03 15:31 | 24C | 0.94782 | 4.1138e-06 | 2.3961e-05 | 0.94784 | 32.9568 | 0.00015985
LSSL2010#165 | CB-17 | 1159 |  | 2011-02-03 15:33 | 24C | 0.94859 | 5.1485e-06 | 2.4021e-05 | 0.94861 | 32.9867 | 0.00019999
LSSL2010#166 | CB-17 | 1160 |  | 2011-02-03 15:36 | 24C | 0.94058 | 4.1859e-06 | 2.3961e-05 | 0.94061 | 32.6748 | 0.00016257
LSSL2010#167 | CB-17 | 1161 |  | 2011-02-03 15:38 | 24C | 0.93386 | 4.6353e-06 | 2.3961e-05 | 0.93388 | 32.4132 | 0.00018175
LSSL2010#168 | CB-17 | 1162 |  | 2011-02-03 15:41 | 24C | 0.91953 | 2.7757e-06 | 2.3961e-05 | 0.91955 | 31.8564 | 0.00010815
LSSL2010#169 | CB-17 | 1163 |  | 2011-02-03 15:43 | 24C | 0.89152 | 5.8894e-06 | 2.3961e-05 | 0.89154 | 30.7729 | 0.00022498
LSSL2010#170 | CB-17 | 1164 |  | 2011-02-03 15:45 | 24C | 0.80607 | 4.1603e-06 | 2.4021e-05 | 0.8061 | 27.5018 | 0.00015864
LSSL2010#171 | CB-17 | 1165 |  | 2011-02-03 15:47 | 24C | 0.75903 | 3.2516e-06 | 2.408e-05 | 0.75906 | 25.7237 | 0.00012176
LSSL2010#172 | reference sample | CB-3 | 6 | 2011-02-03 15:50 | 24C | 0.99885 | 2.566e-06 | 2.3961e-05 | 0.99888 | 34.9559 | 0.00010155 | 614-615
LSSL2010#173 | CB-54 | 1550 |  | 2011-02-03 15:57 | 24C | 0.98046 | 8.5219e-06 | 2.4021e-05 | 0.98049 | 34.2332 | 0.00033563
LSSL2010#174 | CB-54 | 1551 |  | 2011-02-03 15:59 | 24C | 0.97575 | 1.7599e-06 | 2.4021e-05 | 0.97578 | 34.0487 | 6.8803e-05
LSSL2010#175 | CB-54 | 1552 |  | 2011-02-03 16:01 | 24C | 0.96858 | 4.4373e-06 | 2.4021e-05 | 0.9686 | 33.7676 | 0.0001706
LSSL2010#176 | CB-54 | 1553 |  | 2011-02-03 16:04 | 24C | 0.95153 | 2.7425e-06 | 2.4021e-05 | 0.95155 | 33.1014 | 0.00010447
LSSL2010#177 | CB-54 | 1554 |  | 2011-02-03 16:06 | 24C | 0.93691 | 5.1439e-06 | 2.3961e-05 | 0.93693 | 32.5316 | 0.00020148
LSSL2010#178 | CB-54 | 1555 |  | 2011-02-03 16:09 | 24C | 0.92335 | 2.2394e-06 | 2.4021e-05 | 0.92338 | 32.0049 | 8.5299e-05
LSSL2010#179 | CB-54 | 1556 |  | 2011-02-03 16:11 | 24C | 0.90128 | 4.5128e-06 | 2.3961e-05 | 0.90131 | 31.15 | 0.00017531
LSSL2010#180 | CB-54 | 1557 |  | 2011-02-03 16:13 | 24C | 0.86233 | 3.6104e-06 | 2.3961e-05 | 0.86235 | 29.6494 | 0.00013702
LSSL2010#181 | CB-54 | 1558 |  | 2011-02-03 16:15 | 24C | 0.83805 | 4.1137e-06 | 2.4021e-05 | 0.83807 | 28.7197 | 0.00015439
LSSL2010#182 | CB-54 | 1559 |  | 2011-02-03 16:17 | 24C | 0.83802 | 4.9101e-06 | 2.3901e-05 | 0.83805 | 28.7187 | 0.00018646
LSSL2010#183 | reference sample | CB-3 | 7 | 2011-02-03 16:20 | 24C | 0.99886 | 4.1135e-06 | 2.3961e-05 | 0.99888 | 34.956 | 0.00016194 | 614-615
LSSL2010#184 | PP-6 | 1238 |  | 2011-02-03 16:25 | 24C | 0.99886 | 5.4855e-06 | 2.3961e-05 | 0.99888 | 34.9561 | 0.00021651
LSSL2010#185 | PP-6 | 1239 |  | 2011-02-03 16:27 | 24C | 0.9988 | 7.1553e-06 | 2.408e-05 | 0.99883 | 34.9539 | 0.00027977
LSSL2010#186 | PP-6 | 1240 |  | 2011-02-03 16:29 | 24C | 0.9985 | 4.7172e-06 | 2.408e-05 | 0.99852 | 34.9419 | 0.00018654
LSSL2010#187 | PP-6 | 1241 |  | 2011-02-03 16:31 | 24C | 0.99793 | 1.4046e-06 | 2.4021e-05 | 0.99795 | 34.9194 | 5.4326e-05
LSSL2010#188 | PP-6 | 1242 |  | 2011-02-03 16:34 | 24C | 0.99707 | 5.6675e-06 | 2.4021e-05 | 0.99709 | 34.8855 | 0.00022339
LSSL2010#189 | PP-6 | 1243 |  | 2011-02-03 16:36 | 24C | 0.99677 | 2.7425e-06 | 2.3961e-05 | 0.99679 | 34.8736 | 0.00010761
LSSL2010#190 | PP-6 | 1244 |  | 2011-02-03 16:39 | 24C | 0.99639 | 7.6436e-06 | 2.408e-05 | 0.99641 | 34.8588 | 0.00030295
LSSL2010#191 | PP-6 | 1245 |  | 2011-02-03 16:41 | 24C | 0.99607 | 5.3481e-06 | 2.3961e-05 | 0.99609 | 34.8462 | 0.00021009
LSSL2010#192 | PP-6 | 1246 |  | 2011-02-03 16:43 | 24C | 0.99571 | 5.168e-06 | 2.3961e-05 | 0.99574 | 34.8323 | 0.00020516
LSSL2010#193 | PP-6 | 1247 |  | 2011-02-03 16:46 | 24C | 0.99569 | 7.6003e-06 | 2.408e-05 | 0.99571 | 34.8314 | 0.00030124
LSSL2010#194 | PP-6 | 1248 |  | 2011-02-03 16:48 | 24C | 0.99441 | 4.8959e-06 | 2.408e-05 | 0.99443 | 34.781 | 0.00019429
LSSL2010#195 | PP-6 | 1249 |  | 2011-02-03 16:51 | 24C | 0.9913 | 7.654e-06 | 2.4021e-05 | 0.99132 | 34.6587 | 0.00029778
LSSL2010#196 | PP-6 | 1250 |  | 2011-02-03 16:54 | 24C | 0.98542 | 7.353e-06 | 2.408e-05 | 0.98544 | 34.4278 | 0.00028867
LSSL2010#197 | PP-6 | 1251 |  | 2011-02-03 16:56 | 24C | 0.98173 | 5.1678e-06 | 2.4021e-05 | 0.98175 | 34.2829 | 0.00020097
LSSL2010#198 | PP-6 | 1252 |  | 2011-02-03 16:58 | 24C | 0.97569 | 4.3839e-06 | 2.408e-05 | 0.97572 | 34.0462 | 0.00017062
LSSL2010#199 | PP-6 | 1253 |  | 2011-02-03 17:16 | 24C | 0.95422 | 7.3696e-06 | 2.3961e-05 | 0.95425 | 33.2066 | 0.00029013
LSSL2010#200 | PP-6 | 1254 |  | 2011-02-03 17:19 | 24C | 0.94663 | 5.9189e-06 | 2.4021e-05 | 0.94665 | 32.9103 | 0.00022985
LSSL2010#201 | PP-6 | 1255 |  | 2011-02-03 17:21 | 24C | 0.9401 | 1.0298e-05 | 2.4021e-05 | 0.94013 | 32.6561 | 0.00040387
LSSL2010#202 | PP-6 | 1256 |  | 2011-02-03 17:24 | 24C | 0.93264 | 3.6105e-06 | 2.3961e-05 | 0.93266 | 32.3657 | 0.00013946
LSSL2010#203 | PP-6 | 1257 |  | 2011-02-03 17:26 | 24C | 0.92143 | 7.4723e-06 | 2.4021e-05 | 0.92145 | 31.9302 | 0.00029207
LSSL2010#204 | PP-6 | 1258 |  | 2011-02-03 17:29 | 24C | 0.91297 | 8.4891e-06 | 2.3961e-05 | 0.913 | 31.6024 | 0.00032934
LSSL2010#205 | PP-6 | 1259 |  | 2011-02-03 17:31 | 24C | 0.89228 | 7.4405e-06 | 2.3961e-05 | 0.8923 | 30.802 | 0.00028508
LSSL2010#206 | PP-6 | 1260 |  | 2011-02-03 17:34 | 24C | 0.80638 | 5.3357e-06 | 2.3961e-05 | 0.8064 | 27.5136 | 0.00020341
LSSL2010#207 | PP-6 | 1261 |  | 2011-02-03 17:37 | 24C | 0.78035 | 2.2259e-06 | 2.4021e-05 | 0.78037 | 26.5275 | 8.4873e-05
LSSL2010#208 | reference sample | CB-3 | 8 | 2011-02-03 17:40 | 24C | 0.99886 | 4.7303e-06 | 2.408e-05 | 0.99888 | 34.956 | 0.00018726 | 614-615
Calibration#3 |  | K15 |  | 2011-02-04 09:20 | 24C | 0.99996 | 2.7425e-06 | 8.1062e-06 | 0.99997
LSSL2010#209 | CB-51 | 1478 |  | 2011-02-04 09:28 | 24C | 0.99878 | 4.6354e-06 | 8.1062e-06 | 0.99879 | 34.9525 | 0.0001827
LSSL2010#210 | CB-51 | 1479 |  | 2011-02-04 09:30 | 24C | 0.99859 | 3.2148e-06 | 7.8678e-06 | 0.9986 | 34.9448 | 0.00012555
| CB-51 | 1480 |  | salinity value was determined using SALINOM.EXE Program | [5 empty] | 34.9168 |  | power bump!! #1480 did not get recorded - communications lost- reason unknown-
LSSL2010#211 | CB-51 | 1481 |  | 2011-02-04 09:50 | 24C | 0.99746 | 4.7428e-06 | 8.1062e-06 | 0.99747 | 34.9003 | 0.00018599
LSSL2010#212 | CB-51 | 1482 |  | 2011-02-04 09:55 | 24C | 0.99713 | 3.3596e-06 | 7.8678e-06 | 0.99713 | 34.8872 | 0.00013211
LSSL2010#213 | CB-51 | 1483 |  | 2011-02-04 09:58 | 24C | 0.99679 | 3.6105e-06 | 8.1062e-06 | 0.9968 | 34.874 | 0.00014493
LSSL2010#214 | CB-51 | 1484 |  | 2011-02-04 10:02 | 24C | 0.99634 | 5.5831e-06 | 7.987e-06 | 0.99635 | 34.8563 | 0.00021838
LSSL2010#215 | CB-51 | 1485 |  | 2011-02-04 10:26 | 24C | 0.99599 | 8.936e-06 | 7.987e-06 | 0.996 | 34.8424 | 0.00035184
LSSL2010#216 | CB-51 | 1486 |  | 2011-02-04 10:28 | 24C | 0.99548 | 3.6104e-06 | 7.987e-06 | 0.99549 | 34.8226 | 0.00013822
LSSL2010#217 | CB-51 | 1487 |  | 2011-02-04 10:30 | 24C | 0.99479 | 2.4777e-06 | 8.0466e-06 | 0.9948 | 34.7955 | 9.4938e-05
| CB-51 | 1488 |  | NO SAMPLE | [7 empty] | no sample 1488
LSSL2010#218 | CB-51 | 1489 |  | 2011-02-04 10:32 | 24C | 0.99086 | 8.5641e-06 | 7.987e-06 | 0.99086 | 34.6407 | 0.00033732
LSSL2010#219 | CB-51 | 1490 |  | 2011-02-04 10:35 | 24C | 0.98615 | 2.7425e-06 | 7.987e-06 | 0.98615 | 34.4557 | 0.00010659
LSSL2010#220 | CB-51 | 1491 |  | 2011-02-04 10:37 | 24C | 0.97917 | 5.4684e-06 | 7.987e-06 | 0.97918 | 34.1819 | 0.00021546
LSSL2010#221 | CB-51 | 1492 |  | 2011-02-04 10:39 | 24C | 0.96748 | 2.4908e-06 | 7.987e-06 | 0.96748 | 33.7239 | 9.869e-05
LSSL2010#222 | CB-51 | 1493 |  | 2011-02-04 10:41 | 24C | 0.95556 | 3.7839e-06 | 7.987e-06 | 0.95557 | 33.2582 | 0.00014715
LSSL2010#223 | CB-51 | 1494 |  | 2011-02-04 10:43 | 24C | 0.95306 | 4.6227e-06 | 8.0466e-06 | 0.95307 | 33.1604 | 0.00018375
LSSL2010#224 | CB-51 | 1495 |  | 2011-02-04 10:45 | 24C | 0.94591 | 2.7425e-06 | 7.987e-06 | 0.94592 | 32.8819 | 0.00010656
LSSL2010#225 | CB-51 | 1496 |  | 2011-02-04 10:47 | 24C | 0.93389 | 2.89e-06 | 7.987e-06 | 0.9339 | 32.4136 | 0.00011207
LSSL2010#226 | CB-51 | 1497 |  | 2011-02-04 10:49 | 24C | 0.9245 | 4.1858e-06 | 7.987e-06 | 0.92451 | 32.049 | 0.00016335
LSSL2010#227 | CB-51 | 1498 |  | 2011-02-04 10:52 | 24C | 0.90438 | 7.9878e-06 | 8.0466e-06 | 0.90439 | 31.2691 | 0.00030599
LSSL2010#228 | CB-51 | 1499 |  | 2011-02-04 10:54 | 24C | 0.88398 | 4.6675e-06 | 7.987e-06 | 0.88399 | 30.4814 | 0.00018152
LSSL2010#229 | CB-51 | 1500 |  | 2011-02-04 10:57 | 24C | 0.83764 | 4.3337e-06 | 7.987e-06 | 0.83765 | 28.7036 | 0.00016464
LSSL2010#230 | CB-50 | 1454 |  | 2011-02-04 11:05 | 24C | 0.99888 | 2.4911e-06 | 7.9274e-06 | 0.99889 | 34.9562 | 0.00010059
LSSL2010#231 | CB-50 | 1455 |  | 2011-02-04 11:08 | 24C | 0.99885 | 8.4423e-06 | 7.987e-06 | 0.99886 | 34.9551 | 0.00033122
LSSL2010#232 | CB-50 | 1456 |  | 2011-02-04 11:10 | 24C | 0.99857 | 2.7097e-06 | 8.0466e-06 | 0.99858 | 34.9441 | 0.00010865
LSSL2010#233 | CB-50 | 1457 |  | 2011-02-04 11:12 | 24C | 0.9979 | 4.9134e-06 | 7.987e-06 | 0.99791 | 34.9177 | 0.00019232
LSSL2010#234 | CB-50 | 1458 |  | 2011-02-04 11:14 | 24C | 0.9971 | 4.3839e-06 | 7.987e-06 | 0.99711 | 34.8862 | 0.00017276
LSSL2010#235 | CB-50 | 1459 |  | 2011-02-04 11:17 | 24C | 0.99676 | 2.4915e-06 | 7.9274e-06 | 0.99676 | 34.8727 | 9.4938e-05
LSSL2010#236 | CB-50 | 1460 |  | 2011-02-04 11:19 | 24C | 0.99633 | 3.8654e-06 | 7.8678e-06 | 0.99634 | 34.8558 | 0.00015205
LSSL2010#237 | CB-50 | 1461 |  | 2011-02-04 11:21 | 24C | 0.996 | 2.4909e-06 | 8.0466e-06 | 0.996 | 34.8428 | 9.9543e-05
LSSL2010#238 | CB-50 | 1462 |  | 2011-02-04 11:24 | 24C | 0.99566 | 6.6505e-06 | 7.987e-06 | 0.99567 | 34.8296 | 0.00025921
LSSL2010#239 | CB-50 | 1463 |  | 2011-02-04 11:27 | 24C | 0.99557 | 4.2257e-06 | 8.1062e-06 | 0.99558 | 34.826 | 0.00016483
LSSL2010#240 | CB-50 | 1464 |  | 2011-02-04 11:29 | 24C | 0.99432 | 4.6041e-06 | 7.987e-06 | 0.99433 | 34.7768 | 0.00018322
LSSL2010#241 | CB-50 | 1465 |  | 2011-02-04 11:32 | 24C | 0.99169 | 6.5495e-06 | 7.987e-06 | 0.99169 | 34.6733 | 0.00025751
LSSL2010#242 | CB-50 | 1466 |  | 2011-02-04 11:34 | 24C | 0.98962 | 7.2313e-06 | 8.0466e-06 | 0.98963 | 34.5923 | 0.00028765
LSSL2010#243 | CB-50 | 1467 |  | 2011-02-04 11:37 | 24C | 0.98613 | 6.3303e-06 | 7.987e-06 | 0.98614 | 34.455 | 0.0002493
LSSL2010#244 | CB-50 | 1468 |  | 2011-02-04 11:39 | 24C | 0.98166 | 5.8911e-06 | 7.987e-06 | 0.98167 | 34.2796 | 0.00022967
LSSL2010#245 | CB-50 | 1469 |  | 2011-02-04 11:41 | 24C | 0.95771 | 3.2151e-06 | 7.987e-06 | 0.95771 | 33.3419 | 0.000126
LSSL2010#246 | CB-50 | 1470 |  | 2011-02-04 11:43 | 24C | 0.94495 | 5.6569e-06 | 8.0466e-06 | 0.94495 | 32.8441 | 0.00021857
LSSL2010#247 | CB-50 | 1471 |  | 2011-02-04 11:46 | 24C | 0.94535 | 6.4249e-06 | 7.987e-06 | 0.94536 | 32.86 | 0.00025058
LSSL2010#248 | CB-50 | 1472 |  | 2011-02-04 11:48 | 24C | 0.93539 | 2.4916e-06 | 8.0466e-06 | 0.93539 | 32.4719 | 9.4885e-05
LSSL2010#249 | CB-50 | 1473 |  | 2011-02-04 11:50 | 24C | 0.92528 | 1.7702e-06 | 8.0466e-06 | 0.92529 | 32.0793 | 6.7435e-05
LSSL2010#250 | CB-50 | 1474 |  | 2011-02-04 11:53 | 24C | 0.90412 | 6.8246e-06 | 7.987e-06 | 0.90413 | 31.259 | 0.00026221
LSSL2010#251 | CB-50 | 1475 |  | 2011-02-04 11:55 | 24C | 0.87374 | 5.4532e-06 | 7.987e-06 | 0.87374 | 30.0872 | 0.00020899
LSSL2010#252 | CB-50 | 1476 |  | 2011-02-04 11:57 | 24C | 0.8198 | 7.9034e-06 | 7.987e-06 | 0.81981 | 28.0232 | 0.00030019
LSSL2010#253 | CB-50 | 1477 |  | 2011-02-04 11:59 | 24C | 0.798 | 6.5629e-06 | 7.9274e-06 | 0.79801 | 27.1951 | 0.00024928
LSSL2010#254 | reference sample | CB-3 | 8 | 2011-02-04 12:03 | 24C | 0.99914 | 6.2796e-06 | 8.0466e-06 | 0.99915 | 34.9666 | 0.0002493 | 614-615
LSSL2010#255 | CB-21 | 1406 |  | 2011-02-04 12:15 | 24C | 0.99888 | 1.1064e-06 | 7.987e-06 | 0.99889 | 34.9563 | 4.1799e-05
LSSL2010#256 | CB-21 | 1407 |  | 2011-02-04 12:17 | 24C | 0.99887 | 1.1203e-06 | 7.9274e-06 | 0.99888 | 34.9558 | 4.5217e-05
LSSL2010#257 | CB-21 | 1408 |  | 2011-02-04 12:19 | 24C | 0.99879 | 5.8609e-06 | 8.1062e-06 | 0.9988 | 34.9529 | 0.00023274
LSSL2010#258 | CB-21 | 1409 |  | 2011-02-04 12:21 | 24C | 0.99852 | 1.3712e-06 | 7.9274e-06 | 0.99852 | 34.9419 | 5.4351e-05
LSSL2010#259 | CB-21 | 1410 |  | 2011-02-04 12:24 | 24C | 0.99785 | 7.7497e-06 | 7.8678e-06 | 0.99786 | 34.9157 | 0.00030526
LSSL2010#260 | CB-21 | 1411 |  | 2011-02-04 12:26 | 24C | 0.99701 | 4.6353e-06 | 7.987e-06 | 0.99702 | 34.8828 | 0.00018273
LSSL2010#261 | CB-21 | 1412 |  | 2011-02-04 12:29 | 24C | 0.99663 | 2.4908e-06 | 7.9274e-06 | 0.99663 | 34.8675 | 9.7029e-05
LSSL2010#262 | CB-21 | 1413 |  | 2011-02-04 12:31 | 24C | 0.99616 | 8.198e-06 | 7.9274e-06 | 0.99617 | 34.8494 | 0.000324
LSSL2010#263 | CB-21 | 1414 |  | 2011-02-04 12:33 | 24C | 0.99577 | 9.3319e-06 | 7.987e-06 | 0.99578 | 34.8338 | 0.00036908
LSSL2010#264 | CB-21 | 1415 |  | 2011-02-04 12:36 | 24C | 0.99572 | 5.865e-06 | 7.987e-06 | 0.99573 | 34.832 | 0.00023196
LSSL2010#265 | CB-21 | 1416 |  | 2011-02-04 12:38 | 24C | 0.9951 | 8.2053e-06 | 7.987e-06 | 0.99511 | 34.8077 | 0.00032017
LSSL2010#266 | CB-21 | 1417 |  | 2011-02-04 12:41 | 24C | 0.99369 | 4.4789e-06 | 7.8678e-06 | 0.9937 | 34.7522 | 0.00017574
LSSL2010#267 | CB-21 | 1418 |  | 2011-02-04 12:43 | 24C | 0.98888 | 5.2207e-06 | 7.987e-06 | 0.98889 | 34.563 | 0.00020243
LSSL2010#268 | CB-21 | 1419 |  | 2011-02-04 12:45 | 24C | 0.98118 | 4.1136e-06 | 7.9274e-06 | 0.98119 | 34.2608 | 0.00016064
LSSL2010#269 | CB-21 | 1420 |  | 2011-02-04 12:47 | 24C | 0.95531 | 4.2259e-06 | 7.987e-06 | 0.95532 | 33.2486 | 0.00016771
LSSL2010#270 | CB-21 | 1421 |  | 2011-02-04 12:50 | 24C | 0.95292 | 2.2275e-06 | 7.987e-06 | 0.95293 | 33.1552 | 8.7556e-05
LSSL2010#271 | CB-21 | 1422 |  | 2011-02-04 12:52 | 24C | 0.94599 | 3.4495e-06 | 7.987e-06 | 0.946 | 32.8848 | 0.00013345
LSSL2010#272 | CB-21 | 1423 |  | 2011-02-04 12:55 | 24C | 0.93973 | 8.9806e-06 | 7.987e-06 | 0.93974 | 32.6409 | 0.00034848
LSSL2010#273 | CB-21 | 1424 |  | 2011-02-04 12:57 | 24C | 0.93255 | 5.2334e-06 | 7.987e-06 | 0.93256 | 32.3616 | 0.0002041
LSSL2010#274 | CB-21 | 1425 |  | 2011-02-04 12:59 | 24C | 0.92279 | 5.2332e-06 | 8.1062e-06 | 0.9228 | 31.9824 | 0.00020115
LSSL2010#275 | CB-21 | 1426 |  | 2011-02-04 13:01 | 24C | 0.90316 | 7.5748e-06 | 8.1062e-06 | 0.90317 | 31.2219 | 0.00029238
LSSL2010#276 | CB-21 | 1427 |  | 2011-02-04 13:04 | 24C | 0.87725 | 5.5671e-06 | 7.987e-06 | 0.87726 | 30.2224 | 0.0002137
LSSL2010#277 | CB-21 | 1428 |  | 2011-02-04 13:06 | 24C | 0.77525 | 8.2918e-06 | 7.987e-06 | 0.77526 | 26.3343 | 0.00031316
LSSL2010#278 | CB-21 | 1429 |  | 2011-02-04 13:09 | 24C | 0.74067 | 1.3712e-06 | 7.987e-06 | 0.74068 | 25.0334 | 5.1717e-05
LSSL2010#279 | reference sample | CB-3 | 9 | 2011-02-04 13:11 | 24C | 0.99884 | 5.5281e-06 | 7.987e-06 | 0.99884 | 34.9545 | 0.00021693 | 614-615
LSSL2010#280 | reference sample | CB-3 | 10 | 2011-02-04 13:31 | 24C | 0.99888 | 6.1457e-06 | 7.987e-06 | 0.99889 | 34.9564 | 0.00024314 | 614-615
Calibration#4 |  | K15 |  | 2011-02-07 10:05 | 24C | 0.99993 | 0 | 4.0054e-05 | 0.99997
LSSL2010#281 | CB-40 | 1430 |  | 2011-02-07 10:39 | 24C | 0.99888 | 1.1199e-06 | 3.9995e-05 | 0.99892 | 34.9575 | 4.265e-05
LSSL2010#282 | CB-40 | 1431 |  | 2011-02-07 10:41 | 24C | 0.99882 | 3.2141e-06 | 4.0054e-05 | 0.99886 | 34.9549 | 0.00012709
LSSL2010#283 | CB-40 | 1432 |  | 2011-02-07 10:43 | 24C | 0.99854 | 4.3336e-06 | 4.0054e-05 | 0.99858 | 34.9443 | 0.00017093
LSSL2010#284 | CB-40 | 1433 |  | 2011-02-07 10:46 | 24C | 0.99795 | 3.865e-06 | 4.0054e-05 | 0.99799 | 34.9209 | 0.00015448
LSSL2010#285 | CB-40 | 1434 |  | 2011-02-07 10:48 | 24C | 0.99702 | 1.1199e-06 | 3.9995e-05 | 0.99706 | 34.8841 | 4.5227e-05
LSSL2010#286 | CB-40 | 1435 |  | 2011-02-07 10:50 | 24C | 0.99676 | 1.1199e-06 | 4.0054e-05 | 0.9968 | 34.8742 | 4.4374e-05
LSSL2010#287 | CB-40 | 1436 |  | 2011-02-07 10:52 | 24C | 0.99636 | 3.2142e-06 | 4.0054e-05 | 0.9964 | 34.8584 | 0.00012382
LSSL2010#288 | CB-40 | 1437 |  | 2011-02-07 10:54 | 24C | 0.99609 | 4.2256e-06 | 3.9995e-05 | 0.99613 | 34.8476 | 0.00016944
LSSL2010#289 | CB-40 | 1438 |  | 2011-02-07 10:56 | 24C | 0.9957 | 4.7431e-06 | 4.0054e-05 | 0.99574 | 34.8323 | 0.00018675
LSSL2010#290 | CB-40 | 1439 |  | 2011-02-07 10:58 | 24C | 0.99523 | 5.0782e-06 | 3.9935e-05 | 0.99527 | 34.8138 | 0.00019832
LSSL2010#291 | CB-40 | 1440 |  | 2011-02-07 11:00 | 24C | 0.99479 | 3.2141e-06 | 3.9995e-05 | 0.99483 | 34.7966 | 0.00012382
LSSL2010#292 | CB-40 | 1441 |  | 2011-02-07 11:03 | 24C | 0.99371 | 3.2144e-06 | 4.0054e-05 | 0.99375 | 34.754 | 0.00012599
LSSL2010#293 | CB-40 | 1442 |  | 2011-02-07 11:05 | 24C | 0.99107 | 4.1786e-06 | 3.9935e-05 | 0.99111 | 34.6505 | 0.00016283
LSSL2010#294 | CB-40 | 1443 |  | 2011-02-07 11:07 | 24C | 0.98613 | 4.0096e-06 | 4.0054e-05 | 0.98617 | 34.4565 | 0.00015769
LSSL2010#295 | CB-40 | 1444 |  | 2011-02-07 11:10 | 24C | 0.97888 | 2.4909e-06 | 3.9935e-05 | 0.97892 | 34.172 | 9.869e-05
LSSL2010#296 | CB-40 | 1445 |  | 2011-02-07 11:12 | 24C | 0.96859 | 5.1607e-06 | 3.9995e-05 | 0.96863 | 33.7688 | 0.00019816
LSSL2010#297 | CB-40 | 1446 |  | 2011-02-07 11:14 | 24C | 0.94128 | 3.9742e-06 | 4.0054e-05 | 0.94132 | 32.7025 | 0.00015499
LSSL2010#298 | CB-40 | 1447 |  | 2011-02-07 11:16 | 24C | 0.96293 | 3.5949e-06 | 3.9995e-05 | 0.96297 | 33.5474 | 0.00014165 | Really??
LSSL2010#299 | CB-40 | 1448 |  | 2011-02-07 11:19 | 24C | 0.9422 | 3.2141e-06 | 3.9995e-05 | 0.94224 | 32.7385 | 0.00012244
LSSL2010#300 | CB-40 | 1449 |  | 2011-02-07 11:22 | 24C | 0.93356 | 3.3595e-06 | 4.0054e-05 | 0.9336 | 32.4021 | 0.0001317
LSSL2010#301 | CB-40 | 1450 |  | 2011-02-07 11:24 | 24C | 0.92237 | 3.1253e-06 | 4.0054e-05 | 0.92241 | 31.9674 | 0.00012085
LSSL2010#302 | CB-40 | 1451 |  | 2011-02-07 11:26 | 24C | 0.89603 | 2.89e-06 | 3.9995e-05 | 0.89607 | 30.9478 | 0.00011103
LSSL2010#303 | CB-40 | 1452 |  | 2011-02-07 11:28 | 24C | 0.80332 | 2.8901e-06 | 3.9995e-05 | 0.80336 | 27.3979 | 0.00010975
LSSL2010#304 | CB-40 | 1453 |  | 2011-02-07 11:31 | 24C | 0.76957 | 2.2394e-06 | 3.9995e-05 | 0.76961 | 26.1212 | 8.2774e-05
LSSL2010#305 | reference sample | CB-3 | 11 | 2011-02-07 11:35 | 24C | 0.99884 | 3.9806e-06 | 4.0054e-05 | 0.99888 | 34.956 | 0.00015625 | 614-615
LSSL2010#306 | CB-53 | 1526 |  | 2011-02-07 12:00 | 24C | 0.99679 | 4.189e-06 | 3.9995e-05 | 0.99683 | 34.8753 | 0.00016758
LSSL2010#307 | CB-54 | 1527 |  | 2011-02-07 12:02 | 24C | 0.99543 | 3.4659e-06 | 3.9995e-05 | 0.99547 | 34.8216 | 0.00013699
LSSL2010#308 | CB-55 | 1528 |  | 2011-02-07 12:04 | 24C | 0.99462 | 2.4615e-06 | 3.9995e-05 | 0.99466 | 34.79 | 9.869e-05
LSSL2010#309 | CB-56 | 1529 |  | 2011-02-07 12:06 | 24C | 0.99408 | 3.8647e-06 | 3.9935e-05 | 0.99412 | 34.7686 | 0.00015042
LSSL2010#310 | CB-57 | 1530 |  | 2011-02-07 12:08 | 24C | 0.99263 | 1.1532e-06 | 4.0054e-05 | 0.99267 | 34.7118 | 4.4374e-05
LSSL2010#311 | CB-58 | 1531 |  | 2011-02-07 12:10 | 24C | 0.99028 | 3.131e-06 | 4.0054e-05 | 0.99032 | 34.6193 | 0.00012312
LSSL2010#312 | CB-59 | 1532 |  | 2011-02-07 12:13 | 24C | 0.9884 | 1.1203e-06 | 3.9995e-05 | 0.98844 | 34.5455 | 4.5227e-05
LSSL2010#313 | CB-60 | 1534 |  | 2011-02-07 12:14 | 24C | 0.97954 | 3.623e-06 | 3.9995e-05 | 0.97958 | 34.1979 | 0.00014089
LSSL2010#314 | CB-61 | 1535 |  | 2011-02-07 12:17 | 24C | 0.96907 | 2.8792e-06 | 3.9995e-05 | 0.96911 | 33.7876 | 0.00011402
LSSL2010#315 | CB-62 | 1536 |  | 2011-02-07 12:19 | 24C | 0.95548 | 2.6733e-06 | 4.0054e-05 | 0.95552 | 33.2562 | 0.00010274
LSSL2010#316 | CB-63 | 1537 |  | 2011-02-07 12:21 | 24C | 0.95068 | 2.2398e-06 | 4.0054e-05 | 0.95072 | 33.069 | 8.7048e-05
LSSL2010#317 | CB-64 | 1538 |  | 2011-02-07 12:23 | 24C | 0.94154 | 2.8901e-06 | 3.9995e-05 | 0.94158 | 32.7127 | 0.00011145
LSSL2010#318 | CB-65 | 1539 |  | 2011-02-07 12:25 | 24C | 0.93287 | 2.4747e-06 | 3.9995e-05 | 0.93291 | 32.3751 | 9.4938e-05
LSSL2010#319 | CB-66 | 1540 |  | 2011-02-07 12:27 | 24C | 0.92293 | 2.4915e-06 | 4.0054e-05 | 0.92297 | 31.9892 | 9.3957e-05
LSSL2010#320 | CB-67 | 1541 |  | 2011-02-07 12:29 | 24C | 0.90216 | 3.6105e-06 | 4.0054e-05 | 0.9022 | 31.1845 | 0.00013706
LSSL2010#321 | CB-68 | 1542 |  | 2011-02-07 12:31 | 24C | 0.86296 | 2.2401e-06 | 4.0054e-05 | 0.863 | 29.6744 | 8.7006e-05
LSSL2010#322 | CB-69 | 1544 |  | 2011-02-07 12:33 | 24C | 0.83224 | 4.384e-06 | 4.0054e-05 | 0.83228 | 28.4987 | 0.00016658
LSSL2010#323 | CB-70 | 1544 |  | 2011-02-07 12:36 | 24C | 0.8314 | 2.7426e-06 | 3.9935e-05 | 0.83144 | 28.4664 | 0.00010449
LSSL2010#324 | reference sample | CB-3 | 12 | 2011-02-07 12:43 | 24C | 0.99883 | 2.9035e-06 | 4.0054e-05 | 0.99887 | 34.9554 | 0.00011369 | 614-615
LSSL2010#325 | CB-52 | 1502 |  | 2011-02-07 12:51 | 24C | 0.99727 | 2.4912e-06 | 4.0054e-05 | 0.99731 | 34.894 | 9.869e-05
LSSL2010#326 | CB-52 | 1503 |  | 2011-02-07 12:53 | 24C | 0.99694 | 2.7262e-06 | 4.0054e-05 | 0.99698 | 34.8812 | 0.00010761
LSSL2010#327 | CB-52 | 1504 |  | 2011-02-07 12:55 | 24C | 0.99677 | 3.3308e-06 | 4.0054e-05 | 0.99681 | 34.8746 | 0.00012839
LSSL2010#328 | CB-52 | 1505 |  | 2011-02-07 12:57 | 24C | 0.99655 | 2.7425e-06 | 3.9995e-05 | 0.99659 | 34.8657 | 0.00010865
LSSL2010#329 | CB-52 | 1506 |  | 2011-02-07 12:59 | 24C | 0.9963 | 3.8748e-06 | 4.0054e-05 | 0.99634 | 34.856 | 0.00015161
LSSL2010#330 | CB-52 | 1507 |  | 2011-02-07 13:01 | 24C | 0.99594 | 5.7377e-06 | 3.9995e-05 | 0.99598 | 34.8418 | 0.00022904
LSSL2010#331 | CB-52 | 1508 |  | 2011-02-07 13:03 | 24C | 0.99578 | 3.2144e-06 | 3.9995e-05 | 0.99582 | 34.8356 | 0.00012382
LSSL2010#332 | CB-52 | 1509 |  | 2011-02-07 13:05 | 24C | 0.99544 | 3.3595e-06 | 4.0054e-05 | 0.99548 | 34.8223 | 0.00013414
LSSL2010#333 | CB-52 | 1510 |  | 2011-02-07 13:07 | 24C | 0.99547 | 3.6104e-06 | 3.9995e-05 | 0.99551 | 34.8233 | 0.00014157
LSSL2010#334 | CB-52 | 1511 |  | 2011-02-07 13:09 | 24C | 0.99506 | 1.1536e-06 | 3.9995e-05 | 0.9951 | 34.8074 | 4.2703e-05
LSSL2010#335 | CB-52 | 1512 |  | 2011-02-07 13:11 | 24C | 0.99378 | 6.4718e-06 | 4.0054e-05 | 0.99382 | 34.7567 | 0.00025338
LSSL2010#336 | CB-52 | 1513 |  | 2011-02-07 13:13 | 24C | 0.98935 | 2.2394e-06 | 4.0054e-05 | 0.98939 | 34.5827 | 8.7067e-05
LSSL2010#337 | CB-52 | 1514 |  | 2011-02-07 13:15 | 24C | 0.9861 | 3.865e-06 | 3.9935e-05 | 0.98614 | 34.4551 | 0.00015134
LSSL2010#338 | CB-52 | 1515 |  | 2011-02-07 13:17 | 24C | 0.97456 | 3.1253e-06 | 3.9935e-05 | 0.9746 | 34.0026 | 0.00012
LSSL2010#339 | CB-52 | 1516 |  | 2011-02-07 13:19 | 24C | 0.97496 | 2.2263e-06 | 3.9995e-05 | 0.975 | 34.0181 | 8.6163e-05
LSSL2010#340 | CB-52 | 1517 |  | 2011-02-07 13:21 | 24C | 0.96638 | 6.311e-06 | 3.9995e-05 | 0.96642 | 33.6824 | 0.00024824
LSSL2010#341 | CB-52 | 1518 |  | 2011-02-07 13:24 | 24C | 0.93328 | 3.4662e-06 | 4.0054e-05 | 0.93332 | 32.3911 | 0.00013479
LSSL2010#342 | CB-52 | 1519 |  | 2011-02-07 13:27 | 24C | 0.94519 | 2.7425e-06 | 3.9935e-05 | 0.94523 | 32.8547 | 0.00010975 | Loose cap - loose insert
LSSL2010#343 | CB-52 | 1520 |  | 2011-02-07 13:29 | 24C | 0.93398 | 3.5406e-06 | 3.9935e-05 | 0.93402 | 32.4185 | 0.00013765
LSSL2010#344 | CB-52 | 1521 |  | 2011-02-07 13:31 | 24C | 0.92562 | 2.4909e-06 | 3.9995e-05 | 0.92566 | 32.0934 | 9.802e-05
LSSL2010#345 | CB-52 | 1522 |  | 2011-02-07 13:33 | 24C | 0.8964 | 3.1417e-06 | 3.9935e-05 | 0.89644 | 30.9619 | 0.00012365
LSSL2010#346 | CB-52 | 1523 |  | 2011-02-07 13:35 | 24C | 0.87116 | 2.2259e-06 | 3.9935e-05 | 0.87119 | 29.9891 | 8.5326e-05
LSSL2010#347 | CB-52 | 1524 |  | 2011-02-07 13:37 | 24C | 0.84289 | 3.8434e-06 | 3.9995e-05 | 0.84293 | 28.9056 | 0.00014661
LSSL2010#348 | CB-52 | 1525 |  | 2011-02-07 13:39 | 24C | 0.83498 | 2.2397e-06 | 3.9995e-05 | 0.83502 | 28.6032 | 8.4895e-05
LSSL2010#349 | reference sample | CB-3 | 13 | 2011-02-07 13:42 | 24C | 0.99885 | 3.218e-06 | 3.9935e-05 | 0.99889 | 34.9564 | 0.00012756 | 614-615
LSSL2010#350 | CB-16 | 998 |  | 2011-02-07 15:32 | 24C | 0.99882 | 6.6317e-06 | 4.0054e-05 | 0.99886 | 34.955 | 0.00025894
LSSL2010#351 | CB-16 | 999 |  | 2011-02-07 15:37 | 24C | 0.99886 | 2.0945e-06 | 4.0054e-05 | 0.9989 | 34.9568 | 8.1866e-05
LSSL2010#352 | CB-16 | 1000 |  | 2011-02-07 15:43 | 24C | 0.99875 | 2.2405e-06 | 3.9995e-05 | 0.99879 | 34.9526 | 8.9617e-05
LSSL2010#353 | CB-16 | 1001 |  | 2011-02-07 15:45 | 24C | 0.9985 | 3.1446e-06 | 3.9995e-05 | 0.99854 | 34.9425 | 0.00012628
LSSL2010#354 | CB-16 | 1002 |  | 2011-02-07 15:48 | 24C | 0.99777 | 2.7425e-06 | 3.9935e-05 | 0.99781 | 34.9139 | 0.00010447
LSSL2010#355 | CB-16 | 1003 |  | 2011-02-07 15:50 | 24C | 0.997 | 3.4657e-06 | 4.0054e-05 | 0.99704 | 34.8836 | 0.00013662
LSSL2010#356 | CB-16 | 1004 |  | 2011-02-07 15:53 | 24C | 0.99663 | 3.1415e-06 | 3.9935e-05 | 0.99667 | 34.869 | 0.0001262
LSSL2010#357 | CB-16 | 1005 |  | 2011-02-07 15:55 | 24C | 0.9962 | 3.1253e-06 | 3.9995e-05 | 0.99624 | 34.852 | 0.00012312
LSSL2010#358 | CB-16 | 1006 |  | 2011-02-07 15:57 | 24C | 0.9962 | 3.4352e-06 | 3.9995e-05 | 0.99624 | 34.8521 | 0.00013596
LSSL2010#359 | CB-16 | 1007 |  | 2011-02-07 16:00 | 24C | 0.99574 | 3.6104e-06 | 4.0054e-05 | 0.99578 | 34.8342 | 0.00013753
LSSL2010#360 | CB-16 | 1008 |  | 2011-02-07 16:02 | 24C | 0.99542 | 2.4908e-06 | 4.0054e-05 | 0.99546 | 34.8214 | 9.7495e-05
LSSL2010#361 | CB-16 | 1009 |  | 2011-02-07 16:04 | 24C | 0.99457 | 3.1253e-06 | 4.0054e-05 | 0.99461 | 34.7881 | 0.00012304
LSSL2010#362 | CB-16 | 1010 |  | 2011-02-07 16:06 | 24C | 0.99312 | 4.2739e-06 | 3.9935e-05 | 0.99316 | 34.7308 | 0.00016608
LSSL2010#363 | CB-16 | 1011 |  | 2011-02-07 16:09 | 24C | 0.98546 | 1.4046e-06 | 4.0054e-05 | 0.9855 | 34.4299 | 5.4326e-05
LSSL2010#364 | CB-16 | 1012 |  | 2011-02-07 16:11 | 24C | 0.97902 | 2.8904e-06 | 4.0054e-05 | 0.97906 | 34.1773 | 0.00011604
LSSL2010#365 | CB-16 | 1013 |  | 2011-02-07 16:12 | 24C | 0.96879 | 3.4657e-06 | 4.0054e-05 | 0.96883 | 33.7766 | 0.00013904
LSSL2010#366 | CB-16 | 1014 |  | 2011-02-07 16:14 | 24C | 0.95353 | 2.2725e-06 | 4.0054e-05 | 0.95357 | 33.1801 | 8.7556e-05
LSSL2010#367 | CB-16 | 1015 |  | 2011-02-07 16:17 | 24C | 0.94751 | 2.7425e-06 | 3.9935e-05 | 0.94755 | 32.9453 | 0.00010447
LSSL2010#368 | CB-16 | 1016 |  | 2011-02-07 16:19 | 24C | 0.94036 | 2.4909e-06 | 3.9995e-05 | 0.9404 | 32.6666 | 9.6125e-05
LSSL2010#369 | CB-16 | 1017 |  | 2011-02-07 16:21 | 24C | 0.93371 | 3.4495e-06 | 3.9995e-05 | 0.93375 | 32.4079 | 0.00013341
LSSL2010#370 | CB-16 | 1018 |  | 2011-02-07 16:23 | 24C | 0.92303 | 2.2401e-06 | 3.9995e-05 | 0.92307 | 31.9931 | 8.5326e-05
LSSL2010#371 | CB-16 | 1019 |  | 2011-02-07 16:26 | 24C | 0.90215 | 2.4908e-06 | 4.0054e-05 | 0.90219 | 31.1843 | 9.6509e-05
LSSL2010#372 | CB-16 | 1020 |  | 2011-02-07 16:28 | 24C | 0.83418 | 2.7425e-06 | 3.9935e-05 | 0.83422 | 28.5728 | 0.00010396
LSSL2010#373 | CB-16 | 1021 |  | 2011-02-07 16:30 | 24C | 0.82273 | 1.1529e-06 | 4.0054e-05 | 0.82277 | 28.1358 | 4.2663e-05
LSSL2010#374 | reference sample | CB-3 | 14 | 2011-02-07 16:41 | 24C | 0.99888 | 2.3951e-06 | 4.0054e-05 | 0.99892 | 34.9576 | 9.9652e-05 | 614-615
LSSL2010#375 | CB-18 | 1286 |  | 2011-02-07 16:52 | 24C | 0.9985 | 0 | 4.0054e-05 | 0.99854 | 34.9424 | 0
LSSL2010#376 | CB-18 | 1287 |  | 2011-02-07 16:54 | 24C | 0.9978 | 2.7425e-06 | 3.9935e-05 | 0.99784 | 34.915 | 0.0001097
LSSL2010#377 | CB-18 | 1288 |  | 2011-02-07 16:57 | 24C | 0.99727 | 3.1424e-06 | 4.0054e-05 | 0.99731 | 34.8941 | 0.00012304
LSSL2010#378 | CB-18 | 1289 |  | 2011-02-07 16:59 | 24C | 0.99694 | 3.9984e-06 | 3.9995e-05 | 0.99698 | 34.8811 | 0.00015827
LSSL2010#379 | CB-18 | 1290 |  | 2011-02-07 17:01 | 24C | 0.99654 | 2.8897e-06 | 3.9935e-05 | 0.99658 | 34.8656 | 0.00011603
LSSL2010#380 | CB-18 | 1291 |  | 2011-02-07 17:04 | 24C | 0.99615 | 4.2739e-06 | 3.9935e-05 | 0.99619 | 34.85 | 0.00016798
LSSL2010#381 | CB-18 | 1292 |  | 2011-02-07 17:06 | 24C | 0.99593 | 2.4908e-06 | 4.0054e-05 | 0.99597 | 34.8413 | 9.87e-05
LSSL2010#382 | CB-18 | 1293 |  | 2011-02-07 17:08 | 24C | 0.99573 | 5.1203e-06 | 3.9935e-05 | 0.99577 | 34.8335 | 0.00019874
LSSL2010#383 | CB-18 | 1294 |  | 2011-02-07 17:10 | 24C | 0.99573 | 4.3837e-06 | 3.9935e-05 | 0.99577 | 34.8337 | 0.00017502
LSSL2010#384 | CB-18 | 1295 |  | 2011-02-07 17:12 | 24C | 0.99496 | 3.974e-06 | 4.0054e-05 | 0.995 | 34.8031 | 0.00015598
LSSL2010#385 | CB-18 | 1296 |  | 2011-02-07 17:14 | 24C | 0.99468 | 3.4657e-06 | 4.0054e-05 | 0.99472 | 34.7922 | 0.00013549
LSSL2010#386 | CB-18 | 1297 |  | 2011-02-07 17:17 | 24C | 0.99312 | 4.1137e-06 | 3.9935e-05 | 0.99316 | 34.7309 | 0.00016506
LSSL2010#387 | CB-18 | 1298 |  | 2011-02-07 17:19 | 24C | 0.98535 | 2.7425e-06 | 3.9935e-05 | 0.98539 | 34.4257 | 0.0001097
LSSL2010#388 | CB-18 | 1299 |  | 2011-02-07 17:21 | 24C | 0.97819 | 2.8877e-06 | 3.9995e-05 | 0.97823 | 34.1448 | 0.00011227
LSSL2010#389 | CB-18 | 1300 |  | 2011-02-07 17:22 | 24C | 0.96647 | 3.4657e-06 | 3.9935e-05 | 0.96651 | 33.6859 | 0.00013596
LSSL2010#390 | CB-18 | 1301 |  | 2011-02-07 17:24 | 24C | 0.95398 | 4.0973e-06 | 3.9995e-05 | 0.95402 | 33.1976 | 0.00015879
LSSL2010#391 | CB-18 | 1302 |  | 2011-02-07 17:26 | 24C | 0.94855 | 5.2338e-06 | 3.9935e-05 | 0.94859 | 32.9858 | 0.00020294
LSSL2010#392 | CB-18 | 1303 |  | 2011-02-07 17:28 | 24C | 0.94086 | 6.3007e-06 | 4.0054e-05 | 0.9409 | 32.6863 | 0.00024377
LSSL2010#393 | CB-18 | 1304 |  | 2011-02-07 17:30 | 24C | 0.93457 | 2.7425e-06 | 4.0054e-05 | 0.93461 | 32.4414 | 0.00010761
LSSL2010#394 | CB-18 | 1305 |  | 2011-02-07 17:32 | 24C | 0.92105 | 1.4046e-06 | 3.9995e-05 | 0.92109 | 31.9162 | 5.2783e-05
LSSL2010#395 | CB-18 | 1306 |  | 2011-02-07 17:34 | 24C | 0.88772 | 2.4921e-06 | 4.0054e-05 | 0.88776 | 30.6267 | 9.4895e-05
LSSL2010#396 | CB-18 | 1307 |  | 2011-02-07 17:36 | 24C | 0.88086 | 2.2127e-06 | 3.9995e-05 | 0.8809 | 30.3625 | 8.5299e-05
LSSL2010#397 | CB-18 | 1308 |  | 2011-02-07 17:38 | 24C | 0.77695 | 2.2401e-06 | 3.9935e-05 | 0.77699 | 26.3996 | 8.5313e-05
LSSL2010#398 | CB-18 | 1309 |  | 2011-02-07 17:40 | 24C | 0.74564 | 2.4908e-06 | 4.0054e-05 | 0.74568 | 25.2209 | 9.394e-05
LSSL2010#399 |  | Ref CB-3 | 15 | 2011-02-07 17:42 | 24C | 0.9988 | 3.0364e-06 | 3.9995e-05 | 0.99884 | 34.9542 | 0.00011696 | 614-615
Calibration#5 |  | Cal |  | 2011-02-09 09:34 | 24C | 0.99997 | 2.7426e-06 | 1.9073e-06 | 0.99997 |  |  | Adjusted Rs to 5.54
LSSL2010#400 | CB-19 | 1310 |  | 2011-02-09 09:42 | 24C | 0.99883 | 1.1196e-06 | 1.967e-06 | 0.99883 | 34.954 | 4.5227e-05
LSSL2010#401 | CB-19 | 1311 |  | 2011-02-09 09:44 | 24C | 0.9988 | 3.0087e-06 | 2.0266e-06 | 0.9988 | 34.9528 | 0.0001183
LSSL2010#402 | CB-19 | 1312 |  | 2011-02-09 09:47 | 24C | 0.99869 | 2.2401e-06 | 1.967e-06 | 0.99869 | 34.9486 | 8.7077e-05
LSSL2010#403 | CB-19 | 1313 |  | 2011-02-09 09:50 | 24C | 0.99838 | 3.359e-06 | 2.0266e-06 | 0.99839 | 34.9364 | 0.00013469
LSSL2010#404 | CB-19 | 1314 |  | 2011-02-09 09:52 | 24C | 0.99774 | 4.4858e-06 | 2.0266e-06 | 0.99774 | 34.9112 | 0.00017536
LSSL2010#405 | CB-19 | 1315 |  | 2011-02-09 09:54 | 24C | 0.99683 | 3.1414e-06 | 1.967e-06 | 0.99683 | 34.8752 | 0.00012381
LSSL2010#406 | CB-19 | 1316 |  | 2011-02-09 09:56 | 24C | 0.99648 | 1.4046e-06 | 1.967e-06 | 0.99648 | 34.8614 | 5.4326e-05
LSSL2010#407 | CB-19 | 1317 |  | 2011-02-09 09:58 | 24C | 0.99616 | 2.4744e-06 | 2.0266e-06 | 0.99616 | 34.849 | 9.9792e-05
LSSL2010#408 | CB-19 | 1318 |  | 2011-02-09 10:01 | 24C | 0.9957 | 2.7432e-06 | 2.0266e-06 | 0.9957 | 34.8309 | 0.00010656
LSSL2010#409 | CB-19 | 1319 |  | 2011-02-09 10:03 | 24C | 0.99574 | 3.2141e-06 | 2.0266e-06 | 0.99575 | 34.8326 | 0.00012429
LSSL2010#410 | CB-19 | 1320 |  | 2011-02-09 10:04 | 24C | 0.9947 | 2.7425e-06 | 1.9073e-06 | 0.9947 | 34.7916 | 0.00011074
LSSL2010#411 | CB-19 | 1321 |  | 2011-02-09 10:07 | 24C | 0.99439 | 3.332e-08 | 2.0266e-06 | 0.99439 | 34.7795 | 0
LSSL2010#412 | CB-19 | 1322 |  | 2011-02-09 10:35 | 24C | 0.9926 | 2.8765e-06 | 2.0266e-06 | 0.9926 | 34.709 | 0.00011401
LSSL2010#413 | CB-19 | 1323 |  | 2011-02-09 10:37 | 24C | 0.98553 | 3.3594e-06 | 1.967e-06 | 0.98553 | 34.4312 | 0.00013308
LSSL2010#414 | CB-19 | 1324 |  | 2011-02-09 10:39 | 24C | 0.9796 | 3.2144e-06 | 1.9073e-06 | 0.9796 | 34.1987 | 0.000126
LSSL2010#415 | CB-19 | 1325 |  | 2011-02-09 10:41 | 24C | 0.96935 | 2.8904e-06 | 2.0266e-06 | 0.96935 | 33.797 | 0.00011519
LSSL2010#416 | CB-19 | 1326 |  | 2011-02-09 10:44 | 24C | 0.9569 | 2.2396e-06 | 1.9073e-06 | 0.9569 | 33.3101 | 8.7223e-05
LSSL2010#417 | CB-19 | 1327 |  | 2011-02-09 10:46 | 24C | 0.95563 | 4.2256e-06 | 2.0266e-06 | 0.95563 | 33.2607 | 0.00016598
LSSL2010#418 | CB-19 | 1328 |  | 2011-02-09 10:48 | 24C | 0.94142 | 2.7757e-06 | 2.0266e-06 | 0.94142 | 32.7065 | 0.00010656
LSSL2010#419 | CB-19 | 1329 |  | 2011-02-09 10:50 | 24C | 0.9399 | 3.2141e-06 | 1.967e-06 | 0.9399 | 32.6473 | 0.00012617
LSSL2010#420 | CB-19 | 1330 |  | 2011-02-09 10:52 | 24C | 0.92322 | 2.4909e-06 | 2.0266e-06 | 0.92322 | 31.999 | 9.9639e-05
LSSL2010#421 | CB-19 | 1331 |  | 2011-02-09 10:54 | 24C | 0.87199 | 1.1199e-06 | 2.0266e-06 | 0.87199 | 30.0198 | 4.4356e-05
LSSL2010#422 | CB-19 | 1332 |  | 2011-02-09 10:57 | 24C | 0.75973 | 5.4724e-06 | 2.0266e-06 | 0.75973 | 25.7492 | 0.00020379
LSSL2010#423 | CB-19 | 1333 |  | 2011-02-09 10:59 | 24C | 0.73492 | 1.3551e-06 | 1.967e-06 | 0.73492 | 24.8176 | 5.2235e-05
LSSL2010#424 | reference sample | CB-6 | 16 | 2011-02-09 11:06 | 24C | 0.99888 | 6.019e-06 | 1.967e-06 | 0.99888 | 34.956 | 0.00023831 | 1334-1335
LSSL2010#425 | CB-51 | 1478 | d | 2011-02-09 11:14 | 24C | 0.99872 | 3.0085e-06 | 2.0266e-06 | 0.99872 | 34.9497 | 0.00011691
LSSL2010#426 | CB-40 | 1433 | d | 2011-02-09 11:18 | 24C | 0.99786 | 4.2856e-06 | 1.967e-06 | 0.99786 | 34.9157 | 0.00016832
LSSL2010#427 | CB-50 | 1458 | d | 2011-02-09 11:22 | 24C | 0.99707 | 2.4909e-06 | 2.0266e-06 | 0.99707 | 34.8846 | 9.4938e-05
LSSL2010#428 | CB-52 | 1506 | d | 2011-02-09 11:25 | 24C | 0.99627 | 1.1529e-06 | 1.9073e-06 | 0.99628 | 34.8535 | 4.5227e-05
LSSL2010#429 | CB-40 | 1440 | d | 2011-02-09 11:27 | 24C | 0.99473 | 2.4908e-06 | 2.0266e-06 | 0.99473 | 34.7926 | 0.00010059
LSSL2010#430 | CB-53 | 1528 | d | 2011-02-09 11:30 | 24C | 0.9945 | 3.3588e-06 | 2.0266e-06 | 0.9945 | 34.7838 | 0.00013341
LSSL2010#431 | CB-52 | 1512 | d | 2011-02-09 11:33 | 24C | 0.99375 | 2.4915e-06 | 2.0266e-06 | 0.99376 | 34.7544 | 9.4885e-05
LSSL2010#432 | CB-51 | 1489 | d | 2011-02-09 11:35 | 24C | 0.99082 | 3.1418e-06 | 1.967e-06 | 0.99082 | 34.639 | 0.00012312
LSSL2010#433 | CB-50 | 1467 | d | 2011-02-09 11:37 | 24C | 0.98604 | 2.2396e-06 | 1.9073e-06 | 0.98604 | 34.4513 | 8.8203e-05
LSSL2010#434 | CB-54 | 1551 | d | 2011-02-09 11:40 | 24C | 0.9757 | 1.7597e-06 | 2.0266e-06 | 0.9757 | 34.0458 | 6.7435e-05
LSSL2010#435 | CB-52 | 1518 | d | 2011-02-09 11:42 | 24C | 0.93325 | 1.4046e-06 | 1.9073e-06 | 0.93325 | 32.3885 | 5.2235e-05
LSSL2010#436 | CB-15 | 1134 | d | 2011-02-09 11:44 | 24C | 0.95343 | 1.7026e-05 | 1.9073e-06 | 0.95343 | 33.1749 | 0.0006637
LSSL2010#437 | CB-53 | 1536 | d | 2011-02-09 11:46 | 24C | 0.95471 | 2.4911e-06 | 1.967e-06 | 0.95471 | 33.2247 | 9.4938e-05
LSSL2010#438 | CB-50 | 1471 | d | 2011-02-09 11:49 | 24C | 0.94526 | 2.8901e-06 | 2.0862e-06 | 0.94526 | 32.8561 | 0.00011213
LSSL2010#439 | CB-53 | 1542 | d | 2011-02-09 11:51 | 24C | 0.86055 | 2.2394e-06 | 1.967e-06 | 0.86055 | 29.5804 | 8.4887e-05
LSSL2010#440 | CB-54 | 1557 | d | 2011-02-09 11:53 | 24C | 0.86093 | 2.2126e-06 | 2.1458e-06 | 0.86093 | 29.5951 | 8.8712e-05
LSSL2010#441 | CB-51 | 1500 | d | 2011-02-09 11:56 | 24C | 0.83791 | 3.6105e-06 | 1.967e-06 | 0.83791 | 28.7137 | 0.00013841
LSSL2010#442 | reference sample | CB-6 | 17 | 2011-02-09 11:59 | 24C | 0.99887 | 1.7994e-06 | 2.0266e-06 | 0.99887 | 34.9556 | 7.2001e-05 | 1334-1335
LSSL2010#443 | CB-17 | 1143 | d | 2011-02-09 12:12 | 24C | 0.99889 | 2.8792e-06 | 2.0266e-06 | 0.99889 | 34.9563 | 0.00011317
LSSL2010#444 | CB-21 | 1410 | d | 2011-02-09 12:15 | 24C | 0.99785 | 3.4658e-06 | 1.967e-06 | 0.99785 | 34.9155 | 0.00013387
LSSL2010#445 | CB-17 | 1148 | d | 2011-02-09 12:17 | 24C | 0.9966 | 2.4908e-06 | 2.0266e-06 | 0.99661 | 34.8665 | 9.6642e-05
LSSL2010#446 | PP-7 | 1196 | d | 2011-02-09 12:20 | 24C | 0.9967 | 4.9165e-06 | 2.0266e-06 | 0.9967 | 34.8702 | 0.00019397
LSSL2010#447 | CB-15 | 1127 | d | 2011-02-09 12:22 | 24C | 0.99574 | 1.3713e-06 | 1.967e-06 | 0.99574 | 34.8323 | 5.4326e-05
LSSL2010#448 | CB-15 | 1129 | d | 2011-02-09 12:24 | 24C | 0.99429 | 2.4909e-06 | 2.0266e-06 | 0.9943 | 34.7756 | 0.00010437
LSSL2010#449 | CB-15 | 1076 | d | 2011-02-09 12:30 | 24C | 0.97279 | 2.4909e-06 | 1.9073e-06 | 0.97279 | 33.9317 | 9.7029e-05
LSSL2010#450 | CB-40 | 1447 | d | 2011-02-09 12:33 | 24C | 0.96254 | 1.1196e-06 | 2.0266e-06 | 0.96254 | 33.5307 | 4.265e-05
LSSL2010#451 | CB-15 | 1085 | d | 2011-02-09 12:35 | 24C | 0.93015 | 2.7262e-06 | 2.0266e-06 | 0.93015 | 32.2681 | 0.00010659
LSSL2010#452 | CB-21 | 1425 | d | 2011-02-09 12:38 | 24C | 0.92289 | 4.9162e-06 | 2.0266e-06 | 0.92289 | 31.986 | 0.00018936
LSSL2010#453 | CB-17 | 1162 | d | 2011-02-09 12:41 | 24C | 0.91951 | 3.4666e-06 | 1.967e-06 | 0.91951 | 31.855 | 0.00013693
LSSL2010#454 | CB21 | 1429 | d | 2011-02-09 12:43 | 24C | 0.74066 | 2.2391e-06 | 1.967e-06 | 0.74066 | 25.0329 | 8.872e-05
LSSL2010#455 | reference sample | CB-6 | 18 | 2011-02-09 12:47 | 24C | 0.9989 | 3.4657e-06 | 2.0266e-06 | 0.9989 | 34.9568 | 0.0001355 | 1334-1335
LSSL2010#456 | CB-70 | 806 | d | 2011-02-09 12:55 | 24C | 0.99892 | 2.2397e-06 | 2.0266e-06 | 0.99892 | 34.9577 | 8.7013e-05
LSSL2010#457 | CB-70 | 810 | d | 2011-02-09 12:57 | 24C | 0.99751 | 1.1529e-06 | 1.9073e-06 | 0.99751 | 34.9021 | 4.265e-05
LSSL2010#458 | CB-8 | 739 | d | 2011-02-09 12:59 | 24C | 0.99675 | 7.2738e-06 | 1.967e-06 | 0.99675 | 34.8723 | 0.00028709
LSSL2010#459 | CB-11 | 859 | d | 2011-02-09 13:01 | 24C | 0.99677 | 3.1145e-06 | 2.0862e-06 | 0.99678 | 34.8731 | 0.00011967
LSSL2010#460 | CB-70 | 811 | d | 2011-02-09 13:03 | 24C | 0.99677 | 2.8661e-06 | 2.0266e-06 | 0.99678 | 34.8731 | 0.00011014
LSSL2010#461 | CB-10 | 789 | d | 2011-02-09 13:06 | 24C | 0.9963 | 3.7354e-06 | 2.0862e-06 | 0.9963 | 34.8545 | 0.00014444
LSSL2010#462 | CB-9 | 909 | d | 2011-02-09 13:08 | 24C | 0.99624 | 3.9994e-06 | 2.0266e-06 | 0.99625 | 34.8523 | 0.00015603
LSSL2010#463 | CB-8 | 743 | d | 2011-02-09 13:10 | 24C | 0.99582 | 2.2394e-06 | 2.0266e-06 | 0.99582 | 34.8356 | 8.793e-05
LSSL2010#464 | CB-7 | 719 | d | 2011-02-09 14:49 | 24C | 0.99568 | 4.1469e-06 | 2.1458e-06 | 0.99568 | 34.8301 | 0.00016043
LSSL2010#465 | CB-11 | 863 | d | 2011-02-09 14:52 | 24C | 0.99573 | 3.1418e-06 | 1.967e-06 | 0.99574 | 34.8322 | 0.00012381
LSSL2010#466 | CB-70 | 830 | d | 2011-02-09 14:54 | 24C | 0.99461 | 3.4657e-06 | 2.0266e-06 | 0.99462 | 34.7882 | 0.00013904
LSSL2010#467 | CB-70 | 836 | d | 2011-02-09 14:56 | 24C | 0.98182 | 3.4658e-06 | 2.0266e-06 | 0.98182 | 34.2856 | 0.00013595
LSSL2010#468 | CB-10 | 795 | d | 2011-02-09 14:58 | 24C | 0.98012 | 3.6105e-06 | 2.0266e-06 | 0.98012 | 34.2189 | 0.0001408
LSSL2010#469 | CB70 | 819 | d | 2011-02-09 15:00 | 24C | 0.97841 | 2.2394e-06 | 1.9073e-06 | 0.97841 | 34.1518 | 8.7024e-05
LSSL2010#470 | CB-11 | 870 | d | 2011-02-09 15:02 | 24C | 0.95334 | 2.7425e-06 | 1.9073e-06 | 0.95334 | 33.171 | 0.00010761
LSSL2010#471 | CB-70 | 821 | d | 2011-02-09 15:04 | 24C | 0.95214 | 3.4359e-06 | 2.0862e-06 | 0.95214 | 33.1243 | 0.00013203
LSSL2010#472 | CB-8 | 750 | d | 2011-02-09 15:07 | 24C | 0.94997 | 3.6104e-06 | 2.0266e-06 | 0.94997 | 33.0396 | 0.00013925
LSSL2010#473 | CB-7 | 726 | d | 2011-02-09 15:10 | 24C | 0.9563 | 2.4915e-06 | 2.0266e-06 | 0.9563 | 33.2868 | 9.4981e-05
LSSL2010#474 | CB-70 | 845 | d | 2011-02-09 15:12 | 24C | 0.93661 | 2.4916e-06 | 2.0266e-06 | 0.93661 | 32.5192 | 9.6976e-05
LSSL2010#475 | CB10 | 800 | d | 2011-02-09 15:14 | 24C | 0.93642 | 3.1314e-06 | 1.967e-06 | 0.93642 | 32.5118 | 0.00012312
LSSL2010#476 | reference sample | CB-6 | 19 | 2011-02-09 15:17 | 24C | 0.9989 | 5.0527e-06 | 2.0266e-06 | 0.9989 | 34.9566 | 0.00019929 | 1334-1335
LSSL2010#477 | CB-6 | 1334 | d | 2011-02-09 15:39 | 24C | 0.99885 | 2.4915e-06 | 1.967e-06 | 0.99885 | 34.9549 | 0.00010059
LSSL2010#478 | CB-19 | 1312 | d | 2011-02-09 15:42 | 24C | 0.99884 | 2.0804e-06 | 2.0862e-06 | 0.99884 | 34.9544 | 8.1621e-05
LSSL2010#479 | PP-6 | 1240 | d | 2011-02-09 15:44 | 24C | 0.99853 | 2.2126e-06 | 2.0862e-06 | 0.99853 | 34.9422 | 8.8728e-05
LSSL2010#480 | CB-DW | 1264 | d | 2011-02-09 15:46 | 24C | 0.99792 | 1.3882e-06 | 2.0862e-06 | 0.99792 | 34.9182 | 5.2278e-05
LSSL2010#481 | CB-DW | 1265 | d | 2011-02-09 15:48 | 24C | 0.99708 | 0 | 2.0862e-06 | 0.99708 | 34.8853 | 0
LSSL2010#482 | CB-18 | 1289 | d | 2011-02-09 15:50 | 24C | 0.99695 | 5.2144e-06 | 2.0266e-06 | 0.99695 | 34.88 | 0.00020293
LSSL2010#483 | CB-18 | 1293 | d | 2011-02-09 15:53 | 24C | 0.99575 | 2.4908e-06 | 2.0266e-06 | 0.99576 | 34.833 | 9.9734e-05
LSSL2010#484 | CB-19 | 1319 | d | 2011-02-09 15:55 | 24C | 0.99573 | 2.7264e-06 | 2.0266e-06 | 0.99573 | 34.8319 | 0.00010344
LSSL2010#485 | PP-7 | 1204 | d | 2011-02-09 15:58 | 24C | 0.97918 | 2.7425e-06 | 2.0266e-06 | 0.97918 | 34.182 | 0.00010865
LSSL2010#486 | CB-21 | 1386 | d | 2011-02-09 16:00 | 24C | 0.97834 | 3.131e-06 | 2.0266e-06 | 0.97834 | 34.1491 | 0.00012172
LSSL2010#487 | CB-DW | 1275 | d | 2011-02-09 16:03 | 24C | 0.97482 | 2.4911e-06 | 2.0266e-06 | 0.97482 | 34.0113 | 9.869e-05
LSSL2010#488 | CB-18 | 1300 | d | 2011-02-09 16:05 | 24C | 0.96649 | 3.1905e-06 | 2.0266e-06 | 0.96649 | 33.6851 | 0.00012553
LSSL2010#489 | PP-6 | 1253 | d | 2011-02-09 16:07 | 24C | 0.95439 | 2.4909e-06 | 2.0266e-06 | 0.9544 | 33.2124 | 9.4981e-05
LSSL2010#490 | CB-DW | 1277 | d | 2011-02-09 16:09 | 24C | 0.95025 | 2.866e-06 | 2.0862e-06 | 0.95025 | 33.0508 | 0.00011317
LSSL2010#491 | CB-6 | 1351 | d | 2011-02-09 16:12 | 24C | 0.93939 | 3.1418e-06 | 1.9073e-06 | 0.9394 | 32.6276 | 0.00012486
LSSL2010#492 | CB-DW | 1279 | d | 2011-02-09 16:14 | 24C | 0.93916 | 1.1203e-06 | 2.0862e-06 | 0.93916 | 32.6184 | 4.4364e-05
LSSL2010#493 | PP-6 | 1256 | d | 2011-02-09 16:16 | 24C | 0.93244 | 3.1417e-06 | 2.0266e-06 | 0.93245 | 32.3572 | 0.00012313
LSSL2010#494 | CB-DW | 1281 | d | 2011-02-09 16:19 | 24C | 0.91977 | 1.1196e-06 | 2.0266e-06 | 0.91977 | 31.8649 | 4.3729e-05
LSSL2010#495 | CB-21 | 1401 | d | 2011-02-09 16:21 | 24C | 0.88377 | 3.1418e-06 | 1.967e-06 | 0.88377 | 30.473 | 0.00012365
LSSL2010#496 | PP-7 | 1213 | d | 2011-02-09 16:23 | 24C | 0.77691 | 1.1196e-06 | 1.967e-06 | 0.77691 | 26.3966 | 4.3716e-05
LSSL2010#497 | CB-DW | 1285 | d | 2011-02-09 16:26 | 24C | 0.76381 | 0 | 2.0266e-06 | 0.76381 | 25.9028 | 1.0662e-06
LSSL2010#498 | CB-21 | 1405 | d | 2011-02-09 16:29 | 24C | 0.74217 | 2.2135e-06 | 2.0862e-06 | 0.74217 | 25.0894 | 8.5299e-05
LSSL2010#499 | CB-19 | 1333 | d | 2011-02-09 16:31 | 24C | 0.73494 | 3.3588e-06 | 1.967e-06 | 0.73494 | 24.8184 | 0.0001271
LSSL2010#500 | reference sample | CB-6 | 20 | 2011-02-09 16:36 | 24C | 0.99891 | 2.4912e-06 | 1.967e-06 | 0.99891 | 34.9572 | 9.6609e-05 | 1334-1335
LSSL2010#501 | CB-13 | 1047 | d | 2011-02-09 16:49 | 24C | 0.99887 | 2.8543e-06 | 1.967e-06 | 0.99887 | 34.9556 | 0.00011192
LSSL2010#502 | CB-16 | 1003 | d | 2011-02-09 16:57 | 24C | 0.997 | 2.8907e-06 | 2.1458e-06 | 0.997 | 34.8821 | 0.00011519
LSSL2010#503 | CB-12 | 957 | d | 2011-02-09 16:59 | 24C | 0.99432 | 2.4908e-06 | 2.0266e-06 | 0.99432 | 34.7767 | 9.4885e-05
LSSL2010#504 | CB-13 | 1055 | d | 2011-02-09 17:02 | 24C | 0.99643 | 2.4908e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 9.9543e-05
LSSL2010#505 | CB-16 | 1007 | d | 2011-02-09 17:04 | 24C | 0.99575 | 1.3712e-06 | 1.9073e-06 | 0.99576 | 34.833 | 5.5371e-05
LSSL2010#506 | CB-15 | 1070 | d | 2011-02-09 17:06 | 24C | 0.99425 | 2.7425e-06 | 1.9073e-06 | 0.99425 | 34.7739 | 0.00010663
LSSL2010#507 | CB-12a | 974 | d | 2011-02-09 17:09 | 24C | 0.99344 | 2.4751e-06 | 2.0266e-06 | 0.99344 | 34.7418 | 9.9792e-05
LSSL2010#508 | CB-12 | 963 | d | 2011-02-09 17:12 | 24C | 0.9826 | 1.4046e-06 | 1.967e-06 | 0.9826 | 34.3163 | 5.2235e-05
LSSL2010#509 | CB-9 | 915 | d | 2011-02-09 17:14 | 24C | 0.97714 | 2.89e-06 | 1.967e-06 | 0.97715 | 34.1023 | 0.00011317
LSSL2010#510 | CB-12a | 980 | d | 2011-02-09 17:17 | 24C | 0.97401 | 1.1536e-06 | 2.0266e-06 | 0.97401 | 33.9794 | 4.2703e-05
LSSL2010#511 | CB-16 | 1014 | d | 2011-02-09 17:18 | 24C | 0.95357 | 3.6119e-06 | 2.1458e-06 | 0.95357 | 33.1801 | 0.00013908
LSSL2010#512 | CB-12 | 968 | d | 2011-02-09 17:21 | 24C | 0.93818 | 2.0945e-06 | 1.967e-06 | 0.93818 | 32.5802 | 8.1861e-05
LSSL2010#513 | CB-9 | 920 | d | 2011-02-09 17:23 | 24C | 0.93825 | 2.4908e-06 | 1.967e-06 | 0.93825 | 32.583 | 9.869e-05
LSSL2010#514 | CB-12a | 989 | d | 2011-02-09 17:26 | 24C | 0.93005 | 3.1424e-06 | 2.0266e-06 | 0.93005 | 32.264 | 0.00012176
LSSL2010#515 | CB-9 | 923 | d | 2011-02-09 17:28 | 24C | 0.89184 | 3.2141e-06 | 2.0862e-06 | 0.89185 | 30.7846 | 0.00012629
LSSL2010#516 | CB-13 | 1068 | d | 2011-02-09 17:30 | 24C | 0.83527 | 3.6104e-06 | 2.0266e-06 | 0.83527 | 28.6127 | 0.00013717
LSSL2010#517 | reference sample | CB-6 | 21 | 2011-02-09 17:32 | 24C | 0.99888 | 1.6614e-06 | 1.967e-06 | 0.99888 | 34.9561 | 6.5793e-05 | 1334-1335
Calibration#6 |  | K15 |  | 2011-02-10 11:03 | 24C | 0.99995 | 2.2391e-06 | 2.0981e-05 | 0.99997
Calibration#7 |  | K15 |  | 2011-02-10 11:17 | 24C | 0.99997 | 4.4784e-06 | 2.0266e-06 | 0.99997
LSSL2010#518 | CB-13 | 1022 |  | 2011-02-10 11:22 | 24C | 0.9969 | 2.2391e-06 | 2.0266e-06 | 0.9969 | 34.8779 | 8.7024e-05
LSSL2010#519 | CB-13 | 1023 |  | 2011-02-10 11:24 | 24C | 0.99688 | 2.4909e-06 | 2.0862e-06 | 0.99689 | 34.8774 | 9.869e-05
LSSL2010#520 | CB-13 | 1024 |  | 2011-02-10 11:27 | 24C | 0.99676 | 3.2144e-06 | 1.967e-06 | 0.99676 | 34.8724 | 0.00012553
LSSL2010#521 | CB-13 | 1025 |  | 2011-02-10 11:30 | 24C | 0.99685 | 2.7425e-06 | 1.967e-06 | 0.99685 | 34.876 | 0.00010865
LSSL2010#522 | CB-13 | 1026 |  | 2011-02-10 11:32 | 24C | 0.99655 | 1.1396e-06 | 2.0862e-06 | 0.99655 | 34.8643 | 4.2703e-05
LSSL2010#523 | CB-13 | 1027 |  | 2011-02-10 11:35 | 24C | 0.99653 | 2.0947e-06 | 2.0862e-06 | 0.99653 | 34.8636 | 8.1174e-05
LSSL2010#524 | CB-13 | 1028 |  | 2011-02-10 11:37 | 24C | 0.99637 | 2.2391e-06 | 1.967e-06 | 0.99637 | 34.8572 | 8.5352e-05
LSSL2010#525 | CB-13 | 1029 |  | 2011-02-10 11:39 | 24C | 0.99643 | 3.6105e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 0.00014239
LSSL2010#526 | CB-13 | 1030 |  | 2011-02-10 11:41 | 24C | 0.99617 | 2.2394e-06 | 2.1458e-06 | 0.99618 | 34.8496 | 8.959e-05
LSSL2010#527 | CB-13 | 1031 |  | 2011-02-10 11:43 | 24C | 0.99615 | 3.344e-06 | 2.0862e-06 | 0.99616 | 34.8488 | 0.00013404
LSSL2010#528 | CB-13 | 1032 |  | 2011-02-10 11:45 | 24C | 0.9961 | 2.7426e-06 | 1.967e-06 | 0.9961 | 34.8465 | 0.00010865
LSSL2010#529 | CB-13 | 1033 |  | 2011-02-10 11:47 | 24C | 0.99617 | 2.8897e-06 | 2.0862e-06 | 0.99618 | 34.8496 | 0.00011517
LSSL2010#530 | CB-13 | 1034 |  | 2011-02-10 11:50 | 24C | 0.99486 | 2.4908e-06 | 2.0266e-06 | 0.99486 | 34.7979 | 9.557e-05
LSSL2010#531 | CB-13 | 1035 |  | 2011-02-10 11:52 | 24C | 0.99486 | 5.635e-06 | 1.967e-06 | 0.99486 | 34.7979 | 0.00021988
LSSL2010#532 | CB-13 | 1037 |  | 2011-02-10 11:54 | 24C | 0.99495 | 2.4908e-06 | 1.967e-06 | 0.99496 | 34.8016 | 9.7246e-05
LSSL2010#533 | CB-13 | 1037 |  | 2011-02-10 11:56 | 24C | 0.99504 | 2.2391e-06 | 1.967e-06 | 0.99504 | 34.8048 | 8.7866e-05
LSSL2010#534 | CB-13 | 1038 |  | 2011-02-10 11:59 | 24C | 0.98992 | 2.8795e-06 | 2.0266e-06 | 0.98993 | 34.6038 | 0.00011316
LSSL2010#535 | CB-13 | 1039 |  | 2011-02-10 12:01 | 24C | 0.98993 | 2.8661e-06 | 2.0266e-06 | 0.98993 | 34.6041 | 0.00011012
LSSL2010#536 | CB-13 | 1040 |  | 2011-02-10 12:03 | 24C | 0.98969 | 2.4912e-06 | 2.0266e-06 | 0.98969 | 34.5946 | 9.6599e-05
LSSL2010#537 | CB-13 | 1041 |  | 2011-02-10 12:05 | 24C | 0.9897 | 1.1196e-06 | 2.0266e-06 | 0.9897 | 34.595 | 4.265e-05
LSSL2010#538 | CB-13 | 1042 |  | 2011-02-10 12:08 | 24C | 0.79746 | 2.2391e-06 | 1.967e-06 | 0.79746 | 27.1742 | 8.4455e-05
LSSL2010#539 | CB-13 | 1043 |  | 2011-02-10 12:10 | 24C | 0.79753 | 2.7426e-06 | 2.0862e-06 | 0.79753 | 27.1768 | 0.00010449
LSSL2010#540 | CB-13 | 1044 |  | 2011-02-10 12:12 | 24C | 0.79755 | 2.4908e-06 | 2.0266e-06 | 0.79755 | 27.1776 | 9.4895e-05
LSSL2010#541 | CB-13 | 1045 |  | 2011-02-10 12:14 | 24C | 0.79755 | 2.2259e-06 | 2.0862e-06 | 0.79755 | 27.1778 | 8.2741e-05
LSSL2010#542 | reference sample | CB-6 | 22 | 2011-02-10 12:16 | 24C | 0.99892 | 3.0409e-06 | 1.9073e-06 | 0.99892 | 34.9575 | 0.00012299
LSSL2010#543 | PP-7 | 1190 |  | 2011-02-10 12:55 | 24C | 0.99886 | 3.466e-06 | 2.0266e-06 | 0.99886 | 34.9552 | 0.0001355
LSSL2010#544 | PP-7 | 1191 |  | 2011-02-10 12:59 | 24C | 0.99887 | 3.4658e-06 | 1.967e-06 | 0.99887 | 34.9557 | 0.0001334
LSSL2010#545 | PP-7 | 1192 |  | 2011-02-10 13:01 | 24C | 0.9988 | 1.3712e-06 | 2.0266e-06 | 0.9988 | 34.9529 | 5.4326e-05
LSSL2010#546 | PP-7 | 1193 |  | 2011-02-10 13:04 | 24C | 0.99855 | 3.466e-06 | 1.967e-06 | 0.99855 | 34.9428 | 0.00013822
LSSL2010#547 | PP-7 | 1194 |  | 2011-02-10 13:05 | 24C | 0.99789 | 3.4657e-06 | 1.967e-06 | 0.99789 | 34.9171 | 0.00013595
LSSL2010#548 | PP-7 | 1195 |  | 2011-02-10 13:09 | 24C | 0.99702 | 3.4553e-06 | 1.967e-06 | 0.99703 | 34.883 | 0.00013731
LSSL2010#549 | PP-7 | 1196 |  | 2011-02-10 13:11 | 24C | 0.99667 | 3.1414e-06 | 1.967e-06 | 0.99667 | 34.8689 | 0.00012312
LSSL2010#550 | PP-7 | 1197 |  | 2011-02-10 13:13 | 24C | 0.99634 | 1.1532e-06 | 2.1458e-06 | 0.99635 | 34.8562 | 4.4374e-05
LSSL2010#551 | PP-7 | 1198 |  | 2011-02-10 13:15 | 24C | 0.99616 | 3.3304e-06 | 2.0266e-06 | 0.99616 | 34.8489 | 0.0001349
LSSL2010#552 | PP-7 | 1199 |  | 2011-02-10 13:17 | 24C | 0.99568 | 2.4777e-06 | 1.9073e-06 | 0.99568 | 34.8301 | 9.9759e-05
LSSL2010#553 | PP-7 | 1200 |  | 2011-02-10 13:19 | 24C | 0.99537 | 2.7425e-06 | 2.0266e-06 | 0.99537 | 34.818 | 0.00010865
LSSL2010#554 | PP-7 | 1201 |  | 2011-02-10 13:21 | 24C | 0.9944 | 2.7264e-06 | 2.0266e-06 | 0.9944 | 34.7798 | 0.00010656
LSSL2010#555 | PP-7 | 1202 |  | 2011-02-10 13:23 | 24C | 0.98973 | 3.6105e-06 | 2.0266e-06 | 0.98973 | 34.5963 | 0.00014356
LSSL2010#556 | PP-7 | 1203 |  | 2011-02-10 13:25 | 24C | 0.98361 | 2.7426e-06 | 2.0266e-06 | 0.98362 | 34.356 | 0.00010451
LSSL2010#557 | PP-7 | 1204 |  | 2011-02-10 13:27 | 24C | 0.9791 | 3.3594e-06 | 2.0266e-06 | 0.9791 | 34.1788 | 0.00013117
LSSL2010#558 | PP-7 | 1205 |  | 2011-02-10 13:30 | 24C | 0.96774 | 3.2141e-06 | 2.0266e-06 | 0.96774 | 33.7339 | 0.00012528
LSSL2010#559 | PP-7 | 1206 |  | 2011-02-10 13:32 | 24C | 0.95254 | 2.4908e-06 | 2.0266e-06 | 0.95254 | 33.1401 | 9.4885e-05
LSSL2010#560 | PP-7 | 1207 |  | 2011-02-10 13:34 | 24C | 0.94337 | 2.2401e-06 | 2.0862e-06 | 0.94337 | 32.7825 | 8.7876e-05
LSSL2010#561 | PP-7 | 1208 |  | 2011-02-10 13:36 | 24C | 0.93792 | 2.226e-06 | 2.0266e-06 | 0.93792 | 32.5701 | 8.7876e-05
LSSL2010#562 | PP-7 | 1209 |  | 2011-02-10 13:38 | 24C | 0.93129 | 4.9004e-06 | 2.0266e-06 | 0.9313 | 32.3125 | 0.00018835
LSSL2010#563 | PP-7 | 1210 |  | 2011-02-10 13:40 | 24C | 0.91919 | 3.4657e-06 | 2.0266e-06 | 0.91919 | 31.8425 | 0.00013331
LSSL2010#564 | PP-7 | 1211 |  | 2011-02-10 13:46 | 24C | 0.83636 | 3.1311e-06 | 1.967e-06 | 0.83636 | 28.6543 | 0.00011968
LSSL2010#565 | PP-7 | 1212 |  | 2011-02-10 13:48 | 24C | 0.82699 | 2.7425e-06 | 2.0266e-06 | 0.82699 | 28.2969 | 0.00010395
LSSL2010#566 | PP-7 | 1213 |  | 2011-02-10 13:50 | 24C | 0.77688 | 2.4908e-06 | 2.0266e-06 | 0.77688 | 26.3955 | 9.1847e-05
LSSL2010#567 | reference sample | CB-6 | 23 | 2011-02-10 13:53 | 24C | 0.99897 | 2.9667e-06 | 2.0862e-06 | 0.99897 | 34.9596 | 0.00011473
LSSL2010#568 | CB-15 | 1070 |  | 2011-02-10 13:59 | 24C | 0.99427 | 2.4909e-06 | 1.967e-06 | 0.99427 | 34.7745 | 9.869e-05
LSSL2010#569 | CB-15 | 1071 |  | 2011-02-10 14:02 | 24C | 0.9917 | 3.2144e-06 | 1.967e-06 | 0.9917 | 34.6734 | 0.00012623
LSSL2010#570 | CB-15 | 1072 |  | 2011-02-10 14:04 | 24C | 0.9896 | 3.3587e-06 | 1.9073e-06 | 0.9896 | 34.5912 | 0.00012927
LSSL2010#571 | CB-15 | 1073 |  | 2011-02-10 14:06 | 24C | 0.98732 | 2.128e-06 | 1.9073e-06 | 0.98732 | 34.5016 | 8.1621e-05
LSSL2010#572 | CB-15 | 1074 |  | 2011-02-10 14:08 | 24C | 0.9843 | 2.4909e-06 | 1.967e-06 | 0.9843 | 34.3829 | 9.87e-05
LSSL2010#573 | CB-15 | 1075 |  | 2011-02-10 14:13 | 24C | 0.98072 | 2.2355e-05 | 2.0266e-06 | 0.98072 | 34.2426 | 0.000875
LSSL2010#574 | CB-15 | 1076 |  | 2011-02-10 14:59 | 24C | 0.9727 | 3.1415e-06 | 1.967e-06 | 0.9727 | 33.9281 | 0.00012621
LSSL2010#575 | CB-15 | 1077 |  | 2011-02-10 15:02 | 24C | 0.96704 | 2.8897e-06 | 2.0266e-06 | 0.96704 | 33.7065 | 0.00011386
LSSL2010#576 | CB-15 | 1078 |  | 2011-02-10 15:04 | 24C | 0.9613 | 3.2145e-06 | 1.967e-06 | 0.9613 | 33.482 | 0.000126
LSSL2010#577 | CB-15 | 1079 |  | 2011-02-10 15:07 | 24C | 0.95605 | 1.1196e-06 | 2.0266e-06 | 0.95605 | 33.2769 | 4.2703e-05
LSSL2010#578 | CB-15 | 1080 |  | 2011-02-10 15:09 | 24C | 0.95287 | 2.7425e-06 | 2.0266e-06 | 0.95287 | 33.1527 | 0.00010343
LSSL2010#579 | CB-15 | 1081 |  | 2011-02-10 15:11 | 24C | 0.94946 | 2.4908e-06 | 2.0266e-06 | 0.94946 | 33.0199 | 9.87e-05
LSSL2010#580 | CB-15 | 1082 |  | 2011-02-10 15:14 | 24C | 0.9466 | 2.8897e-06 | 1.967e-06 | 0.9466 | 32.9082 | 0.00011145
LSSL2010#581 | CB-15 | 1083 |  | 2011-02-10 15:16 | 24C | 0.94145 | 2.4912e-06 | 1.967e-06 | 0.94145 | 32.7076 | 9.4885e-05
LSSL2010#582 | CB-15 | 1084 |  | 2011-02-10 15:18 | 24C | 0.93589 | 2.8904e-06 | 1.9073e-06 | 0.93589 | 32.4911 | 0.00011403
LSSL2010#583 | CB-15 | 1085 |  | 2011-02-10 15:34 | 24C | 0.93012 | 2.4911e-06 | 2.0862e-06 | 0.93012 | 32.2668 | 9.9553e-05
LSSL2010#584 | CB-15 | 1086 |  | 2011-02-10 15:36 | 24C | 0.9283 | 3.1148e-06 | 2.0266e-06 | 0.92831 | 32.1963 | 0.00012313
LSSL2010#585 | CB-15 | 1087 |  | 2011-02-10 15:38 | 24C | 0.92934 | 3.3591e-06 | 1.9073e-06 | 0.92935 | 32.2367 | 0.00012795
LSSL2010#586 | CB-15 | 1088 |  | 2011-02-10 15:41 | 24C | 0.91925 | 2.7425e-06 | 2.0266e-06 | 0.91925 | 31.8449 | 0.00010447
LSSL2010#587 | CB-15 | 1089 |  | 2011-02-10 15:42 | 24C | 0.91104 | 4.2073e-06 | 2.0266e-06 | 0.91104 | 31.5266 | 0.00016433
LSSL2010#588 | CB-15 | 1090 |  | 2011-02-10 15:45 | 24C | 0.88018 | 4.1136e-06 | 2.0266e-06 | 0.88018 | 30.3349 | 0.00015711
LSSL2010#589 | CB-15 | 1091 |  | 2011-02-10 15:47 | 24C | 0.86349 | 3.9638e-06 | 2.0862e-06 | 0.86349 | 29.6931 | 0.0001517
LSSL2010#590 | CB-15 | 1092 |  | 2011-02-10 15:49 | 24C | 0.83638 | 4.2256e-06 | 2.0266e-06 | 0.83638 | 28.6552 | 0.00016216
LSSL2010#591 | CB-15 | 1093 |  | 2011-02-10 15:51 | 24C | 0.79389 | 3.2141e-06 | 2.0266e-06 | 0.79389 | 27.039 | 0.0001238
LSSL2010#592 | reference sample | CB-6 | 24 | 2011-02-10 15:54 | 24C | 0.99893 | 1.6611e-06 | 2.0266e-06 | 0.99894 | 34.9581 | 6.4399e-05
LSSL2010#593 | CB-12a | 974 |  | 2011-02-10 15:59 | 24C | 0.99343 | 3.3459e-06 | 2.0862e-06 | 0.99343 | 34.7417 | 0.00013202
LSSL2010#594 | CB-12a | 975 |  | 2011-02-10 16:01 | 24C | 0.99238 | 4.189e-06 | 1.9073e-06 | 0.99238 | 34.7003 | 0.00016533
LSSL2010#595 | CB-12a | 976 |  | 2011-02-10 16:03 | 24C | 0.98988 | 2.4776e-06 | 2.0266e-06 | 0.98988 | 34.6022 | 9.6976e-05
LSSL2010#596 | CB-12a | 977 |  | 2011-02-10 16:05 | 24C | 0.98898 | 4.3836e-06 | 2.0266e-06 | 0.98898 | 34.5668 | 0.00017346
LSSL2010#597 | CB-12a | 978 |  | 2011-02-10 16:07 | 24C | 0.98419 | 2.2394e-06 | 1.967e-06 | 0.98419 | 34.3787 | 8.7067e-05
LSSL2010#598 | CB-12a | 979 |  | 2011-02-10 16:09 | 24C | 0.98092 | 4.2164e-06 | 1.967e-06 | 0.98092 | 34.2505 | 0.00016491
LSSL2010#599 | CB-12a | 980 |  | 2011-02-10 16:11 | 24C | 0.97357 | 3.8756e-06 | 2.0862e-06 | 0.97358 | 33.9624 | 0.0001485
LSSL2010#600 | CB-12a | 981 |  | 2011-02-10 16:13 | 24C | 0.96653 | 3.4352e-06 | 2.0862e-06 | 0.96653 | 33.6867 | 0.0001341
LSSL2010#601 | CB-12a | 982 |  | 2011-02-10 16:16 | 24C | 0.96233 | 2.89e-06 | 2.0266e-06 | 0.96233 | 33.5225 | 0.00011316
LSSL2010#602 | CB-12a | 983 |  | 2011-02-10 16:18 | 24C | 0.95714 | 6.664e-08 | 2.0266e-06 | 0.95714 | 33.3197 | 0
LSSL2010#603 | CB-12a | 984 |  | 2011-02-10 16:20 | 24C | 0.9541 | 3.6105e-06 | 2.0266e-06 | 0.9541 | 33.2007 | 0.00014113
LSSL2010#604 | CB-12a | 985 |  | 2011-02-10 16:22 | 24C | 0.95102 | 3.8747e-06 | 2.0862e-06 | 0.95103 | 33.0808 | 0.00015161
LSSL2010#605 | CB-12a | 986 |  | 2011-02-10 16:24 | 24C | 0.94902 | 4.6041e-06 | 2.0266e-06 | 0.94903 | 33.0028 | 0.00018033
LSSL2010#606 | CB-12a | 987 |  | 2011-02-10 16:26 | 24C | 0.94293 | 4.9162e-06 | 2.0266e-06 | 0.94293 | 32.7654 | 0.00019131
LSSL2010#607 | CB-12a | 988 |  | 2011-02-10 16:27 | 24C | 0.93753 | 2.7425e-06 | 2.0266e-06 | 0.93753 | 32.555 | 0.00010865
LSSL2010#608 | CB-12a | 989 |  | 2011-02-10 16:29 | 24C | 0.93018 | 3.4657e-06 | 2.0266e-06 | 0.93018 | 32.2692 | 0.0001355
LSSL2010#609 | CB-12a | 990 |  | 2011-02-10 16:31 | 24C | 0.92547 | 4.0099e-06 | 2.1458e-06 | 0.92547 | 32.0862 | 0.00015276
LSSL2010#610 | CB-12a | 991 |  | 2011-02-10 16:33 | 24C | 0.92221 | 3.5406e-06 | 2.0266e-06 | 0.92221 | 31.9598 | 0.00013625
LSSL2010#611 | CB-12a | 992 |  | 2011-02-10 16:35 | 24C | 0.91164 | 1.7702e-06 | 1.967e-06 | 0.91164 | 31.5499 | 6.9141e-05
LSSL2010#612 | CB-12a | 993 |  | 2011-02-10 16:37 | 24C | 0.90407 | 3.3591e-06 | 2.0862e-06 | 0.90408 | 31.257 | 0.00012893
LSSL2010#613 | CB-12a | 994 |  | 2011-02-10 16:39 | 24C | 0.89295 | 3.1417e-06 | 1.967e-06 | 0.89296 | 30.8274 | 0.00012086
LSSL2010#614 | CB-12a | 995 |  | 2011-02-10 16:42 | 24C | 0.87458 | 3.4515e-06 | 2.0862e-06 | 0.87458 | 30.1194 | 0.00013157
LSSL2010#615 | CB-12a | 996 |  | 2011-02-10 16:44 | 24C | 0.8196 | 2.2394e-06 | 1.967e-06 | 0.81961 | 28.0155 | 8.5313e-05
LSSL2010#616 | CB-12a | 997 |  | 2011-02-10 16:46 | 24C | 0.78066 | 2.4908e-06 | 2.0862e-06 | 0.78067 | 26.5386 | 9.4478e-05
LSSL2010#617 | reference sample | CB-6 | 25 | 2011-02-10 16:48 | 24C | 0.99887 | 4.276e-06 | 2.0862e-06 | 0.99888 | 34.9558 | 0.00016853
LSSL2010#618 | CB-21 | 1382 |  | 2011-02-10 16:54 | 24C | 0.99053 | 3.3587e-06 | 2.0266e-06 | 0.99054 | 34.6278 | 0.00012951
LSSL2010#619 | CB-21 | 1383 |  | 2011-02-10 16:57 | 24C | 0.98679 | 4.6353e-06 | 2.0266e-06 | 0.9868 | 34.4809 | 0.00018285
LSSL2010#620 | CB-21 | 1384 |  | 2011-02-10 16:59 | 24C | 0.98339 | 2.2263e-06 | 2.0266e-06 | 0.98339 | 34.3473 | 8.9601e-05
LSSL2010#621 | CB-21 | 1385 |  | 2011-02-10 17:01 | 24C | 0.98126 | 2.7426e-06 | 1.967e-06 | 0.98126 | 34.2636 | 0.00010447
LSSL2010#622 | CB-21 | 1386 |  | 2011-02-10 17:03 | 24C | 0.97856 | 3.2012e-06 | 2.0266e-06 | 0.97856 | 34.1578 | 0.00012091
LSSL2010#623 | CB-21 | 1387 |  | 2011-02-10 17:05 | 24C | 0.9744 | 4.3336e-06 | 2.0266e-06 | 0.9744 | 33.9947 | 0.00016843
LSSL2010#624 | CB-21 | 1388 |  | 2011-02-10 17:08 | 24C | 0.96983 | 4.2113e-06 | 2.0862e-06 | 0.96983 | 33.8158 | 0.00016236
LSSL2010#625 | CB-21 | 1389 |  | 2011-02-10 17:10 | 24C | 0.96326 | 3.323e-06 | 1.967e-06 | 0.96327 | 33.5589 | 0.00013002
LSSL2010#626 | CB-21 | 1390 |  | 2011-02-10 17:12 | 24C | 0.95706 | 2.0947e-06 | 2.0862e-06 | 0.95706 | 33.3165 | 8.2541e-05
LSSL2010#627 | CB-21 | 1391 |  | 2011-02-10 17:14 | 24C | 0.95116 | 2.4908e-06 | 2.0266e-06 | 0.95116 | 33.0862 | 9.7029e-05
LSSL2010#628 | CB-21 | 1392 |  | 2011-02-10 17:16 | 24C | 0.9475 | 3.8622e-06 | 2.0266e-06 | 0.9475 | 32.9435 | 0.00014836
LSSL2010#629 | CB-21 | 1393 |  | 2011-02-10 17:19 | 24C | 0.94466 | 2.4911e-06 | 1.967e-06 | 0.94466 | 32.8327 | 9.869e-05
LSSL2010#630 | CB-21 | 1394 |  | 2011-02-10 17:21 | 24C | 0.94203 | 2.7425e-06 | 1.9073e-06 | 0.94203 | 32.7303 | 0.00010447
LSSL2010#631 | CB-21 | 1395 |  | 2011-02-10 17:23 | 24C | 0.93603 | 3.8538e-06 | 2.0266e-06 | 0.93603 | 32.4967 | 0.00015042
LSSL2010#632 | CB-21 | 1396 |  | 2011-02-10 17:25 | 24C | 0.9308 | 4.6144e-06 | 1.9073e-06 | 0.9308 | 32.2934 | 0.00018034
LSSL2010#633 | CB-21 | 1397 |  | 2011-02-10 17:27 | 24C | 0.92361 | 4.409e-06 | 2.0862e-06 | 0.92361 | 32.014 | 0.00017061
LSSL2010#634 | CB-21 | 1398 |  | 2011-02-10 17:29 | 24C | 0.91814 | 2.4909e-06 | 2.0266e-06 | 0.91814 | 31.8017 | 9.8682e-05
LSSL2010#635 | CB-21 | 1399 |  | 2011-02-10 17:31 | 24C | 0.90898 | 1.3713e-06 | 2.0266e-06 | 0.90899 | 31.447 | 5.3301e-05
LSSL2010#636 | CB-21 | 1400 |  | 2011-02-10 17:34 | 24C | 0.89644 | 2.4754e-06 | 2.1458e-06 | 0.89644 | 30.962 | 9.6496e-05
LSSL2010#637 | CB-21 | 1401 |  | 2011-02-10 17:36 | 24C | 0.88541 | 2.4909e-06 | 1.967e-06 | 0.88541 | 30.5363 | 9.4885e-05
LSSL2010#638 | CB-21 | 1402 |  | 2011-02-10 17:38 | 24C | 0.85779 | 3.8623e-06 | 2.0862e-06 | 0.85779 | 29.4746 | 0.00014921
LSSL2010#639 | CB-21 | 1403 |  | 2011-02-10 17:40 | 24C | 0.83299 | 3.6105e-06 | 2.0266e-06 | 0.83299 | 28.5257 | 0.00013967
LSSL2010#640 | CB-21 | 1404 |  | 2011-02-10 17:42 | 24C | 0.78216 | 2.2392e-06 | 2.0266e-06 | 0.78216 | 26.5951 | 8.5306e-05
LSSL2010#641 | CB-21 | 1405 |  | 2011-02-10 17:45 | 24C | 0.74258 | 2.8897e-06 | 2.0266e-06 | 0.74259 | 25.105 | 0.00010553
LSSL2010#642 | reference sample | CB-6 | 26 | 2011-02-10 17:48 | 24C | 0.99887 | 3.2247e-06 | 1.9073e-06 | 0.99888 | 34.9557 | 0.00012608
Calibration#8 |  | K15 |  | 2011-02-11 09:25 | 24C | 0.99997 | 0 | 5.0664e-06 | 0.99997
LSSL2010#643 | PP-7 | 1166 |  | 2011-02-11 09:29 | 24C | 0.99716 | 1.4046e-06 | 5.0068e-06 | 0.99717 | 34.8885 | 5.5371e-05
LSSL2010#644 | PP-7 | 1167 |  | 2011-02-11 09:32 | 24C | 0.99715 | 1.515e-06 | 4.8876e-06 | 0.99716 | 34.8882 | 5.8044e-05
LSSL2010#645 | PP-7 | 1168 |  | 2011-02-11 09:36 | 24C | 0.99713 | 3.191e-06 | 5.0068e-06 | 0.99713 | 34.8871 | 0.00012429
LSSL2010#646 | PP-7 | 1169 |  | 2011-02-11 09:38 | 24C | 0.99705 | 3.1414e-06 | 5.0068e-06 | 0.99706 | 34.8842 | 0.00012621
LSSL2010#647 | PP-7 | 1170 |  | 2011-02-11 09:40 | 24C | 0.99689 | 3.2141e-06 | 5.0068e-06 | 0.9969 | 34.8779 | 0.00012432
LSSL2010#648 | PP-7 | 1171 |  | 2011-02-11 09:42 | 24C | 0.99673 | 3.579e-06 | 5.0068e-06 | 0.99674 | 34.8716 | 0.00013979
LSSL2010#649 | PP-7 | 1172 |  | 2011-02-11 09:44 | 24C | 0.99668 | 3.4657e-06 | 4.9472e-06 | 0.99668 | 34.8695 | 0.00013203
LSSL2010#650 | PP-7 | 1173 |  | 2011-02-11 09:46 | 24C | 0.99705 | 2.1278e-06 | 5.0068e-06 | 0.99705 | 34.8839 | 8.1638e-05
LSSL2010#651 | PP-7 | 1174 |  | 2011-02-11 09:49 | 24C | 0.99662 | 1.3713e-06 | 5.0068e-06 | 0.99663 | 34.8673 | 5.5371e-05
LSSL2010#652 | PP-7 | 1175 |  | 2011-02-11 09:51 | 24C | 0.99638 | 2.4908e-06 | 4.9472e-06 | 0.99638 | 34.8576 | 9.6599e-05
LSSL2010#653 | PP-7 | 1176 |  | 2011-02-11 09:54 | 24C | 0.99638 | 3.1535e-06 | 5.0068e-06 | 0.99639 | 34.8579 | 0.00012228
LSSL2010#654 | PP-7 | 1177 |  | 2011-02-11 09:57 | 24C | 0.99642 | 2.4925e-06 | 4.9472e-06 | 0.99642 | 34.8592 | 9.869e-05
LSSL2010#655 | PP-7 | 1178 |  | 2011-02-11 09:59 | 24C | 0.99577 | 3.6104e-06 | 5.0068e-06 | 0.99578 | 34.8339 | 0.00014762
LSSL2010#656 | PP-7 | 1179 |  | 2011-02-11 10:02 | 24C | 0.99538 | 3.2141e-06 | 5.0664e-06 | 0.99539 | 34.8186 | 0.00012624
LSSL2010#657 | PP-7 | 1180 |  | 2011-02-11 10:04 | 24C | 0.99543 | 2.4911e-06 | 5.0664e-06 | 0.99543 | 34.8204 | 9.6433e-05
LSSL2010#658 | PP-7 | 1181 |  | 2011-02-11 10:06 | 24C | 0.9954 | 2.4909e-06 | 4.9472e-06 | 0.9954 | 34.8191 | 9.4885e-05
LSSL2010#659 | PP-7 | 1182 |  | 2011-02-11 10:09 | 24C | 0.99542 | 2.4908e-06 | 5.0664e-06 | 0.99543 | 34.8201 | 0.0001006
LSSL2010#660 | PP-7 | 1183 |  | 2011-02-11 10:10 | 24C | 0.99067 | 2.7425e-06 | 4.8876e-06 | 0.99067 | 34.6331 | 0.00010659
LSSL2010#661 | PP-7 | 1184 |  | 2011-02-11 10:12 | 24C | 0.99064 | 3.8647e-06 | 5.0068e-06 | 0.99064 | 34.632 | 0.00015066
LSSL2010#662 | PP-7 | 1185 |  | 2011-02-11 10:29 | 24C | 0.99068 | 3.8542e-06 | 5.0068e-06 | 0.99068 | 34.6336 | 0.00015136
LSSL2010#663 | PP-7 | 1186 |  | 2011-02-11 10:31 | 24C | 0.99038 | 4.3336e-06 | 4.9472e-06 | 0.99039 | 34.6219 | 0.00016864
LSSL2010#664 | PP-7 | 1187 |  | 2011-02-11 10:33 | 24C | 0.94388 | 3.6081e-06 | 5.0068e-06 | 0.94389 | 32.8026 | 0.00013993
LSSL2010#665 | PP-7 | 1188 |  | 2011-02-11 10:35 | 24C | 0.8738 | 4.336e-06 | 4.9472e-06 | 0.8738 | 30.0895 | 0.0001679
LSSL2010#666 | PP-7 | 1189 |  | 2011-02-11 10:37 | 24C | 0.77705 | 3.1414e-06 | 5.0068e-06 | 0.77705 | 26.4021 | 0.00011629
LSSL2010#667 | reference sample | CB-6 | 27 | 2011-02-11 10:39 | 24C | 0.99891 | 3.3298e-06 | 4.828e-06 | 0.99891 | 34.9571 | 0.00012925
LSSL2010#668 | CB-15 | 1094 |  | 2011-02-11 10:47 | 24C | 0.99702 | 1.3713e-06 | 4.8876e-06 | 0.99702 | 34.8828 | 5.2235e-05
LSSL2010#669 | CB-15 | 1095 |  | 2011-02-11 10:49 | 24C | 0.99692 | 3.3595e-06 | 4.9472e-06 | 0.99693 | 34.8792 | 0.00013116
LSSL2010#670 | CB-15 | 1096 |  | 2011-02-11 10:51 | 24C | 0.99702 | 5.4532e-06 | 5.0068e-06 | 0.99703 | 34.8831 | 0.00021569
LSSL2010#671 | CB-15 | 1097 |  | 2011-02-11 10:53 | 24C | 0.99701 | 4.9165e-06 | 5.0068e-06 | 0.99702 | 34.8827 | 0.00019367
LSSL2010#672 | CB-15 | 1098 |  | 2011-02-11 10:55 | 24C | 0.99656 | 2.7425e-06 | 5.0068e-06 | 0.99656 | 34.8648 | 0.00010657
LSSL2010#673 | CB-15 | 1099 |  | 2011-02-11 10:58 | 24C | 0.9966 | 2.2267e-06 | 4.9472e-06 | 0.9966 | 34.8662 | 8.7013e-05
LSSL2010#674 | CB-15 | 1100 |  | 2011-02-11 10:59 | 24C | 0.99655 | 3.4389e-06 | 5.0068e-06 | 0.99656 | 34.8646 | 0.00013478
LSSL2010#675 | CB-15 | 1101 |  | 2011-02-11 11:02 | 24C | 0.99659 | 4.6353e-06 | 5.0068e-06 | 0.99659 | 34.8659 | 0.00017966
LSSL2010#676 | CB-15 | 1102 |  | 2011-02-11 11:04 | 24C | 0.99617 | 3.9742e-06 | 5.0068e-06 | 0.99618 | 34.8497 | 0.0001555
LSSL2010#677 | CB-15 | 1103 |  | 2011-02-11 11:06 | 24C | 0.99611 | 1.1532e-06 | 4.8876e-06 | 0.99612 | 34.8472 | 4.4374e-05
LSSL2010#678 | CB-15 | 1104 |  | 2011-02-11 11:07 | 24C | 0.99617 | 4.4783e-06 | 4.8876e-06 | 0.99617 | 34.8493 | 0.00017404
LSSL2010#679 | CB-15 | 1105 |  | 2011-02-11 11:09 | 24C | 0.99622 | 3.1421e-06 | 4.8876e-06 | 0.99623 | 34.8516 | 0.00012383
LSSL2010#680 | CB-15 | 1106 |  | 2011-02-11 11:11 | 24C | 0.99534 | 3.6111e-06 | 4.8876e-06 | 0.99534 | 34.8167 | 0.00014174
LSSL2010#681 | CB-15 | 1107 |  | 2011-02-11 11:13 | 24C | 0.99538 | 3.865e-06 | 5.0068e-06 | 0.99539 | 34.8185 | 0.00015069
LSSL2010#682 | CB-15 | 1108 |  | 2011-02-11 11:15 | 24C | 0.99542 | 3.865e-06 | 5.0068e-06 | 0.99543 | 34.8202 | 0.00015069
LSSL2010#683 | CB-15 | 1109 |  | 2011-02-11 11:18 | 24C | 0.99543 | 3.1417e-06 | 5.126e-06 | 0.99544 | 34.8205 | 0.00012202
LSSL2010#684 | CB-15 | 1110 |  | 2011-02-11 11:20 | 24C | 0.99009 | 2.4911e-06 | 5.0664e-06 | 0.99009 | 34.6104 | 9.869e-05
LSSL2010#685 | CB-15 | 1111 |  | 2011-02-11 11:22 | 24C | 0.9907 | 2.4908e-06 | 5.0068e-06 | 0.9907 | 34.6344 | 9.7462e-05
LSSL2010#686 | CB-15 | 1112 |  | 2011-02-11 11:24 | 24C | 0.98999 | 3.1418e-06 | 5.0068e-06 | 0.99 | 34.6066 | 0.00012421
LSSL2010#687 | CB-15 | 1113 |  | 2011-02-11 11:26 | 24C | 0.98976 | 4.0093e-06 | 5.0664e-06 | 0.98977 | 34.5977 | 0.00015704
LSSL2010#688 | CB-15 | 1114 |  | 2011-02-11 11:28 | 24C | 0.79146 | 3.1415e-06 | 5.0068e-06 | 0.79146 | 26.947 | 0.00011629
LSSL2010#689 | CB-15 | 1115 |  | 2011-02-11 11:30 | 24C | 0.78969 | 2.7425e-06 | 5.0068e-06 | 0.78969 | 26.8799 | 9.8205e-05
LSSL2010#690 | CB-15 | 1116 |  | 2011-02-11 11:33 | 24C | 0.78938 | 2.2391e-06 | 5.0068e-06 | 0.78938 | 26.8682 | 8.4895e-05
LSSL2010#691 | CB-15 | 1117 |  | 2011-02-11 11:35 | 24C | 0.78937 | 2.8897e-06 | 5.0068e-06 | 0.78937 | 26.8679 | 0.00010594
LSSL2010#692 | reference sample | CB-6 | 28 | 2011-02-11 11:40 | 24C | 0.99889 | 2.0943e-06 | 5.0664e-06 | 0.9989 | 34.9566 | 8.2542e-05
LSSL2010#693 | LOOP | 96 |  | 2011-02-11 11:50 | 24C | 0.7901 | 3.4657e-06 | 5.126e-06 | 0.79011 | 26.8956 | 0.00013117
LSSL2010#694 | LOOP | 100 |  | 2011-02-11 11:52 | 24C | 0.76471 | 3.1414e-06 | 5.0068e-06 | 0.76471 | 25.9367 | 0.00011969
LSSL2010#695 | LOOP | 102 |  | 2011-02-11 11:54 | 24C | 0.76467 | 3.2145e-06 | 5.0068e-06 | 0.76468 | 25.9353 | 0.00011784
LSSL2010#696 | LOOP | 107 |  | 2011-02-11 11:56 | 24C | 0.77809 | 1.3712e-06 | 5.0068e-06 | 0.77809 | 26.4413 | 5.2783e-05
LSSL2010#697 | LOOP | 107 |  | 2011-02-11 11:58 | 24C | 0.78529 | 2.4908e-06 | 5.0068e-06 | 0.7853 | 26.7136 | 8.962e-05
LSSL2010#698 | LOOP | 110 |  | 2011-02-11 12:00 | 24C | 0.8168 | 2.4915e-06 | 5.126e-06 | 0.81681 | 27.909 | 8.9222e-05
LSSL2010#699 | LOOP | 115 |  | 2011-02-11 12:02 | 24C | 0.75796 | 2.2391e-06 | 5.0068e-06 | 0.75796 | 25.6825 | 8.2749e-05
LSSL2010#700 | LOOP | 117 |  | 2011-02-11 12:05 | 24C | 0.75627 | 2.7425e-06 | 5.0068e-06 | 0.75627 | 25.6189 | 0.00010135
LSSL2010#701 | LOOP | 122 |  | 2011-02-11 12:06 | 24C | 0.72351 | 3.854e-06 | 5.0068e-06 | 0.72352 | 24.3913 | 0.00014262
LSSL2010#702 | LOOP | 126 |  | 2011-02-11 12:09 | 24C | 0.73684 | 2.4908e-06 | 5.0068e-06 | 0.73685 | 24.8898 | 9.4459e-05
LSSL2010#703 | LOOP | 128 |  | 2011-02-11 12:11 | 24C | 0.76751 | 2.7426e-06 | 5.0068e-06 | 0.76752 | 26.0424 | 0.00010448
LSSL2010#704 | LOOP | 131 |  | 2011-02-11 12:14 | 24C | 0.79058 | 2.7425e-06 | 5.0068e-06 | 0.79059 | 26.9138 | 9.9246e-05
LSSL2010#705 | LOOP | 134 |  | 2011-02-11 12:16 | 24C | 0.87179 | 2.0945e-06 | 5.0068e-06 | 0.8718 | 30.0123 | 7.9449e-05
LSSL2010#706 | reference sample | CB-6 | 29 | 2011-02-11 12:19 | 24C | 0.99892 | 2.8297e-06 | 4.8876e-06 | 0.99892 | 34.9576 | 0.00011337

## Autosal-Data for Merging
Salinometer ID | Station | Sample ID (bottle label) | Duplicates | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments | Diff
Sample ID | Station | Bottle Label |  | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
|  | 719 | [10 empty] | 2011-02-09 14:49 | 24C | 0.99568 | 4.1469e-06 | 2.1458e-06 | 0.99568 | 34.8301 | 0.00016043 |  | -34.8301
|  | 726 | [10 empty] | 2011-02-09 15:10 | 24C | 0.9563 | 2.4915e-06 | 2.0266e-06 | 0.9563 | 33.2868 | 9.4981e-05 |  | -33.2868
|  | 739 | [10 empty] | 2011-02-09 12:59 | 24C | 0.99675 | 7.2738e-06 | 1.967e-06 | 0.99675 | 34.8723 | 0.00028709 |  | -34.8723
|  | 743 | [10 empty] | 2011-02-09 13:10 | 24C | 0.99582 | 2.2394e-06 | 2.0266e-06 | 0.99582 | 34.8356 | 8.793e-05 |  | -34.8356
|  | 750 | [10 empty] | 2011-02-09 15:07 | 24C | 0.94997 | 3.6104e-06 | 2.0266e-06 | 0.94997 | 33.0396 | 0.00013925 |  | -33.0396
|  | 789 | [10 empty] | 2011-02-09 13:06 | 24C | 0.9963 | 3.7354e-06 | 2.0862e-06 | 0.9963 | 34.8545 | 0.00014444 |  | -34.8545
|  | 795 | [10 empty] | 2011-02-09 14:58 | 24C | 0.98012 | 3.6105e-06 | 2.0266e-06 | 0.98012 | 34.2189 | 0.0001408 |  | -34.2189
|  | 800 | [10 empty] | 2011-02-09 15:14 | 24C | 0.93642 | 3.1314e-06 | 1.967e-06 | 0.93642 | 32.5118 | 0.00012312 |  | -32.5118
|  | 806 | [10 empty] | 2011-02-09 12:55 | 24C | 0.99892 | 2.2397e-06 | 2.0266e-06 | 0.99892 | 34.9577 | 8.7013e-05 |  | -34.9577
|  | 810 | [10 empty] | 2011-02-09 12:57 | 24C | 0.99751 | 1.1529e-06 | 1.9073e-06 | 0.99751 | 34.9021 | 4.265e-05 |  | -34.9021
|  | 811 | [10 empty] | 2011-02-09 13:03 | 24C | 0.99677 | 2.8661e-06 | 2.0266e-06 | 0.99678 | 34.8731 | 0.00011014 |  | -34.8731
|  | 819 | [10 empty] | 2011-02-09 15:00 | 24C | 0.97841 | 2.2394e-06 | 1.9073e-06 | 0.97841 | 34.1518 | 8.7024e-05 |  | -34.1518
|  | 821 | [10 empty] | 2011-02-09 15:04 | 24C | 0.95214 | 3.4359e-06 | 2.0862e-06 | 0.95214 | 33.1243 | 0.00013203 |  | -33.1243
|  | 830 | [10 empty] | 2011-02-09 14:54 | 24C | 0.99461 | 3.4657e-06 | 2.0266e-06 | 0.99462 | 34.7882 | 0.00013904 |  | -34.7882
|  | 836 | [10 empty] | 2011-02-09 14:56 | 24C | 0.98182 | 3.4658e-06 | 2.0266e-06 | 0.98182 | 34.2856 | 0.00013595 |  | -34.2856
|  | 845 | [10 empty] | 2011-02-09 15:12 | 24C | 0.93661 | 2.4916e-06 | 2.0266e-06 | 0.93661 | 32.5192 | 9.6976e-05 |  | -32.5192
|  | 859 | [10 empty] | 2011-02-09 13:01 | 24C | 0.99677 | 3.1145e-06 | 2.0862e-06 | 0.99678 | 34.8731 | 0.00011967 |  | -34.8731
|  | 863 | [10 empty] | 2011-02-09 14:52 | 24C | 0.99573 | 3.1418e-06 | 1.967e-06 | 0.99574 | 34.8322 | 0.00012381 |  | -34.8322
|  | 870 | [10 empty] | 2011-02-09 15:02 | 24C | 0.95334 | 2.7425e-06 | 1.9073e-06 | 0.95334 | 33.171 | 0.00010761 |  | -33.171
|  | 909 | [10 empty] | 2011-02-09 13:08 | 24C | 0.99624 | 3.9994e-06 | 2.0266e-06 | 0.99625 | 34.8523 | 0.00015603 |  | -34.8523
|  | 915 | [10 empty] | 2011-02-09 17:14 | 24C | 0.97714 | 2.89e-06 | 1.967e-06 | 0.97715 | 34.1023 | 0.00011317 |  | -34.1023
|  | 920 | [10 empty] | 2011-02-09 17:23 | 24C | 0.93825 | 2.4908e-06 | 1.967e-06 | 0.93825 | 32.583 | 9.869e-05 |  | -32.583
|  | 923 | [10 empty] | 2011-02-09 17:28 | 24C | 0.89184 | 3.2141e-06 | 2.0862e-06 | 0.89185 | 30.7846 | 0.00012629 |  | -30.7846
|  | 957 | [10 empty] | 2011-02-09 16:59 | 24C | 0.99432 | 2.4908e-06 | 2.0266e-06 | 0.99432 | 34.7767 | 9.4885e-05 |  | -34.7767
|  | 963 | [10 empty] | 2011-02-09 17:12 | 24C | 0.9826 | 1.4046e-06 | 1.967e-06 | 0.9826 | 34.3163 | 5.2235e-05 |  | -34.3163
|  | 968 | [10 empty] | 2011-02-09 17:21 | 24C | 0.93818 | 2.0945e-06 | 1.967e-06 | 0.93818 | 32.5802 | 8.1861e-05 |  | -32.5802
0
LSSL2010#593 | CB-12a | 974 |  | 2011-02-10 15:59 | 24C | 0.99343 | 3.3459e-06 | 2.0862e-06 | 0.99343 | 34.7417 | 0.00013202 |  | 2011-02-09 17:09 | 24C | 0.99344 | 2.4751e-06 | 2.0266e-06 | 0.99344 | 34.7418 | 9.9792e-05 | 0 | -0.00013351
LSSL2010#594 | CB-12a | 975 |  | 2011-02-10 16:01 | 24C | 0.99238 | 4.189e-06 | 1.9073e-06 | 0.99238 | 34.7003 | 0.00016533 | [10 empty] | 34.7003
LSSL2010#595 | CB-12a | 976 |  | 2011-02-10 16:03 | 24C | 0.98988 | 2.4776e-06 | 2.0266e-06 | 0.98988 | 34.6022 | 9.6976e-05 | [10 empty] | 34.6022
LSSL2010#596 | CB-12a | 977 |  | 2011-02-10 16:05 | 24C | 0.98898 | 4.3836e-06 | 2.0266e-06 | 0.98898 | 34.5668 | 0.00017346 | [10 empty] | 34.5668
LSSL2010#597 | CB-12a | 978 |  | 2011-02-10 16:07 | 24C | 0.98419 | 2.2394e-06 | 1.967e-06 | 0.98419 | 34.3787 | 8.7067e-05 | [10 empty] | 34.3787
LSSL2010#598 | CB-12a | 979 |  | 2011-02-10 16:09 | 24C | 0.98092 | 4.2164e-06 | 1.967e-06 | 0.98092 | 34.2505 | 0.00016491 | [10 empty] | 34.2505
LSSL2010#599 | CB-12a | 980 |  | 2011-02-10 16:11 | 24C | 0.97357 | 3.8756e-06 | 2.0862e-06 | 0.97358 | 33.9624 | 0.0001485 |  | 2011-02-09 17:17 | 24C | 0.97401 | 1.1536e-06 | 2.0266e-06 | 0.97401 | 33.9794 | 4.2703e-05 | 0 | -0.016998
LSSL2010#600 | CB-12a | 981 |  | 2011-02-10 16:13 | 24C | 0.96653 | 3.4352e-06 | 2.0862e-06 | 0.96653 | 33.6867 | 0.0001341 | [10 empty] | 33.6867
LSSL2010#601 | CB-12a | 982 |  | 2011-02-10 16:16 | 24C | 0.96233 | 2.89e-06 | 2.0266e-06 | 0.96233 | 33.5225 | 0.00011316 | [10 empty] | 33.5225
LSSL2010#602 | CB-12a | 983 |  | 2011-02-10 16:18 | 24C | 0.95714 | 6.664e-08 | 2.0266e-06 | 0.95714 | 33.3197 | 0 | [10 empty] | 33.3197
LSSL2010#603 | CB-12a | 984 |  | 2011-02-10 16:20 | 24C | 0.9541 | 3.6105e-06 | 2.0266e-06 | 0.9541 | 33.2007 | 0.00014113 | [10 empty] | 33.2007
LSSL2010#604 | CB-12a | 985 |  | 2011-02-10 16:22 | 24C | 0.95102 | 3.8747e-06 | 2.0862e-06 | 0.95103 | 33.0808 | 0.00015161 | [10 empty] | 33.0808
LSSL2010#605 | CB-12a | 986 |  | 2011-02-10 16:24 | 24C | 0.94902 | 4.6041e-06 | 2.0266e-06 | 0.94903 | 33.0028 | 0.00018033 | [10 empty] | 33.0028
LSSL2010#606 | CB-12a | 987 |  | 2011-02-10 16:26 | 24C | 0.94293 | 4.9162e-06 | 2.0266e-06 | 0.94293 | 32.7654 | 0.00019131 | [10 empty] | 32.7654
LSSL2010#607 | CB-12a | 988 |  | 2011-02-10 16:27 | 24C | 0.93753 | 2.7425e-06 | 2.0266e-06 | 0.93753 | 32.555 | 0.00010865 | [10 empty] | 32.555
LSSL2010#608 | CB-12a | 989 |  | 2011-02-10 16:29 | 24C | 0.93018 | 3.4657e-06 | 2.0266e-06 | 0.93018 | 32.2692 | 0.0001355 |  | 2011-02-09 17:26 | 24C | 0.93005 | 3.1424e-06 | 2.0266e-06 | 0.93005 | 32.264 | 0.00012176 | 0 | 0.005188
LSSL2010#609 | CB-12a | 990 |  | 2011-02-10 16:31 | 24C | 0.92547 | 4.0099e-06 | 2.1458e-06 | 0.92547 | 32.0862 | 0.00015276 | [10 empty] | 32.0862
LSSL2010#610 | CB-12a | 991 |  | 2011-02-10 16:33 | 24C | 0.92221 | 3.5406e-06 | 2.0266e-06 | 0.92221 | 31.9598 | 0.00013625 | [10 empty] | 31.9598
LSSL2010#611 | CB-12a | 992 |  | 2011-02-10 16:35 | 24C | 0.91164 | 1.7702e-06 | 1.967e-06 | 0.91164 | 31.5499 | 6.9141e-05 | [10 empty] | 31.5499
LSSL2010#612 | CB-12a | 993 |  | 2011-02-10 16:37 | 24C | 0.90407 | 3.3591e-06 | 2.0862e-06 | 0.90408 | 31.257 | 0.00012893 | [10 empty] | 31.257
LSSL2010#613 | CB-12a | 994 |  | 2011-02-10 16:39 | 24C | 0.89295 | 3.1417e-06 | 1.967e-06 | 0.89296 | 30.8274 | 0.00012086 | [10 empty] | 30.8274
LSSL2010#614 | CB-12a | 995 |  | 2011-02-10 16:42 | 24C | 0.87458 | 3.4515e-06 | 2.0862e-06 | 0.87458 | 30.1194 | 0.00013157 | [10 empty] | 30.1194
LSSL2010#615 | CB-12a | 996 |  | 2011-02-10 16:44 | 24C | 0.8196 | 2.2394e-06 | 1.967e-06 | 0.81961 | 28.0155 | 8.5313e-05 | [10 empty] | 28.0155
LSSL2010#616 | CB-12a | 997 |  | 2011-02-10 16:46 | 24C | 0.78066 | 2.4908e-06 | 2.0862e-06 | 0.78067 | 26.5386 | 9.4478e-05 | [10 empty] | 26.5386
LSSL2010#350 | CB-16 | 998 |  | 2011-02-07 15:32 | 24C | 0.99882 | 6.6317e-06 | 4.0054e-05 | 0.99886 | 34.955 | 0.00025894 | [10 empty] | 34.955
LSSL2010#351 | CB-16 | 999 |  | 2011-02-07 15:37 | 24C | 0.99886 | 2.0945e-06 | 4.0054e-05 | 0.9989 | 34.9568 | 8.1866e-05 | [10 empty] | 34.9568
LSSL2010#352 | CB-16 | 1000 |  | 2011-02-07 15:43 | 24C | 0.99875 | 2.2405e-06 | 3.9995e-05 | 0.99879 | 34.9526 | 8.9617e-05 | [10 empty] | 34.9526
LSSL2010#353 | CB-16 | 1001 |  | 2011-02-07 15:45 | 24C | 0.9985 | 3.1446e-06 | 3.9995e-05 | 0.99854 | 34.9425 | 0.00012628 | [10 empty] | 34.9425
LSSL2010#354 | CB-16 | 1002 |  | 2011-02-07 15:48 | 24C | 0.99777 | 2.7425e-06 | 3.9935e-05 | 0.99781 | 34.9139 | 0.00010447 | [10 empty] | 34.9139
LSSL2010#355 | CB-16 | 1003 |  | 2011-02-07 15:50 | 24C | 0.997 | 3.4657e-06 | 4.0054e-05 | 0.99704 | 34.8836 | 0.00013662 |  | 2011-02-09 16:57 | 24C | 0.997 | 2.8907e-06 | 2.1458e-06 | 0.997 | 34.8821 | 0.00011519 | 0 | 0.0014534
LSSL2010#356 | CB-16 | 1004 |  | 2011-02-07 15:53 | 24C | 0.99663 | 3.1415e-06 | 3.9935e-05 | 0.99667 | 34.869 | 0.0001262 | [10 empty] | 34.869
LSSL2010#357 | CB-16 | 1005 |  | 2011-02-07 15:55 | 24C | 0.9962 | 3.1253e-06 | 3.9995e-05 | 0.99624 | 34.852 | 0.00012312 | [10 empty] | 34.852
LSSL2010#358 | CB-16 | 1006 |  | 2011-02-07 15:57 | 24C | 0.9962 | 3.4352e-06 | 3.9995e-05 | 0.99624 | 34.8521 | 0.00013596 | [10 empty] | 34.8521
LSSL2010#359 | CB-16 | 1007 |  | 2011-02-07 16:00 | 24C | 0.99574 | 3.6104e-06 | 4.0054e-05 | 0.99578 | 34.8342 | 0.00013753 |  | 2011-02-09 17:04 | 24C | 0.99575 | 1.3712e-06 | 1.9073e-06 | 0.99576 | 34.833 | 5.5371e-05 | 0 | 0.0011597
LSSL2010#360 | CB-16 | 1008 |  | 2011-02-07 16:02 | 24C | 0.99542 | 2.4908e-06 | 4.0054e-05 | 0.99546 | 34.8214 | 9.7495e-05 | [10 empty] | 34.8214
LSSL2010#361 | CB-16 | 1009 |  | 2011-02-07 16:04 | 24C | 0.99457 | 3.1253e-06 | 4.0054e-05 | 0.99461 | 34.7881 | 0.00012304 | [10 empty] | 34.7881
LSSL2010#362 | CB-16 | 1010 |  | 2011-02-07 16:06 | 24C | 0.99312 | 4.2739e-06 | 3.9935e-05 | 0.99316 | 34.7308 | 0.00016608 | [10 empty] | 34.7308
LSSL2010#363 | CB-16 | 1011 |  | 2011-02-07 16:09 | 24C | 0.98546 | 1.4046e-06 | 4.0054e-05 | 0.9855 | 34.4299 | 5.4326e-05 | [10 empty] | 34.4299
LSSL2010#364 | CB-16 | 1012 |  | 2011-02-07 16:11 | 24C | 0.97902 | 2.8904e-06 | 4.0054e-05 | 0.97906 | 34.1773 | 0.00011604 | [10 empty] | 34.1773
LSSL2010#365 | CB-16 | 1013 |  | 2011-02-07 16:12 | 24C | 0.96879 | 3.4657e-06 | 4.0054e-05 | 0.96883 | 33.7766 | 0.00013904 | [10 empty] | 33.7766
LSSL2010#366 | CB-16 | 1014 |  | 2011-02-07 16:14 | 24C | 0.95353 | 2.2725e-06 | 4.0054e-05 | 0.95357 | 33.1801 | 8.7556e-05 |  | 2011-02-09 17:18 | 24C | 0.95357 | 3.6119e-06 | 2.1458e-06 | 0.95357 | 33.1801 | 0.00013908 | 0 | 3.8147e-06
LSSL2010#367 | CB-16 | 1015 |  | 2011-02-07 16:17 | 24C | 0.94751 | 2.7425e-06 | 3.9935e-05 | 0.94755 | 32.9453 | 0.00010447 | [10 empty] | 32.9453
LSSL2010#368 | CB-16 | 1016 |  | 2011-02-07 16:19 | 24C | 0.94036 | 2.4909e-06 | 3.9995e-05 | 0.9404 | 32.6666 | 9.6125e-05 | [10 empty] | 32.6666
LSSL2010#369 | CB-16 | 1017 |  | 2011-02-07 16:21 | 24C | 0.93371 | 3.4495e-06 | 3.9995e-05 | 0.93375 | 32.4079 | 0.00013341 | [10 empty] | 32.4079
LSSL2010#370 | CB-16 | 1018 |  | 2011-02-07 16:23 | 24C | 0.92303 | 2.2401e-06 | 3.9995e-05 | 0.92307 | 31.9931 | 8.5326e-05 | [10 empty] | 31.9931
LSSL2010#371 | CB-16 | 1019 |  | 2011-02-07 16:26 | 24C | 0.90215 | 2.4908e-06 | 4.0054e-05 | 0.90219 | 31.1843 | 9.6509e-05 | [10 empty] | 31.1843
LSSL2010#372 | CB-16 | 1020 |  | 2011-02-07 16:28 | 24C | 0.83418 | 2.7425e-06 | 3.9935e-05 | 0.83422 | 28.5728 | 0.00010396 | [10 empty] | 28.5728
LSSL2010#373 | CB-16 | 1021 |  | 2011-02-07 16:30 | 24C | 0.82273 | 1.1529e-06 | 4.0054e-05 | 0.82277 | 28.1358 | 4.2663e-05 | [10 empty] | 28.1358
LSSL2010#518 | CB-13 | 1022 |  | 2011-02-10 11:22 | 24C | 0.9969 | 2.2391e-06 | 2.0266e-06 | 0.9969 | 34.8779 | 8.7024e-05 | [10 empty] | 34.8779
LSSL2010#519 | CB-13 | 1023 |  | 2011-02-10 11:24 | 24C | 0.99688 | 2.4909e-06 | 2.0862e-06 | 0.99689 | 34.8774 | 9.869e-05 | [10 empty] | 34.8774
LSSL2010#520 | CB-13 | 1024 |  | 2011-02-10 11:27 | 24C | 0.99676 | 3.2144e-06 | 1.967e-06 | 0.99676 | 34.8724 | 0.00012553 | [10 empty] | 34.8724
LSSL2010#521 | CB-13 | 1025 |  | 2011-02-10 11:30 | 24C | 0.99685 | 2.7425e-06 | 1.967e-06 | 0.99685 | 34.876 | 0.00010865 | [10 empty] | 34.876
LSSL2010#522 | CB-13 | 1026 |  | 2011-02-10 11:32 | 24C | 0.99655 | 1.1396e-06 | 2.0862e-06 | 0.99655 | 34.8643 | 4.2703e-05 | [10 empty] | 34.8643
LSSL2010#523 | CB-13 | 1027 |  | 2011-02-10 11:35 | 24C | 0.99653 | 2.0947e-06 | 2.0862e-06 | 0.99653 | 34.8636 | 8.1174e-05 | [10 empty] | 34.8636
LSSL2010#524 | CB-13 | 1028 |  | 2011-02-10 11:37 | 24C | 0.99637 | 2.2391e-06 | 1.967e-06 | 0.99637 | 34.8572 | 8.5352e-05 | [10 empty] | 34.8572
LSSL2010#525 | CB-13 | 1029 |  | 2011-02-10 11:39 | 24C | 0.99643 | 3.6105e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 0.00014239 | [10 empty] | 34.8596
LSSL2010#526 | CB-13 | 1030 |  | 2011-02-10 11:41 | 24C | 0.99617 | 2.2394e-06 | 2.1458e-06 | 0.99618 | 34.8496 | 8.959e-05 | [10 empty] | 34.8496
LSSL2010#527 | CB-13 | 1031 |  | 2011-02-10 11:43 | 24C | 0.99615 | 3.344e-06 | 2.0862e-06 | 0.99616 | 34.8488 | 0.00013404 | [10 empty] | 34.8488
LSSL2010#528 | CB-13 | 1032 |  | 2011-02-10 11:45 | 24C | 0.9961 | 2.7426e-06 | 1.967e-06 | 0.9961 | 34.8465 | 0.00010865 | [10 empty] | 34.8465
LSSL2010#529 | CB-13 | 1033 |  | 2011-02-10 11:47 | 24C | 0.99617 | 2.8897e-06 | 2.0862e-06 | 0.99618 | 34.8496 | 0.00011517 | [10 empty] | 34.8496
LSSL2010#530 | CB-13 | 1034 |  | 2011-02-10 11:50 | 24C | 0.99486 | 2.4908e-06 | 2.0266e-06 | 0.99486 | 34.7979 | 9.557e-05 | [10 empty] | 34.7979
LSSL2010#531 | CB-13 | 1035 |  | 2011-02-10 11:52 | 24C | 0.99486 | 5.635e-06 | 1.967e-06 | 0.99486 | 34.7979 | 0.00021988 | [10 empty] | 34.7979
LSSL2010#532 | CB-13 | 1036 |  | 2011-02-10 11:54 | 24C | 0.99495 | 2.4908e-06 | 1.967e-06 | 0.99496 | 34.8016 | 9.7246e-05 | SZ- Changed from 1037 to 1036 | [9 empty] | 34.8016
LSSL2010#533 | CB-13 | 1037 |  | 2011-02-10 11:56 | 24C | 0.99504 | 2.2391e-06 | 1.967e-06 | 0.99504 | 34.8048 | 8.7866e-05 | [10 empty] | 34.8048
LSSL2010#534 | CB-13 | 1038 |  | 2011-02-10 11:59 | 24C | 0.98992 | 2.8795e-06 | 2.0266e-06 | 0.98993 | 34.6038 | 0.00011316 | [10 empty] | 34.6038
LSSL2010#535 | CB-13 | 1039 |  | 2011-02-10 12:01 | 24C | 0.98993 | 2.8661e-06 | 2.0266e-06 | 0.98993 | 34.6041 | 0.00011012 | [10 empty] | 34.6041
LSSL2010#536 | CB-13 | 1040 |  | 2011-02-10 12:03 | 24C | 0.98969 | 2.4912e-06 | 2.0266e-06 | 0.98969 | 34.5946 | 9.6599e-05 | [10 empty] | 34.5946
LSSL2010#537 | CB-13 | 1041 |  | 2011-02-10 12:05 | 24C | 0.9897 | 1.1196e-06 | 2.0266e-06 | 0.9897 | 34.595 | 4.265e-05 | [10 empty] | 34.595
LSSL2010#538 | CB-13 | 1042 |  | 2011-02-10 12:08 | 24C | 0.79746 | 2.2391e-06 | 1.967e-06 | 0.79746 | 27.1742 | 8.4455e-05 | [10 empty] | 27.1742
LSSL2010#539 | CB-13 | 1043 |  | 2011-02-10 12:10 | 24C | 0.79753 | 2.7426e-06 | 2.0862e-06 | 0.79753 | 27.1768 | 0.00010449 | [10 empty] | 27.1768
LSSL2010#540 | CB-13 | 1044 |  | 2011-02-10 12:12 | 24C | 0.79755 | 2.4908e-06 | 2.0266e-06 | 0.79755 | 27.1776 | 9.4895e-05 | [10 empty] | 27.1776
LSSL2010#541 | CB-13 | 1045 |  | 2011-02-10 12:14 | 24C | 0.79755 | 2.2259e-06 | 2.0862e-06 | 0.79755 | 27.1778 | 8.2741e-05 | [10 empty] | 27.1778
LSSL2010#99 | CB-13 | 1046 |  | 2011-02-03 11:12 | 24C | 0.99891 | 2.8901e-06 | 2.3961e-05 | 0.99893 | 34.9579 | 0.00011213 | [10 empty] | 34.9579
LSSL2010#100 | CB-13 | 1047 |  | 2011-02-03 11:14 | 24C | 0.9989 | 5.3052e-06 | 2.4021e-05 | 0.99892 | 34.9575 | 0.00020517 |  | 2011-02-09 16:49 | 24C | 0.99887 | 2.8543e-06 | 1.967e-06 | 0.99887 | 34.9556 | 0.00011192 | 0 | 0.0019302
LSSL2010#101 | CB-13 | 1048 |  | 2011-02-03 11:16 | 24C | 0.99891 | 3.8619e-06 | 2.4021e-05 | 0.99894 | 34.9581 | 0.00015532 | [10 empty] | 34.9581
LSSL2010#102 | CB-13 | 1049 |  | 2011-02-03 11:19 | 24C | 0.9985 | 3.1415e-06 | 2.3961e-05 | 0.99853 | 34.942 | 0.00012313 | [10 empty] | 34.942
LSSL2010#103 | CB-13 | 1050 |  | 2011-02-03 11:21 | 24C | 0.99782 | 2.7425e-06 | 2.3961e-05 | 0.99784 | 34.9152 | 0.0001097 | [10 empty] | 34.9152
LSSL2010#104 | CB-13 | 1051 |  | 2011-02-03 11:23 | 24C | 0.99692 | 4.9037e-06 | 2.4021e-05 | 0.99694 | 34.8796 | 0.00019225 | [10 empty] | 34.8796
LSSL2010#105 | CB-13 | 1052 |  | 2011-02-03 11:25 | 24C | 0.99658 | 9.3855e-06 | 2.4021e-05 | 0.9966 | 34.8662 | 0.00036876 | [10 empty] | 34.8662
LSSL2010#106 | CB-13 | 1053 |  | 2011-02-03 11:28 | 24C | 0.99611 | 5.1551e-06 | 2.408e-05 | 0.99614 | 34.8481 | 0.00020366 | [10 empty] | 34.8481
LSSL2010#107 | CB-13 | 1054 |  | 2011-02-03 11:30 | 24C | 0.99602 | 6.8489e-06 | 2.3961e-05 | 0.99605 | 34.8444 | 0.00027 | [10 empty] | 34.8444
LSSL2010#108 | CB-13 | 1055 |  | 2011-02-03 11:32 | 24C | 0.99587 | 6.3301e-06 | 2.3961e-05 | 0.9959 | 34.8385 | 0.00024835 |  | 2011-02-09 17:02 | 24C | 0.99643 | 2.4908e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 9.9543e-05 | 0 | -0.021126
LSSL2010#109 | CB-13 | 1056 |  | 2011-02-03 11:35 | 24C | 0.99508 | 7.4722e-06 | 2.4021e-05 | 0.9951 | 34.8072 | 0.00029601 | [10 empty] | 34.8072
LSSL2010#110 | CB-13 | 1057 |  | 2011-02-03 11:37 | 24C | 0.99387 | 2.7301e-06 | 2.4021e-05 | 0.9939 | 34.7599 | 0.00010708 | [10 empty] | 34.7599
LSSL2010#111 | CB-13 | 1058 |  | 2011-02-03 11:39 | 24C | 0.98949 | 2.7426e-06 | 2.408e-05 | 0.98952 | 34.5877 | 0.00010241 | [10 empty] | 34.5877
LSSL2010#112 | CB-13 | 1059 |  | 2011-02-03 11:42 | 24C | 0.98177 | 5.4721e-06 | 2.3961e-05 | 0.98179 | 34.2844 | 0.0002165 | [10 empty] | 34.2844
LSSL2010#113 | CB-13 | 1060 |  | 2011-02-03 11:44 | 24C | 0.97318 | 3.6104e-06 | 2.3961e-05 | 0.9732 | 33.9478 | 0.00014182 | [10 empty] | 33.9478
LSSL2010#114 | CB-13 | 1061 |  | 2011-02-03 11:47 | 24C | 0.95881 | 4.6063e-06 | 2.3961e-05 | 0.95883 | 33.3857 | 0.00017979 | [10 empty] | 33.3857
LSSL2010#115 | CB-13 | 1062 |  | 2011-02-03 11:49 | 24C | 0.95393 | 3.8756e-06 | 2.3901e-05 | 0.95395 | 33.195 | 0.00014951 | [10 empty] | 33.195
LSSL2010#116 | CB-13 | 1063 |  | 2011-02-03 11:52 | 24C | 0.94382 | 3.392e-06 | 2.4021e-05 | 0.94385 | 32.801 | 0.00013165 | [10 empty] | 32.801
LSSL2010#117 | CB-13 | 1064 |  | 2011-02-03 11:54 | 24C | 0.93725 | 3.8437e-06 | 2.3901e-05 | 0.93728 | 32.5452 | 0.00014723 | [10 empty] | 32.5452
LSSL2010#118 | CB-13 | 1065 |  | 2011-02-03 11:57 | 24C | 0.93194 | 5.7638e-06 | 2.3961e-05 | 0.93197 | 32.3387 | 0.00022378 | [10 empty] | 32.3387
LSSL2010#119 | CB-13 | 1066 |  | 2011-02-03 11:59 | 24C | 0.91967 | 6.3446e-06 | 2.408e-05 | 0.91969 | 31.8618 | 0.00024698 | [10 empty] | 31.8618
LSSL2010#120 | CB-13 | 1067 |  | 2011-02-03 12:01 | 24C | 0.88833 | 7.1397e-06 | 2.3961e-05 | 0.88835 | 30.6498 | 0.00027429 | [10 empty] | 30.6498
LSSL2010#121 | CB-13 | 1068 |  | 2011-02-03 12:03 | 24C | 0.83546 | 8.0422e-06 | 2.3961e-05 | 0.83548 | 28.6207 | 0.00031011 |  | 2011-02-09 17:30 | 24C | 0.83527 | 3.6104e-06 | 2.0266e-06 | 0.83527 | 28.6127 | 0.00013717 | 0 | 0.0080605
LSSL2010#122 | CB-13 | 1069 |  | 2011-02-03 12:06 | 24C | 0.79683 | 1.3713e-06 | 2.4021e-05 | 0.79686 | 27.1512 | 5.2235e-05 | [10 empty] | 27.1512
LSSL2010#568 | CB-15 | 1070 |  | 2011-02-10 13:59 | 24C | 0.99427 | 2.4909e-06 | 1.967e-06 | 0.99427 | 34.7745 | 9.869e-05 |  | 2011-02-09 17:06 | 24C | 0.99425 | 2.7425e-06 | 1.9073e-06 | 0.99425 | 34.7739 | 0.00010663 | 0 | 0.00064087
LSSL2010#569 | CB-15 | 1071 |  | 2011-02-10 14:02 | 24C | 0.9917 | 3.2144e-06 | 1.967e-06 | 0.9917 | 34.6734 | 0.00012623 | [10 empty] | 34.6734
LSSL2010#570 | CB-15 | 1072 |  | 2011-02-10 14:04 | 24C | 0.9896 | 3.3587e-06 | 1.9073e-06 | 0.9896 | 34.5912 | 0.00012927 | [10 empty] | 34.5912
LSSL2010#571 | CB-15 | 1073 |  | 2011-02-10 14:06 | 24C | 0.98732 | 2.128e-06 | 1.9073e-06 | 0.98732 | 34.5016 | 8.1621e-05 | [10 empty] | 34.5016
LSSL2010#572 | CB-15 | 1074 |  | 2011-02-10 14:08 | 24C | 0.9843 | 2.4909e-06 | 1.967e-06 | 0.9843 | 34.3829 | 9.87e-05 | [10 empty] | 34.3829
LSSL2010#573 | CB-15 | 1075 |  | 2011-02-10 14:13 | 24C | 0.98072 | 2.2355e-05 | 2.0266e-06 | 0.98072 | 34.2426 | 0.000875 | [10 empty] | 34.2426
LSSL2010#574 | CB-15 | 1076 |  | 2011-02-10 14:59 | 24C | 0.9727 | 3.1415e-06 | 1.967e-06 | 0.9727 | 33.9281 | 0.00012621 |  | 2011-02-09 12:30 | 24C | 0.97279 | 2.4909e-06 | 1.9073e-06 | 0.97279 | 33.9317 | 9.7029e-05 | 0 | -0.0035477
LSSL2010#575 | CB-15 | 1077 |  | 2011-02-10 15:02 | 24C | 0.96704 | 2.8897e-06 | 2.0266e-06 | 0.96704 | 33.7065 | 0.00011386 | [10 empty] | 33.7065
LSSL2010#576 | CB-15 | 1078 |  | 2011-02-10 15:04 | 24C | 0.9613 | 3.2145e-06 | 1.967e-06 | 0.9613 | 33.482 | 0.000126 | [10 empty] | 33.482
LSSL2010#577 | CB-15 | 1079 |  | 2011-02-10 15:07 | 24C | 0.95605 | 1.1196e-06 | 2.0266e-06 | 0.95605 | 33.2769 | 4.2703e-05 | [10 empty] | 33.2769
LSSL2010#578 | CB-15 | 1080 |  | 2011-02-10 15:09 | 24C | 0.95287 | 2.7425e-06 | 2.0266e-06 | 0.95287 | 33.1527 | 0.00010343 | [10 empty] | 33.1527
LSSL2010#579 | CB-15 | 1081 |  | 2011-02-10 15:11 | 24C | 0.94946 | 2.4908e-06 | 2.0266e-06 | 0.94946 | 33.0199 | 9.87e-05 | [10 empty] | 33.0199
LSSL2010#580 | CB-15 | 1082 |  | 2011-02-10 15:14 | 24C | 0.9466 | 2.8897e-06 | 1.967e-06 | 0.9466 | 32.9082 | 0.00011145 | [10 empty] | 32.9082
LSSL2010#581 | CB-15 | 1083 |  | 2011-02-10 15:16 | 24C | 0.94145 | 2.4912e-06 | 1.967e-06 | 0.94145 | 32.7076 | 9.4885e-05 | [10 empty] | 32.7076
LSSL2010#582 | CB-15 | 1084 |  | 2011-02-10 15:18 | 24C | 0.93589 | 2.8904e-06 | 1.9073e-06 | 0.93589 | 32.4911 | 0.00011403 | [10 empty] | 32.4911
LSSL2010#583 | CB-15 | 1085 |  | 2011-02-10 15:34 | 24C | 0.93012 | 2.4911e-06 | 2.0862e-06 | 0.93012 | 32.2668 | 9.9553e-05 |  | 2011-02-09 12:35 | 24C | 0.93015 | 2.7262e-06 | 2.0266e-06 | 0.93015 | 32.2681 | 0.00010659 | 0 | -0.0012817
LSSL2010#584 | CB-15 | 1086 |  | 2011-02-10 15:36 | 24C | 0.9283 | 3.1148e-06 | 2.0266e-06 | 0.92831 | 32.1963 | 0.00012313 | [10 empty] | 32.1963
LSSL2010#585 | CB-15 | 1087 |  | 2011-02-10 15:38 | 24C | 0.92934 | 3.3591e-06 | 1.9073e-06 | 0.92935 | 32.2367 | 0.00012795 | [10 empty] | 32.2367
LSSL2010#586 | CB-15 | 1088 |  | 2011-02-10 15:41 | 24C | 0.91925 | 2.7425e-06 | 2.0266e-06 | 0.91925 | 31.8449 | 0.00010447 | [10 empty] | 31.8449
LSSL2010#587 | CB-15 | 1089 |  | 2011-02-10 15:42 | 24C | 0.91104 | 4.2073e-06 | 2.0266e-06 | 0.91104 | 31.5266 | 0.00016433 | [10 empty] | 31.5266
LSSL2010#588 | CB-15 | 1090 |  | 2011-02-10 15:45 | 24C | 0.88018 | 4.1136e-06 | 2.0266e-06 | 0.88018 | 30.3349 | 0.00015711 | [10 empty] | 30.3349
LSSL2010#589 | CB-15 | 1091 |  | 2011-02-10 15:47 | 24C | 0.86349 | 3.9638e-06 | 2.0862e-06 | 0.86349 | 29.6931 | 0.0001517 | [10 empty] | 29.6931
LSSL2010#590 | CB-15 | 1092 |  | 2011-02-10 15:49 | 24C | 0.83638 | 4.2256e-06 | 2.0266e-06 | 0.83638 | 28.6552 | 0.00016216 | [10 empty] | 28.6552
LSSL2010#591 | CB-15 | 1093 |  | 2011-02-10 15:51 | 24C | 0.79389 | 3.2141e-06 | 2.0266e-06 | 0.79389 | 27.039 | 0.0001238 | [10 empty] | 27.039
LSSL2010#668 | CB-15 | 1094 |  | 2011-02-11 10:47 | 24C | 0.99702 | 1.3713e-06 | 4.8876e-06 | 0.99702 | 34.8828 | 5.2235e-05 | [10 empty] | 34.8828
LSSL2010#669 | CB-15 | 1095 |  | 2011-02-11 10:49 | 24C | 0.99692 | 3.3595e-06 | 4.9472e-06 | 0.99693 | 34.8792 | 0.00013116 | [10 empty] | 34.8792
LSSL2010#670 | CB-15 | 1096 |  | 2011-02-11 10:51 | 24C | 0.99702 | 5.4532e-06 | 5.0068e-06 | 0.99703 | 34.8831 | 0.00021569 | [10 empty] | 34.8831
LSSL2010#671 | CB-15 | 1097 |  | 2011-02-11 10:53 | 24C | 0.99701 | 4.9165e-06 | 5.0068e-06 | 0.99702 | 34.8827 | 0.00019367 | [10 empty] | 34.8827
LSSL2010#672 | CB-15 | 1098 |  | 2011-02-11 10:55 | 24C | 0.99656 | 2.7425e-06 | 5.0068e-06 | 0.99656 | 34.8648 | 0.00010657 | [10 empty] | 34.8648
LSSL2010#673 | CB-15 | 1099 |  | 2011-02-11 10:58 | 24C | 0.9966 | 2.2267e-06 | 4.9472e-06 | 0.9966 | 34.8662 | 8.7013e-05 | [10 empty] | 34.8662
LSSL2010#674 | CB-15 | 1100 |  | 2011-02-11 10:59 | 24C | 0.99655 | 3.4389e-06 | 5.0068e-06 | 0.99656 | 34.8646 | 0.00013478 | [10 empty] | 34.8646
LSSL2010#675 | CB-15 | 1101 |  | 2011-02-11 11:02 | 24C | 0.99659 | 4.6353e-06 | 5.0068e-06 | 0.99659 | 34.8659 | 0.00017966 | [10 empty] | 34.8659
LSSL2010#676 | CB-15 | 1102 |  | 2011-02-11 11:04 | 24C | 0.99617 | 3.9742e-06 | 5.0068e-06 | 0.99618 | 34.8497 | 0.0001555 | [10 empty] | 34.8497
LSSL2010#677 | CB-15 | 1103 |  | 2011-02-11 11:06 | 24C | 0.99611 | 1.1532e-06 | 4.8876e-06 | 0.99612 | 34.8472 | 4.4374e-05 | [10 empty] | 34.8472
LSSL2010#678 | CB-15 | 1104 |  | 2011-02-11 11:07 | 24C | 0.99617 | 4.4783e-06 | 4.8876e-06 | 0.99617 | 34.8493 | 0.00017404 | [10 empty] | 34.8493
LSSL2010#679 | CB-15 | 1105 |  | 2011-02-11 11:09 | 24C | 0.99622 | 3.1421e-06 | 4.8876e-06 | 0.99623 | 34.8516 | 0.00012383 | [10 empty] | 34.8516
LSSL2010#680 | CB-15 | 1106 |  | 2011-02-11 11:11 | 24C | 0.99534 | 3.6111e-06 | 4.8876e-06 | 0.99534 | 34.8167 | 0.00014174 | [10 empty] | 34.8167
LSSL2010#681 | CB-15 | 1107 |  | 2011-02-11 11:13 | 24C | 0.99538 | 3.865e-06 | 5.0068e-06 | 0.99539 | 34.8185 | 0.00015069 | [10 empty] | 34.8185
LSSL2010#682 | CB-15 | 1108 |  | 2011-02-11 11:15 | 24C | 0.99542 | 3.865e-06 | 5.0068e-06 | 0.99543 | 34.8202 | 0.00015069 | [10 empty] | 34.8202
LSSL2010#683 | CB-15 | 1109 |  | 2011-02-11 11:18 | 24C | 0.99543 | 3.1417e-06 | 5.126e-06 | 0.99544 | 34.8205 | 0.00012202 | [10 empty] | 34.8205
LSSL2010#684 | CB-15 | 1110 |  | 2011-02-11 11:20 | 24C | 0.99009 | 2.4911e-06 | 5.0664e-06 | 0.99009 | 34.6104 | 9.869e-05 | [10 empty] | 34.6104
LSSL2010#685 | CB-15 | 1111 |  | 2011-02-11 11:22 | 24C | 0.9907 | 2.4908e-06 | 5.0068e-06 | 0.9907 | 34.6344 | 9.7462e-05 | [10 empty] | 34.6344
LSSL2010#686 | CB-15 | 1112 |  | 2011-02-11 11:24 | 24C | 0.98999 | 3.1418e-06 | 5.0068e-06 | 0.99 | 34.6066 | 0.00012421 | [10 empty] | 34.6066
LSSL2010#687 | CB-15 | 1113 |  | 2011-02-11 11:26 | 24C | 0.98976 | 4.0093e-06 | 5.0664e-06 | 0.98977 | 34.5977 | 0.00015704 | [10 empty] | 34.5977
LSSL2010#688 | CB-15 | 1114 |  | 2011-02-11 11:28 | 24C | 0.79146 | 3.1415e-06 | 5.0068e-06 | 0.79146 | 26.947 | 0.00011629 | [10 empty] | 26.947
LSSL2010#689 | CB-15 | 1115 |  | 2011-02-11 11:30 | 24C | 0.78969 | 2.7425e-06 | 5.0068e-06 | 0.78969 | 26.8799 | 9.8205e-05 | [10 empty] | 26.8799
LSSL2010#690 | CB-15 | 1116 |  | 2011-02-11 11:33 | 24C | 0.78938 | 2.2391e-06 | 5.0068e-06 | 0.78938 | 26.8682 | 8.4895e-05 | [10 empty] | 26.8682
LSSL2010#691 | CB-15 | 1117 |  | 2011-02-11 11:35 | 24C | 0.78937 | 2.8897e-06 | 5.0068e-06 | 0.78937 | 26.8679 | 0.00010594 | [10 empty] | 26.8679
LSSL2010#74 | CB-15 | 1118 |  | 2011-02-03 09:32 | 24C | 0.9988 | 5.871e-06 | 2.408e-05 | 0.99882 | 34.9536 | 0.00023096 | [10 empty] | 34.9536
LSSL2010#75 | CB-15 | 1119 |  | 2011-02-03 09:35 | 24C | 0.99863 | 3.4658e-06 | 2.4021e-05 | 0.99866 | 34.9471 | 0.00013341 | [10 empty] | 34.9471
LSSL2010#76 | CB-15 | 1120 |  | 2011-02-03 09:37 | 24C | 0.99875 | 4.3223e-06 | 2.3961e-05 | 0.99878 | 34.9519 | 0.00017063 | [10 empty] | 34.9519
LSSL2010#77 | CB-15 | 1121 |  | 2011-02-03 09:39 | 24C | 0.99847 | 5.9206e-06 | 2.408e-05 | 0.99849 | 34.9406 | 0.00023374 | [10 empty] | 34.9406
LSSL2010#78 | CB-15 | 1122 |  | 2011-02-03 09:42 | 24C | 0.9978 | 4.4373e-06 | 2.408e-05 | 0.99782 | 34.9143 | 0.00017504 | [10 empty] | 34.9143
LSSL2010#79 | CB-15 | 1123 |  | 2011-02-03 09:45 | 24C | 0.99695 | 3.5949e-06 | 2.3961e-05 | 0.99697 | 34.8808 | 0.00014346 | [10 empty] | 34.8808
LSSL2010#80 | CB-15 | 1124 |  | 2011-02-03 09:47 | 24C | 0.99664 | 3.6104e-06 | 2.3961e-05 | 0.99666 | 34.8686 | 0.00014379 | [10 empty] | 34.8686
LSSL2010#81 | CB-15 | 1125 |  | 2011-02-03 09:49 | 24C | 0.99629 | 6.5199e-06 | 2.3961e-05 | 0.99632 | 34.855 | 0.00025501 | [10 empty] | 34.855
LSSL2010#82 | CB-15 | 1126 |  | 2011-02-03 09:51 | 24C | 0.99623 | 2.2396e-06 | 2.3961e-05 | 0.99625 | 34.8525 | 8.9681e-05 | [10 empty] | 34.8525
LSSL2010#83 | CB-15 | 1127 |  | 2011-02-03 09:54 | 24C | 0.99575 | 2.8901e-06 | 2.3901e-05 | 0.99577 | 34.8337 | 0.00011606 |  | 2011-02-09 12:22 | 24C | 0.99574 | 1.3713e-06 | 1.967e-06 | 0.99574 | 34.8323 | 5.4326e-05 | 0 | 0.0014343
LSSL2010#84 | CB-15 | 1128 |  | 2011-02-03 09:56 | 24C | 0.99545 | 2.4908e-06 | 2.3961e-05 | 0.99547 | 34.822 | 9.869e-05 | [10 empty] | 34.822
LSSL2010#85 | CB-15 | 1129 |  | 2011-02-03 09:58 | 24C | 0.9943 | 1.3713e-06 | 2.3961e-05 | 0.99433 | 34.7768 | 5.2235e-05 |  | 2011-02-09 12:24 | 24C | 0.99429 | 2.4909e-06 | 2.0266e-06 | 0.9943 | 34.7756 | 0.00010437 | 0 | 0.0011787
LSSL2010#86 | CB-15 | 1130 |  | 2011-02-03 10:00 | 24C | 0.98943 | 6.6227e-06 | 2.4021e-05 | 0.98946 | 34.5854 | 0.00025908 | [10 empty] | 34.5854
LSSL2010#87 | CB-15 | 1131 |  | 2011-02-03 10:03 | 24C | 0.9761 | 2.4744e-06 | 2.3961e-05 | 0.97613 | 34.0623 | 9.6652e-05 | [10 empty] | 34.0623
LSSL2010#88 | CB-15 | 1132 |  | 2011-02-03 10:05 | 24C | 0.97646 | 3.781e-06 | 2.4021e-05 | 0.97648 | 34.0763 | 0.00014929 | [10 empty] | 34.0763
LSSL2010#89 | CB-15 | 1133 |  | 2011-02-03 10:07 | 24C | 0.962 | 5.8084e-06 | 2.3901e-05 | 0.96203 | 33.5105 | 0.00022729 | [10 empty] | 33.5105
LSSL2010#90 | CB-15 | 1134 |  | 2011-02-03 10:10 | 24C | 0.95338 | 2.2392e-06 | 2.3961e-05 | 0.95341 | 33.1737 | 8.7015e-05 |  | 2011-02-09 11:44 | 24C | 0.95343 | 1.7026e-05 | 1.9073e-06 | 0.95343 | 33.1749 | 0.0006637 | 0 | -0.0011559
LSSL2010#91 | CB-15 | 1135 |  | 2011-02-03 10:12 | 24C | 0.94763 | 9.0252e-06 | 2.3961e-05 | 0.94765 | 32.9494 | 0.00035108 | [10 empty] | 32.9494
LSSL2010#92 | CB-15 | 1136 |  | 2011-02-03 10:15 | 24C | 0.9388 | 7.1241e-06 | 2.4021e-05 | 0.93882 | 32.6053 | 0.00027659 | [10 empty] | 32.6053
LSSL2010#93 | CB-15 | 1137 |  | 2011-02-03 10:17 | 24C | 0.93319 | 6.8045e-06 | 2.4021e-05 | 0.93321 | 32.3871 | 0.00026278 | [10 empty] | 32.3871
LSSL2010#94 | CB-15 | 1138 |  | 2011-02-03 10:19 | 24C | 0.92299 | 4.7007e-06 | 2.3901e-05 | 0.92302 | 31.9909 | 0.00018663 | [10 empty] | 31.9909
LSSL2010#95 | CB-15 | 1139 |  | 2011-02-03 10:22 | 24C | 0.86631 | 6.2687e-06 | 2.408e-05 | 0.86634 | 29.8025 | 0.00024276 | [10 empty] | 29.8025
LSSL2010#96 | CB-15 | 1140 |  | 2011-02-03 10:25 | 24C | 0.83407 | 3.8489e-06 | 2.3961e-05 | 0.83409 | 28.5676 | 0.00014713 | [10 empty] | 28.5676
LSSL2010#97 | CB-15 | 1141 |  | 2011-02-03 10:27 | 24C | 0.7848 | 8.9936e-06 | 2.3901e-05 | 0.78482 | 26.6956 | 0.00033962 | [10 empty] | 26.6956
LSSL2010#149 | CB-17 | 1142 |  | 2011-02-03 13:45 | 24C | 0.99886 | 7.8373e-06 | 2.3961e-05 | 0.99888 | 34.9561 | 0.00030983 | [10 empty] | 34.9561
LSSL2010#150 | CB-17 | 1143 |  | 2011-02-03 13:47 | 24C | 0.99886 | 4.2256e-06 | 2.4021e-05 | 0.99888 | 34.956 | 0.00016938 |  | 2011-02-09 12:12 | 24C | 0.99889 | 2.8792e-06 | 2.0266e-06 | 0.99889 | 34.9563 | 0.00011317 | 0 | -0.00031281
LSSL2010#151 | CB-17 | 1144 |  | 2011-02-03 13:49 | 24C | 0.99876 | 3.8623e-06 | 2.3961e-05 | 0.99879 | 34.9523 | 0.00015198 | [10 empty] | 34.9523
LSSL2010#152 | CB-17 | 1145 |  | 2011-02-03 13:52 | 24C | 0.99848 | 7.277e-06 | 2.4021e-05 | 0.9985 | 34.941 | 0.00028506 | [10 empty] | 34.941
LSSL2010#153 | CB-17 | 1146 |  | 2011-02-03 13:54 | 24C | 0.9977 | 4.1602e-06 | 2.3961e-05 | 0.99773 | 34.9105 | 0.00016507 | [10 empty] | 34.9105
LSSL2010#154 | CB-17 | 1147 |  | 2011-02-03 13:57 | 24C | 0.99696 | 6.6094e-06 | 2.408e-05 | 0.99698 | 34.8813 | 0.00025993 | [10 empty] | 34.8813
LSSL2010#155 | CB-17 | 1148 |  | 2011-02-03 13:59 | 24C | 0.99672 | 5.1679e-06 | 2.408e-05 | 0.99674 | 34.8719 | 0.000204 |  | 2011-02-09 12:17 | 24C | 0.9966 | 2.4908e-06 | 2.0266e-06 | 0.99661 | 34.8665 | 9.6642e-05 | SZ- Correct there was no sample 1149 taken | 0.0054207
LSSL2010#156 | CB-17 | 1150 |  | 2011-02-03 14:02 | 24C | 0.9959 | 5.346e-06 | 2.4021e-05 | 0.99592 | 34.8395 | 0.00020888 | [10 empty] | 34.8395
LSSL2010#157 | CB-17 | 1151 |  | 2011-02-03 14:04 | 24C | 0.99564 | 5.6518e-06 | 2.3961e-05 | 0.99566 | 34.8293 | 0.00022341 | [10 empty] | 34.8293
LSSL2010#158 | CB-17 | 1152 |  | 2011-02-03 15:15 | 24C | 0.99524 | 7.8752e-06 | 2.3961e-05 | 0.99527 | 34.8137 | 0.00030593 | [10 empty] | 34.8137
LSSL2010#159 | CB-17 | 1153 |  | 2011-02-03 15:18 | 24C | 0.99453 | 7.1644e-06 | 2.3901e-05 | 0.99455 | 34.7858 | 0.00028045 | [10 empty] | 34.7858
LSSL2010#160 | CB-17 | 1154 |  | 2011-02-03 15:20 | 24C | 0.99344 | 4.9004e-06 | 2.3961e-05 | 0.99346 | 34.7429 | 0.0001893 | [10 empty] | 34.7429
LSSL2010#161 | CB-17 | 1155 |  | 2011-02-03 15:23 | 24C | 0.98595 | 2.2398e-06 | 2.4021e-05 | 0.98597 | 34.4485 | 9.0443e-05 | [10 empty] | 34.4485
LSSL2010#162 | CB-17 | 1156 |  | 2011-02-03 15:25 | 24C | 0.97925 | 2.8897e-06 | 2.3961e-05 | 0.97928 | 34.1858 | 0.00011403 | [10 empty] | 34.1858
LSSL2010#163 | CB-17 | 1157 |  | 2011-02-03 15:28 | 24C | 0.96571 | 2.4909e-06 | 2.4021e-05 | 0.96574 | 33.6555 | 9.3081e-05 | [10 empty] | 33.6555
LSSL2010#164 | CB-17 | 1158 |  | 2011-02-03 15:31 | 24C | 0.94782 | 4.1138e-06 | 2.3961e-05 | 0.94784 | 32.9568 | 0.00015985 | [10 empty] | 32.9568
LSSL2010#165 | CB-17 | 1159 |  | 2011-02-03 15:33 | 24C | 0.94859 | 5.1485e-06 | 2.4021e-05 | 0.94861 | 32.9867 | 0.00019999 | [10 empty] | 32.9867
LSSL2010#166 | CB-17 | 1160 |  | 2011-02-03 15:36 | 24C | 0.94058 | 4.1859e-06 | 2.3961e-05 | 0.94061 | 32.6748 | 0.00016257 | [10 empty] | 32.6748
LSSL2010#167 | CB-17 | 1161 |  | 2011-02-03 15:38 | 24C | 0.93386 | 4.6353e-06 | 2.3961e-05 | 0.93388 | 32.4132 | 0.00018175 | [10 empty] | 32.4132
LSSL2010#168 | CB-17 | 1162 |  | 2011-02-03 15:41 | 24C | 0.91953 | 2.7757e-06 | 2.3961e-05 | 0.91955 | 31.8564 | 0.00010815 |  | 2011-02-09 12:41 | 24C | 0.91951 | 3.4666e-06 | 1.967e-06 | 0.91951 | 31.855 | 0.00013693 | 0 | 0.0013962
LSSL2010#169 | CB-17 | 1163 |  | 2011-02-03 15:43 | 24C | 0.89152 | 5.8894e-06 | 2.3961e-05 | 0.89154 | 30.7729 | 0.00022498 | [10 empty] | 30.7729
LSSL2010#170 | CB-17 | 1164 |  | 2011-02-03 15:45 | 24C | 0.80607 | 4.1603e-06 | 2.4021e-05 | 0.8061 | 27.5018 | 0.00015864 | [10 empty] | 27.5018
LSSL2010#171 | CB-17 | 1165 |  | 2011-02-03 15:47 | 24C | 0.75903 | 3.2516e-06 | 2.408e-05 | 0.75906 | 25.7237 | 0.00012176 | [10 empty] | 25.7237
LSSL2010#643 | PP-7 | 1166 |  | 2011-02-11 09:29 | 24C | 0.99716 | 1.4046e-06 | 5.0068e-06 | 0.99717 | 34.8885 | 5.5371e-05 | [10 empty] | 34.8885
LSSL2010#644 | PP-7 | 1167 |  | 2011-02-11 09:32 | 24C | 0.99715 | 1.515e-06 | 4.8876e-06 | 0.99716 | 34.8882 | 5.8044e-05 | [10 empty] | 34.8882
LSSL2010#645 | PP-7 | 1168 |  | 2011-02-11 09:36 | 24C | 0.99713 | 3.191e-06 | 5.0068e-06 | 0.99713 | 34.8871 | 0.00012429 | [10 empty] | 34.8871
LSSL2010#646 | PP-7 | 1169 |  | 2011-02-11 09:38 | 24C | 0.99705 | 3.1414e-06 | 5.0068e-06 | 0.99706 | 34.8842 | 0.00012621 | [10 empty] | 34.8842
LSSL2010#647 | PP-7 | 1170 |  | 2011-02-11 09:40 | 24C | 0.99689 | 3.2141e-06 | 5.0068e-06 | 0.9969 | 34.8779 | 0.00012432 | [10 empty] | 34.8779
LSSL2010#648 | PP-7 | 1171 |  | 2011-02-11 09:42 | 24C | 0.99673 | 3.579e-06 | 5.0068e-06 | 0.99674 | 34.8716 | 0.00013979 | [10 empty] | 34.8716
LSSL2010#649 | PP-7 | 1172 |  | 2011-02-11 09:44 | 24C | 0.99668 | 3.4657e-06 | 4.9472e-06 | 0.99668 | 34.8695 | 0.00013203 | [10 empty] | 34.8695
LSSL2010#650 | PP-7 | 1173 |  | 2011-02-11 09:46 | 24C | 0.99705 | 2.1278e-06 | 5.0068e-06 | 0.99705 | 34.8839 | 8.1638e-05 | [10 empty] | 34.8839
LSSL2010#651 | PP-7 | 1174 |  | 2011-02-11 09:49 | 24C | 0.99662 | 1.3713e-06 | 5.0068e-06 | 0.99663 | 34.8673 | 5.5371e-05 | [10 empty] | 34.8673
LSSL2010#652 | PP-7 | 1175 |  | 2011-02-11 09:51 | 24C | 0.99638 | 2.4908e-06 | 4.9472e-06 | 0.99638 | 34.8576 | 9.6599e-05 | [10 empty] | 34.8576
LSSL2010#653 | PP-7 | 1176 |  | 2011-02-11 09:54 | 24C | 0.99638 | 3.1535e-06 | 5.0068e-06 | 0.99639 | 34.8579 | 0.00012228 | [10 empty] | 34.8579
LSSL2010#654 | PP-7 | 1177 |  | 2011-02-11 09:57 | 24C | 0.99642 | 2.4925e-06 | 4.9472e-06 | 0.99642 | 34.8592 | 9.869e-05 | [10 empty] | 34.8592
LSSL2010#655 | PP-7 | 1178 |  | 2011-02-11 09:59 | 24C | 0.99577 | 3.6104e-06 | 5.0068e-06 | 0.99578 | 34.8339 | 0.00014762 | [10 empty] | 34.8339
LSSL2010#656 | PP-7 | 1179 |  | 2011-02-11 10:02 | 24C | 0.99538 | 3.2141e-06 | 5.0664e-06 | 0.99539 | 34.8186 | 0.00012624 | [10 empty] | 34.8186
LSSL2010#657 | PP-7 | 1180 |  | 2011-02-11 10:04 | 24C | 0.99543 | 2.4911e-06 | 5.0664e-06 | 0.99543 | 34.8204 | 9.6433e-05 | [10 empty] | 34.8204
LSSL2010#658 | PP-7 | 1181 |  | 2011-02-11 10:06 | 24C | 0.9954 | 2.4909e-06 | 4.9472e-06 | 0.9954 | 34.8191 | 9.4885e-05 | [10 empty] | 34.8191
LSSL2010#659 | PP-7 | 1182 |  | 2011-02-11 10:09 | 24C | 0.99542 | 2.4908e-06 | 5.0664e-06 | 0.99543 | 34.8201 | 0.0001006 | [10 empty] | 34.8201
LSSL2010#660 | PP-7 | 1183 |  | 2011-02-11 10:10 | 24C | 0.99067 | 2.7425e-06 | 4.8876e-06 | 0.99067 | 34.6331 | 0.00010659 | [10 empty] | 34.6331
LSSL2010#661 | PP-7 | 1184 |  | 2011-02-11 10:12 | 24C | 0.99064 | 3.8647e-06 | 5.0068e-06 | 0.99064 | 34.632 | 0.00015066 | [10 empty] | 34.632
LSSL2010#662 | PP-7 | 1185 |  | 2011-02-11 10:29 | 24C | 0.99068 | 3.8542e-06 | 5.0068e-06 | 0.99068 | 34.6336 | 0.00015136 | [10 empty] | 34.6336
LSSL2010#663 | PP-7 | 1186 |  | 2011-02-11 10:31 | 24C | 0.99038 | 4.3336e-06 | 4.9472e-06 | 0.99039 | 34.6219 | 0.00016864 | [10 empty] | 34.6219
LSSL2010#664 | PP-7 | 1187 |  | 2011-02-11 10:33 | 24C | 0.94388 | 3.6081e-06 | 5.0068e-06 | 0.94389 | 32.8026 | 0.00013993 | [10 empty] | 32.8026
LSSL2010#665 | PP-7 | 1188 |  | 2011-02-11 10:35 | 24C | 0.8738 | 4.336e-06 | 4.9472e-06 | 0.8738 | 30.0895 | 0.0001679 | [10 empty] | 30.0895
LSSL2010#666 | PP-7 | 1189 |  | 2011-02-11 10:37 | 24C | 0.77705 | 3.1414e-06 | 5.0068e-06 | 0.77705 | 26.4021 | 0.00011629 | [10 empty] | 26.4021
LSSL2010#543 | PP-7 | 1190 |  | 2011-02-10 12:55 | 24C | 0.99886 | 3.466e-06 | 2.0266e-06 | 0.99886 | 34.9552 | 0.0001355 | [10 empty] | 34.9552
LSSL2010#544 | PP-7 | 1191 |  | 2011-02-10 12:59 | 24C | 0.99887 | 3.4658e-06 | 1.967e-06 | 0.99887 | 34.9557 | 0.0001334 | [10 empty] | 34.9557
LSSL2010#545 | PP-7 | 1192 |  | 2011-02-10 13:01 | 24C | 0.9988 | 1.3712e-06 | 2.0266e-06 | 0.9988 | 34.9529 | 5.4326e-05 | [10 empty] | 34.9529
LSSL2010#546 | PP-7 | 1193 |  | 2011-02-10 13:04 | 24C | 0.99855 | 3.466e-06 | 1.967e-06 | 0.99855 | 34.9428 | 0.00013822 | [10 empty] | 34.9428
LSSL2010#547 | PP-7 | 1194 |  | 2011-02-10 13:05 | 24C | 0.99789 | 3.4657e-06 | 1.967e-06 | 0.99789 | 34.9171 | 0.00013595 | [10 empty] | 34.9171
LSSL2010#548 | PP-7 | 1195 |  | 2011-02-10 13:09 | 24C | 0.99702 | 3.4553e-06 | 1.967e-06 | 0.99703 | 34.883 | 0.00013731 | [10 empty] | 34.883
LSSL2010#549 | PP-7 | 1196 |  | 2011-02-10 13:11 | 24C | 0.99667 | 3.1414e-06 | 1.967e-06 | 0.99667 | 34.8689 | 0.00012312 |  | 2011-02-09 12:20 | 24C | 0.9967 | 4.9165e-06 | 2.0266e-06 | 0.9967 | 34.8702 | 0.00019397 | 0 | -0.0013008
LSSL2010#550 | PP-7 | 1197 |  | 2011-02-10 13:13 | 24C | 0.99634 | 1.1532e-06 | 2.1458e-06 | 0.99635 | 34.8562 | 4.4374e-05 | [10 empty] | 34.8562
LSSL2010#551 | PP-7 | 1198 |  | 2011-02-10 13:15 | 24C | 0.99616 | 3.3304e-06 | 2.0266e-06 | 0.99616 | 34.8489 | 0.0001349 | [10 empty] | 34.8489
LSSL2010#552 | PP-7 | 1199 |  | 2011-02-10 13:17 | 24C | 0.99568 | 2.4777e-06 | 1.9073e-06 | 0.99568 | 34.8301 | 9.9759e-05 | [10 empty] | 34.8301
LSSL2010#553 | PP-7 | 1200 |  | 2011-02-10 13:19 | 24C | 0.99537 | 2.7425e-06 | 2.0266e-06 | 0.99537 | 34.818 | 0.00010865 | [10 empty] | 34.818
LSSL2010#554 | PP-7 | 1201 |  | 2011-02-10 13:21 | 24C | 0.9944 | 2.7264e-06 | 2.0266e-06 | 0.9944 | 34.7798 | 0.00010656 | [10 empty] | 34.7798
LSSL2010#555 | PP-7 | 1202 |  | 2011-02-10 13:23 | 24C | 0.98973 | 3.6105e-06 | 2.0266e-06 | 0.98973 | 34.5963 | 0.00014356 | [10 empty] | 34.5963
LSSL2010#556 | PP-7 | 1203 |  | 2011-02-10 13:25 | 24C | 0.98361 | 2.7426e-06 | 2.0266e-06 | 0.98362 | 34.356 | 0.00010451 | [10 empty] | 34.356
LSSL2010#557 | PP-7 | 1204 |  | 2011-02-10 13:27 | 24C | 0.9791 | 3.3594e-06 | 2.0266e-06 | 0.9791 | 34.1788 | 0.00013117 |  | 2011-02-09 15:58 | 24C | 0.97918 | 2.7425e-06 | 2.0266e-06 | 0.97918 | 34.182 | 0.00010865 | 0 | -0.003231
LSSL2010#558 | PP-7 | 1205 |  | 2011-02-10 13:30 | 24C | 0.96774 | 3.2141e-06 | 2.0266e-06 | 0.96774 | 33.7339 | 0.00012528 | [10 empty] | 33.7339
LSSL2010#559 | PP-7 | 1206 |  | 2011-02-10 13:32 | 24C | 0.95254 | 2.4908e-06 | 2.0266e-06 | 0.95254 | 33.1401 | 9.4885e-05 | [10 empty] | 33.1401
LSSL2010#560 | PP-7 | 1207 |  | 2011-02-10 13:34 | 24C | 0.94337 | 2.2401e-06 | 2.0862e-06 | 0.94337 | 32.7825 | 8.7876e-05 | [10 empty] | 32.7825
LSSL2010#561 | PP-7 | 1208 |  | 2011-02-10 13:36 | 24C | 0.93792 | 2.226e-06 | 2.0266e-06 | 0.93792 | 32.5701 | 8.7876e-05 | [10 empty] | 32.5701
LSSL2010#562 | PP-7 | 1209 |  | 2011-02-10 13:38 | 24C | 0.93129 | 4.9004e-06 | 2.0266e-06 | 0.9313 | 32.3125 | 0.00018835 | [10 empty] | 32.3125
LSSL2010#563 | PP-7 | 1210 |  | 2011-02-10 13:40 | 24C | 0.91919 | 3.4657e-06 | 2.0266e-06 | 0.91919 | 31.8425 | 0.00013331 | [10 empty] | 31.8425
LSSL2010#564 | PP-7 | 1211 |  | 2011-02-10 13:46 | 24C | 0.83636 | 3.1311e-06 | 1.967e-06 | 0.83636 | 28.6543 | 0.00011968 | [10 empty] | 28.6543
LSSL2010#565 | PP-7 | 1212 |  | 2011-02-10 13:48 | 24C | 0.82699 | 2.7425e-06 | 2.0266e-06 | 0.82699 | 28.2969 | 0.00010395 | [10 empty] | 28.2969
LSSL2010#566 | PP-7 | 1213 |  | 2011-02-10 13:50 | 24C | 0.77688 | 2.4908e-06 | 2.0266e-06 | 0.77688 | 26.3955 | 9.1847e-05 |  | 2011-02-09 16:23 | 24C | 0.77691 | 1.1196e-06 | 1.967e-06 | 0.77691 | 26.3966 | 4.3716e-05 | 0 | -0.0011349
LSSL2010#124 | PP-6 | 1214 |  | 2011-02-03 12:29 | 24C | 0.99707 | 2.5749e-06 | 2.3961e-05 | 0.9971 | 34.8857 | 9.9229e-05 | [10 empty] | 34.8857
LSSL2010#125 | PP-6 | 1215 |  | 2011-02-03 12:34 | 24C | 0.99709 | 5.1558e-06 | 2.3961e-05 | 0.99711 | 34.8863 | 0.00020151 | [10 empty] | 34.8863
LSSL2010#126 | PP-6 | 1216 |  | 2011-02-03 12:37 | 24C | 0.99706 | 5.168e-06 | 2.4021e-05 | 0.99708 | 34.8852 | 0.00019939 | [10 empty] | 34.8852
LSSL2010#127 | PP-6 | 1217 |  | 2011-02-03 12:39 | 24C | 0.99707 | 1.3712e-06 | 2.3961e-05 | 0.9971 | 34.8857 | 5.4326e-05 | [10 empty] | 34.8857
LSSL2010#128 | PP-6 | 1218 |  | 2011-02-03 12:42 | 24C | 0.99672 | 6.9617e-06 | 2.4021e-05 | 0.99674 | 34.8719 | 0.00027607 | [10 empty] | 34.8719
LSSL2010#129 | PP-6 | 1219 |  | 2011-02-03 12:44 | 24C | 0.99653 | 5.5668e-06 | 2.408e-05 | 0.99656 | 34.8646 | 0.00021841 | [10 empty] | 34.8646
LSSL2010#130 | PP-6 | 1220 |  | 2011-02-03 12:46 | 24C | 0.99679 | 1.0181e-05 | 2.4021e-05 | 0.99681 | 34.8746 | 0.00040166 | [10 empty] | 34.8746
LSSL2010#131 | PP-6 | 1221 |  | 2011-02-03 12:48 | 24C | 0.9967 | 3.3588e-06 | 2.4021e-05 | 0.99672 | 34.871 | 0.00012971 | [10 empty] | 34.871
LSSL2010#132 | PP-6 | 1222 |  | 2011-02-03 12:51 | 24C | 0.99627 | 2.2394e-06 | 2.408e-05 | 0.99629 | 34.854 | 8.7024e-05 | [10 empty] | 34.854
LSSL2010#133 | PP-6 | 1223 |  | 2011-02-03 12:53 | 24C | 0.99637 | 4.1735e-06 | 2.408e-05 | 0.9964 | 34.8582 | 0.00016273 | [10 empty] | 34.8582
LSSL2010#134 | PP-6 | 1224 |  | 2011-02-03 12:56 | 24C | 0.99637 | 5.2931e-06 | 2.3961e-05 | 0.99639 | 34.8582 | 0.00020805 | [10 empty] | 34.8582
LSSL2010#135 | PP-6 | 1225 |  | 2011-02-03 13:01 | 24C | 0.99637 | 5.2035e-06 | 2.3961e-05 | 0.9964 | 34.8582 | 0.0002015 | [10 empty] | 34.8582
LSSL2010#136 | PP-6 | 1226 |  | 2011-02-03 13:03 | 24C | 0.99565 | 5.5668e-06 | 2.408e-05 | 0.99567 | 34.8298 | 0.00021849 | [10 empty] | 34.8298
LSSL2010#137 | PP-6 | 1227 |  | 2011-02-03 13:05 | 24C | 0.99573 | 7.3523e-06 | 2.4021e-05 | 0.99575 | 34.8328 | 0.00028976 | [10 empty] | 34.8328
LSSL2010#138 | PP-6 | 1228 |  | 2011-02-03 13:08 | 24C | 0.9958 | 4.9029e-06 | 2.3961e-05 | 0.99583 | 34.8358 | 0.00019366 | [10 empty] | 34.8358
LSSL2010#139 | PP-6 | 1229 |  | 2011-02-03 13:10 | 24C | 0.99565 | 4.9162e-06 | 2.4021e-05 | 0.99567 | 34.8296 | 0.0001954 | [10 empty] | 34.8296
LSSL2010#140 | PP-6 | 1230 |  | 2011-02-03 13:12 | 24C | 0.99154 | 7.4722e-06 | 2.4021e-05 | 0.99157 | 34.6683 | 0.00028939 | [10 empty] | 34.6683
LSSL2010#141 | PP-6 | 1231 |  | 2011-02-03 13:15 | 24C | 0.99148 | 8.3443e-06 | 2.3961e-05 | 0.99151 | 34.666 | 0.00032832 | [10 empty] | 34.666
LSSL2010#142 | PP-6 | 1232 |  | 2011-02-03 13:17 | 24C | 0.99141 | 4.6066e-06 | 2.3961e-05 | 0.99144 | 34.6632 | 0.00017824 | [10 empty] | 34.6632
LSSL2010#143 | PP-6 | 1233 |  | 2011-02-03 13:19 | 24C | 0.99125 | 3.854e-06 | 2.3961e-05 | 0.99127 | 34.6568 | 0.00015227 | [10 empty] | 34.6568
LSSL2010#144 | PP-6 | 1234 |  | 2011-02-03 13:22 | 24C | 0.78013 | 4.6354e-06 | 2.4021e-05 | 0.78015 | 26.5191 | 0.00017619 | [10 empty] | 26.5191
LSSL2010#145 | PP-6 | 1235 |  | 2011-02-03 13:25 | 24C | 0.78006 | 3.6105e-06 | 2.3961e-05 | 0.78009 | 26.5167 | 0.00013716 | [10 empty] | 26.5167
LSSL2010#146 | PP-6 | 1236 |  | 2011-02-03 13:28 | 24C | 0.78005 | 9.5635e-06 | 2.4021e-05 | 0.78008 | 26.5163 | 0.00035717 | [10 empty] | 26.5163
LSSL2010#147 | PP-6 | 1237 |  | 2011-02-03 13:30 | 24C | 0.78004 | 5.4694e-06 | 2.3901e-05 | 0.78007 | 26.516 | 0.00020784 | [10 empty] | 26.516
LSSL2010#184 | PP-6 | 1238 |  | 2011-02-03 16:25 | 24C | 0.99886 | 5.4855e-06 | 2.3961e-05 | 0.99888 | 34.9561 | 0.00021651 | [10 empty] | 34.9561
LSSL2010#185 | PP-6 | 1239 |  | 2011-02-03 16:27 | 24C | 0.9988 | 7.1553e-06 | 2.408e-05 | 0.99883 | 34.9539 | 0.00027977 | [10 empty] | 34.9539
LSSL2010#186 | PP-6 | 1240 |  | 2011-02-03 16:29 | 24C | 0.9985 | 4.7172e-06 | 2.408e-05 | 0.99852 | 34.9419 | 0.00018654 |  | 2011-02-09 15:44 | 24C | 0.99853 | 2.2126e-06 | 2.0862e-06 | 0.99853 | 34.9422 | 8.8728e-05 | 0 | -0.00028229
LSSL2010#187 | PP-6 | 1241 |  | 2011-02-03 16:31 | 24C | 0.99793 | 1.4046e-06 | 2.4021e-05 | 0.99795 | 34.9194 | 5.4326e-05 | [10 empty] | 34.9194
LSSL2010#188 | PP-6 | 1242 |  | 2011-02-03 16:34 | 24C | 0.99707 | 5.6675e-06 | 2.4021e-05 | 0.99709 | 34.8855 | 0.00022339 | [10 empty] | 34.8855
LSSL2010#189 | PP-6 | 1243 |  | 2011-02-03 16:36 | 24C | 0.99677 | 2.7425e-06 | 2.3961e-05 | 0.99679 | 34.8736 | 0.00010761 | [10 empty] | 34.8736
LSSL2010#190 | PP-6 | 1244 |  | 2011-02-03 16:39 | 24C | 0.99639 | 7.6436e-06 | 2.408e-05 | 0.99641 | 34.8588 | 0.00030295 | [10 empty] | 34.8588
LSSL2010#191 | PP-6 | 1245 |  | 2011-02-03 16:41 | 24C | 0.99607 | 5.3481e-06 | 2.3961e-05 | 0.99609 | 34.8462 | 0.00021009 | [10 empty] | 34.8462
LSSL2010#192 | PP-6 | 1246 |  | 2011-02-03 16:43 | 24C | 0.99571 | 5.168e-06 | 2.3961e-05 | 0.99574 | 34.8323 | 0.00020516 | [10 empty] | 34.8323
LSSL2010#193 | PP-6 | 1247 |  | 2011-02-03 16:46 | 24C | 0.99569 | 7.6003e-06 | 2.408e-05 | 0.99571 | 34.8314 | 0.00030124 | [10 empty] | 34.8314
LSSL2010#194 | PP-6 | 1248 |  | 2011-02-03 16:48 | 24C | 0.99441 | 4.8959e-06 | 2.408e-05 | 0.99443 | 34.781 | 0.00019429 | [10 empty] | 34.781
LSSL2010#195 | PP-6 | 1249 |  | 2011-02-03 16:51 | 24C | 0.9913 | 7.654e-06 | 2.4021e-05 | 0.99132 | 34.6587 | 0.00029778 | [10 empty] | 34.6587
LSSL2010#196 | PP-6 | 1250 |  | 2011-02-03 16:54 | 24C | 0.98542 | 7.353e-06 | 2.408e-05 | 0.98544 | 34.4278 | 0.00028867 | [10 empty] | 34.4278
LSSL2010#197 | PP-6 | 1251 |  | 2011-02-03 16:56 | 24C | 0.98173 | 5.1678e-06 | 2.4021e-05 | 0.98175 | 34.2829 | 0.00020097 | [10 empty] | 34.2829
LSSL2010#198 | PP-6 | 1252 |  | 2011-02-03 16:58 | 24C | 0.97569 | 4.3839e-06 | 2.408e-05 | 0.97572 | 34.0462 | 0.00017062 | [10 empty] | 34.0462
LSSL2010#199 | PP-6 | 1253 |  | 2011-02-03 17:16 | 24C | 0.95422 | 7.3696e-06 | 2.3961e-05 | 0.95425 | 33.2066 | 0.00029013 |  | 2011-02-09 16:07 | 24C | 0.95439 | 2.4909e-06 | 2.0266e-06 | 0.9544 | 33.2124 | 9.4981e-05 | 0 | -0.005806
LSSL2010#200 | PP-6 | 1254 |  | 2011-02-03 17:19 | 24C | 0.94663 | 5.9189e-06 | 2.4021e-05 | 0.94665 | 32.9103 | 0.00022985 | [10 empty] | 32.9103
LSSL2010#201 | PP-6 | 1255 |  | 2011-02-03 17:21 | 24C | 0.9401 | 1.0298e-05 | 2.4021e-05 | 0.94013 | 32.6561 | 0.00040387 | [10 empty] | 32.6561
LSSL2010#202 | PP-6 | 1256 |  | 2011-02-03 17:24 | 24C | 0.93264 | 3.6105e-06 | 2.3961e-05 | 0.93266 | 32.3657 | 0.00013946 |  | 2011-02-09 16:16 | 24C | 0.93244 | 3.1417e-06 | 2.0266e-06 | 0.93245 | 32.3572 | 0.00012313 | 0 | 0.0084839
LSSL2010#203 | PP-6 | 1257 |  | 2011-02-03 17:26 | 24C | 0.92143 | 7.4723e-06 | 2.4021e-05 | 0.92145 | 31.9302 | 0.00029207 | [10 empty] | 31.9302
LSSL2010#204 | PP-6 | 1258 |  | 2011-02-03 17:29 | 24C | 0.91297 | 8.4891e-06 | 2.3961e-05 | 0.913 | 31.6024 | 0.00032934 | [10 empty] | 31.6024
LSSL2010#205 | PP-6 | 1259 |  | 2011-02-03 17:31 | 24C | 0.89228 | 7.4405e-06 | 2.3961e-05 | 0.8923 | 30.802 | 0.00028508 | [10 empty] | 30.802
LSSL2010#206 | PP-6 | 1260 |  | 2011-02-03 17:34 | 24C | 0.80638 | 5.3357e-06 | 2.3961e-05 | 0.8064 | 27.5136 | 0.00020341 | [10 empty] | 27.5136
LSSL2010#207 | PP-6 | 1261 |  | 2011-02-03 17:37 | 24C | 0.78035 | 2.2259e-06 | 2.4021e-05 | 0.78037 | 26.5275 | 8.4873e-05 | [10 empty] | 26.5275
LSSL2010#1 | CB-DW | 1262 |  | 2011-02-02 11:30 | 24C | 0.99858 | 3.1311e-06 | 4.6015e-05 | 0.99862 | 34.9458 | 0.00012618 | [10 empty] | 34.9458
LSSL2010#2 | CB-DW | 1263 |  | 2011-02-02 11:32 | 24C | 0.9983 | 3.8749e-06 | 4.6015e-05 | 0.99835 | 34.935 | 0.00014763 | [10 empty] | 34.935
LSSL2010#3 | CB-DW | 1264 |  | 2011-02-02 11:35 | 24C | 0.99788 | 6.0358e-06 | 4.6015e-05 | 0.99793 | 34.9186 | 0.00023836 |  | 2011-02-09 15:46 | 24C | 0.99792 | 1.3882e-06 | 2.0862e-06 | 0.99792 | 34.9182 | 5.2278e-05 | 0 | 0.00041962
LSSL2010#4 | CB-DW | 1265 |  | 2011-02-02 11:38 | 24C | 0.99705 | 3.4659e-06 | 4.6074e-05 | 0.9971 | 34.8858 | 0.00013799 |  | 2011-02-09 15:48 | 24C | 0.99708 | 0 | 2.0862e-06 | 0.99708 | 34.8853 | 0 | 0 | 0.00053406
LSSL2010#5 | CB-DW | 1266 |  | 2011-02-02 11:40 | 24C | 0.99664 | 5.0211e-06 | 4.5896e-05 | 0.99668 | 34.8695 | 0.00019505 | [10 empty] | 34.8695
LSSL2010#6 | CB-DW | 1267 |  | 2011-02-02 11:43 | 24C | 0.9963 | 2.7423e-06 | 4.6015e-05 | 0.99634 | 34.856 | 0.00010656 | [10 empty] | 34.856
LSSL2010#7 | CB-DW | 1268 |  | 2011-02-02 11:46 | 24C | 0.9961 | 2.4909e-06 | 4.6015e-05 | 0.99615 | 34.8485 | 9.9734e-05 | [10 empty] | 34.8485
LSSL2010#8 | CB-DW | 1269 |  | 2011-02-02 11:48 | 24C | 0.99553 | 6.2135e-06 | 4.6015e-05 | 0.99558 | 34.8259 | 0.00024024 | [10 empty] | 34.8259
LSSL2010#9 | CB-DW | 1270 |  | 2011-02-02 11:51 | 24C | 0.99564 | 7.1714e-06 | 4.5955e-05 | 0.99569 | 34.8302 | 0.00028287 | [10 empty] | 34.8302
LSSL2010#10 | CB-DW | 1271 |  | 2011-02-02 11:53 | 24C | 0.99532 | 5.5828e-06 | 4.6015e-05 | 0.99537 | 34.8178 | 0.00021731 | [10 empty] | 34.8178
LSSL2010#11 | CB-DW | 1272 |  | 2011-02-02 11:56 | 24C | 0.99442 | 4.4786e-06 | 4.6015e-05 | 0.99446 | 34.7822 | 0.00017744 | [10 empty] | 34.7822
LSSL2010#12 | CB-DW | 1273 |  | 2011-02-02 11:59 | 24C | 0.99201 | 2.7425e-06 | 4.6015e-05 | 0.99206 | 34.6876 | 0.00010447 | [10 empty] | 34.6876
LSSL2010#13 | CB-DW | 1274 |  | 2011-02-02 12:02 | 24C | 0.98578 | 3.3587e-06 | 4.6015e-05 | 0.98583 | 34.4429 | 0.00013108 | [10 empty] | 34.4429
LSSL2010#14 | CB-DW | 1275 |  | 2011-02-02 12:05 | 24C | 0.97479 | 4.9007e-06 | 4.6074e-05 | 0.97483 | 34.0117 | 0.00019277 |  | 2011-02-09 16:03 | 24C | 0.97482 | 2.4911e-06 | 2.0266e-06 | 0.97482 | 34.0113 | 9.869e-05 | 0 | 0.00037003
LSSL2010#15 | CB-DW | 1276 |  | 2011-02-02 12:07 | 24C | 0.96762 | 3.6235e-06 | 4.6015e-05 | 0.96766 | 33.7309 | 0.00014164 | [10 empty] | 33.7309
LSSL2010#16 | CB-DW | 1277 |  | 2011-02-02 12:10 | 24C | 0.95007 | 4.8959e-06 | 4.5955e-05 | 0.95012 | 33.0455 | 0.0001943 |  | 2011-02-09 16:09 | 24C | 0.95025 | 2.866e-06 | 2.0862e-06 | 0.95025 | 33.0508 | 0.00011317 | 0 | -0.0052261
LSSL2010#17 | CB-DW | 1278 |  | 2011-02-02 12:13 | 24C | 0.9466 | 4.7277e-06 | 4.6015e-05 | 0.94665 | 32.9101 | 0.00018258 | [10 empty] | 32.9101
LSSL2010#18 | CB-DW | 1279 |  | 2011-02-02 12:15 | 24C | 0.93913 | 3.4658e-06 | 4.6015e-05 | 0.93918 | 32.6191 | 0.00013386 |  | 2011-02-09 16:14 | 24C | 0.93916 | 1.1203e-06 | 2.0862e-06 | 0.93916 | 32.6184 | 4.4364e-05 | 0 | 0.00066376
LSSL2010#19 | CB-DW | 1280 |  | 2011-02-02 12:18 | 24C | 0.92996 | 4.3843e-06 | 4.6015e-05 | 0.93 | 32.2623 | 0.0001726 | [10 empty] | 32.2623
LSSL2010#20 | CB-DW | 1281 |  | 2011-02-02 12:21 | 24C | 0.91984 | 6.1978e-06 | 4.5955e-05 | 0.91989 | 31.8695 | 0.00024055 |  | 2011-02-09 16:19 | 24C | 0.91977 | 1.1196e-06 | 2.0266e-06 | 0.91977 | 31.8649 | 4.3729e-05 | 0 | 0.0045662
LSSL2010#21 | CB-DW | 1282 |  | 2011-02-02 12:24 | 24C | 0.90726 | 4.1856e-06 | 4.6015e-05 | 0.9073 | 31.3818 | 0.0001651 | [10 empty] | 31.3818
LSSL2010#22 | CB-DW | 1283 |  | 2011-02-02 12:26 | 24C | 0.83034 | 3.3834e-06 | 4.6015e-05 | 0.83039 | 28.4264 | 0.00012852 | [10 empty] | 28.4264
LSSL2010#23 | CB-DW | 1284 |  | 2011-02-02 12:29 | 24C | 0.77006 | 4.9815e-06 | 4.6015e-05 | 0.77011 | 26.1399 | 0.00018882 | [10 empty] | 26.1399
LSSL2010#24 | CB-DW | 1285 |  | 2011-02-02 12:32 | 24C | 0.76375 | 2.7301e-06 | 4.5896e-05 | 0.7638 | 25.9022 | 0.00010253 |  | 2011-02-09 16:26 | 24C | 0.76381 | 0 | 2.0266e-06 | 0.76381 | 25.9028 | 1.0662e-06 | 0 | -0.00068665
LSSL2010#375 | CB-18 | 1286 |  | 2011-02-07 16:52 | 24C | 0.9985 | 0 | 4.0054e-05 | 0.99854 | 34.9424 | 0 | [10 empty] | 34.9424
LSSL2010#376 | CB-18 | 1287 |  | 2011-02-07 16:54 | 24C | 0.9978 | 2.7425e-06 | 3.9935e-05 | 0.99784 | 34.915 | 0.0001097 | [10 empty] | 34.915
LSSL2010#377 | CB-18 | 1288 |  | 2011-02-07 16:57 | 24C | 0.99727 | 3.1424e-06 | 4.0054e-05 | 0.99731 | 34.8941 | 0.00012304 | [10 empty] | 34.8941
LSSL2010#378 | CB-18 | 1289 |  | 2011-02-07 16:59 | 24C | 0.99694 | 3.9984e-06 | 3.9995e-05 | 0.99698 | 34.8811 | 0.00015827 |  | 2011-02-09 15:50 | 24C | 0.99695 | 5.2144e-06 | 2.0266e-06 | 0.99695 | 34.88 | 0.00020293 | 0 | 0.0010986
LSSL2010#379 | CB-18 | 1290 |  | 2011-02-07 17:01 | 24C | 0.99654 | 2.8897e-06 | 3.9935e-05 | 0.99658 | 34.8656 | 0.00011603 | [10 empty] | 34.8656
LSSL2010#380 | CB-18 | 1291 |  | 2011-02-07 17:04 | 24C | 0.99615 | 4.2739e-06 | 3.9935e-05 | 0.99619 | 34.85 | 0.00016798 | [10 empty] | 34.85
LSSL2010#381 | CB-18 | 1292 |  | 2011-02-07 17:06 | 24C | 0.99593 | 2.4908e-06 | 4.0054e-05 | 0.99597 | 34.8413 | 9.87e-05 | [10 empty] | 34.8413
LSSL2010#382 | CB-18 | 1293 |  | 2011-02-07 17:08 | 24C | 0.99573 | 5.1203e-06 | 3.9935e-05 | 0.99577 | 34.8335 | 0.00019874 |  | 2011-02-09 15:53 | 24C | 0.99575 | 2.4908e-06 | 2.0266e-06 | 0.99576 | 34.833 | 9.9734e-05 | 0 | 0.00050735
LSSL2010#383 | CB-18 | 1294 |  | 2011-02-07 17:10 | 24C | 0.99573 | 4.3837e-06 | 3.9935e-05 | 0.99577 | 34.8337 | 0.00017502 | [10 empty] | 34.8337
LSSL2010#384 | CB-18 | 1295 |  | 2011-02-07 17:12 | 24C | 0.99496 | 3.974e-06 | 4.0054e-05 | 0.995 | 34.8031 | 0.00015598 | [10 empty] | 34.8031
LSSL2010#385 | CB-18 | 1296 |  | 2011-02-07 17:14 | 24C | 0.99468 | 3.4657e-06 | 4.0054e-05 | 0.99472 | 34.7922 | 0.00013549 | [10 empty] | 34.7922
LSSL2010#386 | CB-18 | 1297 |  | 2011-02-07 17:17 | 24C | 0.99312 | 4.1137e-06 | 3.9935e-05 | 0.99316 | 34.7309 | 0.00016506 | [10 empty] | 34.7309
LSSL2010#387 | CB-18 | 1298 |  | 2011-02-07 17:19 | 24C | 0.98535 | 2.7425e-06 | 3.9935e-05 | 0.98539 | 34.4257 | 0.0001097 | [10 empty] | 34.4257
LSSL2010#388 | CB-18 | 1299 |  | 2011-02-07 17:21 | 24C | 0.97819 | 2.8877e-06 | 3.9995e-05 | 0.97823 | 34.1448 | 0.00011227 | [10 empty] | 34.1448
LSSL2010#389 | CB-18 | 1300 |  | 2011-02-07 17:22 | 24C | 0.96647 | 3.4657e-06 | 3.9935e-05 | 0.96651 | 33.6859 | 0.00013596 |  | 2011-02-09 16:05 | 24C | 0.96649 | 3.1905e-06 | 2.0266e-06 | 0.96649 | 33.6851 | 0.00012553 | 0 | 0.0007782
LSSL2010#390 | CB-18 | 1301 |  | 2011-02-07 17:24 | 24C | 0.95398 | 4.0973e-06 | 3.9995e-05 | 0.95402 | 33.1976 | 0.00015879 | [10 empty] | 33.1976
LSSL2010#391 | CB-18 | 1302 |  | 2011-02-07 17:26 | 24C | 0.94855 | 5.2338e-06 | 3.9935e-05 | 0.94859 | 32.9858 | 0.00020294 | [10 empty] | 32.9858
LSSL2010#392 | CB-18 | 1303 |  | 2011-02-07 17:28 | 24C | 0.94086 | 6.3007e-06 | 4.0054e-05 | 0.9409 | 32.6863 | 0.00024377 | [10 empty] | 32.6863
LSSL2010#393 | CB-18 | 1304 |  | 2011-02-07 17:30 | 24C | 0.93457 | 2.7425e-06 | 4.0054e-05 | 0.93461 | 32.4414 | 0.00010761 | [10 empty] | 32.4414
LSSL2010#394 | CB-18 | 1305 |  | 2011-02-07 17:32 | 24C | 0.92105 | 1.4046e-06 | 3.9995e-05 | 0.92109 | 31.9162 | 5.2783e-05 | [10 empty] | 31.9162
LSSL2010#395 | CB-18 | 1306 |  | 2011-02-07 17:34 | 24C | 0.88772 | 2.4921e-06 | 4.0054e-05 | 0.88776 | 30.6267 | 9.4895e-05 | [10 empty] | 30.6267
LSSL2010#396 | CB-18 | 1307 |  | 2011-02-07 17:36 | 24C | 0.88086 | 2.2127e-06 | 3.9995e-05 | 0.8809 | 30.3625 | 8.5299e-05 | [10 empty] | 30.3625
LSSL2010#397 | CB-18 | 1308 |  | 2011-02-07 17:38 | 24C | 0.77695 | 2.2401e-06 | 3.9935e-05 | 0.77699 | 26.3996 | 8.5313e-05 | [10 empty] | 26.3996
LSSL2010#398 | CB-18 | 1309 |  | 2011-02-07 17:40 | 24C | 0.74564 | 2.4908e-06 | 4.0054e-05 | 0.74568 | 25.2209 | 9.394e-05 | [10 empty] | 25.2209
LSSL2010#400 | CB-19 | 1310 |  | 2011-02-09 09:42 | 24C | 0.99883 | 1.1196e-06 | 1.967e-06 | 0.99883 | 34.954 | 4.5227e-05 | [10 empty] | 34.954
LSSL2010#401 | CB-19 | 1311 |  | 2011-02-09 09:44 | 24C | 0.9988 | 3.0087e-06 | 2.0266e-06 | 0.9988 | 34.9528 | 0.0001183 | [10 empty] | 34.9528
LSSL2010#402 | CB-19 | 1312 |  | 2011-02-09 09:47 | 24C | 0.99869 | 2.2401e-06 | 1.967e-06 | 0.99869 | 34.9486 | 8.7077e-05 |  | 2011-02-09 15:42 | 24C | 0.99884 | 2.0804e-06 | 2.0862e-06 | 0.99884 | 34.9544 | 8.1621e-05 | 0 | -0.0057755
LSSL2010#403 | CB-19 | 1313 |  | 2011-02-09 09:50 | 24C | 0.99838 | 3.359e-06 | 2.0266e-06 | 0.99839 | 34.9364 | 0.00013469 | [10 empty] | 34.9364
LSSL2010#404 | CB-19 | 1314 |  | 2011-02-09 09:52 | 24C | 0.99774 | 4.4858e-06 | 2.0266e-06 | 0.99774 | 34.9112 | 0.00017536 | [10 empty] | 34.9112
LSSL2010#405 | CB-19 | 1315 |  | 2011-02-09 09:54 | 24C | 0.99683 | 3.1414e-06 | 1.967e-06 | 0.99683 | 34.8752 | 0.00012381 | [10 empty] | 34.8752
LSSL2010#406 | CB-19 | 1316 |  | 2011-02-09 09:56 | 24C | 0.99648 | 1.4046e-06 | 1.967e-06 | 0.99648 | 34.8614 | 5.4326e-05 | [10 empty] | 34.8614
LSSL2010#407 | CB-19 | 1317 |  | 2011-02-09 09:58 | 24C | 0.99616 | 2.4744e-06 | 2.0266e-06 | 0.99616 | 34.849 | 9.9792e-05 | [10 empty] | 34.849
LSSL2010#408 | CB-19 | 1318 |  | 2011-02-09 10:01 | 24C | 0.9957 | 2.7432e-06 | 2.0266e-06 | 0.9957 | 34.8309 | 0.00010656 | [10 empty] | 34.8309
LSSL2010#409 | CB-19 | 1319 |  | 2011-02-09 10:03 | 24C | 0.99574 | 3.2141e-06 | 2.0266e-06 | 0.99575 | 34.8326 | 0.00012429 |  | 2011-02-09 15:55 | 24C | 0.99573 | 2.7264e-06 | 2.0266e-06 | 0.99573 | 34.8319 | 0.00010344 | 0 | 0.0007019
LSSL2010#410 | CB-19 | 1320 |  | 2011-02-09 10:04 | 24C | 0.9947 | 2.7425e-06 | 1.9073e-06 | 0.9947 | 34.7916 | 0.00011074 | [10 empty] | 34.7916
LSSL2010#411 | CB-19 | 1321 |  | 2011-02-09 10:07 | 24C | 0.99439 | 3.332e-08 | 2.0266e-06 | 0.99439 | 34.7795 | 0 | [10 empty] | 34.7795
LSSL2010#412 | CB-19 | 1322 |  | 2011-02-09 10:35 | 24C | 0.9926 | 2.8765e-06 | 2.0266e-06 | 0.9926 | 34.709 | 0.00011401 | [10 empty] | 34.709
LSSL2010#413 | CB-19 | 1323 |  | 2011-02-09 10:37 | 24C | 0.98553 | 3.3594e-06 | 1.967e-06 | 0.98553 | 34.4312 | 0.00013308 | [10 empty] | 34.4312
LSSL2010#414 | CB-19 | 1324 |  | 2011-02-09 10:39 | 24C | 0.9796 | 3.2144e-06 | 1.9073e-06 | 0.9796 | 34.1987 | 0.000126 | [10 empty] | 34.1987
LSSL2010#415 | CB-19 | 1325 |  | 2011-02-09 10:41 | 24C | 0.96935 | 2.8904e-06 | 2.0266e-06 | 0.96935 | 33.797 | 0.00011519 | [10 empty] | 33.797
LSSL2010#416 | CB-19 | 1326 |  | 2011-02-09 10:44 | 24C | 0.9569 | 2.2396e-06 | 1.9073e-06 | 0.9569 | 33.3101 | 8.7223e-05 | [10 empty] | 33.3101
LSSL2010#417 | CB-19 | 1327 |  | 2011-02-09 10:46 | 24C | 0.95563 | 4.2256e-06 | 2.0266e-06 | 0.95563 | 33.2607 | 0.00016598 | [10 empty] | 33.2607
LSSL2010#418 | CB-19 | 1328 |  | 2011-02-09 10:48 | 24C | 0.94142 | 2.7757e-06 | 2.0266e-06 | 0.94142 | 32.7065 | 0.00010656 | [10 empty] | 32.7065
LSSL2010#419 | CB-19 | 1329 |  | 2011-02-09 10:50 | 24C | 0.9399 | 3.2141e-06 | 1.967e-06 | 0.9399 | 32.6473 | 0.00012617 | [10 empty] | 32.6473
LSSL2010#420 | CB-19 | 1330 |  | 2011-02-09 10:52 | 24C | 0.92322 | 2.4909e-06 | 2.0266e-06 | 0.92322 | 31.999 | 9.9639e-05 | [10 empty] | 31.999
LSSL2010#421 | CB-19 | 1331 |  | 2011-02-09 10:54 | 24C | 0.87199 | 1.1199e-06 | 2.0266e-06 | 0.87199 | 30.0198 | 4.4356e-05 | [10 empty] | 30.0198
LSSL2010#422 | CB-19 | 1332 |  | 2011-02-09 10:57 | 24C | 0.75973 | 5.4724e-06 | 2.0266e-06 | 0.75973 | 25.7492 | 0.00020379 | [10 empty] | 25.7492
LSSL2010#423 | CB-19 | 1333 |  | 2011-02-09 10:59 | 24C | 0.73492 | 1.3551e-06 | 1.967e-06 | 0.73492 | 24.8176 | 5.2235e-05 |  | 2011-02-09 16:31 | 24C | 0.73494 | 3.3588e-06 | 1.967e-06 | 0.73494 | 24.8184 | 0.0001271 | 0 | -0.00078201
LSSL2010#27 | CB-6 | 1334 |  | 2011-02-02 14:03 | 24C | 0.99886 | 3.264e-06 | 4.5955e-05 | 0.99891 | 34.9571 | 0.00012912 |  | 2011-02-09 15:39 | 24C | 0.99885 | 2.4915e-06 | 1.967e-06 | 0.99885 | 34.9549 | 0.00010059 | 0 | 0.0021858
LSSL2010#28 | CB-6 | 1335 |  | 2011-02-02 14:06 | 24C | 0.99886 | 5.6675e-06 | 4.5955e-05 | 0.9989 | 34.9568 | 0.00022572 | [10 empty] | 34.9568
LSSL2010#29 | CB-6 | 1336 |  | 2011-02-02 14:09 | 24C | 0.99927 | 2.7426e-06 | 4.6015e-05 | 0.99931 | 34.973 | 0.00010456 | [10 empty] | 34.973
LSSL2010#30 | CB-6 | 1337 |  | 2011-02-02 14:14 | 24C | 0.99853 | 3.8402e-06 | 4.6015e-05 | 0.99857 | 34.9439 | 0.00014859 | [10 empty] | 34.9439
LSSL2010#31 | CB-6 | 1338 |  | 2011-02-02 14:17 | 24C | 0.99771 | 1.3713e-06 | 4.6015e-05 | 0.99776 | 34.9118 | 5.4351e-05 | [10 empty] | 34.9118
LSSL2010#32 | CB-6 | 1339 |  | 2011-02-02 14:20 | 24C | 0.99689 | 3.8748e-06 | 4.6015e-05 | 0.99694 | 34.8794 | 0.00015047 | [10 empty] | 34.8794
LSSL2010#33 | CB-6 | 1340 |  | 2011-02-02 14:22 | 24C | 0.99647 | 8.4334e-06 | 4.6015e-05 | 0.99652 | 34.8631 | 0.00033289 | [10 empty] | 34.8631
LSSL2010#34 | CB-6 | 1341 |  | 2011-02-02 14:24 | 24C | 0.99615 | 5.5979e-06 | 4.6015e-05 | 0.99619 | 34.8502 | 0.00022175 | [10 empty] | 34.8502
LSSL2010#35 | CB-6 | 1342 |  | 2011-02-02 14:27 | 24C | 0.99577 | 4.5027e-06 | 4.6015e-05 | 0.99581 | 34.8353 | 0.00017851 | [10 empty] | 34.8353
LSSL2010#36 | CB-6 | 1343 |  | 2011-02-02 14:30 | 24C | 0.99565 | 1.3713e-06 | 4.6015e-05 | 0.9957 | 34.8308 | 5.2235e-05 | [10 empty] | 34.8308
LSSL2010#37 | CB-6 | 1344 |  | 2011-02-02 14:33 | 24C | 0.99469 | 5.4847e-06 | 4.6074e-05 | 0.99473 | 34.7928 | 0.00021364 | [10 empty] | 34.7928
LSSL2010#38 | CB-6 | 1345 |  | 2011-02-02 14:35 | 24C | 0.99286 | 7.7817e-06 | 4.6015e-05 | 0.99291 | 34.7211 | 0.00030835 | [10 empty] | 34.7211
LSSL2010#39 | CB-6 | 1346 |  | 2011-02-02 14:38 | 24C | 0.98572 | 4.5125e-06 | 4.6015e-05 | 0.98577 | 34.4405 | 0.00017485 | [10 empty] | 34.4405
LSSL2010#40 | CB-6 | 1347 |  | 2011-02-02 14:40 | 24C | 0.96722 | 2.1278e-06 | 4.6015e-05 | 0.96726 | 33.7153 | 8.1169e-05 | [10 empty] | 33.7153
LSSL2010#41 | CB-6 | 1348 |  | 2011-02-02 14:42 | 24C | 0.96711 | 6.6064e-06 | 4.6074e-05 | 0.96716 | 33.7111 | 0.00025526 | [10 empty] | 33.7111
LSSL2010#42 | CB-6 | 1349 |  | 2011-02-02 14:45 | 24C | 0.95239 | 3.3587e-06 | 4.6015e-05 | 0.95243 | 33.1358 | 0.00013383 | [10 empty] | 33.1358
LSSL2010#43 | CB-6 | 1350 |  | 2011-02-02 14:48 | 24C | 0.94684 | 8.6176e-06 | 4.6015e-05 | 0.94689 | 32.9196 | 0.00033553 | [10 empty] | 32.9196
LSSL2010#44 | CB-6 | 1351 |  | 2011-02-02 14:52 | 24C | 0.9394 | 6.2131e-06 | 4.5896e-05 | 0.93944 | 32.6294 | 0.00024166 |  | 2011-02-09 16:12 | 24C | 0.93939 | 3.1418e-06 | 1.9073e-06 | 0.9394 | 32.6276 | 0.00012486 | 0 | 0.0017471
LSSL2010#45 | CB-6 | 1352 |  | 2011-02-02 14:55 | 24C | 0.93189 | 3.0662e-06 | 4.6015e-05 | 0.93194 | 32.3375 | 0.00011836 | [10 empty] | 32.3375
LSSL2010#46 | CB-6 | 1353 |  | 2011-02-02 14:57 | 24C | 0.92048 | 3.862e-06 | 4.5896e-05 | 0.92052 | 31.8941 | 0.00015135 | SZ- Correct, there is gap btw 1353 and 1357 | [9 empty] | 31.8941
LSSL2010#47 | CB-6 | 1357 |  | 2011-02-02 15:01 | 24C | 0.73233 | 2.8904e-06 | 4.5896e-05 | 0.73238 | 24.7226 | 0.00010975 | [10 empty] | 24.7226
LSSL2010#49 | CB-21 | 1358 |  | 2011-02-02 15:17 | 24C | 0.99705 | 3.3588e-06 | 4.5955e-05 | 0.9971 | 34.8858 | 0.00013233 | [10 empty] | 34.8858
LSSL2010#50 | CB-21 | 1359 |  | 2011-02-02 15:20 | 24C | 0.99701 | 4.7299e-06 | 4.6015e-05 | 0.99705 | 34.8841 | 0.00018569 | [10 empty] | 34.8841
LSSL2010#51 | CB-21 | 1360 |  | 2011-02-02 15:23 | 24C | 0.99694 | 6.1012e-06 | 4.6074e-05 | 0.99698 | 34.8813 | 0.00023971 | [10 empty] | 34.8813
LSSL2010#52 | CB-21 | 1361 |  | 2011-02-02 15:25 | 24C | 0.997 | 6.3002e-06 | 4.6074e-05 | 0.99705 | 34.8839 | 0.00024981 | [10 empty] | 34.8839
LSSL2010#53 | CB-21 | 1362 |  | 2011-02-02 15:27 | 24C | 0.99656 | 3.8648e-06 | 4.6015e-05 | 0.99661 | 34.8666 | 0.00015243 | [10 empty] | 34.8666
LSSL2010#54 | CB-21 | 1363 |  | 2011-02-02 15:30 | 24C | 0.99652 | 4.5125e-06 | 4.6015e-05 | 0.99657 | 34.8651 | 0.00017433 | [10 empty] | 34.8651
LSSL2010#55 | CB-21 | 1364 |  | 2011-02-02 15:32 | 24C | 0.99657 | 4.1136e-06 | 4.6015e-05 | 0.99662 | 34.8668 | 0.00016089 | [10 empty] | 34.8668
LSSL2010#56 | CB-21 | 1365 |  | 2011-02-02 15:35 | 24C | 0.99657 | 7.0607e-06 | 4.5955e-05 | 0.99662 | 34.8669 | 0.0002774 | [10 empty] | 34.8669
LSSL2010#57 | CB-21 | 1366 |  | 2011-02-02 15:37 | 24C | 0.99612 | 5.5575e-06 | 4.6015e-05 | 0.99617 | 34.8492 | 0.00022085 | [10 empty] | 34.8492
LSSL2010#58 | CB-21 | 1367 |  | 2011-02-02 15:40 | 24C | 0.99609 | 2.4908e-06 | 4.6074e-05 | 0.99614 | 34.8479 | 9.6976e-05 | [10 empty] | 34.8479
LSSL2010#59 | CB-21 | 1368 |  | 2011-02-02 15:44 | 24C | 0.99619 | 3.8619e-06 | 4.5955e-05 | 0.99623 | 34.8517 | 0.00015179 | [10 empty] | 34.8517
LSSL2010#60 | CB-21 | 1369 |  | 2011-02-02 15:46 | 24C | 0.99621 | 3.6085e-06 | 4.6015e-05 | 0.99625 | 34.8526 | 0.00014196 | [10 empty] | 34.8526
LSSL2010#61 | CB-21 | 1370 |  | 2011-02-02 15:49 | 24C | 0.99528 | 3.8323e-06 | 4.6015e-05 | 0.99533 | 34.8162 | 0.00014921 | [10 empty] | 34.8162
LSSL2010#62 | CB-21 | 1371 |  | 2011-02-02 15:51 | 24C | 0.99529 | 5.8623e-06 | 4.6015e-05 | 0.99534 | 34.8166 | 0.00023023 | [10 empty] | 34.8166
LSSL2010#63 | CB-21 | 1372 |  | 2011-02-02 15:53 | 24C | 0.99573 | 3.9989e-06 | 4.6015e-05 | 0.99577 | 34.8336 | 0.00015782 | [10 empty] | 34.8336
LSSL2010#64 | CB-21 | 1373 |  | 2011-02-02 15:57 | 24C | 0.99542 | 4.3587e-06 | 4.5955e-05 | 0.99546 | 34.8215 | 0.00016984 | [10 empty] | 34.8215
LSSL2010#65 | CB-21 | 1374 |  | 2011-02-02 16:01 | 24C | 0.98885 | 8.3446e-06 | 4.5896e-05 | 0.9889 | 34.5635 | 0.00032792 | [10 empty] | 34.5635
LSSL2010#66 | CB-21 | 1375 |  | 2011-02-02 16:05 | 24C | 0.98885 | 5.4849e-06 | 4.6134e-05 | 0.9889 | 34.5635 | 0.00021388 | [10 empty] | 34.5635
LSSL2010#67 | CB-21 | 1376 |  | 2011-02-02 16:09 | 24C | 0.98889 | 7.3776e-06 | 4.5955e-05 | 0.98894 | 34.565 | 0.00029054 | [10 empty] | 34.565
LSSL2010#68 | CB-21 | 1377 |  | 2011-02-02 16:11 | 24C | 0.98891 | 6.8046e-06 | 4.5896e-05 | 0.98895 | 34.5656 | 0.00027202 | [10 empty] | 34.5656
LSSL2010#69 | CB-21 | 1378 |  | 2011-02-02 16:14 | 24C | 0.74173 | 2.4776e-06 | 4.6015e-05 | 0.74177 | 25.0744 | 9.2326e-05 | [10 empty] | 25.0744
LSSL2010#70 | CB-21 | 1379 |  | 2011-02-02 16:17 | 24C | 0.74127 | 6.7174e-06 | 4.6015e-05 | 0.74131 | 25.0573 | 0.00024994 | [10 empty] | 25.0573
LSSL2010#71 | CB-21 | 1380 |  | 2011-02-02 16:19 | 24C | 0.74123 | 3.8619e-06 | 4.6015e-05 | 0.74128 | 25.0559 | 0.00014712 | [10 empty] | 25.0559
LSSL2010#72 | CB-21 | 1381 |  | 2011-02-02 16:22 | 24C | 0.74119 | 4.9029e-06 | 4.6015e-05 | 0.74124 | 25.0545 | 0.00018162 | [10 empty] | 25.0545
LSSL2010#618 | CB-21 | 1382 |  | 2011-02-10 16:54 | 24C | 0.99053 | 3.3587e-06 | 2.0266e-06 | 0.99054 | 34.6278 | 0.00012951 | [10 empty] | 34.6278
LSSL2010#619 | CB-21 | 1383 |  | 2011-02-10 16:57 | 24C | 0.98679 | 4.6353e-06 | 2.0266e-06 | 0.9868 | 34.4809 | 0.00018285 | [10 empty] | 34.4809
LSSL2010#620 | CB-21 | 1384 |  | 2011-02-10 16:59 | 24C | 0.98339 | 2.2263e-06 | 2.0266e-06 | 0.98339 | 34.3473 | 8.9601e-05 | [10 empty] | 34.3473
LSSL2010#621 | CB-21 | 1385 |  | 2011-02-10 17:01 | 24C | 0.98126 | 2.7426e-06 | 1.967e-06 | 0.98126 | 34.2636 | 0.00010447 | [10 empty] | 34.2636
LSSL2010#622 | CB-21 | 1386 |  | 2011-02-10 17:03 | 24C | 0.97856 | 3.2012e-06 | 2.0266e-06 | 0.97856 | 34.1578 | 0.00012091 |  | 2011-02-09 16:00 | 24C | 0.97834 | 3.131e-06 | 2.0266e-06 | 0.97834 | 34.1491 | 0.00012172 | 0 | 0.0086899
LSSL2010#623 | CB-21 | 1387 |  | 2011-02-10 17:05 | 24C | 0.9744 | 4.3336e-06 | 2.0266e-06 | 0.9744 | 33.9947 | 0.00016843 | [10 empty] | 33.9947
LSSL2010#624 | CB-21 | 1388 |  | 2011-02-10 17:08 | 24C | 0.96983 | 4.2113e-06 | 2.0862e-06 | 0.96983 | 33.8158 | 0.00016236 | [10 empty] | 33.8158
LSSL2010#625 | CB-21 | 1389 |  | 2011-02-10 17:10 | 24C | 0.96326 | 3.323e-06 | 1.967e-06 | 0.96327 | 33.5589 | 0.00013002 | [10 empty] | 33.5589
LSSL2010#626 | CB-21 | 1390 |  | 2011-02-10 17:12 | 24C | 0.95706 | 2.0947e-06 | 2.0862e-06 | 0.95706 | 33.3165 | 8.2541e-05 | [10 empty] | 33.3165
LSSL2010#627 | CB-21 | 1391 |  | 2011-02-10 17:14 | 24C | 0.95116 | 2.4908e-06 | 2.0266e-06 | 0.95116 | 33.0862 | 9.7029e-05 | [10 empty] | 33.0862
LSSL2010#628 | CB-21 | 1392 |  | 2011-02-10 17:16 | 24C | 0.9475 | 3.8622e-06 | 2.0266e-06 | 0.9475 | 32.9435 | 0.00014836 | [10 empty] | 32.9435
LSSL2010#629 | CB-21 | 1393 |  | 2011-02-10 17:19 | 24C | 0.94466 | 2.4911e-06 | 1.967e-06 | 0.94466 | 32.8327 | 9.869e-05 | [10 empty] | 32.8327
LSSL2010#630 | CB-21 | 1394 |  | 2011-02-10 17:21 | 24C | 0.94203 | 2.7425e-06 | 1.9073e-06 | 0.94203 | 32.7303 | 0.00010447 | [10 empty] | 32.7303
LSSL2010#631 | CB-21 | 1395 |  | 2011-02-10 17:23 | 24C | 0.93603 | 3.8538e-06 | 2.0266e-06 | 0.93603 | 32.4967 | 0.00015042 | [10 empty] | 32.4967
LSSL2010#632 | CB-21 | 1396 |  | 2011-02-10 17:25 | 24C | 0.9308 | 4.6144e-06 | 1.9073e-06 | 0.9308 | 32.2934 | 0.00018034 | [10 empty] | 32.2934
LSSL2010#633 | CB-21 | 1397 |  | 2011-02-10 17:27 | 24C | 0.92361 | 4.409e-06 | 2.0862e-06 | 0.92361 | 32.014 | 0.00017061 | [10 empty] | 32.014
LSSL2010#634 | CB-21 | 1398 |  | 2011-02-10 17:29 | 24C | 0.91814 | 2.4909e-06 | 2.0266e-06 | 0.91814 | 31.8017 | 9.8682e-05 | [10 empty] | 31.8017
LSSL2010#635 | CB-21 | 1399 |  | 2011-02-10 17:31 | 24C | 0.90898 | 1.3713e-06 | 2.0266e-06 | 0.90899 | 31.447 | 5.3301e-05 | [10 empty] | 31.447
LSSL2010#636 | CB-21 | 1400 |  | 2011-02-10 17:34 | 24C | 0.89644 | 2.4754e-06 | 2.1458e-06 | 0.89644 | 30.962 | 9.6496e-05 | [10 empty] | 30.962
LSSL2010#637 | CB-21 | 1401 |  | 2011-02-10 17:36 | 24C | 0.88541 | 2.4909e-06 | 1.967e-06 | 0.88541 | 30.5363 | 9.4885e-05 |  | 2011-02-09 16:21 | 24C | 0.88377 | 3.1418e-06 | 1.967e-06 | 0.88377 | 30.473 | 0.00012365 | 0 | 0.063282
LSSL2010#638 | CB-21 | 1402 |  | 2011-02-10 17:38 | 24C | 0.85779 | 3.8623e-06 | 2.0862e-06 | 0.85779 | 29.4746 | 0.00014921 | [10 empty] | 29.4746
LSSL2010#639 | CB-21 | 1403 |  | 2011-02-10 17:40 | 24C | 0.83299 | 3.6105e-06 | 2.0266e-06 | 0.83299 | 28.5257 | 0.00013967 | [10 empty] | 28.5257
LSSL2010#640 | CB-21 | 1404 |  | 2011-02-10 17:42 | 24C | 0.78216 | 2.2392e-06 | 2.0266e-06 | 0.78216 | 26.5951 | 8.5306e-05 | [10 empty] | 26.5951
LSSL2010#641 | CB-21 | 1405 |  | 2011-02-10 17:45 | 24C | 0.74258 | 2.8897e-06 | 2.0266e-06 | 0.74259 | 25.105 | 0.00010553 |  | 2011-02-09 16:29 | 24C | 0.74217 | 2.2135e-06 | 2.0862e-06 | 0.74217 | 25.0894 | 8.5299e-05 | 0 | 0.01564
LSSL2010#255 | CB-21 | 1406 |  | 2011-02-04 12:15 | 24C | 0.99888 | 1.1064e-06 | 7.987e-06 | 0.99889 | 34.9563 | 4.1799e-05 | [10 empty] | 34.9563
LSSL2010#256 | CB-21 | 1407 |  | 2011-02-04 12:17 | 24C | 0.99887 | 1.1203e-06 | 7.9274e-06 | 0.99888 | 34.9558 | 4.5217e-05 | [10 empty] | 34.9558
LSSL2010#257 | CB-21 | 1408 |  | 2011-02-04 12:19 | 24C | 0.99879 | 5.8609e-06 | 8.1062e-06 | 0.9988 | 34.9529 | 0.00023274 | [10 empty] | 34.9529
LSSL2010#258 | CB-21 | 1409 |  | 2011-02-04 12:21 | 24C | 0.99852 | 1.3712e-06 | 7.9274e-06 | 0.99852 | 34.9419 | 5.4351e-05 | [10 empty] | 34.9419
LSSL2010#259 | CB-21 | 1410 |  | 2011-02-04 12:24 | 24C | 0.99785 | 7.7497e-06 | 7.8678e-06 | 0.99786 | 34.9157 | 0.00030526 |  | 2011-02-09 12:15 | 24C | 0.99785 | 3.4658e-06 | 1.967e-06 | 0.99785 | 34.9155 | 0.00013387 | 0 | 0.00025177
LSSL2010#260 | CB-21 | 1411 |  | 2011-02-04 12:26 | 24C | 0.99701 | 4.6353e-06 | 7.987e-06 | 0.99702 | 34.8828 | 0.00018273 | [10 empty] | 34.8828
LSSL2010#261 | CB-21 | 1412 |  | 2011-02-04 12:29 | 24C | 0.99663 | 2.4908e-06 | 7.9274e-06 | 0.99663 | 34.8675 | 9.7029e-05 | [10 empty] | 34.8675
LSSL2010#262 | CB-21 | 1413 |  | 2011-02-04 12:31 | 24C | 0.99616 | 8.198e-06 | 7.9274e-06 | 0.99617 | 34.8494 | 0.000324 | [10 empty] | 34.8494
LSSL2010#263 | CB-21 | 1414 |  | 2011-02-04 12:33 | 24C | 0.99577 | 9.3319e-06 | 7.987e-06 | 0.99578 | 34.8338 | 0.00036908 | [10 empty] | 34.8338
LSSL2010#264 | CB-21 | 1415 |  | 2011-02-04 12:36 | 24C | 0.99572 | 5.865e-06 | 7.987e-06 | 0.99573 | 34.832 | 0.00023196 | [10 empty] | 34.832
LSSL2010#265 | CB-21 | 1416 |  | 2011-02-04 12:38 | 24C | 0.9951 | 8.2053e-06 | 7.987e-06 | 0.99511 | 34.8077 | 0.00032017 | [10 empty] | 34.8077
LSSL2010#266 | CB-21 | 1417 |  | 2011-02-04 12:41 | 24C | 0.99369 | 4.4789e-06 | 7.8678e-06 | 0.9937 | 34.7522 | 0.00017574 | [10 empty] | 34.7522
LSSL2010#267 | CB-21 | 1418 |  | 2011-02-04 12:43 | 24C | 0.98888 | 5.2207e-06 | 7.987e-06 | 0.98889 | 34.563 | 0.00020243 | [10 empty] | 34.563
LSSL2010#268 | CB-21 | 1419 |  | 2011-02-04 12:45 | 24C | 0.98118 | 4.1136e-06 | 7.9274e-06 | 0.98119 | 34.2608 | 0.00016064 | [10 empty] | 34.2608
LSSL2010#269 | CB-21 | 1420 |  | 2011-02-04 12:47 | 24C | 0.95531 | 4.2259e-06 | 7.987e-06 | 0.95532 | 33.2486 | 0.00016771 | [10 empty] | 33.2486
LSSL2010#270 | CB-21 | 1421 |  | 2011-02-04 12:50 | 24C | 0.95292 | 2.2275e-06 | 7.987e-06 | 0.95293 | 33.1552 | 8.7556e-05 | [10 empty] | 33.1552
LSSL2010#271 | CB-21 | 1422 |  | 2011-02-04 12:52 | 24C | 0.94599 | 3.4495e-06 | 7.987e-06 | 0.946 | 32.8848 | 0.00013345 | [10 empty] | 32.8848
LSSL2010#272 | CB-21 | 1423 |  | 2011-02-04 12:55 | 24C | 0.93973 | 8.9806e-06 | 7.987e-06 | 0.93974 | 32.6409 | 0.00034848 | [10 empty] | 32.6409
LSSL2010#273 | CB-21 | 1424 |  | 2011-02-04 12:57 | 24C | 0.93255 | 5.2334e-06 | 7.987e-06 | 0.93256 | 32.3616 | 0.0002041 | [10 empty] | 32.3616
LSSL2010#274 | CB-21 | 1425 |  | 2011-02-04 12:59 | 24C | 0.92279 | 5.2332e-06 | 8.1062e-06 | 0.9228 | 31.9824 | 0.00020115 |  | 2011-02-09 12:38 | 24C | 0.92289 | 4.9162e-06 | 2.0266e-06 | 0.92289 | 31.986 | 0.00018936 | 0 | -0.0035915
LSSL2010#275 | CB-21 | 1426 |  | 2011-02-04 13:01 | 24C | 0.90316 | 7.5748e-06 | 8.1062e-06 | 0.90317 | 31.2219 | 0.00029238 | [10 empty] | 31.2219
LSSL2010#276 | CB-21 | 1427 |  | 2011-02-04 13:04 | 24C | 0.87725 | 5.5671e-06 | 7.987e-06 | 0.87726 | 30.2224 | 0.0002137 | [10 empty] | 30.2224
LSSL2010#277 | CB-21 | 1428 |  | 2011-02-04 13:06 | 24C | 0.77525 | 8.2918e-06 | 7.987e-06 | 0.77526 | 26.3343 | 0.00031316 | [10 empty] | 26.3343
LSSL2010#278 | CB-21 | 1429 |  | 2011-02-04 13:09 | 24C | 0.74067 | 1.3712e-06 | 7.987e-06 | 0.74068 | 25.0334 | 5.1717e-05 |  | 2011-02-09 12:43 | 24C | 0.74066 | 2.2391e-06 | 1.967e-06 | 0.74066 | 25.0329 | 8.872e-05 | 0 | 0.00047302
LSSL2010#281 | CB-40 | 1430 |  | 2011-02-07 10:39 | 24C | 0.99888 | 1.1199e-06 | 3.9995e-05 | 0.99892 | 34.9575 | 4.265e-05 | [10 empty] | 34.9575
LSSL2010#282 | CB-40 | 1431 |  | 2011-02-07 10:41 | 24C | 0.99882 | 3.2141e-06 | 4.0054e-05 | 0.99886 | 34.9549 | 0.00012709 | [10 empty] | 34.9549
LSSL2010#283 | CB-40 | 1432 |  | 2011-02-07 10:43 | 24C | 0.99854 | 4.3336e-06 | 4.0054e-05 | 0.99858 | 34.9443 | 0.00017093 | [10 empty] | 34.9443
LSSL2010#284 | CB-40 | 1433 |  | 2011-02-07 10:46 | 24C | 0.99795 | 3.865e-06 | 4.0054e-05 | 0.99799 | 34.9209 | 0.00015448 |  | 2011-02-09 11:18 | 24C | 0.99786 | 4.2856e-06 | 1.967e-06 | 0.99786 | 34.9157 | 0.00016832 | 0 | 0.0051727
LSSL2010#285 | CB-40 | 1434 |  | 2011-02-07 10:48 | 24C | 0.99702 | 1.1199e-06 | 3.9995e-05 | 0.99706 | 34.8841 | 4.5227e-05 | [10 empty] | 34.8841
LSSL2010#286 | CB-40 | 1435 |  | 2011-02-07 10:50 | 24C | 0.99676 | 1.1199e-06 | 4.0054e-05 | 0.9968 | 34.8742 | 4.4374e-05 | [10 empty] | 34.8742
LSSL2010#287 | CB-40 | 1436 |  | 2011-02-07 10:52 | 24C | 0.99636 | 3.2142e-06 | 4.0054e-05 | 0.9964 | 34.8584 | 0.00012382 | [10 empty] | 34.8584
LSSL2010#288 | CB-40 | 1437 |  | 2011-02-07 10:54 | 24C | 0.99609 | 4.2256e-06 | 3.9995e-05 | 0.99613 | 34.8476 | 0.00016944 | [10 empty] | 34.8476
LSSL2010#289 | CB-40 | 1438 |  | 2011-02-07 10:56 | 24C | 0.9957 | 4.7431e-06 | 4.0054e-05 | 0.99574 | 34.8323 | 0.00018675 | [10 empty] | 34.8323
LSSL2010#290 | CB-40 | 1439 |  | 2011-02-07 10:58 | 24C | 0.99523 | 5.0782e-06 | 3.9935e-05 | 0.99527 | 34.8138 | 0.00019832 | [10 empty] | 34.8138
LSSL2010#291 | CB-40 | 1440 |  | 2011-02-07 11:00 | 24C | 0.99479 | 3.2141e-06 | 3.9995e-05 | 0.99483 | 34.7966 | 0.00012382 |  | 2011-02-09 11:27 | 24C | 0.99473 | 2.4908e-06 | 2.0266e-06 | 0.99473 | 34.7926 | 0.00010059 | 0 | 0.0039291
LSSL2010#292 | CB-40 | 1441 |  | 2011-02-07 11:03 | 24C | 0.99371 | 3.2144e-06 | 4.0054e-05 | 0.99375 | 34.754 | 0.00012599 | [10 empty] | 34.754
LSSL2010#293 | CB-40 | 1442 |  | 2011-02-07 11:05 | 24C | 0.99107 | 4.1786e-06 | 3.9935e-05 | 0.99111 | 34.6505 | 0.00016283 | [10 empty] | 34.6505
LSSL2010#294 | CB-40 | 1443 |  | 2011-02-07 11:07 | 24C | 0.98613 | 4.0096e-06 | 4.0054e-05 | 0.98617 | 34.4565 | 0.00015769 | [10 empty] | 34.4565
LSSL2010#295 | CB-40 | 1444 |  | 2011-02-07 11:10 | 24C | 0.97888 | 2.4909e-06 | 3.9935e-05 | 0.97892 | 34.172 | 9.869e-05 | [10 empty] | 34.172
LSSL2010#296 | CB-40 | 1445 |  | 2011-02-07 11:12 | 24C | 0.96859 | 5.1607e-06 | 3.9995e-05 | 0.96863 | 33.7688 | 0.00019816 | [10 empty] | 33.7688
LSSL2010#297 | CB-40 | 1446 |  | 2011-02-07 11:14 | 24C | 0.94128 | 3.9742e-06 | 4.0054e-05 | 0.94132 | 32.7025 | 0.00015499 | [10 empty] | 32.7025
LSSL2010#298 | CB-40 | 1447 |  | 2011-02-07 11:16 | 24C | 0.96293 | 3.5949e-06 | 3.9995e-05 | 0.96297 | 33.5474 | 0.00014165 | Really?? | 2011-02-09 12:33 | 24C | 0.96254 | 1.1196e-06 | 2.0266e-06 | 0.96254 | 33.5307 | 4.265e-05 | 0 | 0.016678
LSSL2010#299 | CB-40 | 1448 |  | 2011-02-07 11:19 | 24C | 0.9422 | 3.2141e-06 | 3.9995e-05 | 0.94224 | 32.7385 | 0.00012244 | [10 empty] | 32.7385
LSSL2010#300 | CB-40 | 1449 |  | 2011-02-07 11:22 | 24C | 0.93356 | 3.3595e-06 | 4.0054e-05 | 0.9336 | 32.4021 | 0.0001317 | [10 empty] | 32.4021
LSSL2010#301 | CB-40 | 1450 |  | 2011-02-07 11:24 | 24C | 0.92237 | 3.1253e-06 | 4.0054e-05 | 0.92241 | 31.9674 | 0.00012085 | [10 empty] | 31.9674
LSSL2010#302 | CB-40 | 1451 |  | 2011-02-07 11:26 | 24C | 0.89603 | 2.89e-06 | 3.9995e-05 | 0.89607 | 30.9478 | 0.00011103 | [10 empty] | 30.9478
LSSL2010#303 | CB-40 | 1452 |  | 2011-02-07 11:28 | 24C | 0.80332 | 2.8901e-06 | 3.9995e-05 | 0.80336 | 27.3979 | 0.00010975 | [10 empty] | 27.3979
LSSL2010#304 | CB-40 | 1453 |  | 2011-02-07 11:31 | 24C | 0.76957 | 2.2394e-06 | 3.9995e-05 | 0.76961 | 26.1212 | 8.2774e-05 | [10 empty] | 26.1212
LSSL2010#230 | CB-50 | 1454 |  | 2011-02-04 11:05 | 24C | 0.99888 | 2.4911e-06 | 7.9274e-06 | 0.99889 | 34.9562 | 0.00010059 | [10 empty] | 34.9562
LSSL2010#231 | CB-50 | 1455 |  | 2011-02-04 11:08 | 24C | 0.99885 | 8.4423e-06 | 7.987e-06 | 0.99886 | 34.9551 | 0.00033122 | [10 empty] | 34.9551
LSSL2010#232 | CB-50 | 1456 |  | 2011-02-04 11:10 | 24C | 0.99857 | 2.7097e-06 | 8.0466e-06 | 0.99858 | 34.9441 | 0.00010865 | [10 empty] | 34.9441
LSSL2010#233 | CB-50 | 1457 |  | 2011-02-04 11:12 | 24C | 0.9979 | 4.9134e-06 | 7.987e-06 | 0.99791 | 34.9177 | 0.00019232 | [10 empty] | 34.9177
LSSL2010#234 | CB-50 | 1458 |  | 2011-02-04 11:14 | 24C | 0.9971 | 4.3839e-06 | 7.987e-06 | 0.99711 | 34.8862 | 0.00017276 |  | 2011-02-09 11:22 | 24C | 0.99707 | 2.4909e-06 | 2.0266e-06 | 0.99707 | 34.8846 | 9.4938e-05 | 0 | 0.0015564
LSSL2010#235 | CB-50 | 1459 |  | 2011-02-04 11:17 | 24C | 0.99676 | 2.4915e-06 | 7.9274e-06 | 0.99676 | 34.8727 | 9.4938e-05 | [10 empty] | 34.8727
LSSL2010#236 | CB-50 | 1460 |  | 2011-02-04 11:19 | 24C | 0.99633 | 3.8654e-06 | 7.8678e-06 | 0.99634 | 34.8558 | 0.00015205 | [10 empty] | 34.8558
LSSL2010#237 | CB-50 | 1461 |  | 2011-02-04 11:21 | 24C | 0.996 | 2.4909e-06 | 8.0466e-06 | 0.996 | 34.8428 | 9.9543e-05 | [10 empty] | 34.8428
LSSL2010#238 | CB-50 | 1462 |  | 2011-02-04 11:24 | 24C | 0.99566 | 6.6505e-06 | 7.987e-06 | 0.99567 | 34.8296 | 0.00025921 | [10 empty] | 34.8296
LSSL2010#239 | CB-50 | 1463 |  | 2011-02-04 11:27 | 24C | 0.99557 | 4.2257e-06 | 8.1062e-06 | 0.99558 | 34.826 | 0.00016483 | [10 empty] | 34.826
LSSL2010#240 | CB-50 | 1464 |  | 2011-02-04 11:29 | 24C | 0.99432 | 4.6041e-06 | 7.987e-06 | 0.99433 | 34.7768 | 0.00018322 | [10 empty] | 34.7768
LSSL2010#241 | CB-50 | 1465 |  | 2011-02-04 11:32 | 24C | 0.99169 | 6.5495e-06 | 7.987e-06 | 0.99169 | 34.6733 | 0.00025751 | [10 empty] | 34.6733
LSSL2010#242 | CB-50 | 1466 |  | 2011-02-04 11:34 | 24C | 0.98962 | 7.2313e-06 | 8.0466e-06 | 0.98963 | 34.5923 | 0.00028765 | [10 empty] | 34.5923
LSSL2010#243 | CB-50 | 1467 |  | 2011-02-04 11:37 | 24C | 0.98613 | 6.3303e-06 | 7.987e-06 | 0.98614 | 34.455 | 0.0002493 |  | 2011-02-09 11:37 | 24C | 0.98604 | 2.2396e-06 | 1.9073e-06 | 0.98604 | 34.4513 | 8.8203e-05 | 0 | 0.0036469
LSSL2010#244 | CB-50 | 1468 |  | 2011-02-04 11:39 | 24C | 0.98166 | 5.8911e-06 | 7.987e-06 | 0.98167 | 34.2796 | 0.00022967 | [10 empty] | 34.2796
LSSL2010#245 | CB-50 | 1469 |  | 2011-02-04 11:41 | 24C | 0.95771 | 3.2151e-06 | 7.987e-06 | 0.95771 | 33.3419 | 0.000126 | [10 empty] | 33.3419
LSSL2010#246 | CB-50 | 1470 |  | 2011-02-04 11:43 | 24C | 0.94495 | 5.6569e-06 | 8.0466e-06 | 0.94495 | 32.8441 | 0.00021857 | [10 empty] | 32.8441
LSSL2010#247 | CB-50 | 1471 |  | 2011-02-04 11:46 | 24C | 0.94535 | 6.4249e-06 | 7.987e-06 | 0.94536 | 32.86 | 0.00025058 |  | 2011-02-09 11:49 | 24C | 0.94526 | 2.8901e-06 | 2.0862e-06 | 0.94526 | 32.8561 | 0.00011213 | 0 | 0.0038948
LSSL2010#248 | CB-50 | 1472 |  | 2011-02-04 11:48 | 24C | 0.93539 | 2.4916e-06 | 8.0466e-06 | 0.93539 | 32.4719 | 9.4885e-05 | [10 empty] | 32.4719
LSSL2010#249 | CB-50 | 1473 |  | 2011-02-04 11:50 | 24C | 0.92528 | 1.7702e-06 | 8.0466e-06 | 0.92529 | 32.0793 | 6.7435e-05 | [10 empty] | 32.0793
LSSL2010#250 | CB-50 | 1474 |  | 2011-02-04 11:53 | 24C | 0.90412 | 6.8246e-06 | 7.987e-06 | 0.90413 | 31.259 | 0.00026221 | [10 empty] | 31.259
LSSL2010#251 | CB-50 | 1475 |  | 2011-02-04 11:55 | 24C | 0.87374 | 5.4532e-06 | 7.987e-06 | 0.87374 | 30.0872 | 0.00020899 | [10 empty] | 30.0872
LSSL2010#252 | CB-50 | 1476 |  | 2011-02-04 11:57 | 24C | 0.8198 | 7.9034e-06 | 7.987e-06 | 0.81981 | 28.0232 | 0.00030019 | [10 empty] | 28.0232
LSSL2010#253 | CB-50 | 1477 |  | 2011-02-04 11:59 | 24C | 0.798 | 6.5629e-06 | 7.9274e-06 | 0.79801 | 27.1951 | 0.00024928 | [10 empty] | 27.1951
LSSL2010#209 | CB-51 | 1478 |  | 2011-02-04 09:28 | 24C | 0.99878 | 4.6354e-06 | 8.1062e-06 | 0.99879 | 34.9525 | 0.0001827 |  | 2011-02-09 11:14 | 24C | 0.99872 | 3.0085e-06 | 2.0266e-06 | 0.99872 | 34.9497 | 0.00011691 | 0 | 0.0027733
LSSL2010#210 | CB-51 | 1479 |  | 2011-02-04 09:30 | 24C | 0.99859 | 3.2148e-06 | 7.8678e-06 | 0.9986 | 34.9448 | 0.00012555 | [10 empty] | 34.9448
| CB-51 | 1480 |  | salinity value was determined using SALINOM.EXE Program | [5 empty] | 34.9168 |  | power bump!! #1480 did not get recorded - communications lost- reason unknown- | [9 empty] | 34.9168
LSSL2010#211 | CB-51 | 1481 |  | 2011-02-04 09:50 | 24C | 0.99746 | 4.7428e-06 | 8.1062e-06 | 0.99747 | 34.9003 | 0.00018599 | [10 empty] | 34.9003
LSSL2010#212 | CB-51 | 1482 |  | 2011-02-04 09:55 | 24C | 0.99713 | 3.3596e-06 | 7.8678e-06 | 0.99713 | 34.8872 | 0.00013211 | [10 empty] | 34.8872
LSSL2010#213 | CB-51 | 1483 |  | 2011-02-04 09:58 | 24C | 0.99679 | 3.6105e-06 | 8.1062e-06 | 0.9968 | 34.874 | 0.00014493 | [10 empty] | 34.874
LSSL2010#214 | CB-51 | 1484 |  | 2011-02-04 10:02 | 24C | 0.99634 | 5.5831e-06 | 7.987e-06 | 0.99635 | 34.8563 | 0.00021838 | [10 empty] | 34.8563
LSSL2010#215 | CB-51 | 1485 |  | 2011-02-04 10:26 | 24C | 0.99599 | 8.936e-06 | 7.987e-06 | 0.996 | 34.8424 | 0.00035184 | [10 empty] | 34.8424
LSSL2010#216 | CB-51 | 1486 |  | 2011-02-04 10:28 | 24C | 0.99548 | 3.6104e-06 | 7.987e-06 | 0.99549 | 34.8226 | 0.00013822 | [10 empty] | 34.8226
LSSL2010#217 | CB-51 | 1487 |  | 2011-02-04 10:30 | 24C | 0.99479 | 2.4777e-06 | 8.0466e-06 | 0.9948 | 34.7955 | 9.4938e-05 | [10 empty] | 34.7955
LSSL2010#218 | CB-51 | 1489 |  | 2011-02-04 10:32 | 24C | 0.99086 | 8.5641e-06 | 7.987e-06 | 0.99086 | 34.6407 | 0.00033732 |  | 2011-02-09 11:35 | 24C | 0.99082 | 3.1418e-06 | 1.967e-06 | 0.99082 | 34.639 | 0.00012312 | 0 | 0.0017128
LSSL2010#219 | CB-51 | 1490 |  | 2011-02-04 10:35 | 24C | 0.98615 | 2.7425e-06 | 7.987e-06 | 0.98615 | 34.4557 | 0.00010659 | [10 empty] | 34.4557
LSSL2010#220 | CB-51 | 1491 |  | 2011-02-04 10:37 | 24C | 0.97917 | 5.4684e-06 | 7.987e-06 | 0.97918 | 34.1819 | 0.00021546 | [10 empty] | 34.1819
LSSL2010#221 | CB-51 | 1492 |  | 2011-02-04 10:39 | 24C | 0.96748 | 2.4908e-06 | 7.987e-06 | 0.96748 | 33.7239 | 9.869e-05 | [10 empty] | 33.7239
LSSL2010#222 | CB-51 | 1493 |  | 2011-02-04 10:41 | 24C | 0.95556 | 3.7839e-06 | 7.987e-06 | 0.95557 | 33.2582 | 0.00014715 | [10 empty] | 33.2582
LSSL2010#223 | CB-51 | 1494 |  | 2011-02-04 10:43 | 24C | 0.95306 | 4.6227e-06 | 8.0466e-06 | 0.95307 | 33.1604 | 0.00018375 | [10 empty] | 33.1604
LSSL2010#224 | CB-51 | 1495 |  | 2011-02-04 10:45 | 24C | 0.94591 | 2.7425e-06 | 7.987e-06 | 0.94592 | 32.8819 | 0.00010656 | [10 empty] | 32.8819
LSSL2010#225 | CB-51 | 1496 |  | 2011-02-04 10:47 | 24C | 0.93389 | 2.89e-06 | 7.987e-06 | 0.9339 | 32.4136 | 0.00011207 | [10 empty] | 32.4136
LSSL2010#226 | CB-51 | 1497 |  | 2011-02-04 10:49 | 24C | 0.9245 | 4.1858e-06 | 7.987e-06 | 0.92451 | 32.049 | 0.00016335 | [10 empty] | 32.049
LSSL2010#227 | CB-51 | 1498 |  | 2011-02-04 10:52 | 24C | 0.90438 | 7.9878e-06 | 8.0466e-06 | 0.90439 | 31.2691 | 0.00030599 | [10 empty] | 31.2691
LSSL2010#228 | CB-51 | 1499 |  | 2011-02-04 10:54 | 24C | 0.88398 | 4.6675e-06 | 7.987e-06 | 0.88399 | 30.4814 | 0.00018152 | [10 empty] | 30.4814
LSSL2010#229 | CB-51 | 1500 |  | 2011-02-04 10:57 | 24C | 0.83764 | 4.3337e-06 | 7.987e-06 | 0.83765 | 28.7036 | 0.00016464 |  | 2011-02-09 11:56 | 24C | 0.83791 | 3.6105e-06 | 1.967e-06 | 0.83791 | 28.7137 | 0.00013841 | 0 | -0.010147
LSSL2010#325 | CB-52 | 1502 |  | 2011-02-07 12:51 | 24C | 0.99727 | 2.4912e-06 | 4.0054e-05 | 0.99731 | 34.894 | 9.869e-05 | SZ- Correct, there was no 1501 niskin | [9 empty] | 34.894
LSSL2010#326 | CB-52 | 1503 |  | 2011-02-07 12:53 | 24C | 0.99694 | 2.7262e-06 | 4.0054e-05 | 0.99698 | 34.8812 | 0.00010761 | [10 empty] | 34.8812
LSSL2010#327 | CB-52 | 1504 |  | 2011-02-07 12:55 | 24C | 0.99677 | 3.3308e-06 | 4.0054e-05 | 0.99681 | 34.8746 | 0.00012839 | [10 empty] | 34.8746
LSSL2010#328 | CB-52 | 1505 |  | 2011-02-07 12:57 | 24C | 0.99655 | 2.7425e-06 | 3.9995e-05 | 0.99659 | 34.8657 | 0.00010865 | [10 empty] | 34.8657
LSSL2010#329 | CB-52 | 1506 |  | 2011-02-07 12:59 | 24C | 0.9963 | 3.8748e-06 | 4.0054e-05 | 0.99634 | 34.856 | 0.00015161 |  | 2011-02-09 11:25 | 24C | 0.99627 | 1.1529e-06 | 1.9073e-06 | 0.99628 | 34.8535 | 4.5227e-05 | 0 | 0.0025291
LSSL2010#330 | CB-52 | 1507 |  | 2011-02-07 13:01 | 24C | 0.99594 | 5.7377e-06 | 3.9995e-05 | 0.99598 | 34.8418 | 0.00022904 | [10 empty] | 34.8418
LSSL2010#331 | CB-52 | 1508 |  | 2011-02-07 13:03 | 24C | 0.99578 | 3.2144e-06 | 3.9995e-05 | 0.99582 | 34.8356 | 0.00012382 | [10 empty] | 34.8356
LSSL2010#332 | CB-52 | 1509 |  | 2011-02-07 13:05 | 24C | 0.99544 | 3.3595e-06 | 4.0054e-05 | 0.99548 | 34.8223 | 0.00013414 | [10 empty] | 34.8223
LSSL2010#333 | CB-52 | 1510 |  | 2011-02-07 13:07 | 24C | 0.99547 | 3.6104e-06 | 3.9995e-05 | 0.99551 | 34.8233 | 0.00014157 | [10 empty] | 34.8233
LSSL2010#334 | CB-52 | 1511 |  | 2011-02-07 13:09 | 24C | 0.99506 | 1.1536e-06 | 3.9995e-05 | 0.9951 | 34.8074 | 4.2703e-05 | [10 empty] | 34.8074
LSSL2010#335 | CB-52 | 1512 |  | 2011-02-07 13:11 | 24C | 0.99378 | 6.4718e-06 | 4.0054e-05 | 0.99382 | 34.7567 | 0.00025338 |  | 2011-02-09 11:33 | 24C | 0.99375 | 2.4915e-06 | 2.0266e-06 | 0.99376 | 34.7544 | 9.4885e-05 | 0 | 0.0023613
LSSL2010#336 | CB-52 | 1513 |  | 2011-02-07 13:13 | 24C | 0.98935 | 2.2394e-06 | 4.0054e-05 | 0.98939 | 34.5827 | 8.7067e-05 | [10 empty] | 34.5827
LSSL2010#337 | CB-52 | 1514 |  | 2011-02-07 13:15 | 24C | 0.9861 | 3.865e-06 | 3.9935e-05 | 0.98614 | 34.4551 | 0.00015134 | [10 empty] | 34.4551
LSSL2010#338 | CB-52 | 1515 |  | 2011-02-07 13:17 | 24C | 0.97456 | 3.1253e-06 | 3.9935e-05 | 0.9746 | 34.0026 | 0.00012 | [10 empty] | 34.0026
LSSL2010#339 | CB-52 | 1516 |  | 2011-02-07 13:19 | 24C | 0.97496 | 2.2263e-06 | 3.9995e-05 | 0.975 | 34.0181 | 8.6163e-05 | [10 empty] | 34.0181
LSSL2010#340 | CB-52 | 1517 |  | 2011-02-07 13:21 | 24C | 0.96638 | 6.311e-06 | 3.9995e-05 | 0.96642 | 33.6824 | 0.00024824 | [10 empty] | 33.6824
LSSL2010#341 | CB-52 | 1518 |  | 2011-02-07 13:24 | 24C | 0.93328 | 3.4662e-06 | 4.0054e-05 | 0.93332 | 32.3911 | 0.00013479 |  | 2011-02-09 11:42 | 24C | 0.93325 | 1.4046e-06 | 1.9073e-06 | 0.93325 | 32.3885 | 5.2235e-05 | 0 | 0.0025291
LSSL2010#342 | CB-52 | 1519 |  | 2011-02-07 13:27 | 24C | 0.94519 | 2.7425e-06 | 3.9935e-05 | 0.94523 | 32.8547 | 0.00010975 | Loose cap - loose insert | [9 empty] | 32.8547
LSSL2010#343 | CB-52 | 1520 |  | 2011-02-07 13:29 | 24C | 0.93398 | 3.5406e-06 | 3.9935e-05 | 0.93402 | 32.4185 | 0.00013765 | [10 empty] | 32.4185
LSSL2010#344 | CB-52 | 1521 |  | 2011-02-07 13:31 | 24C | 0.92562 | 2.4909e-06 | 3.9995e-05 | 0.92566 | 32.0934 | 9.802e-05 | [10 empty] | 32.0934
LSSL2010#345 | CB-52 | 1522 |  | 2011-02-07 13:33 | 24C | 0.8964 | 3.1417e-06 | 3.9935e-05 | 0.89644 | 30.9619 | 0.00012365 | [10 empty] | 30.9619
LSSL2010#346 | CB-52 | 1523 |  | 2011-02-07 13:35 | 24C | 0.87116 | 2.2259e-06 | 3.9935e-05 | 0.87119 | 29.9891 | 8.5326e-05 | [10 empty] | 29.9891
LSSL2010#347 | CB-52 | 1524 |  | 2011-02-07 13:37 | 24C | 0.84289 | 3.8434e-06 | 3.9995e-05 | 0.84293 | 28.9056 | 0.00014661 | [10 empty] | 28.9056
LSSL2010#348 | CB-52 | 1525 |  | 2011-02-07 13:39 | 24C | 0.83498 | 2.2397e-06 | 3.9995e-05 | 0.83502 | 28.6032 | 8.4895e-05 | [10 empty] | 28.6032
LSSL2010#306 | CB-53 | 1526 |  | 2011-02-07 12:00 | 24C | 0.99679 | 4.189e-06 | 3.9995e-05 | 0.99683 | 34.8753 | 0.00016758 | [10 empty] | 34.8753
LSSL2010#307 | CB-54 | 1527 |  | 2011-02-07 12:02 | 24C | 0.99543 | 3.4659e-06 | 3.9995e-05 | 0.99547 | 34.8216 | 0.00013699 | [10 empty] | 34.8216
LSSL2010#308 | CB-55 | 1528 |  | 2011-02-07 12:04 | 24C | 0.99462 | 2.4615e-06 | 3.9995e-05 | 0.99466 | 34.79 | 9.869e-05 |  | 2011-02-09 11:30 | 24C | 0.9945 | 3.3588e-06 | 2.0266e-06 | 0.9945 | 34.7838 | 0.00013341 | 0 | 0.0061569
LSSL2010#309 | CB-56 | 1529 |  | 2011-02-07 12:06 | 24C | 0.99408 | 3.8647e-06 | 3.9935e-05 | 0.99412 | 34.7686 | 0.00015042 | [10 empty] | 34.7686
LSSL2010#310 | CB-57 | 1530 |  | 2011-02-07 12:08 | 24C | 0.99263 | 1.1532e-06 | 4.0054e-05 | 0.99267 | 34.7118 | 4.4374e-05 | [10 empty] | 34.7118
LSSL2010#311 | CB-58 | 1531 |  | 2011-02-07 12:10 | 24C | 0.99028 | 3.131e-06 | 4.0054e-05 | 0.99032 | 34.6193 | 0.00012312 | [10 empty] | 34.6193
LSSL2010#312 | CB-59 | 1532 |  | 2011-02-07 12:13 | 24C | 0.9884 | 1.1203e-06 | 3.9995e-05 | 0.98844 | 34.5455 | 4.5227e-05 | SZ- Correct, 1533 was a misfire | [9 empty] | 34.5455
LSSL2010#313 | CB-60 | 1534 |  | 2011-02-07 12:14 | 24C | 0.97954 | 3.623e-06 | 3.9995e-05 | 0.97958 | 34.1979 | 0.00014089 | [10 empty] | 34.1979
LSSL2010#314 | CB-61 | 1535 |  | 2011-02-07 12:17 | 24C | 0.96907 | 2.8792e-06 | 3.9995e-05 | 0.96911 | 33.7876 | 0.00011402 | [10 empty] | 33.7876
LSSL2010#315 | CB-62 | 1536 |  | 2011-02-07 12:19 | 24C | 0.95548 | 2.6733e-06 | 4.0054e-05 | 0.95552 | 33.2562 | 0.00010274 |  | 2011-02-09 11:46 | 24C | 0.95471 | 2.4911e-06 | 1.967e-06 | 0.95471 | 33.2247 | 9.4938e-05 | 0 | 0.031494
LSSL2010#316 | CB-63 | 1537 |  | 2011-02-07 12:21 | 24C | 0.95068 | 2.2398e-06 | 4.0054e-05 | 0.95072 | 33.069 | 8.7048e-05 | [10 empty] | 33.069
LSSL2010#317 | CB-64 | 1538 |  | 2011-02-07 12:23 | 24C | 0.94154 | 2.8901e-06 | 3.9995e-05 | 0.94158 | 32.7127 | 0.00011145 | [10 empty] | 32.7127
LSSL2010#318 | CB-65 | 1539 |  | 2011-02-07 12:25 | 24C | 0.93287 | 2.4747e-06 | 3.9995e-05 | 0.93291 | 32.3751 | 9.4938e-05 | [10 empty] | 32.3751
LSSL2010#319 | CB-66 | 1540 |  | 2011-02-07 12:27 | 24C | 0.92293 | 2.4915e-06 | 4.0054e-05 | 0.92297 | 31.9892 | 9.3957e-05 | [10 empty] | 31.9892
LSSL2010#320 | CB-67 | 1541 |  | 2011-02-07 12:29 | 24C | 0.90216 | 3.6105e-06 | 4.0054e-05 | 0.9022 | 31.1845 | 0.00013706 | [10 empty] | 31.1845
LSSL2010#321 | CB-68 | 1542 |  | 2011-02-07 12:31 | 24C | 0.86296 | 2.2401e-06 | 4.0054e-05 | 0.863 | 29.6744 | 8.7006e-05 |  | 2011-02-09 11:51 | 24C | 0.86055 | 2.2394e-06 | 1.967e-06 | 0.86055 | 29.5804 | 8.4887e-05 | 0 | 0.094038
LSSL2010#322 | CB-69 | 1543 |  | 2011-02-07 12:33 | 24C | 0.83224 | 4.384e-06 | 4.0054e-05 | 0.83228 | 28.4987 | 0.00016658 | SZ- Changed from 1544 to 1543 based on elog | [9 empty] | 28.4987
LSSL2010#323 | CB-70 | 1544 |  | 2011-02-07 12:36 | 24C | 0.8314 | 2.7426e-06 | 3.9935e-05 | 0.83144 | 28.4664 | 0.00010449 | [10 empty] | 28.4664
LSSL2010#173 | CB-54 | 1550 |  | 2011-02-03 15:57 | 24C | 0.98046 | 8.5219e-06 | 2.4021e-05 | 0.98049 | 34.2332 | 0.00033563 | SZ- Correct, gap from 1544 to 1550 | [9 empty] | 34.2332
LSSL2010#174 | CB-54 | 1551 |  | 2011-02-03 15:59 | 24C | 0.97575 | 1.7599e-06 | 2.4021e-05 | 0.97578 | 34.0487 | 6.8803e-05 |  | 2011-02-09 11:40 | 24C | 0.9757 | 1.7597e-06 | 2.0266e-06 | 0.9757 | 34.0458 | 6.7435e-05 | 0 | 0.0029182
LSSL2010#175 | CB-54 | 1552 |  | 2011-02-03 16:01 | 24C | 0.96858 | 4.4373e-06 | 2.4021e-05 | 0.9686 | 33.7676 | 0.0001706 | [10 empty] | 33.7676
LSSL2010#176 | CB-54 | 1553 |  | 2011-02-03 16:04 | 24C | 0.95153 | 2.7425e-06 | 2.4021e-05 | 0.95155 | 33.1014 | 0.00010447 | [10 empty] | 33.1014
LSSL2010#177 | CB-54 | 1554 |  | 2011-02-03 16:06 | 24C | 0.93691 | 5.1439e-06 | 2.3961e-05 | 0.93693 | 32.5316 | 0.00020148 | [10 empty] | 32.5316
LSSL2010#178 | CB-54 | 1555 |  | 2011-02-03 16:09 | 24C | 0.92335 | 2.2394e-06 | 2.4021e-05 | 0.92338 | 32.0049 | 8.5299e-05 | [10 empty] | 32.0049
LSSL2010#179 | CB-54 | 1556 |  | 2011-02-03 16:11 | 24C | 0.90128 | 4.5128e-06 | 2.3961e-05 | 0.90131 | 31.15 | 0.00017531 | [10 empty] | 31.15
LSSL2010#180 | CB-54 | 1557 |  | 2011-02-03 16:13 | 24C | 0.86233 | 3.6104e-06 | 2.3961e-05 | 0.86235 | 29.6494 | 0.00013702 |  | 2011-02-09 11:53 | 24C | 0.86093 | 2.2126e-06 | 2.1458e-06 | 0.86093 | 29.5951 | 8.8712e-05 | 0 | 0.054281
LSSL2010#181 | CB-54 | 1558 |  | 2011-02-03 16:15 | 24C | 0.83805 | 4.1137e-06 | 2.4021e-05 | 0.83807 | 28.7197 | 0.00015439 | [10 empty] | 28.7197
LSSL2010#182 | CB-54 | 1559 |  | 2011-02-03 16:17 | 24C | 0.83802 | 4.9101e-06 | 2.3901e-05 | 0.83805 | 28.7187 | 0.00018646 | [10 empty] | 28.7187
0
0
0
LSSL2010#693 | LOOP | 96 |  | 2011-02-11 11:50 | 24C | 0.7901 | 3.4657e-06 | 5.126e-06 | 0.79011 | 26.8956 | 0.00013117 | [10 empty] | 26.8956
LSSL2010#694 | LOOP | 100 |  | 2011-02-11 11:52 | 24C | 0.76471 | 3.1414e-06 | 5.0068e-06 | 0.76471 | 25.9367 | 0.00011969 | [10 empty] | 25.9367
LSSL2010#695 | LOOP | 102 |  | 2011-02-11 11:54 | 24C | 0.76467 | 3.2145e-06 | 5.0068e-06 | 0.76468 | 25.9353 | 0.00011784 | [10 empty] | 25.9353
LSSL2010#696 | LOOP | 105 |  | 2011-02-11 11:56 | 24C | 0.77809 | 1.3712e-06 | 5.0068e-06 | 0.77809 | 26.4413 | 5.2783e-05 | SZ- Changed from 105 to 107 | [9 empty] | 26.4413
LSSL2010#697 | LOOP | 107 |  | 2011-02-11 11:58 | 24C | 0.78529 | 2.4908e-06 | 5.0068e-06 | 0.7853 | 26.7136 | 8.962e-05 | [10 empty] | 26.7136
LSSL2010#698 | LOOP | 110 |  | 2011-02-11 12:00 | 24C | 0.8168 | 2.4915e-06 | 5.126e-06 | 0.81681 | 27.909 | 8.9222e-05 | [10 empty] | 27.909
LSSL2010#699 | LOOP | 115 |  | 2011-02-11 12:02 | 24C | 0.75796 | 2.2391e-06 | 5.0068e-06 | 0.75796 | 25.6825 | 8.2749e-05 | [10 empty] | 25.6825
LSSL2010#700 | LOOP | 117 |  | 2011-02-11 12:05 | 24C | 0.75627 | 2.7425e-06 | 5.0068e-06 | 0.75627 | 25.6189 | 0.00010135 | [10 empty] | 25.6189
LSSL2010#701 | LOOP | 122 |  | 2011-02-11 12:06 | 24C | 0.72351 | 3.854e-06 | 5.0068e-06 | 0.72352 | 24.3913 | 0.00014262 | [10 empty] | 24.3913
LSSL2010#702 | LOOP | 126 |  | 2011-02-11 12:09 | 24C | 0.73684 | 2.4908e-06 | 5.0068e-06 | 0.73685 | 24.8898 | 9.4459e-05 | [10 empty] | 24.8898
LSSL2010#703 | LOOP | 128 |  | 2011-02-11 12:11 | 24C | 0.76751 | 2.7426e-06 | 5.0068e-06 | 0.76752 | 26.0424 | 0.00010448 | [10 empty] | 26.0424
LSSL2010#704 | LOOP | 131 |  | 2011-02-11 12:14 | 24C | 0.79058 | 2.7425e-06 | 5.0068e-06 | 0.79059 | 26.9138 | 9.9246e-05 | [10 empty] | 26.9138
LSSL2010#705 | LOOP | 134 |  | 2011-02-11 12:16 | 24C | 0.87179 | 2.0945e-06 | 5.0068e-06 | 0.8718 | 30.0123 | 7.9449e-05 | [10 empty] | 30.0123
These have already been dealt with - ignore
LSSL2010#464 | CB-7 | 719 | d | 2011-02-09 14:49 | 24C | 0.99568 | 4.1469e-06 | 2.1458e-06 | 0.99568 | 34.8301 | 0.00016043
LSSL2010#473 | CB-7 | 726 | d | 2011-02-09 15:10 | 24C | 0.9563 | 2.4915e-06 | 2.0266e-06 | 0.9563 | 33.2868 | 9.4981e-05
LSSL2010#458 | CB-8 | 739 | d | 2011-02-09 12:59 | 24C | 0.99675 | 7.2738e-06 | 1.967e-06 | 0.99675 | 34.8723 | 0.00028709
LSSL2010#463 | CB-8 | 743 | d | 2011-02-09 13:10 | 24C | 0.99582 | 2.2394e-06 | 2.0266e-06 | 0.99582 | 34.8356 | 8.793e-05
LSSL2010#472 | CB-8 | 750 | d | 2011-02-09 15:07 | 24C | 0.94997 | 3.6104e-06 | 2.0266e-06 | 0.94997 | 33.0396 | 0.00013925
LSSL2010#461 | CB-10 | 789 | d | 2011-02-09 13:06 | 24C | 0.9963 | 3.7354e-06 | 2.0862e-06 | 0.9963 | 34.8545 | 0.00014444
LSSL2010#468 | CB-10 | 795 | d | 2011-02-09 14:58 | 24C | 0.98012 | 3.6105e-06 | 2.0266e-06 | 0.98012 | 34.2189 | 0.0001408
LSSL2010#475 | CB10 | 800 | d | 2011-02-09 15:14 | 24C | 0.93642 | 3.1314e-06 | 1.967e-06 | 0.93642 | 32.5118 | 0.00012312
LSSL2010#456 | CB-70 | 806 | d | 2011-02-09 12:55 | 24C | 0.99892 | 2.2397e-06 | 2.0266e-06 | 0.99892 | 34.9577 | 8.7013e-05
LSSL2010#457 | CB-70 | 810 | d | 2011-02-09 12:57 | 24C | 0.99751 | 1.1529e-06 | 1.9073e-06 | 0.99751 | 34.9021 | 4.265e-05
LSSL2010#460 | CB-70 | 811 | d | 2011-02-09 13:03 | 24C | 0.99677 | 2.8661e-06 | 2.0266e-06 | 0.99678 | 34.8731 | 0.00011014
LSSL2010#469 | CB70 | 819 | d | 2011-02-09 15:00 | 24C | 0.97841 | 2.2394e-06 | 1.9073e-06 | 0.97841 | 34.1518 | 8.7024e-05
LSSL2010#471 | CB-70 | 821 | d | 2011-02-09 15:04 | 24C | 0.95214 | 3.4359e-06 | 2.0862e-06 | 0.95214 | 33.1243 | 0.00013203
LSSL2010#466 | CB-70 | 830 | d | 2011-02-09 14:54 | 24C | 0.99461 | 3.4657e-06 | 2.0266e-06 | 0.99462 | 34.7882 | 0.00013904
LSSL2010#467 | CB-70 | 836 | d | 2011-02-09 14:56 | 24C | 0.98182 | 3.4658e-06 | 2.0266e-06 | 0.98182 | 34.2856 | 0.00013595
LSSL2010#474 | CB-70 | 845 | d | 2011-02-09 15:12 | 24C | 0.93661 | 2.4916e-06 | 2.0266e-06 | 0.93661 | 32.5192 | 9.6976e-05
LSSL2010#459 | CB-11 | 859 | d | 2011-02-09 13:01 | 24C | 0.99677 | 3.1145e-06 | 2.0862e-06 | 0.99678 | 34.8731 | 0.00011967
LSSL2010#465 | CB-11 | 863 | d | 2011-02-09 14:52 | 24C | 0.99573 | 3.1418e-06 | 1.967e-06 | 0.99574 | 34.8322 | 0.00012381
LSSL2010#470 | CB-11 | 870 | d | 2011-02-09 15:02 | 24C | 0.95334 | 2.7425e-06 | 1.9073e-06 | 0.95334 | 33.171 | 0.00010761
LSSL2010#462 | CB-9 | 909 | d | 2011-02-09 13:08 | 24C | 0.99624 | 3.9994e-06 | 2.0266e-06 | 0.99625 | 34.8523 | 0.00015603
LSSL2010#509 | CB-9 | 915 | d | 2011-02-09 17:14 | 24C | 0.97714 | 2.89e-06 | 1.967e-06 | 0.97715 | 34.1023 | 0.00011317
LSSL2010#513 | CB-9 | 920 | d | 2011-02-09 17:23 | 24C | 0.93825 | 2.4908e-06 | 1.967e-06 | 0.93825 | 32.583 | 9.869e-05
LSSL2010#515 | CB-9 | 923 | d | 2011-02-09 17:28 | 24C | 0.89184 | 3.2141e-06 | 2.0862e-06 | 0.89185 | 30.7846 | 0.00012629
LSSL2010#503 | CB-12 | 957 | d | 2011-02-09 16:59 | 24C | 0.99432 | 2.4908e-06 | 2.0266e-06 | 0.99432 | 34.7767 | 9.4885e-05
LSSL2010#508 | CB-12 | 963 | d | 2011-02-09 17:12 | 24C | 0.9826 | 1.4046e-06 | 1.967e-06 | 0.9826 | 34.3163 | 5.2235e-05
LSSL2010#512 | CB-12 | 968 | d | 2011-02-09 17:21 | 24C | 0.93818 | 2.0945e-06 | 1.967e-06 | 0.93818 | 32.5802 | 8.1861e-05
LSSL2010#507 | CB-12a | 974 | d | 2011-02-09 17:09 | 24C | 0.99344 | 2.4751e-06 | 2.0266e-06 | 0.99344 | 34.7418 | 9.9792e-05
LSSL2010#510 | CB-12a | 980 | d | 2011-02-09 17:17 | 24C | 0.97401 | 1.1536e-06 | 2.0266e-06 | 0.97401 | 33.9794 | 4.2703e-05
LSSL2010#514 | CB-12a | 989 | d | 2011-02-09 17:26 | 24C | 0.93005 | 3.1424e-06 | 2.0266e-06 | 0.93005 | 32.264 | 0.00012176
LSSL2010#502 | CB-16 | 1003 | d | 2011-02-09 16:57 | 24C | 0.997 | 2.8907e-06 | 2.1458e-06 | 0.997 | 34.8821 | 0.00011519
LSSL2010#505 | CB-16 | 1007 | d | 2011-02-09 17:04 | 24C | 0.99575 | 1.3712e-06 | 1.9073e-06 | 0.99576 | 34.833 | 5.5371e-05
LSSL2010#511 | CB-16 | 1014 | d | 2011-02-09 17:18 | 24C | 0.95357 | 3.6119e-06 | 2.1458e-06 | 0.95357 | 33.1801 | 0.00013908
LSSL2010#501 | CB-13 | 1047 | d | 2011-02-09 16:49 | 24C | 0.99887 | 2.8543e-06 | 1.967e-06 | 0.99887 | 34.9556 | 0.00011192
LSSL2010#504 | CB-13 | 1055 | d | 2011-02-09 17:02 | 24C | 0.99643 | 2.4908e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 9.9543e-05
LSSL2010#516 | CB-13 | 1068 | d | 2011-02-09 17:30 | 24C | 0.83527 | 3.6104e-06 | 2.0266e-06 | 0.83527 | 28.6127 | 0.00013717
LSSL2010#506 | CB-15 | 1070 | d | 2011-02-09 17:06 | 24C | 0.99425 | 2.7425e-06 | 1.9073e-06 | 0.99425 | 34.7739 | 0.00010663
LSSL2010#449 | CB-15 | 1076 | d | 2011-02-09 12:30 | 24C | 0.97279 | 2.4909e-06 | 1.9073e-06 | 0.97279 | 33.9317 | 9.7029e-05
LSSL2010#451 | CB-15 | 1085 | d | 2011-02-09 12:35 | 24C | 0.93015 | 2.7262e-06 | 2.0266e-06 | 0.93015 | 32.2681 | 0.00010659
LSSL2010#447 | CB-15 | 1127 | d | 2011-02-09 12:22 | 24C | 0.99574 | 1.3713e-06 | 1.967e-06 | 0.99574 | 34.8323 | 5.4326e-05
LSSL2010#448 | CB-15 | 1129 | d | 2011-02-09 12:24 | 24C | 0.99429 | 2.4909e-06 | 2.0266e-06 | 0.9943 | 34.7756 | 0.00010437
LSSL2010#436 | CB-15 | 1134 | d | 2011-02-09 11:44 | 24C | 0.95343 | 1.7026e-05 | 1.9073e-06 | 0.95343 | 33.1749 | 0.0006637
LSSL2010#443 | CB-17 | 1143 | d | 2011-02-09 12:12 | 24C | 0.99889 | 2.8792e-06 | 2.0266e-06 | 0.99889 | 34.9563 | 0.00011317
LSSL2010#445 | CB-17 | 1148 | d | 2011-02-09 12:17 | 24C | 0.9966 | 2.4908e-06 | 2.0266e-06 | 0.99661 | 34.8665 | 9.6642e-05 | SZ- Correct there was no sample 1149 taken
LSSL2010#453 | CB-17 | 1162 | d | 2011-02-09 12:41 | 24C | 0.91951 | 3.4666e-06 | 1.967e-06 | 0.91951 | 31.855 | 0.00013693
LSSL2010#446 | PP-7 | 1196 | d | 2011-02-09 12:20 | 24C | 0.9967 | 4.9165e-06 | 2.0266e-06 | 0.9967 | 34.8702 | 0.00019397
LSSL2010#485 | PP-7 | 1204 | d | 2011-02-09 15:58 | 24C | 0.97918 | 2.7425e-06 | 2.0266e-06 | 0.97918 | 34.182 | 0.00010865
LSSL2010#496 | PP-7 | 1213 | d | 2011-02-09 16:23 | 24C | 0.77691 | 1.1196e-06 | 1.967e-06 | 0.77691 | 26.3966 | 4.3716e-05
LSSL2010#479 | PP-6 | 1240 | d | 2011-02-09 15:44 | 24C | 0.99853 | 2.2126e-06 | 2.0862e-06 | 0.99853 | 34.9422 | 8.8728e-05
LSSL2010#489 | PP-6 | 1253 | d | 2011-02-09 16:07 | 24C | 0.95439 | 2.4909e-06 | 2.0266e-06 | 0.9544 | 33.2124 | 9.4981e-05
LSSL2010#493 | PP-6 | 1256 | d | 2011-02-09 16:16 | 24C | 0.93244 | 3.1417e-06 | 2.0266e-06 | 0.93245 | 32.3572 | 0.00012313
LSSL2010#480 | CB-DW | 1264 | d | 2011-02-09 15:46 | 24C | 0.99792 | 1.3882e-06 | 2.0862e-06 | 0.99792 | 34.9182 | 5.2278e-05
LSSL2010#481 | CB-DW | 1265 | d | 2011-02-09 15:48 | 24C | 0.99708 | 0 | 2.0862e-06 | 0.99708 | 34.8853 | 0
LSSL2010#487 | CB-DW | 1275 | d | 2011-02-09 16:03 | 24C | 0.97482 | 2.4911e-06 | 2.0266e-06 | 0.97482 | 34.0113 | 9.869e-05
LSSL2010#490 | CB-DW | 1277 | d | 2011-02-09 16:09 | 24C | 0.95025 | 2.866e-06 | 2.0862e-06 | 0.95025 | 33.0508 | 0.00011317
LSSL2010#492 | CB-DW | 1279 | d | 2011-02-09 16:14 | 24C | 0.93916 | 1.1203e-06 | 2.0862e-06 | 0.93916 | 32.6184 | 4.4364e-05
LSSL2010#494 | CB-DW | 1281 | d | 2011-02-09 16:19 | 24C | 0.91977 | 1.1196e-06 | 2.0266e-06 | 0.91977 | 31.8649 | 4.3729e-05
LSSL2010#497 | CB-DW | 1285 | d | 2011-02-09 16:26 | 24C | 0.76381 | 0 | 2.0266e-06 | 0.76381 | 25.9028 | 1.0662e-06
LSSL2010#482 | CB-18 | 1289 | d | 2011-02-09 15:50 | 24C | 0.99695 | 5.2144e-06 | 2.0266e-06 | 0.99695 | 34.88 | 0.00020293
LSSL2010#483 | CB-18 | 1293 | d | 2011-02-09 15:53 | 24C | 0.99575 | 2.4908e-06 | 2.0266e-06 | 0.99576 | 34.833 | 9.9734e-05
LSSL2010#488 | CB-18 | 1300 | d | 2011-02-09 16:05 | 24C | 0.96649 | 3.1905e-06 | 2.0266e-06 | 0.96649 | 33.6851 | 0.00012553
LSSL2010#478 | CB-19 | 1312 | d | 2011-02-09 15:42 | 24C | 0.99884 | 2.0804e-06 | 2.0862e-06 | 0.99884 | 34.9544 | 8.1621e-05
LSSL2010#484 | CB-19 | 1319 | d | 2011-02-09 15:55 | 24C | 0.99573 | 2.7264e-06 | 2.0266e-06 | 0.99573 | 34.8319 | 0.00010344
LSSL2010#499 | CB-19 | 1333 | d | 2011-02-09 16:31 | 24C | 0.73494 | 3.3588e-06 | 1.967e-06 | 0.73494 | 24.8184 | 0.0001271
LSSL2010#477 | CB-6 | 1334 | d | 2011-02-09 15:39 | 24C | 0.99885 | 2.4915e-06 | 1.967e-06 | 0.99885 | 34.9549 | 0.00010059
LSSL2010#491 | CB-6 | 1351 | d | 2011-02-09 16:12 | 24C | 0.93939 | 3.1418e-06 | 1.9073e-06 | 0.9394 | 32.6276 | 0.00012486
LSSL2010#486 | CB-21 | 1386 | d | 2011-02-09 16:00 | 24C | 0.97834 | 3.131e-06 | 2.0266e-06 | 0.97834 | 34.1491 | 0.00012172
LSSL2010#495 | CB-21 | 1401 | d | 2011-02-09 16:21 | 24C | 0.88377 | 3.1418e-06 | 1.967e-06 | 0.88377 | 30.473 | 0.00012365
LSSL2010#498 | CB-21 | 1405 | d | 2011-02-09 16:29 | 24C | 0.74217 | 2.2135e-06 | 2.0862e-06 | 0.74217 | 25.0894 | 8.5299e-05
LSSL2010#444 | CB-21 | 1410 | d | 2011-02-09 12:15 | 24C | 0.99785 | 3.4658e-06 | 1.967e-06 | 0.99785 | 34.9155 | 0.00013387
LSSL2010#452 | CB-21 | 1425 | d | 2011-02-09 12:38 | 24C | 0.92289 | 4.9162e-06 | 2.0266e-06 | 0.92289 | 31.986 | 0.00018936
LSSL2010#454 | CB21 | 1429 | d | 2011-02-09 12:43 | 24C | 0.74066 | 2.2391e-06 | 1.967e-06 | 0.74066 | 25.0329 | 8.872e-05
LSSL2010#426 | CB-40 | 1433 | d | 2011-02-09 11:18 | 24C | 0.99786 | 4.2856e-06 | 1.967e-06 | 0.99786 | 34.9157 | 0.00016832
LSSL2010#429 | CB-40 | 1440 | d | 2011-02-09 11:27 | 24C | 0.99473 | 2.4908e-06 | 2.0266e-06 | 0.99473 | 34.7926 | 0.00010059
LSSL2010#450 | CB-40 | 1447 | d | 2011-02-09 12:33 | 24C | 0.96254 | 1.1196e-06 | 2.0266e-06 | 0.96254 | 33.5307 | 4.265e-05
LSSL2010#427 | CB-50 | 1458 | d | 2011-02-09 11:22 | 24C | 0.99707 | 2.4909e-06 | 2.0266e-06 | 0.99707 | 34.8846 | 9.4938e-05
LSSL2010#433 | CB-50 | 1467 | d | 2011-02-09 11:37 | 24C | 0.98604 | 2.2396e-06 | 1.9073e-06 | 0.98604 | 34.4513 | 8.8203e-05
LSSL2010#438 | CB-50 | 1471 | d | 2011-02-09 11:49 | 24C | 0.94526 | 2.8901e-06 | 2.0862e-06 | 0.94526 | 32.8561 | 0.00011213
LSSL2010#425 | CB-51 | 1478 | d | 2011-02-09 11:14 | 24C | 0.99872 | 3.0085e-06 | 2.0266e-06 | 0.99872 | 34.9497 | 0.00011691
LSSL2010#432 | CB-51 | 1489 | d | 2011-02-09 11:35 | 24C | 0.99082 | 3.1418e-06 | 1.967e-06 | 0.99082 | 34.639 | 0.00012312
LSSL2010#441 | CB-51 | 1500 | d | 2011-02-09 11:56 | 24C | 0.83791 | 3.6105e-06 | 1.967e-06 | 0.83791 | 28.7137 | 0.00013841
LSSL2010#428 | CB-52 | 1506 | d | 2011-02-09 11:25 | 24C | 0.99627 | 1.1529e-06 | 1.9073e-06 | 0.99628 | 34.8535 | 4.5227e-05
LSSL2010#431 | CB-52 | 1512 | d | 2011-02-09 11:33 | 24C | 0.99375 | 2.4915e-06 | 2.0266e-06 | 0.99376 | 34.7544 | 9.4885e-05
LSSL2010#435 | CB-52 | 1518 | d | 2011-02-09 11:42 | 24C | 0.93325 | 1.4046e-06 | 1.9073e-06 | 0.93325 | 32.3885 | 5.2235e-05
LSSL2010#430 | CB-53 | 1528 | d | 2011-02-09 11:30 | 24C | 0.9945 | 3.3588e-06 | 2.0266e-06 | 0.9945 | 34.7838 | 0.00013341
LSSL2010#437 | CB-53 | 1536 | d | 2011-02-09 11:46 | 24C | 0.95471 | 2.4911e-06 | 1.967e-06 | 0.95471 | 33.2247 | 9.4938e-05
LSSL2010#439 | CB-53 | 1542 | d | 2011-02-09 11:51 | 24C | 0.86055 | 2.2394e-06 | 1.967e-06 | 0.86055 | 29.5804 | 8.4887e-05
LSSL2010#434 | CB-54 | 1551 | d | 2011-02-09 11:40 | 24C | 0.9757 | 1.7597e-06 | 2.0266e-06 | 0.9757 | 34.0458 | 6.7435e-05
LSSL2010#440 | CB-54 | 1557 | d | 2011-02-09 11:53 | 24C | 0.86093 | 2.2126e-06 | 2.1458e-06 | 0.86093 | 29.5951 | 8.8712e-05

## Autosal Data Examplesheet
2010-07 Sal Data | Example data in pale yellow
coloured data is to be entered on final paper in that format
Salinometer ID | Station | Sample ID (bottle label) | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#2 | P151 |  | 2010-09-18 13:08 | 24C | 0.99997 | 2.7426e-06 | 1.9073e-06 | 0.99997
2010-07#23 | CABOS | 23 | 2010-09-18 14:29 | 24C | 0.9972 | 4.6253e-06 | 1.967e-06 | 0.9972 | 34.8897 | 0.00017952
2010-07#24 | CABOS | 24 | 2010-09-18 14:33 | 24C | 0.99698 | 3.2141e-06 | 2.0266e-06 | 0.99699 | 34.8814 | 0.00012599
2010-07#25 | CABOS | 25 | 2010-09-18 14:35 | 24C | 0.99679 | 5.0935e-06 | 2.0266e-06 | 0.99679 | 34.8738 | 0.00020105
2010-07#26 | CABOS | 26 | 2010-09-18 16:01 | 24C | 0.99659 | 3.5943e-06 | 2.0266e-06 | 0.9966 | 34.866 | 0.00013924
2010-07#27 | CABOS | 27 | 2010-09-18 16:05 | 24C | 0.99637 | 3.8544e-06 | 1.967e-06 | 0.99637 | 34.8572 | 0.00015042
2010-07#28 | CABOS | 27 | 2010-09-18 16:07 | 24C | 0.99637 | 3.4659e-06 | 2.0266e-06 | 0.99637 | 34.8573 | 0.00013203
2010-07#46 | 2010-08 LOOP 15 |  | 2010-09-18 22:18 | 24C | 0.73328 | 4.8273e-06 | 2.0862e-06 | 0.73328 | 24.7563 | 0.00018047
2010-07#47 | 2010-08 LOOP 16 |  | 2010-09-18 22:22 | 24C | 0.74613 | 2.7425e-06 | 2.0266e-06 | 0.74614 | 25.2382 | 0.00010447
2010-07#48 | 2010-08 LOOP 17 |  | 2010-09-18 22:24 | 24C | 0.77025 | 2.2241e-06 | 1.967e-06 | 0.77026 | 26.1456 | 8.3074e-05
Calibration#3 | P151 |  | 2010-09-22 01:17 | 24C | 0.99995 | 3.5403e-06 | 1.508e-05 | 0.99997
Calibration#4 | P151 |  | 2010-09-22 01:19 | 24C | 0.99997 | 9.0741e-06 | -1.967e-06 | 0.99997
2010-07#75 | Deep Water Standard (2010-07) | CB-28b | 2010-09-22 03:45 | 24C | 0.99697 | 4.6245e-06 | -2.0266e-06 | 0.99697 | 34.8808 | 0.00018091
2010-07#78 | CB-31b | 45 | 2010-09-22 03:58 | 24C | 0.99878 | 4.1861e-06 | -1.9073e-06 | 0.99878 | 34.952 | 0.00016514
2010-07#79 | CB-31b | 46 | 2010-09-22 04:01 | 24C | 0.99807 | 3.8626e-06 | -1.967e-06 | 0.99807 | 34.9241 | 0.00015407
2010-07#80 | CB-31b | 47 | 2010-09-22 04:05 | 24C | 0.99751 | 6.4383e-06 | -2.0266e-06 | 0.99751 | 34.902 | 0.00025197
2010-07#81 | CB-31b | 48 | 2010-09-22 04:09 | 24C | 0.99711 | 4.3205e-06 | -1.9073e-06 | 0.99711 | 34.8864 | 0.00016818
EXAMPLE SHEET ONLY

## Timeline Summary
Summary of Analysis - One row per Cage
|  |  |  | Starting Standby # |  | Temperature
CASE | Date (d-mon-yr) | Station | Samples Analysed | Starting | Ending | Room | Sample |  | Comments
1 | 2011-02-02 | CB-DW | 1262-1285 | 24+5985 | 24+5984 | 22.9 | -- | 4.5955e-05
2 |  | CB-6 | 1334-1357 | 24+5984 | 24+5984 | 22.6 | -- | 4.5955e-05
3 |  | CB-21 | 1358-1381 | 24+5984 | 24+5984 | 21.4 | 22.7 | 4.5955e-05
4 | 2011-02-03 | CB-15 | 1118-1141 | 24+5984 | 24+5984 | 23.1 | 21.8 | 2.3961e-05
5 |  | CB-13 | 1046-1069 | 24+5985 | 24+5985 | 22.6 | 22.1 | 2.3961e-05
6 |  | PP-6 | 1214-1237 | 24+5985 | 24+5985 | 22.4 | -- | 2.3961e-05
7 |  | CB-17 | 1142-1165 | 24+5985 | 24+5985 | 21.7 | -- | 2.3961e-05
8 |  | CB-54 | 1550-1559 | 24+5985 |  | 21.8 | 21.3 | 2.3961e-05
|  | PP-6 | 1238-1252 |  | 24+5985 |  |  | 2.3961e-05
9 |  | PP-6 | 1253-1261 | 24+5985 | 24+5985 | 22.9 | -- | 2.3961e-05
10 | 2011-02-04 | CB-51 | 1478-1500 | 24+5985 | 24+5986 | 22 | 21.8 | 8.0466e-06
11 |  | CB-50 | 1454-1477 | 24+5986 | 24+5986 | 21.9 | 22.4 | 8.0466e-06
12 |  | CB-21 | 1406-1429 | 24+5986 | 24+5987 | 21.5 | 21.5 | 8.0466e-06
13 | 2011-02-07 | CB-40 | 1430-1453 | 24+5986 | 24+5985 | 21.3 | -- | 4.0054e-05
14 |  | CB-53 | 1526-1544 | 24+5985 | 24+5986 | 22.9 | 21.4 | 4.0054e-05
15 |  | CB-52 | 1502-1525 | 24+5986 | 24+5986 | 22.6 | 20.3 | 4.0054e-05
16 |  | CB-16 | 998-1021 | 24+5986 | 24+5985 | 22.5 | 21 | 4.0054e-05
17 |  | CB-18 | 1286-1309 | 24+5985 | 24+5985 | 20.4 | 20.9 | 4.0054e-05
18 | 2011-02-09 | CB-19 | 1310-1333 | 24+5986 | 24+5987 | 22.8 | 23.3 | 1.967e-06 | Changed Rs knob during Autosal Calibration at start of the day
19 |  | DUPLICATES -SEE WORKSHEET |  | 24+5986 | 24+5987 | 22.9 | 22.3 | 1.967e-06
20 |  | DUPLICATES -SEE WORKSHEET |  | 24+5986 | 24+5987 | 23.1 | 22.3 | 1.967e-06
21 |  | DUPLICATES -SEE WORKSHEET |  | 24+5986 | 24+5987 | 21.8 | 21.4 | 1.967e-06
22 |  | DUPLICATES -SEE WORKSHEET |  | 24+5986 | 24+5986 | 22.9 | 22.6 | 1.967e-06
23 |  | DUPLICATES -SEE WORKSHEET |  | 24+5986 | 24+5987 | 22.9 | 22.1 | 1.967e-06
24 | 2011-02-10 | CB-13 | 1022-1045 | 24+5985 | 24+5985 | 23.6 | 24.1 | 1.967e-06
25 |  | PP-7 | 1190-1213 | 24+5985 | 24+5985 | 22.8 | -- | 1.967e-06
26 |  | CB-15 | 1070-1093 | 24+5986 | 24+5988 | 23 | 21.9 | 1.967e-06
27 |  | CB-12a | 974-997 | 24+5988 | 24+5987 | 23.2 | -- | 1.967e-06
28 |  | CB-21 | 1382-1405 | 24+5988 | 24+5988 | 23.4 | 22.2 | 1.967e-06
29 | 2011-02-11 | PP-7 | 1166-1189 | 24+5987 | 24+5988 | 22.8 | 22.2 | 4.828e-06
30 |  | CB-15 | 1094-1117 | 24+5988 | 24+5989 | 22.5 | 22.2 | 4.828e-06
31 |  | LOOP | SEE WORKSHEET | 24+5988 | 24+5988 | 22.2 | 21.4 | 4.828e-06

## Cal Records
LSSL 2010-07 Salinity |  | Example data in pale yellow
IAPSO STD Water |  | Actual Data for LSSL 2010-07 Salinity in Black
|  | Standby # | Standby # |  |  |  | Uncorrected Ratio | Correction | Adjusted Ratio | IAPSO Standard Seawater Bottle Info.
Date/Time | Sample ID | Before | After | Read Value @ Start Calibration | Read Value @ End Calibration | Zero |  |  |  | K15 | Batch # | Batch Date | Salinity | Comments
2010-09-18 13:07 | Calibration#1 | 24.6099 | 24.609 |  |  | 0 | 0.99996 | 8.1062e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2010-09-18 13:08 | Calibration#2 | 24.609 | 24.609 |  |  | 0 | 0.99997 | 1.9073e-06 | 0.99997 | 0.99997 | P151 | 2012-05-20 | 34.999
2011-02-02 10:19 | Calibration#1 | 24+5985 | 24+5985 |  |  | 1e-05 | 0.99992 | 2.2397e-06 | 0.99997 | 0.99997 | P151 | 2012-05-21 | 34.999 | Rs is 5.56
2011-02-03 09:22 | Calibration#2 | 24+5984 | 24+5984 |  |  | 1e-05 | 0.99995 | 2.2391e-06 | 0.99997 | 0.99997 | P151 | 2012-05-22 | 34.999
2011-02-04 09:20 | Calibration#3 | 24+5985 | 24+5985 |  |  | 1e-05 | 0.99996 | 2.7425e-06 | 0.99997 | 0.99997 | P151 | 2012-05-23 | 34.999
2011-02-07 10:05 | Calibration#4 | 24+5986 | 24+5985 |  |  | 1e-05 | 0.99993 | 0 | 0.99997 | 0.99997 | P151 | 2012-05-24 | 34.999
2011-02-09 09:34 | Calibration#5 | 24+5986 | 24+5987 | 24+5984 | 24+5986 | 1e-05 | 0.99997 | 2.7426e-06 | 0.99997 | 0.99997 | P151 | 2012-05-25 | 34.999 | Adjusted Rs to 5.54
2011-02-10 11:03 | Calibration#6 | 24+5985 | 24+5985 |  |  | 1e-05 | 0.99995 | 2.2391e-06 | 0.99997 | 0.99997 | P151 | 2012-05-26 | 34.999
2011-02-10 11:17 | Calibration#7 | 24+5985 | 24+5985 |  |  | 1e-05 | 0.99997 | 4.4784e-06 | 0.99997 | 0.99997 | P151 | 2012-05-27 | 34.999
2011-02-11 09:25 | Calibration#8 | 24+5987 | 24+5988 |  |  | 1e-05 | 0.99997 | 0 | 0.99997 | 0.99997 | P151 | 2012-05-28 | 34.999

## DeepWaterCheck Records
LSSL 2010-07 Salinity |  | Example data in pale yellow
Deep Water Check
Cast No. | Station | DWS ID | Sample mix | DateTime | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments | Average | STDEV
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 11:43 | 0.99892 | 2.5747e-06 | -1.967e-06 | 0.99892 | 34.9576 | 0.00010077
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 14:09 | 0.99891 | 3.6232e-06 | -2.0862e-06 | 0.99891 | 34.957 | 0.00014261 | DELETED READING 1 - LOWER THAN OTHER 2
| CB-3 | Deep Water Standard (2010-07) |  | 2010-10-05 15:12 | 0.9989 | 4.7529e-06 | -2.1458e-06 | 0.99889 | 34.9565 | 0.00018681 | DELETED READING 1 - LOW
?? | CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-02 12:35 | 0.99882 | 6.5389e-06 | 4.5955e-05 | 0.99886 | 34.9553 | 0.0002569
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-02 12:37 | 0.99885 | 4.3332e-06 | 4.5955e-05 | 0.9989 | 34.9567 | 0.00017044
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-02 15:13 | 0.99894 | 2.7425e-06 | 4.6015e-05 | 0.99899 | 34.9601 | 0.00010659
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-02 16:26 | 0.99884 | 5.084e-06 | 4.5955e-05 | 0.99889 | 34.9563 | 0.00020024
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 10:31 | 0.99884 | 8.9631e-06 | 2.3961e-05 | 0.99886 | 34.9553 | 0.00035389
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 12:08 | 0.99886 | 7.2013e-06 | 2.4021e-05 | 0.99888 | 34.9559 | 0.00028346
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 13:33 | 0.99885 | 4.9731e-06 | 2.408e-05 | 0.99887 | 34.9555 | 0.00019541
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 15:50 | 0.99885 | 2.566e-06 | 2.3961e-05 | 0.99888 | 34.9559 | 0.00010155
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 16:20 | 0.99886 | 4.1135e-06 | 2.3961e-05 | 0.99888 | 34.956 | 0.00016194
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-03 17:40 | 0.99886 | 4.7303e-06 | 2.408e-05 | 0.99888 | 34.956 | 0.00018726
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-04 12:03 | 0.99914 | 6.2796e-06 | 8.0466e-06 | 0.99915 | 34.9666 | 0.0002493
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-04 13:11 | 0.99884 | 5.5281e-06 | 7.987e-06 | 0.99884 | 34.9545 | 0.00021693
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-04 13:31 | 0.99888 | 6.1457e-06 | 7.987e-06 | 0.99889 | 34.9564 | 0.00024314
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-07 11:35 | 0.99884 | 3.9806e-06 | 4.0054e-05 | 0.99888 | 34.956 | 0.00015625
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-07 12:43 | 0.99883 | 2.9035e-06 | 4.0054e-05 | 0.99887 | 34.9554 | 0.00011369
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-07 13:42 | 0.99885 | 3.218e-06 | 3.9935e-05 | 0.99889 | 34.9564 | 0.00012756
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-07 16:41 | 0.99888 | 2.3951e-06 | 4.0054e-05 | 0.99892 | 34.9576 | 9.9652e-05
| CB-3 | Deep Water Standard (2010-07) | 614-615 | 2011-02-07 17:42 | 0.9988 | 3.0364e-06 | 3.9995e-05 | 0.99884 | 34.9542 | 0.00011696
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 11:06 | 0.99888 | 6.019e-06 | 1.967e-06 | 0.99888 | 34.956 | 0.00023831
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 11:59 | 0.99887 | 1.7994e-06 | 2.0266e-06 | 0.99887 | 34.9556 | 7.2001e-05
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 12:47 | 0.9989 | 3.4657e-06 | 2.0266e-06 | 0.9989 | 34.9568 | 0.0001355
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 15:17 | 0.9989 | 5.0527e-06 | 2.0266e-06 | 0.9989 | 34.9566 | 0.00019929
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 16:36 | 0.99891 | 2.4912e-06 | 1.967e-06 | 0.99891 | 34.9572 | 9.6609e-05
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-09 17:32 | 0.99888 | 1.6614e-06 | 1.967e-06 | 0.99888 | 34.9561 | 6.5793e-05
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-10 12:16 | 0.99892 | 3.0409e-06 | 1.9073e-06 | 0.99892 | 34.9575 | 0.00012299
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-10 13:53 | 0.99897 | 2.9667e-06 | 2.0862e-06 | 0.99897 | 34.9596 | 0.00011473
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-10 15:54 | 0.99893 | 1.6611e-06 | 2.0266e-06 | 0.99894 | 34.9581 | 6.4399e-05
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-10 16:48 | 0.99887 | 4.276e-06 | 2.0862e-06 | 0.99888 | 34.9558 | 0.00016853
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-10 17:48 | 0.99887 | 3.2247e-06 | 1.9073e-06 | 0.99888 | 34.9557 | 0.00012608
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-11 10:39 | 0.99891 | 3.3298e-06 | 4.828e-06 | 0.99891 | 34.9571 | 0.00012925
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-11 11:40 | 0.99889 | 2.0943e-06 | 5.0664e-06 | 0.9989 | 34.9566 | 8.2542e-05
| CB-6 | Deep Water Standard (2010-07) | 1334-1335 | 2011-02-11 12:19 | 0.99892 | 2.8297e-06 | 4.8876e-06 | 0.99892 | 34.9576 | 0.00011337
[8 empty] | CB-3 Avg, Stdev | 34.9558 | 0.00081142
[8 empty] | CB-6 Avg, Stdev | 34.9571 | 0.0011262

## Sp
Example data in pale yellow
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded
Cast No. | Station | Bottle | Actual difference (not abs) | Sea/IOS | Sample ID | Salt-1 | Salt-2 | diff | diff^2 | Chauvenet | Accept or Reject |  | QA/QC - Analyst | QA/QC - Data |  | mean diff | sd diff | totN
[5 empty] | 6 | 34.7256 | 34.7255 | 8.0109e-05 | 6.4174e-09 | 0.4585 | ACCEPT |  |  |  | IOS/IOS | 0.0073621 | 0.015882 | 62
[5 empty] | 11 | 34.3057 | 34.3064 | 0.00060272 | 3.6327e-07 | 0.42559 | ACCEPT | SAL Sea/IOS
[5 empty] | 16 | 32.6789 | 32.6781 | 0.00083923 | 7.0431e-07 | 0.4107 | ACCEPT | N tot |  | 25 | Sea/IOS | 0.001696 | 0.0018164 | 25
[12 empty] | N used |  | 23
|  |  |  | Duplicates where Primary run at sea and duplicate run at IOS | [7 empty] | sum of sqrs |  | 4.7109e-05
|  |  | -0.00011826 | Sea/IOS | 719 | 34.83 | 34.8301 | 0.00011826 | 1.3984e-08 | 0.86864 | ACCEPT |  |  | 1.0241e-06
|  |  |  | Sea/IOS | 726 | 33.2797 | 33.2868 | 0.0071106 | 5.0561e-05 | 2.981 | REJECT | Sp |  | 0.001012
|  |  | 0.0014343 | Sea/IOS | 739 | 34.8737 | 34.8723 | 0.0014343 | 2.0573e-06 | 0.14407 | ACCEPT | N discard |  | 2
|  |  | -5.722e-05 | Sea/IOS | 743 | 34.8355 | 34.8356 | 5.722e-05 | 3.2742e-09 | 0.90224 | ACCEPT | STD for compare |  | 0.00077075
|  |  |  | Sea/IOS | 750 | 33.0322 | 33.0396 | 0.007309 | 5.3421e-05 | 3.0902 | REJECT
|  |  | 0.00040054 | Sea/IOS | 789 | 34.8549 | 34.8545 | 0.00040054 | 1.6043e-07 | 0.71323 | ACCEPT
|  |  | 0.0020676 | Sea/IOS | 795 | 34.221 | 34.2189 | 0.0020676 | 4.2748e-06 | 0.20456 | ACCEPT
|  |  | 0.0013466 | Sea/IOS | 800 | 32.5132 | 32.5118 | 0.0013466 | 1.8133e-06 | 0.19238 | ACCEPT
|  |  | -0.00042343 | Sea/IOS | 806 | 34.9572 | 34.9577 | 0.00042343 | 1.7929e-07 | 0.70063 | ACCEPT
|  |  | 0.0011597 | Sea/IOS | 810 | 34.9033 | 34.9021 | 0.0011597 | 1.3448e-06 | 0.29529 | ACCEPT
|  |  | 0.00084305 | Sea/IOS | 811 | 34.8739 | 34.8731 | 0.00084305 | 7.1073e-07 | 0.4696 | ACCEPT
|  |  | 0.001297 | Sea/IOS | 819 | 34.1531 | 34.1518 | 0.001297 | 1.6822e-06 | 0.21968 | ACCEPT | SAL ALL IOS/IOS
|  |  | 0.0013313 | Sea/IOS | 821 | 33.1256 | 33.1243 | 0.0013313 | 1.7724e-06 | 0.20078 | ACCEPT | N tot |  | 62
|  |  | 0.001442 | Sea/IOS | 830 | 34.7896 | 34.7882 | 0.001442 | 2.0792e-06 | 0.13987 | ACCEPT | N used |  | 59
|  |  | 0.0021591 | Sea/IOS | 836 | 34.2878 | 34.2856 | 0.0021591 | 4.6618e-06 | 0.25496 | ACCEPT | sum of sqrs |  | 0.0029532
|  |  | 0.0026398 | Sea/IOS | 845 | 32.5219 | 32.5192 | 0.0026398 | 6.9684e-06 | 0.51959 | ACCEPT |  |  | 2.5027e-05
|  |  | 0.0010147 | Sea/IOS | 859 | 34.8741 | 34.8731 | 0.0010147 | 1.0296e-06 | 0.3751 | ACCEPT | Sp |  | 0.0050027
|  |  | 0.00098419 | Sea/IOS | 863 | 34.8332 | 34.8322 | 0.00098419 | 9.6863e-07 | 0.3919 | ACCEPT | N discard |  | 3
|  |  | 0.0019913 | Sea/IOS | 870 | 33.173 | 33.171 | 0.0019913 | 3.9652e-06 | 0.16256 | ACCEPT | STD for compare |  | 0.0057791
|  |  | 0.00059128 | Sea/IOS | 909 | 34.8529 | 34.8523 | 0.00059128 | 3.4961e-07 | 0.60822 | ACCEPT
|  |  | 0.0012169 | Sea/IOS | 915 | 34.1035 | 34.1023 | 0.0012169 | 1.4808e-06 | 0.26378 | ACCEPT | Critical values for Chauvenet statistic
|  |  | 0.00049973 | Sea/IOS | 920 | 32.5835 | 32.583 | 0.00049973 | 2.4973e-07 | 0.65862 | ACCEPT | n | Chauvenet
|  | AVG |  | Sea/IOS | 923 |  | 30.7846 |  |  |  |  | 3 | 1.37
|  | 0.0011011 | -0.00072861 | Sea/IOS | 957 | 34.776 | 34.7767 | 0.00072861 | 5.3087e-07 | 0.53261 | ACCEPT | 4 | 1.53
|  | STD | 0.0030785 | Sea/IOS | 963 | 34.3194 | 34.3163 | 0.0030785 | 9.4769e-06 | 0.76111 | ACCEPT | 5 | 1.64
|  | 0.00093475 | 0.0011559 | Sea/IOS | 968 | 32.5814 | 32.5802 | 0.0011559 | 1.336e-06 | 0.29739 | ACCEPT | 6 | 1.73
[12 empty] | 7 | 1.8
|  |  |  | Duplicates at IOS | [7 empty] | 8 | 1.87
|  |  | -0.00078201 | IOS/IOS | 1333 | 24.8176 | 24.8184 | 0.00078201 | 6.1154e-07 | 0.41431 | ACCEPT | 9 | 1.91
|  |  | 0.00047302 | IOS/IOS | 1429 | 25.0334 | 25.0329 | 0.00047302 | 2.2375e-07 | 0.43376 | ACCEPT | 10 | 1.97
|  |  | 0.01564 | IOS/IOS | 1405 | 25.105 | 25.0894 | 0.01564 | 0.00024462 | 0.52122 | ACCEPT | 20 | 2.24
|  |  | -0.00068665 | IOS/IOS | 1285 | 25.9022 | 25.9028 | 0.00068665 | 4.7148e-07 | 0.42031 | ACCEPT | 30 | 2.39
|  |  | -0.0011349 | IOS/IOS | 1213 | 26.3955 | 26.3966 | 0.0011349 | 1.2879e-06 | 0.39209 | ACCEPT | 50 | 2.57
|  |  | 0.0080605 | IOS/IOS | 1068 | 28.6207 | 28.6127 | 0.0080605 | 6.4971e-05 | 0.04397 | ACCEPT | 80 | 2.74
|  |  | -0.010147 | IOS/IOS | 1500 | 28.7036 | 28.7137 | 0.010147 | 0.00010296 | 0.17535 | ACCEPT | 100 | 2.81
|  |  |  | IOS/IOS | 1542 | 29.6744 | 29.5804 | 0.094038 | 0.0088431 | 5.4574 | REJECT | 150 | 2.93
|  |  |  | IOS/IOS | 1557 | 29.6494 | 29.5951 | 0.054281 | 0.0029465 | 2.9542 | REJECT | 200 | 3.02
|  |  |  | IOS/IOS | 1401 | 30.5363 | 30.473 | 0.063282 | 0.0040046 | 3.5209 | REJECT
|  |  | 0.0013962 | IOS/IOS | 1162 | 31.8564 | 31.855 | 0.0013962 | 1.9493e-06 | 0.37564 | ACCEPT
|  |  | 0.0045662 | IOS/IOS | 1281 | 31.8695 | 31.8649 | 0.0045662 | 2.085e-05 | 0.17604 | ACCEPT
|  |  | -0.0035915 | IOS/IOS | 1425 | 31.9824 | 31.986 | 0.0035915 | 1.2899e-05 | 0.23741 | ACCEPT | Pooled standard deviation of pairs of samples (Sp) was calculated by:
|  |  | 0.005188 | IOS/IOS | 989 | 32.2692 | 32.264 | 0.005188 | 2.6915e-05 | 0.13689 | ACCEPT |  | Sp = SQRT{sum (d*d)/2k}
|  |  | -0.0012817 | IOS/IOS | 1085 | 32.2668 | 32.2681 | 0.0012817 | 1.6429e-06 | 0.38284 | ACCEPT |  | where k = no. of pairs and d = difference between pairs
|  |  | 0.0084839 | IOS/IOS | 1256 | 32.3657 | 32.3572 | 0.0084839 | 7.1976e-05 | 0.07063 | ACCEPT
|  |  | 0.0025291 | IOS/IOS | 1518 | 32.3911 | 32.3885 | 0.0025291 | 6.3966e-06 | 0.3043 | ACCEPT
|  |  | 0.00066376 | IOS/IOS | 1279 | 32.6191 | 32.6184 | 0.00066376 | 4.4057e-07 | 0.42175 | ACCEPT
|  |  | 0.0017471 | IOS/IOS | 1351 | 32.6294 | 32.6276 | 0.0017471 | 3.0525e-06 | 0.35354 | ACCEPT
|  |  | 0.0038948 | IOS/IOS | 1471 | 32.86 | 32.8561 | 0.0038948 | 1.517e-05 | 0.21831 | ACCEPT
|  |  | -0.0052261 | IOS/IOS | 1277 | 33.0455 | 33.0508 | 0.0052261 | 2.7312e-05 | 0.13449 | ACCEPT
|  |  | -0.0011559 | IOS/IOS | 1134 | 33.1737 | 33.1749 | 0.0011559 | 1.336e-06 | 0.39077 | ACCEPT
|  |  | 3.8147e-06 | IOS/IOS | 1014 | 33.1801 | 33.1801 | 3.8147e-06 | 1.4552e-11 | 0.4633 | ACCEPT
|  |  | -0.005806 | IOS/IOS | 1253 | 33.2066 | 33.2124 | 0.005806 | 3.3709e-05 | 0.097981 | ACCEPT
|  |  | 0.031494 | IOS/IOS | 1536 | 33.2562 | 33.2247 | 0.031494 | 0.00099188 | 1.5194 | ACCEPT
|  |  | 0.016678 | IOS/IOS | 1447 | 33.5474 | 33.5307 | 0.016678 | 0.00027815 | 0.58655 | ACCEPT
|  |  | 0.0007782 | IOS/IOS | 1300 | 33.6859 | 33.6851 | 0.0007782 | 6.0559e-07 | 0.41455 | ACCEPT
|  |  | -0.0035477 | IOS/IOS | 1076 | 33.9281 | 33.9317 | 0.0035477 | 1.2586e-05 | 0.24017 | ACCEPT
|  |  | -0.016998 | IOS/IOS | 980 | 33.9624 | 33.9794 | 0.016998 | 0.00028894 | 0.60673 | ACCEPT
|  |  | 0.00037003 | IOS/IOS | 1275 | 34.0117 | 34.0113 | 0.00037003 | 1.3692e-07 | 0.44025 | ACCEPT
|  |  | 0.0029182 | IOS/IOS | 1551 | 34.0487 | 34.0458 | 0.0029182 | 8.5161e-06 | 0.2798 | ACCEPT
|  |  | 0.0086899 | IOS/IOS | 1386 | 34.1578 | 34.1491 | 0.0086899 | 7.5514e-05 | 0.0836 | ACCEPT
|  |  | -0.003231 | IOS/IOS | 1204 | 34.1788 | 34.182 | 0.003231 | 1.044e-05 | 0.26011 | ACCEPT
|  |  | 0.0036469 | IOS/IOS | 1467 | 34.455 | 34.4513 | 0.0036469 | 1.33e-05 | 0.23393 | ACCEPT
|  |  | 0.0017128 | IOS/IOS | 1489 | 34.6407 | 34.639 | 0.0017128 | 2.9337e-06 | 0.3557 | ACCEPT
|  |  | -0.00013351 | IOS/IOS | 974 | 34.7417 | 34.7418 | 0.00013351 | 1.7826e-08 | 0.45514 | ACCEPT
|  |  | 0.0023613 | IOS/IOS | 1512 | 34.7567 | 34.7544 | 0.0023613 | 5.5757e-06 | 0.31487 | ACCEPT
|  |  | 0.00064087 | IOS/IOS | 1070 | 34.7745 | 34.7739 | 0.00064087 | 4.1071e-07 | 0.42319 | ACCEPT
|  |  | 0.0011787 | IOS/IOS | 1129 | 34.7768 | 34.7756 | 0.0011787 | 1.3894e-06 | 0.38933 | ACCEPT
|  |  | 0.0061569 | IOS/IOS | 1528 | 34.79 | 34.7838 | 0.0061569 | 3.7908e-05 | 0.075883 | ACCEPT
|  |  | 0.0039291 | IOS/IOS | 1440 | 34.7966 | 34.7926 | 0.0039291 | 1.5438e-05 | 0.21615 | ACCEPT
|  |  | 0.0007019 | IOS/IOS | 1319 | 34.8326 | 34.8319 | 0.0007019 | 4.9267e-07 | 0.41935 | ACCEPT
|  |  | 0.0014343 | IOS/IOS | 1127 | 34.8337 | 34.8323 | 0.0014343 | 2.0573e-06 | 0.37323 | ACCEPT
|  |  | 0.0011597 | IOS/IOS | 1007 | 34.8342 | 34.833 | 0.0011597 | 1.3448e-06 | 0.39053 | ACCEPT
|  |  | 0.00050735 | IOS/IOS | 1293 | 34.8335 | 34.833 | 0.00050735 | 2.5741e-07 | 0.4316 | ACCEPT
|  |  | 0.0025291 | IOS/IOS | 1506 | 34.856 | 34.8535 | 0.0025291 | 6.3966e-06 | 0.3043 | ACCEPT
|  |  | -0.021126 | IOS/IOS | 1055 | 34.8385 | 34.8596 | 0.021126 | 0.0004463 | 0.86661 | ACCEPT
|  |  | 0.0054207 | IOS/IOS | 1148 | 34.8719 | 34.8665 | 0.0054207 | 2.9384e-05 | 0.12224 | ACCEPT
|  |  | -0.0013008 | IOS/IOS | 1196 | 34.8689 | 34.8702 | 0.0013008 | 1.6921e-06 | 0.38164 | ACCEPT
|  |  | 0.0010986 | IOS/IOS | 1289 | 34.8811 | 34.88 | 0.0010986 | 1.207e-06 | 0.39437 | ACCEPT
|  |  | 0.0014534 | IOS/IOS | 1003 | 34.8836 | 34.8821 | 0.0014534 | 2.1124e-06 | 0.37203 | ACCEPT
|  |  | 0.0015564 | IOS/IOS | 1458 | 34.8862 | 34.8846 | 0.0015564 | 2.4224e-06 | 0.36555 | ACCEPT
|  |  | 0.00053406 | IOS/IOS | 1265 | 34.8858 | 34.8853 | 0.00053406 | 2.8522e-07 | 0.42992 | ACCEPT
|  |  | 0.00025177 | IOS/IOS | 1410 | 34.9157 | 34.9155 | 0.00025177 | 6.3388e-08 | 0.44769 | ACCEPT
|  |  | 0.0051727 | IOS/IOS | 1433 | 34.9209 | 34.9157 | 0.0051727 | 2.6757e-05 | 0.13785 | ACCEPT
|  |  | 0.00041962 | IOS/IOS | 1264 | 34.9186 | 34.9182 | 0.00041962 | 1.7608e-07 | 0.43712 | ACCEPT
|  |  | -0.00028229 | IOS/IOS | 1240 | 34.9419 | 34.9422 | 0.00028229 | 7.9686e-08 | 0.44577 | ACCEPT
|  |  | 0.0027733 | IOS/IOS | 1478 | 34.9525 | 34.9497 | 0.0027733 | 7.6911e-06 | 0.28893 | ACCEPT
|  | AVG | -0.0057755 | IOS/IOS | 1312 | 34.9486 | 34.9544 | 0.0057755 | 3.3356e-05 | 0.099902 | ACCEPT
|  | 0.0013527 | 0.0021858 | IOS/IOS | 1334 | 34.9571 | 34.9549 | 0.0021858 | 4.7778e-06 | 0.32592 | ACCEPT
|  | STD | 0.0019302 | IOS/IOS | 1047 | 34.9575 | 34.9556 | 0.0019302 | 3.7258e-06 | 0.34201 | ACCEPT
|  | 0.007004 | -0.00031281 | IOS/IOS | 1143 | 34.956 | 34.9563 | 0.00031281 | 9.7847e-08 | 0.44385 | ACCEPT

## Bottom-Water Plots
Cruise | Cast No. | Station Name | Sample No. [All others match to this sample number] | Trip [US (up stop), UN (up no stop), USM (up stop mix) or DN (down no stop)] | CTDPres [dbar] | CTDSalt-2 [PSS-78] | Salt-1 | Salt-2 | Analyst Comment: Sea or IOS | Sort Reference | Salt-1 | Chauvenet | Accept or Reject |  | mean | std | N
2010-07 | 67 | CB-40 | 1430 | UN | 3058.127 | 34.9549 | 34.9575 |  | IOS | 1430 | 34.9575 | 0.6598 | ACCEPT | All | 34.9562 | 0.0018968 | 41
2010-07 | 66 | CB-21 | 1407 | UN | 3058.158 | 34.955 | 34.9558 |  | IOS | 1407 | 34.9558 | 0.24119 | ACCEPT | Remove outlier | 34.9564 | 0.0012233 | 40
2010-07 | 66 | CB-21 | 1406 | UN | 3396.672 | 34.955 | 34.9563 |  | IOS | 1406 | 34.9563 | 0.050425 | ACCEPT
2010-07 | 63 | CB-6 | 1335 | UN | 3057.893 | 34.9554 | 34.9568 |  | IOS | 1335 | 34.9568 | 0.28975 | ACCEPT
2010-07 | 63 | CB-6 | 1334 | UN | 3622.441 | 34.9553 | 34.9571 | 34.9549 | IOS/IOS | 1334 | 34.9571 | 0.4788 | ACCEPT |  | Critical values for Chauvenet statistic
2010-07 | 62 | CB-19 | 1311 | UN | 3058.296 | 34.9552 | 34.9528 |  | IOS | 1311 | 34.9528 | 1.7817 | ACCEPT |  | n | Chauvenet
2010-07 | 62 | CB-19 | 1310 | UN | 3571.218 | 34.9551 | 34.954 |  | IOS | 1310 | 34.954 | 1.1583 | ACCEPT |  | 3 | 1.37
2010-07 | 57 | PP-7 | 1191 | UN | 3057.93 | 34.9556 | 34.9557 |  | IOS | 1191 | 34.9557 | 0.29147 | ACCEPT |  | 4 | 1.53
2010-07 | 57 | PP-7 | 1190 | UN | 3434.121 | 34.9557 | 34.9552 |  | IOS | 1190 | 34.9552 | 0.51872 | ACCEPT |  | 5 | 1.64
2010-07 | 55 | CB-17 | 1143 | UN | 3058.33 | 34.9559 | 34.956 | 34.9563 | IOS/IOS | 1143 | 34.956 | 0.11449 | ACCEPT |  | 6 | 1.73
2010-07 | 55 | CB-17 | 1142 | UN | 3572.255 | 34.9558 | 34.9561 |  | IOS | 1142 | 34.9561 | 0.062198 | ACCEPT |  | 7 | 1.8
2010-07 | 54 | CB-15 | 1119 | UN | 3057.768 | 34.9559 | 34.9471 |  | IOS | 1119 | 34.9471 | 4.8165 | REJECT |  | 8 | 1.87
2010-07 | 54 | CB-15 | 1118 | UN | 3532.327 | 34.956 | 34.9536 |  | IOS | 1118 | 34.9536 | 1.3815 | ACCEPT |  | 9 | 1.91
2010-07 | 51 | CB-13 | 1047 | UN | 3058.473 | 34.9561 | 34.9575 | 34.9556 | IOS/IOS | 1047 | 34.9575 | 0.67387 | ACCEPT |  | 10 | 1.97
2010-07 | 51 | CB-13 | 1046 | UN | 3641.368 | 34.956 | 34.9579 |  | IOS | 1046 | 34.9579 | 0.87097 | ACCEPT |  | 20 | 2.24
2010-07 | 49 | CB-16 | 999 | UN | 3059.213 | 34.9562 | 34.9568 |  | IOS | 999 | 34.9568 | 0.31992 | ACCEPT |  | 30 | 2.39
2010-07 | 49 | CB-16 | 998 | UN | 3623.125 | 34.956 | 34.955 |  | IOS | 998 | 34.955 | 0.63336 | ACCEPT |  | 50 | 2.57
2010-07 | 47 | CB-12 | 951 | UN | 3058.69 | 34.9563 | 34.9575 |  | Sea | 951 | 34.9575 | 0.68594 | ACCEPT |  | 80 | 2.74
2010-07 | 47 | CB-12 | 950 | UN | 3695.499 | 34.9561 | 34.9574 |  | Sea | 950 | 34.9574 | 0.6417 | ACCEPT |  | 100 | 2.81
2010-07 | 44 | CB-9 | 903 | UN | 3058.698 | 34.9562 | 34.957 |  | Sea | 903 | 34.957 | 0.42651 | ACCEPT |  | 150 | 2.93
2010-07 | 44 | CB-9 | 902 | UN | 3688.705 | 34.9562 | 34.9574 |  | Sea | 902 | 34.9574 | 0.63566 | ACCEPT |  | 200 | 3.02
2010-07 | 42 | CB-11 | 855 | UN | 3058.158 | 34.9562 | 34.9571 |  | Sea | 855 | 34.9571 | 0.45667 | ACCEPT
2010-07 | 42 | CB-11 | 854 | UN | 3699.673 | 34.9562 | 34.9568 |  | Sea | 854 | 34.9568 | 0.31388 | ACCEPT
2010-07 | 40 | CB-70 | 807 | UN | 3053.682 | 34.9564 | 34.9575 |  | Sea | 807 | 34.9575 | 0.66382 | ACCEPT
2010-07 | 40 | CB-70 | 806 | USM | 3690.568 | 34.9563 | 34.9572 | 34.9577 | Sea/IOS | 806 | 34.9572 | 0.54315 | ACCEPT
2010-07 | 37 | CB-8 | 735 | UN | 3058.527 | 34.9559 | 34.957 |  | Sea | 735 | 34.957 | 0.43656 | ACCEPT
2010-07 | 37 | CB-8 | 734 | UN | 3696.793 | 34.956 | 34.9555 |  | Sea | 734 | 34.9555 | 0.36387 | ACCEPT
2010-07 | 36 | CB-7 | 711 | UN | 3058.741 | 34.9562 | 34.9573 |  | Sea | 711 | 34.9573 | 0.57734 | ACCEPT
2010-07 | 36 | CB-7 | 710 | UN | 3674.398 | 34.9561 | 34.9575 |  | Sea | 710 | 34.9575 | 0.67186 | ACCEPT
2010-07 | 35 | CB-5 | 687 | UN | 3057.988 | 34.9561 | 34.9582 |  | Sea | 687 | 34.9582 | 1.0459 | ACCEPT
2010-07 | 35 | CB-5 | 686 | UN | 3482.699 | 34.956 | 34.9556 |  | Sea | 686 | 34.9556 | 0.3156 | ACCEPT
2010-07 | 34 | CB-4 | 663 | UN | 3059.316 | 34.956 | 34.9559 |  | Sea | 663 | 34.9559 | 0.17884 | ACCEPT
2010-07 | 34 | CB-4 | 662 | UN | 3673.074 | 34.956 | 34.9564 |  | Sea | 662 | 34.9564 | 0.10271 | ACCEPT
2010-07 | 32 | CB-3 | 615 | UN | 3058.964 | 34.9561 | 34.9577 |  | Sea | 615 | 34.9577 | 0.7865 | ACCEPT
2010-07 | 32 | CB-3 | 614 | UN | 3674.735 | 34.9559 | 34.9573 |  | Sea | 614 | 34.9573 | 0.5693 | ACCEPT
2010-07 | 31 | CB-2 | 591 | UN | 3057.918 | 34.9557 | 34.9561 |  | Sea | 591 | 34.9561 | 0.04812 | ACCEPT
2010-07 | 31 | CB-2 | 590 | UN | 3570.45 | 34.956 | 34.9554 |  | Sea | 590 | 34.9554 | 0.44833 | ACCEPT
2010-07 | 30 | CB-2a | 567 | UN | 3057.648 | 34.9555 | 34.9566 |  | Sea | 567 | 34.9566 | 0.19925 | ACCEPT
2010-07 | 30 | CB-2a | 566 | UN | 3591.904 | 34.9561 | 34.9545 |  | Sea | 566 | 34.9545 | 0.88274 | ACCEPT
2010-07 | 29 | Sta-A | 543 | UN | 3057.724 | 34.9561 | 34.9575 |  | Sea | 543 | 34.9575 | 0.65376 | ACCEPT
2010-07 | 29 | Sta-A | 542 | UN | 3313.695 | 34.9561 | 34.9571 |  | Sea | 542 | 34.9571 | 0.48282 | ACCEPT
Note on J1: DFO Employee: Differences >0.05 are highlighted in red.

## Duplicate plots
Salinometer ID | Station | Sample ID (bottle label) | Duplicates | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments | All Diffs | Diff if less than abs(0.01) | Avg | STD
Sample ID | Station | Bottle Label |  | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
LSSL2010#3 | CB-DW | 1264 |  | 2011-02-02 11:35 | 24C | 0.99788 | 6.0358e-06 | 4.6015e-05 | 0.99793 | 34.9186 | 0.00023836 |  | 2011-02-09 15:46 | 24C | 0.99792 | 1.3882e-06 | 2.0862e-06 | 0.99792 | 34.9182 | 5.2278e-05 | 0 | 0.00041962 | 0.00041962
LSSL2010#4 | CB-DW | 1265 |  | 2011-02-02 11:38 | 24C | 0.99705 | 3.4659e-06 | 4.6074e-05 | 0.9971 | 34.8858 | 0.00013799 |  | 2011-02-09 15:48 | 24C | 0.99708 | 0 | 2.0862e-06 | 0.99708 | 34.8853 | 0 | 0 | 0.00053406 | 0.00053406
LSSL2010#14 | CB-DW | 1275 |  | 2011-02-02 12:05 | 24C | 0.97479 | 4.9007e-06 | 4.6074e-05 | 0.97483 | 34.0117 | 0.00019277 |  | 2011-02-09 16:03 | 24C | 0.97482 | 2.4911e-06 | 2.0266e-06 | 0.97482 | 34.0113 | 9.869e-05 | 0 | 0.00037003 | 0.00037003
LSSL2010#16 | CB-DW | 1277 |  | 2011-02-02 12:10 | 24C | 0.95007 | 4.8959e-06 | 4.5955e-05 | 0.95012 | 33.0455 | 0.0001943 |  | 2011-02-09 16:09 | 24C | 0.95025 | 2.866e-06 | 2.0862e-06 | 0.95025 | 33.0508 | 0.00011317 | 0 | -0.0052261 | -0.0052261
LSSL2010#18 | CB-DW | 1279 |  | 2011-02-02 12:15 | 24C | 0.93913 | 3.4658e-06 | 4.6015e-05 | 0.93918 | 32.6191 | 0.00013386 |  | 2011-02-09 16:14 | 24C | 0.93916 | 1.1203e-06 | 2.0862e-06 | 0.93916 | 32.6184 | 4.4364e-05 | 0 | 0.00066376 | 0.00066376
LSSL2010#20 | CB-DW | 1281 |  | 2011-02-02 12:21 | 24C | 0.91984 | 6.1978e-06 | 4.5955e-05 | 0.91989 | 31.8695 | 0.00024055 |  | 2011-02-09 16:19 | 24C | 0.91977 | 1.1196e-06 | 2.0266e-06 | 0.91977 | 31.8649 | 4.3729e-05 | 0 | 0.0045662 | 0.0045662
LSSL2010#24 | CB-DW | 1285 |  | 2011-02-02 12:32 | 24C | 0.76375 | 2.7301e-06 | 4.5896e-05 | 0.7638 | 25.9022 | 0.00010253 |  | 2011-02-09 16:26 | 24C | 0.76381 | 0 | 2.0266e-06 | 0.76381 | 25.9028 | 1.0662e-06 | 0 | -0.00068665 | -0.00068665
LSSL2010#27 | CB-6 | 1334 |  | 2011-02-02 14:03 | 24C | 0.99886 | 3.264e-06 | 4.5955e-05 | 0.99891 | 34.9571 | 0.00012912 |  | 2011-02-09 15:39 | 24C | 0.99885 | 2.4915e-06 | 1.967e-06 | 0.99885 | 34.9549 | 0.00010059 | 0 | 0.0021858 | 0.0021858
LSSL2010#44 | CB-6 | 1351 |  | 2011-02-02 14:52 | 24C | 0.9394 | 6.2131e-06 | 4.5896e-05 | 0.93944 | 32.6294 | 0.00024166 |  | 2011-02-09 16:12 | 24C | 0.93939 | 3.1418e-06 | 1.9073e-06 | 0.9394 | 32.6276 | 0.00012486 | 0 | 0.0017471 | 0.0017471
LSSL2010#83 | CB-15 | 1127 |  | 2011-02-03 09:54 | 24C | 0.99575 | 2.8901e-06 | 2.3901e-05 | 0.99577 | 34.8337 | 0.00011606 |  | 2011-02-09 12:22 | 24C | 0.99574 | 1.3713e-06 | 1.967e-06 | 0.99574 | 34.8323 | 5.4326e-05 | 0 | 0.0014343 | 0.0014343
LSSL2010#85 | CB-15 | 1129 |  | 2011-02-03 09:58 | 24C | 0.9943 | 1.3713e-06 | 2.3961e-05 | 0.99433 | 34.7768 | 5.2235e-05 |  | 2011-02-09 12:24 | 24C | 0.99429 | 2.4909e-06 | 2.0266e-06 | 0.9943 | 34.7756 | 0.00010437 | 0 | 0.0011787 | 0.0011787
LSSL2010#90 | CB-15 | 1134 |  | 2011-02-03 10:10 | 24C | 0.95338 | 2.2392e-06 | 2.3961e-05 | 0.95341 | 33.1737 | 8.7015e-05 |  | 2011-02-09 11:44 | 24C | 0.95343 | 1.7026e-05 | 1.9073e-06 | 0.95343 | 33.1749 | 0.0006637 | 0 | -0.0011559 | -0.0011559
LSSL2010#100 | CB-13 | 1047 |  | 2011-02-03 11:14 | 24C | 0.9989 | 5.3052e-06 | 2.4021e-05 | 0.99892 | 34.9575 | 0.00020517 |  | 2011-02-09 16:49 | 24C | 0.99887 | 2.8543e-06 | 1.967e-06 | 0.99887 | 34.9556 | 0.00011192 | 0 | 0.0019302 | 0.0019302
LSSL2010#108 | CB-13 | 1055 |  | 2011-02-03 11:32 | 24C | 0.99587 | 6.3301e-06 | 2.3961e-05 | 0.9959 | 34.8385 | 0.00024835 |  | 2011-02-09 17:02 | 24C | 0.99643 | 2.4908e-06 | 2.0266e-06 | 0.99643 | 34.8596 | 9.9543e-05 | 0 | -0.021126
LSSL2010#121 | CB-13 | 1068 |  | 2011-02-03 12:03 | 24C | 0.83546 | 8.0422e-06 | 2.3961e-05 | 0.83548 | 28.6207 | 0.00031011 |  | 2011-02-09 17:30 | 24C | 0.83527 | 3.6104e-06 | 2.0266e-06 | 0.83527 | 28.6127 | 0.00013717 | 0 | 0.0080605 | 0.0080605
LSSL2010#150 | CB-17 | 1143 |  | 2011-02-03 13:47 | 24C | 0.99886 | 4.2256e-06 | 2.4021e-05 | 0.99888 | 34.956 | 0.00016938 |  | 2011-02-09 12:12 | 24C | 0.99889 | 2.8792e-06 | 2.0266e-06 | 0.99889 | 34.9563 | 0.00011317 | 0 | -0.00031281 | -0.00031281
LSSL2010#155 | CB-17 | 1148 |  | 2011-02-03 13:59 | 24C | 0.99672 | 5.1679e-06 | 2.408e-05 | 0.99674 | 34.8719 | 0.000204 |  | 2011-02-09 12:17 | 24C | 0.9966 | 2.4908e-06 | 2.0266e-06 | 0.99661 | 34.8665 | 9.6642e-05 | SZ- Correct there was no sample 1149 taken | 0.0054207 | 0.0054207
LSSL2010#168 | CB-17 | 1162 |  | 2011-02-03 15:41 | 24C | 0.91953 | 2.7757e-06 | 2.3961e-05 | 0.91955 | 31.8564 | 0.00010815 |  | 2011-02-09 12:41 | 24C | 0.91951 | 3.4666e-06 | 1.967e-06 | 0.91951 | 31.855 | 0.00013693 | 0 | 0.0013962 | 0.0013962
LSSL2010#174 | CB-54 | 1551 |  | 2011-02-03 15:59 | 24C | 0.97575 | 1.7599e-06 | 2.4021e-05 | 0.97578 | 34.0487 | 6.8803e-05 |  | 2011-02-09 11:40 | 24C | 0.9757 | 1.7597e-06 | 2.0266e-06 | 0.9757 | 34.0458 | 6.7435e-05 | 0 | 0.0029182 | 0.0029182
LSSL2010#180 | CB-54 | 1557 |  | 2011-02-03 16:13 | 24C | 0.86233 | 3.6104e-06 | 2.3961e-05 | 0.86235 | 29.6494 | 0.00013702 |  | 2011-02-09 11:53 | 24C | 0.86093 | 2.2126e-06 | 2.1458e-06 | 0.86093 | 29.5951 | 8.8712e-05 | 0 | 0.054281
LSSL2010#186 | PP-6 | 1240 |  | 2011-02-03 16:29 | 24C | 0.9985 | 4.7172e-06 | 2.408e-05 | 0.99852 | 34.9419 | 0.00018654 |  | 2011-02-09 15:44 | 24C | 0.99853 | 2.2126e-06 | 2.0862e-06 | 0.99853 | 34.9422 | 8.8728e-05 | 0 | -0.00028229 | -0.00028229
LSSL2010#199 | PP-6 | 1253 |  | 2011-02-03 17:16 | 24C | 0.95422 | 7.3696e-06 | 2.3961e-05 | 0.95425 | 33.2066 | 0.00029013 |  | 2011-02-09 16:07 | 24C | 0.95439 | 2.4909e-06 | 2.0266e-06 | 0.9544 | 33.2124 | 9.4981e-05 | 0 | -0.005806 | -0.005806
LSSL2010#202 | PP-6 | 1256 |  | 2011-02-03 17:24 | 24C | 0.93264 | 3.6105e-06 | 2.3961e-05 | 0.93266 | 32.3657 | 0.00013946 |  | 2011-02-09 16:16 | 24C | 0.93244 | 3.1417e-06 | 2.0266e-06 | 0.93245 | 32.3572 | 0.00012313 | 0 | 0.0084839 | 0.0084839
LSSL2010#209 | CB-51 | 1478 |  | 2011-02-04 09:28 | 24C | 0.99878 | 4.6354e-06 | 8.1062e-06 | 0.99879 | 34.9525 | 0.0001827 |  | 2011-02-09 11:14 | 24C | 0.99872 | 3.0085e-06 | 2.0266e-06 | 0.99872 | 34.9497 | 0.00011691 | 0 | 0.0027733 | 0.0027733
LSSL2010#218 | CB-51 | 1489 |  | 2011-02-04 10:32 | 24C | 0.99086 | 8.5641e-06 | 7.987e-06 | 0.99086 | 34.6407 | 0.00033732 |  | 2011-02-09 11:35 | 24C | 0.99082 | 3.1418e-06 | 1.967e-06 | 0.99082 | 34.639 | 0.00012312 | 0 | 0.0017128 | 0.0017128
LSSL2010#229 | CB-51 | 1500 |  | 2011-02-04 10:57 | 24C | 0.83764 | 4.3337e-06 | 7.987e-06 | 0.83765 | 28.7036 | 0.00016464 |  | 2011-02-09 11:56 | 24C | 0.83791 | 3.6105e-06 | 1.967e-06 | 0.83791 | 28.7137 | 0.00013841 | 0 | -0.010147
LSSL2010#234 | CB-50 | 1458 |  | 2011-02-04 11:14 | 24C | 0.9971 | 4.3839e-06 | 7.987e-06 | 0.99711 | 34.8862 | 0.00017276 |  | 2011-02-09 11:22 | 24C | 0.99707 | 2.4909e-06 | 2.0266e-06 | 0.99707 | 34.8846 | 9.4938e-05 | 0 | 0.0015564 | 0.0015564
LSSL2010#243 | CB-50 | 1467 |  | 2011-02-04 11:37 | 24C | 0.98613 | 6.3303e-06 | 7.987e-06 | 0.98614 | 34.455 | 0.0002493 |  | 2011-02-09 11:37 | 24C | 0.98604 | 2.2396e-06 | 1.9073e-06 | 0.98604 | 34.4513 | 8.8203e-05 | 0 | 0.0036469 | 0.0036469
LSSL2010#247 | CB-50 | 1471 |  | 2011-02-04 11:46 | 24C | 0.94535 | 6.4249e-06 | 7.987e-06 | 0.94536 | 32.86 | 0.00025058 |  | 2011-02-09 11:49 | 24C | 0.94526 | 2.8901e-06 | 2.0862e-06 | 0.94526 | 32.8561 | 0.00011213 | 0 | 0.0038948 | 0.0038948
LSSL2010#259 | CB-21 | 1410 |  | 2011-02-04 12:24 | 24C | 0.99785 | 7.7497e-06 | 7.8678e-06 | 0.99786 | 34.9157 | 0.00030526 |  | 2011-02-09 12:15 | 24C | 0.99785 | 3.4658e-06 | 1.967e-06 | 0.99785 | 34.9155 | 0.00013387 | 0 | 0.00025177 | 0.00025177
LSSL2010#274 | CB-21 | 1425 |  | 2011-02-04 12:59 | 24C | 0.92279 | 5.2332e-06 | 8.1062e-06 | 0.9228 | 31.9824 | 0.00020115 |  | 2011-02-09 12:38 | 24C | 0.92289 | 4.9162e-06 | 2.0266e-06 | 0.92289 | 31.986 | 0.00018936 | 0 | -0.0035915 | -0.0035915
LSSL2010#278 | CB-21 | 1429 |  | 2011-02-04 13:09 | 24C | 0.74067 | 1.3712e-06 | 7.987e-06 | 0.74068 | 25.0334 | 5.1717e-05 |  | 2011-02-09 12:43 | 24C | 0.74066 | 2.2391e-06 | 1.967e-06 | 0.74066 | 25.0329 | 8.872e-05 | 0 | 0.00047302 | 0.00047302
LSSL2010#284 | CB-40 | 1433 |  | 2011-02-07 10:46 | 24C | 0.99795 | 3.865e-06 | 4.0054e-05 | 0.99799 | 34.9209 | 0.00015448 |  | 2011-02-09 11:18 | 24C | 0.99786 | 4.2856e-06 | 1.967e-06 | 0.99786 | 34.9157 | 0.00016832 | 0 | 0.0051727 | 0.0051727
LSSL2010#291 | CB-40 | 1440 |  | 2011-02-07 11:00 | 24C | 0.99479 | 3.2141e-06 | 3.9995e-05 | 0.99483 | 34.7966 | 0.00012382 |  | 2011-02-09 11:27 | 24C | 0.99473 | 2.4908e-06 | 2.0266e-06 | 0.99473 | 34.7926 | 0.00010059 | 0 | 0.0039291 | 0.0039291
LSSL2010#298 | CB-40 | 1447 |  | 2011-02-07 11:16 | 24C | 0.96293 | 3.5949e-06 | 3.9995e-05 | 0.96297 | 33.5474 | 0.00014165 | Really?? | 2011-02-09 12:33 | 24C | 0.96254 | 1.1196e-06 | 2.0266e-06 | 0.96254 | 33.5307 | 4.265e-05 | 0 | 0.016678
LSSL2010#308 | CB-55 | 1528 |  | 2011-02-07 12:04 | 24C | 0.99462 | 2.4615e-06 | 3.9995e-05 | 0.99466 | 34.79 | 9.869e-05 |  | 2011-02-09 11:30 | 24C | 0.9945 | 3.3588e-06 | 2.0266e-06 | 0.9945 | 34.7838 | 0.00013341 | 0 | 0.0061569 | 0.0061569
LSSL2010#315 | CB-62 | 1536 |  | 2011-02-07 12:19 | 24C | 0.95548 | 2.6733e-06 | 4.0054e-05 | 0.95552 | 33.2562 | 0.00010274 |  | 2011-02-09 11:46 | 24C | 0.95471 | 2.4911e-06 | 1.967e-06 | 0.95471 | 33.2247 | 9.4938e-05 | 0 | 0.031494
LSSL2010#321 | CB-68 | 1542 |  | 2011-02-07 12:31 | 24C | 0.86296 | 2.2401e-06 | 4.0054e-05 | 0.863 | 29.6744 | 8.7006e-05 |  | 2011-02-09 11:51 | 24C | 0.86055 | 2.2394e-06 | 1.967e-06 | 0.86055 | 29.5804 | 8.4887e-05 | 0 | 0.094038
LSSL2010#329 | CB-52 | 1506 |  | 2011-02-07 12:59 | 24C | 0.9963 | 3.8748e-06 | 4.0054e-05 | 0.99634 | 34.856 | 0.00015161 |  | 2011-02-09 11:25 | 24C | 0.99627 | 1.1529e-06 | 1.9073e-06 | 0.99628 | 34.8535 | 4.5227e-05 | 0 | 0.0025291 | 0.0025291
LSSL2010#335 | CB-52 | 1512 |  | 2011-02-07 13:11 | 24C | 0.99378 | 6.4718e-06 | 4.0054e-05 | 0.99382 | 34.7567 | 0.00025338 |  | 2011-02-09 11:33 | 24C | 0.99375 | 2.4915e-06 | 2.0266e-06 | 0.99376 | 34.7544 | 9.4885e-05 | 0 | 0.0023613 | 0.0023613
LSSL2010#341 | CB-52 | 1518 |  | 2011-02-07 13:24 | 24C | 0.93328 | 3.4662e-06 | 4.0054e-05 | 0.93332 | 32.3911 | 0.00013479 |  | 2011-02-09 11:42 | 24C | 0.93325 | 1.4046e-06 | 1.9073e-06 | 0.93325 | 32.3885 | 5.2235e-05 | 0 | 0.0025291 | 0.0025291
LSSL2010#355 | CB-16 | 1003 |  | 2011-02-07 15:50 | 24C | 0.997 | 3.4657e-06 | 4.0054e-05 | 0.99704 | 34.8836 | 0.00013662 |  | 2011-02-09 16:57 | 24C | 0.997 | 2.8907e-06 | 2.1458e-06 | 0.997 | 34.8821 | 0.00011519 | 0 | 0.0014534 | 0.0014534
LSSL2010#359 | CB-16 | 1007 |  | 2011-02-07 16:00 | 24C | 0.99574 | 3.6104e-06 | 4.0054e-05 | 0.99578 | 34.8342 | 0.00013753 |  | 2011-02-09 17:04 | 24C | 0.99575 | 1.3712e-06 | 1.9073e-06 | 0.99576 | 34.833 | 5.5371e-05 | 0 | 0.0011597 | 0.0011597
LSSL2010#366 | CB-16 | 1014 |  | 2011-02-07 16:14 | 24C | 0.95353 | 2.2725e-06 | 4.0054e-05 | 0.95357 | 33.1801 | 8.7556e-05 |  | 2011-02-09 17:18 | 24C | 0.95357 | 3.6119e-06 | 2.1458e-06 | 0.95357 | 33.1801 | 0.00013908 | 0 | 3.8147e-06 | 3.8147e-06
LSSL2010#378 | CB-18 | 1289 |  | 2011-02-07 16:59 | 24C | 0.99694 | 3.9984e-06 | 3.9995e-05 | 0.99698 | 34.8811 | 0.00015827 |  | 2011-02-09 15:50 | 24C | 0.99695 | 5.2144e-06 | 2.0266e-06 | 0.99695 | 34.88 | 0.00020293 | 0 | 0.0010986 | 0.0010986
LSSL2010#382 | CB-18 | 1293 |  | 2011-02-07 17:08 | 24C | 0.99573 | 5.1203e-06 | 3.9935e-05 | 0.99577 | 34.8335 | 0.00019874 |  | 2011-02-09 15:53 | 24C | 0.99575 | 2.4908e-06 | 2.0266e-06 | 0.99576 | 34.833 | 9.9734e-05 | 0 | 0.00050735 | 0.00050735
LSSL2010#389 | CB-18 | 1300 |  | 2011-02-07 17:22 | 24C | 0.96647 | 3.4657e-06 | 3.9935e-05 | 0.96651 | 33.6859 | 0.00013596 |  | 2011-02-09 16:05 | 24C | 0.96649 | 3.1905e-06 | 2.0266e-06 | 0.96649 | 33.6851 | 0.00012553 | 0 | 0.0007782 | 0.0007782 | 0.0016155 | 0.0028653
LSSL2010#402 | CB-19 | 1312 |  | 2011-02-09 09:47 | 24C | 0.99869 | 2.2401e-06 | 1.967e-06 | 0.99869 | 34.9486 | 8.7077e-05 |  | 2011-02-09 15:42 | 24C | 0.99884 | 2.0804e-06 | 2.0862e-06 | 0.99884 | 34.9544 | 8.1621e-05 | 0 | -0.0057755 | -0.0057755
LSSL2010#409 | CB-19 | 1319 |  | 2011-02-09 10:03 | 24C | 0.99574 | 3.2141e-06 | 2.0266e-06 | 0.99575 | 34.8326 | 0.00012429 |  | 2011-02-09 15:55 | 24C | 0.99573 | 2.7264e-06 | 2.0266e-06 | 0.99573 | 34.8319 | 0.00010344 | 0 | 0.0007019 | 0.0007019
LSSL2010#423 | CB-19 | 1333 |  | 2011-02-09 10:59 | 24C | 0.73492 | 1.3551e-06 | 1.967e-06 | 0.73492 | 24.8176 | 5.2235e-05 |  | 2011-02-09 16:31 | 24C | 0.73494 | 3.3588e-06 | 1.967e-06 | 0.73494 | 24.8184 | 0.0001271 | 0 | -0.00078201 | -0.00078201
LSSL2010#549 | PP-7 | 1196 |  | 2011-02-10 13:11 | 24C | 0.99667 | 3.1414e-06 | 1.967e-06 | 0.99667 | 34.8689 | 0.00012312 |  | 2011-02-09 12:20 | 24C | 0.9967 | 4.9165e-06 | 2.0266e-06 | 0.9967 | 34.8702 | 0.00019397 | 0 | -0.0013008 | -0.0013008
LSSL2010#557 | PP-7 | 1204 |  | 2011-02-10 13:27 | 24C | 0.9791 | 3.3594e-06 | 2.0266e-06 | 0.9791 | 34.1788 | 0.00013117 |  | 2011-02-09 15:58 | 24C | 0.97918 | 2.7425e-06 | 2.0266e-06 | 0.97918 | 34.182 | 0.00010865 | 0 | -0.003231 | -0.003231
LSSL2010#566 | PP-7 | 1213 |  | 2011-02-10 13:50 | 24C | 0.77688 | 2.4908e-06 | 2.0266e-06 | 0.77688 | 26.3955 | 9.1847e-05 |  | 2011-02-09 16:23 | 24C | 0.77691 | 1.1196e-06 | 1.967e-06 | 0.77691 | 26.3966 | 4.3716e-05 | 0 | -0.0011349 | -0.0011349
LSSL2010#568 | CB-15 | 1070 |  | 2011-02-10 13:59 | 24C | 0.99427 | 2.4909e-06 | 1.967e-06 | 0.99427 | 34.7745 | 9.869e-05 |  | 2011-02-09 17:06 | 24C | 0.99425 | 2.7425e-06 | 1.9073e-06 | 0.99425 | 34.7739 | 0.00010663 | 0 | 0.00064087 | 0.00064087
LSSL2010#574 | CB-15 | 1076 |  | 2011-02-10 14:59 | 24C | 0.9727 | 3.1415e-06 | 1.967e-06 | 0.9727 | 33.9281 | 0.00012621 |  | 2011-02-09 12:30 | 24C | 0.97279 | 2.4909e-06 | 1.9073e-06 | 0.97279 | 33.9317 | 9.7029e-05 | 0 | -0.0035477 | -0.0035477
LSSL2010#583 | CB-15 | 1085 |  | 2011-02-10 15:34 | 24C | 0.93012 | 2.4911e-06 | 2.0862e-06 | 0.93012 | 32.2668 | 9.9553e-05 |  | 2011-02-09 12:35 | 24C | 0.93015 | 2.7262e-06 | 2.0266e-06 | 0.93015 | 32.2681 | 0.00010659 | 0 | -0.0012817 | -0.0012817
LSSL2010#593 | CB-12a | 974 |  | 2011-02-10 15:59 | 24C | 0.99343 | 3.3459e-06 | 2.0862e-06 | 0.99343 | 34.7417 | 0.00013202 |  | 2011-02-09 17:09 | 24C | 0.99344 | 2.4751e-06 | 2.0266e-06 | 0.99344 | 34.7418 | 9.9792e-05 | 0 | -0.00013351 | -0.00013351
LSSL2010#599 | CB-12a | 980 |  | 2011-02-10 16:11 | 24C | 0.97357 | 3.8756e-06 | 2.0862e-06 | 0.97358 | 33.9624 | 0.0001485 |  | 2011-02-09 17:17 | 24C | 0.97401 | 1.1536e-06 | 2.0266e-06 | 0.97401 | 33.9794 | 4.2703e-05 | 0 | -0.016998
LSSL2010#608 | CB-12a | 989 |  | 2011-02-10 16:29 | 24C | 0.93018 | 3.4657e-06 | 2.0266e-06 | 0.93018 | 32.2692 | 0.0001355 |  | 2011-02-09 17:26 | 24C | 0.93005 | 3.1424e-06 | 2.0266e-06 | 0.93005 | 32.264 | 0.00012176 | 0 | 0.005188 | 0.005188
LSSL2010#622 | CB-21 | 1386 |  | 2011-02-10 17:03 | 24C | 0.97856 | 3.2012e-06 | 2.0266e-06 | 0.97856 | 34.1578 | 0.00012091 |  | 2011-02-09 16:00 | 24C | 0.97834 | 3.131e-06 | 2.0266e-06 | 0.97834 | 34.1491 | 0.00012172 | 0 | 0.0086899 | 0.0086899
LSSL2010#637 | CB-21 | 1401 |  | 2011-02-10 17:36 | 24C | 0.88541 | 2.4909e-06 | 1.967e-06 | 0.88541 | 30.5363 | 9.4885e-05 |  | 2011-02-09 16:21 | 24C | 0.88377 | 3.1418e-06 | 1.967e-06 | 0.88377 | 30.473 | 0.00012365 | 0 | 0.063282
LSSL2010#641 | CB-21 | 1405 |  | 2011-02-10 17:45 | 24C | 0.74258 | 2.8897e-06 | 2.0266e-06 | 0.74259 | 25.105 | 0.00010553 |  | 2011-02-09 16:29 | 24C | 0.74217 | 2.2135e-06 | 2.0862e-06 | 0.74217 | 25.0894 | 8.5299e-05 | 0 | 0.01564 |  | -0.00016387 | 0.0038688

## Manual Salt Calc
Worksheet in case any sample needs to be recalcualted
Sample # | Bath temp | Pressure | M | Autosal Display of Cond. Ratio (Flush 1) | Autosal Display of Cond. Ratio (Flush 2) | Autosal Display of Cond. Ratio (Flush 3) | Autosal Display of Cond. Ratio (Flush 4) | Average Conductivity Ratio | Correction to Ratio | Salinity
223 | 24 | 0 | 2 | 1.9596 | 1.9597 | 1.9598 |  | 1.9597 | -9.5367e-07 | #VALUE!
1480 | 24 | 0 | ?? | 1.9957 | 1.9958 | 1.9959 |  | 1.9958 |  | 34.9168
Did not use this to calulate salinity
Mary Steel