## File List
Number of Files= |  | 54
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2010-12-0001.mrgcor1 | 2010-12-0001.samcor1 | 12 | 14 | 2010-12-0001.mrgcor1 & 2010-12-0001.samcor1
2 | 2010-12-0002.mrgcor1 | 2010-12-0002.samcor1 | 16 | 17 | 2010-12-0002.mrgcor1 & 2010-12-0002.samcor1
3 | 2010-12-0003.mrgcor1 | 2010-12-0003.samcor1 | 19 | 20 | 2010-12-0003.mrgcor1 & 2010-12-0003.samcor1
4 | 2010-12-0005.mrgcor1 | 2010-12-0005.samcor1 | 22 | 23 | 2010-12-0005.mrgcor1 & 2010-12-0005.samcor1
5 | 2010-12-0009.mrgcor1 | 2010-12-0009.samcor1 | 25 | 26 | 2010-12-0009.mrgcor1 & 2010-12-0009.samcor1
6 | 2010-12-0022.mrgcor1 | 2010-12-0022.samcor1 | 28 | 30 | 2010-12-0022.mrgcor1 & 2010-12-0022.samcor1
7 | 2010-12-0024.mrgcor1 | 2010-12-0024.samcor1 | 32 | 34 | 2010-12-0024.mrgcor1 & 2010-12-0024.samcor1
8 | 2010-12-0026.mrgcor1 | 2010-12-0026.samcor1 | 36 | 38 | 2010-12-0026.mrgcor1 & 2010-12-0026.samcor1
9 | 2010-12-0028.mrgcor1 | 2010-12-0028.samcor1 | 40 | 42 | 2010-12-0028.mrgcor1 & 2010-12-0028.samcor1
10 | 2010-12-0030.mrgcor1 | 2010-12-0030.samcor1 | 44 | 46 | 2010-12-0030.mrgcor1 & 2010-12-0030.samcor1
11 | 2010-12-0031.mrgcor1 | 2010-12-0031.samcor1 | 48 | 48 | 2010-12-0031.mrgcor1 & 2010-12-0031.samcor1
12 | 2010-12-0032.mrgcor1 | 2010-12-0032.samcor1 | 50 | 50 | 2010-12-0032.mrgcor1 & 2010-12-0032.samcor1
13 | 2010-12-0033.mrgcor1 | 2010-12-0033.samcor1 | 52 | 52 | 2010-12-0033.mrgcor1 & 2010-12-0033.samcor1
14 | 2010-12-0037.mrgcor1 | 2010-12-0037.samcor1 | 54 | 56 | 2010-12-0037.mrgcor1 & 2010-12-0037.samcor1
15 | 2010-12-0039.mrgcor1 | 2010-12-0039.samcor1 | 58 | 60 | 2010-12-0039.mrgcor1 & 2010-12-0039.samcor1
16 | 2010-12-0043.mrgcor1 | 2010-12-0043.samcor1 | 62 | 64 | 2010-12-0043.mrgcor1 & 2010-12-0043.samcor1
17 | 2010-12-0046.mrgcor1 | 2010-12-0046.samcor1 | 66 | 68 | 2010-12-0046.mrgcor1 & 2010-12-0046.samcor1
18 | 2010-12-0050.mrgcor1 | 2010-12-0050.samcor1 | 70 | 72 | 2010-12-0050.mrgcor1 & 2010-12-0050.samcor1
19 | 2010-12-0051.mrgcor1 | 2010-12-0051.samcor1 | 74 | 74 | 2010-12-0051.mrgcor1 & 2010-12-0051.samcor1
20 | 2010-12-0052.mrgcor1 | 2010-12-0052.samcor1 | 76 | 77 | 2010-12-0052.mrgcor1 & 2010-12-0052.samcor1
21 | 2010-12-0053.mrgcor1 | 2010-12-0053.samcor1 | 79 | 81 | 2010-12-0053.mrgcor1 & 2010-12-0053.samcor1
22 | 2010-12-0054.mrgcor1 | 2010-12-0054.samcor1 | 83 | 83 | 2010-12-0054.mrgcor1 & 2010-12-0054.samcor1
23 | 2010-12-0055.mrgcor1 | 2010-12-0055.samcor1 | 85 | 85 | 2010-12-0055.mrgcor1 & 2010-12-0055.samcor1
24 | 2010-12-0058.mrgcor1 | 2010-12-0058.samcor1 | 87 | 87 | 2010-12-0058.mrgcor1 & 2010-12-0058.samcor1
25 | 2010-12-0061.mrgcor1 | 2010-12-0061.samcor1 | 89 | 91 | 2010-12-0061.mrgcor1 & 2010-12-0061.samcor1
26 | 2010-12-0063.mrgcor1 | 2010-12-0063.samcor1 | 93 | 95 | 2010-12-0063.mrgcor1 & 2010-12-0063.samcor1
27 | 2010-12-0064.mrgcor1 | 2010-12-0064.samcor1 | 97 | 99 | 2010-12-0064.mrgcor1 & 2010-12-0064.samcor1
28 | 2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 101 | 119 | 2010-12-0067.mrgcor1 & 2010-12-0067.samcor1
29 | 2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 125 | 2010-12-0069.mrgcor1 & 2010-12-0069.samcor1
30 | 2010-12-0070.mrgcor1 | 2010-12-0070.samcor1 | 127 | 127 | 2010-12-0070.mrgcor1 & 2010-12-0070.samcor1
31 | 2010-12-0074.mrgcor1 | 2010-12-0074.samcor1 | 129 | 131 | 2010-12-0074.mrgcor1 & 2010-12-0074.samcor1
32 | 2010-12-0076.mrgcor1 | 2010-12-0076.samcor1 | 133 | 135 | 2010-12-0076.mrgcor1 & 2010-12-0076.samcor1
33 | 2010-12-0079.mrgcor1 | 2010-12-0079.samcor1 | 137 | 139 | 2010-12-0079.mrgcor1 & 2010-12-0079.samcor1
34 | 2010-12-0081.mrgcor1 | 2010-12-0081.samcor1 | 141 | 141 | 2010-12-0081.mrgcor1 & 2010-12-0081.samcor1
35 | 2010-12-0085.mrgcor1 | 2010-12-0085.samcor1 | 143 | 145 | 2010-12-0085.mrgcor1 & 2010-12-0085.samcor1
36 | 2010-12-0088.mrgcor1 | 2010-12-0088.samcor1 | 147 | 149 | 2010-12-0088.mrgcor1 & 2010-12-0088.samcor1
37 | 2010-12-0090.mrgcor1 | 2010-12-0090.samcor1 | 151 | 151 | 2010-12-0090.mrgcor1 & 2010-12-0090.samcor1
38 | 2010-12-0091.mrgcor1 | 2010-12-0091.samcor1 | 153 | 155 | 2010-12-0091.mrgcor1 & 2010-12-0091.samcor1
39 | 2010-12-0099.mrgcor1 | 2010-12-0099.samcor1 | 157 | 159 | 2010-12-0099.mrgcor1 & 2010-12-0099.samcor1
40 | 2010-12-0104.mrgcor1 | 2010-12-0104.samcor1 | 161 | 163 | 2010-12-0104.mrgcor1 & 2010-12-0104.samcor1
41 | 2010-12-0108.mrgcor1 | 2010-12-0108.samcor1 | 165 | 167 | 2010-12-0108.mrgcor1 & 2010-12-0108.samcor1
42 | 2010-12-0116.mrgcor1 | 2010-12-0116.samcor1 | 169 | 171 | 2010-12-0116.mrgcor1 & 2010-12-0116.samcor1
43 | 2010-12-0122.mrgcor1 | 2010-12-0122.samcor1 | 173 | 175 | 2010-12-0122.mrgcor1 & 2010-12-0122.samcor1
44 | 2010-12-0126.mrgcor1 | 2010-12-0126.samcor1 | 177 | 177 | 2010-12-0126.mrgcor1 & 2010-12-0126.samcor1
45 | 2010-12-0127.mrgcor1 | 2010-12-0127.samcor1 | 179 | 179 | 2010-12-0127.mrgcor1 & 2010-12-0127.samcor1
46 | 2010-12-0128.mrgcor1 | 2010-12-0128.samcor1 | 181 | 181 | 2010-12-0128.mrgcor1 & 2010-12-0128.samcor1
47 | 2010-12-0129.mrgcor1 | 2010-12-0129.samcor1 | 183 | 183 | 2010-12-0129.mrgcor1 & 2010-12-0129.samcor1
48 | 2010-12-0130.mrgcor1 | 2010-12-0130.samcor1 | 185 | 185 | 2010-12-0130.mrgcor1 & 2010-12-0130.samcor1
49 | 2010-12-0148.mrgcor1 | 2010-12-0148.samcor1 | 187 | 187 | 2010-12-0148.mrgcor1 & 2010-12-0148.samcor1
50 | 2010-12-0149.mrgcor1 | 2010-12-0149.samcor1 | 189 | 189 | 2010-12-0149.mrgcor1 & 2010-12-0149.samcor1
51 | 2010-12-0151.mrgcor1 | 2010-12-0151.samcor1 | 191 | 191 | 2010-12-0151.mrgcor1 & 2010-12-0151.samcor1
52 | 2010-12-0152.mrgcor1 | 2010-12-0152.samcor1 | 193 | 193 | 2010-12-0152.mrgcor1 & 2010-12-0152.samcor1
53 | 2010-12-0154.mrgcor1 | 2010-12-0154.samcor1 | 195 | 195 | 2010-12-0154.mrgcor1 & 2010-12-0154.samcor1
54 | 2010-12-0155.mrgcor1 | 2010-12-0155.samcor1 | 197 | 197 | 2010-12-0155.mrgcor1 & 2010-12-0155.samcor1

## (Sal1 - Sal_Bot)
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.02449 |  | X | CALC Y
m= | 5.7438e-07 | 5.2878e-07 |  | Std Dev of Resid |  | 0.0018177 |  | 101.056 | 0.00047639
b= | 0.00041834 | 0.00040425 |  | Resid Sum of Squares |  | 0.00015529 |  | 2001.93 | 0.0015682
# of Values Extracted = |  |  | 133
# of Values Used in Fit = |  |  | 49
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 200.896 | 31.2685 | 0.00019071 | 0.000597 | 0.00053374 | -6.3264e-05 |  | 200.896 | 0.000597
1 | YES | 176.086 | 31.2572 | 0.00022002 | -0.0021057 | 0.00051949 | 0.0026252 |  | 176.086 | -0.0021057
1 | NO | 6.1381 | 29.4314 | 0.012534 | 0.0031796 | 0.00042187 | -0.0027577 |  | 125.859 | 0.0020142
2 | NO | 21.2537 | 31.9083 | 0.00030931 | 0.00030518 | 0.00043055 | 0.00012538 |  | 401.027 | 0.00024414
2 | NO | 5.6697 | 31.5926 | 0.003441 | 0.01507 | 0.0004216 | -0.014648 |  | 300.392 | 0.0013466
3 | NO | 30.3238 | 32.0315 | 0.00091539 | -8.7738e-05 | 0.00043576 | 0.0005235 |  | 1000.96 | 0.0028343
3 | NO | 5.9804 | 31.5608 | 0.00082051 | -0.0066032 | 0.00042178 | 0.007025 |  | 750.491 | 0.0022507
4 | NO | 74.8465 | 32.5186 | 0.006668 | 0.011192 | 0.00046134 | -0.010731 |  | 1000.81 | 0.002842
4 | NO | 6.1404 | 31.5587 | 0.00030457 | -0.00020981 | 0.00042187 | 0.00063168 |  | 750.721 | 0.00039673
5 | NO | 77.1555 | 32.9755 | 0.0084633 | 0.018879 | 0.00046266 | -0.018416 |  | 500.646 | 0.0018654
5 | NO | 6.5184 | 31.7429 | 0.0016025 | 0.0065708 | 0.00042209 | -0.0061487 |  | 401.081 | 0.0020561
6 | NO | 126.97 | 33.5207 | 0.0018135 | 0.025761 | 0.00049127 | -0.025269 |  | 175.825 | 0.0015373
6 | NO | 101.076 | 33.3987 | 0.00067391 | 0.0087624 | 0.0004764 | -0.008286 |  | 151.921 | 0.0027657
6 | NO | 6.3452 | 31.8573 | 0.00026521 | -0.00061226 | 0.00042199 | 0.0010342 |  | 125.848 | 0.0021286
7 | NO | 101.005 | 33.5109 | 0.0006562 | 0.0076218 | 0.00047636 | -0.0071454 |  | 101.462 | -0.0022392
7 | NO | 76.0012 | 33.0877 | 0.014548 | 0.068588 | 0.000462 | -0.068126 |  | 102.244 | 0.0027084
7 | NO | 5.7881 | 31.9249 | 0.0011202 | 0.00019264 | 0.00042167 | 0.00022903 |  | 502.011 | 0.00096512
8 | NO | 176.741 | 33.8574 | 0.00059256 | -0.006916 | 0.00051986 | 0.0074359 |  | 402.288 | -7.2479e-05
8 | YES | 125.859 | 33.6631 | 0.00046028 | 0.0020142 | 0.00049064 | -0.0015235 |  | 751.968 | 0.001442
8 | NO | 21.4336 | 31.901 | 0.00035395 | -0.0099678 | 0.00043066 | 0.010398 |  | 500.941 | 0.00069427
9 | YES | 401.027 | 34.0606 | 0.00025356 | 0.00024414 | 0.00064869 | 0.00040455 |  | 1499.58 | 0.0011253
9 | YES | 300.392 | 33.9823 | 0.00085293 | 0.0013466 | 0.00059088 | -0.00075571 |  | 1250.25 | 0.0021095
9 | NO | 6.442 | 31.9741 | 0.00073878 | -0.00021934 | 0.00042205 | 0.00064139 |  | 999.776 | 0.0017242
10 | YES | 1000.96 | 34.3953 | 0.00026701 | 0.0028343 | 0.00099328 | -0.001841 |  | 749.8 | 0.0020447
10 | YES | 750.491 | 34.2849 | 0.00033635 | 0.0022507 | 0.00084941 | -0.0014013 |  | 500.963 | 0.00027084
10 | NO | 5.9333 | 31.9646 | 0.0002237 | -0.00034904 | 0.00042175 | 0.0007708 |  | 401.139 | 0.0018234
11 | NO | 6.2321 | 32.3143 | 0.00021629 | 0.0010147 | 0.00042192 | -0.00059279 |  | 301.457 | 0.0032844
12 | NO | 6.2509 | 32.3666 | 0.002317 | 0.0022163 | 0.00042194 | -0.0017944 |  | 249.801 | 0.0018959
13 | NO | 5.966 | 32.3225 | 0.00062429 | 9.9182e-05 | 0.00042177 | 0.00032259 |  | 200.849 | 0.0014343
14 | YES | 1000.81 | 34.4059 | 0.00029707 | 0.002842 | 0.00099319 | -0.0018488 |  | 175.561 | 0.0026398
14 | YES | 750.721 | 34.2878 | 0.00033198 | 0.00039673 | 0.00084955 | 0.00045282 |  | 101.056 | -0.002121
14 | NO | 6.1049 | 32.1149 | 0.0008836 | -0.0012627 | 0.00042185 | 0.0016845 |  | 1751.97 | 0.00020218
15 | YES | 500.646 | 34.124 | 0.00020015 | 0.0018654 | 0.00070591 | -0.0011595 |  | 1501.23 | 0.0013046
15 | YES | 401.081 | 34.0732 | 0.00026341 | 0.0020561 | 0.00064872 | -0.0014074 |  | 501.929 | 0.0037689
15 | NO | 5.9434 | 31.9971 | 0.00030598 | 0.00045776 | 0.00042176 | -3.6005e-05 |  | 125.721 | -0.0043793
16 | YES | 175.825 | 33.8813 | 0.00047635 | 0.0015373 | 0.00051934 | -0.001018 |  | 751.975 | -0.00090027
16 | YES | 151.921 | 33.8161 | 0.00051674 | 0.0027657 | 0.00050561 | -0.0022601 |  | 500.686 | -0.001194
16 | NO | 6.1782 | 31.9223 | 0.00043003 | -0.00074005 | 0.00042189 | 0.0011619 |  | 101.313 | -0.0045967
17 | NO | 76.7712 | 32.9022 | 0.0076515 | 0.011845 | 0.00046244 | -0.011382 |  | 152.77 | 0.0016136
17 | NO | 51.4401 | 32.5613 | 0.00066571 | 0.0020523 | 0.00044789 | -0.0016044 |  | 127.022 | -0.00044632
17 | NO | 6.06 | 31.9061 | 0.0011017 | -6.485e-05 | 0.00042183 | 0.00048668 |  | 1001.73 | 0.0012474
18 | YES | 125.848 | 32.8958 | 0.00041762 | 0.0021286 | 0.00049063 | -0.001638 |  | 751.719 | 0.00024796
18 | YES | 101.462 | 32.774 | 0.00022162 | -0.0022392 | 0.00047662 | 0.0027159 |  | 2001.93 | 0.00054169
18 | NO | 5.9945 | 31.1543 | 0.0024349 | -0.00038147 | 0.00042179 | 0.00080326 |  | 1503.6 | 0.00098419
19 | NO | 6.5841 | 31.266 | 0.00052424 | -1.3742 | 0.00042213 | 1.3746 |  | 1500.45 | 0.00083542
20 | NO | 51.1952 | 32.2118 | 0.0050313 | 0.0087433 | 0.00044775 | -0.0082955 |  | 752.575 | -0.0019455
20 | NO | 6.2821 | 31.4458 | 0.00037688 | -0.0012302 | 0.00042195 | 0.0016522 |  | 502.033 | -0.0016594
21 | NO | 76.5439 | 32.3277 | 0.0071398 | -0.0036735 | 0.00046231 | 0.0041359 |  | 176.465 | -0.00027847
21 | NO | 51.962 | 32.0465 | 0.00061213 | -0.0029068 | 0.00044819 | 0.003355 |  | 150.874 | 0.0011406
21 | NO | 6.5195 | 31.1328 | 0.0013684 | -0.0046921 | 0.00042209 | 0.0051142
22 | NO | 6.8571 | 31.42 | 0.00046083 | -0.0068912 | 0.00042228 | 0.0073135
23 | NO | 5.9884 | 31.4795 | 0.00027412 | -0.0022545 | 0.00042178 | 0.0026763
24 | NO | 5.6775 | 31.5928 | 0.0010654 | -0.0014248 | 0.00042161 | 0.0018464
25 | YES | 102.244 | 33.3581 | 0.002151 | 0.0027084 | 0.00047707 | -0.0022314
25 | NO | 76.2694 | 33.1024 | 0.00039038 | 0.41437 | 0.00046215 | -0.41391
25 | NO | 5.553 | 31.8793 | 0.00020813 | -0.0011578 | 0.00042153 | 0.0015793
26 | YES | 502.011 | 34.1361 | 0.00030548 | 0.00096512 | 0.00070669 | -0.00025843
26 | YES | 402.288 | 34.0564 | 0.00025961 | -7.2479e-05 | 0.00064941 | 0.00072189
26 | NO | 6.3492 | 31.8714 | 0.00021834 | -0.00076866 | 0.00042199 | 0.0011907
27 | YES | 751.968 | 34.2825 | 0.00021833 | 0.001442 | 0.00085026 | -0.0005917
27 | YES | 500.941 | 34.1191 | 0.00040943 | 0.00069427 | 0.00070608 | 1.1801e-05
27 | NO | 6.1653 | 31.9342 | 0.0002148 | -0.001358 | 0.00042189 | 0.0017799
28 | YES | 1499.58 | 34.5285 | 0.00025359 | 0.0011253 | 0.0012797 | 0.00015434
28 | YES | 1250.25 | 34.4711 | 0.00036415 | 0.0021095 | 0.0011365 | -0.00097306
28 | YES | 999.776 | 34.39 | 0.00040117 | 0.0017242 | 0.0009926 | -0.00073164
28 | YES | 749.8 | 34.2697 | 0.00024069 | 0.0020447 | 0.00084902 | -0.0011957
28 | YES | 500.963 | 34.1086 | 0.00057253 | 0.00027084 | 0.00070609 | 0.00043525
28 | YES | 401.139 | 34.0244 | 0.00087598 | 0.0018234 | 0.00064875 | -0.0011747
28 | YES | 301.457 | 33.99 | 0.00060918 | 0.0032844 | 0.0005915 | -0.002693
28 | YES | 249.801 | 33.956 | 0.00022707 | 0.0018959 | 0.00056183 | -0.0013341
28 | YES | 200.849 | 33.9003 | 0.00030609 | 0.0014343 | 0.00053371 | -0.00090062
28 | YES | 175.561 | 33.864 | 0.00052059 | 0.0026398 | 0.00051918 | -0.0021206
28 | NO | 151.047 | 33.7944 | 0.00061519 | -0.0090027 | 0.0005051 | 0.0095078
28 | NO | 125.502 | 33.603 | 0.0038594 | 0.0098534 | 0.00049043 | -0.0093629
28 | YES | 101.056 | 33.3073 | 0.0007446 | -0.002121 | 0.00047639 | 0.0025974
28 | NO | 75.6507 | 32.719 | 0.0089461 | 0.0085258 | 0.0004618 | -0.0080641
28 | NO | 50.3247 | 32.2847 | 0.0019878 | -0.015697 | 0.00044725 | 0.016145
28 | NO | 31.5045 | 32.1437 | 0.0027933 | 0.00085831 | 0.00043644 | -0.00042187
28 | NO | 21.4633 | 32.0468 | 0.0009124 | -0.00072861 | 0.00043067 | 0.0011593
28 | NO | 11.3507 | 32.0437 | 0.00034083 | -0.00068665 | 0.00042486 | 0.0011115
28 | NO | 5.4783 | 32.0437 | 0.00019597 | -0.0011215 | 0.00042149 | 0.001543
29 | YES | 1751.97 | 34.5616 | 0.00023278 | 0.00020218 | 0.0014246 | 0.0012225
29 | YES | 1501.23 | 34.5153 | 0.00025822 | 0.0013046 | 0.0012806 | -2.4005e-05
29 | YES | 501.929 | 34.0661 | 0.00030188 | 0.0037689 | 0.00070664 | -0.0030623
29 | YES | 125.721 | 33.6286 | 0.0010708 | -0.0043793 | 0.00049056 | 0.0048698
29 | NO | 6.2288 | 31.9769 | 0.000242 | -0.00022507 | 0.00042192 | 0.00064699
30 | NO | 6.0448 | 31.9679 | 0.00025891 | 0.00015449 | 0.00042182 | 0.00026732
31 | YES | 751.975 | 34.2718 | 0.00023102 | -0.00090027 | 0.00085027 | 0.0017505
31 | YES | 500.686 | 34.1156 | 0.00030875 | -0.001194 | 0.00070593 | 0.0018999
31 | NO | 6.1856 | 31.9334 | 0.00086033 | 0.0099754 | 0.0004219 | -0.0095535
32 | YES | 101.313 | 33.5703 | 0.00079388 | -0.0045967 | 0.00047654 | 0.0050732
32 | NO | 76.1829 | 32.9463 | 0.0018186 | 0.020668 | 0.0004621 | -0.020206
32 | NO | 5.7854 | 31.5534 | 0.0061522 | -0.011061 | 0.00042167 | 0.011482
33 | NO | 76.4155 | 32.7994 | 0.00067846 | -0.0035744 | 0.00046224 | 0.0040366
33 | NO | 51.4991 | 32.4094 | 0.0010915 | 0.0084305 | 0.00044793 | -0.0079826
33 | NO | 6.6133 | 31.4556 | 0.00066376 | 0.0084305 | 0.00042214 | -0.0080083
34 | NO | 5.915 | 31.4197 | 0.0024731 | -0.00015449 | 0.00042174 | 0.00057624
35 | YES | 152.77 | 33.6721 | 0.00026808 | 0.0016136 | 0.00050609 | -0.0011075
35 | YES | 127.022 | 33.6218 | 0.0010691 | -0.00044632 | 0.0004913 | 0.00093762
35 | NO | 6.6552 | 31.7449 | 0.00041603 | -0.00022125 | 0.00042217 | 0.00064342
36 | YES | 1001.73 | 34.3897 | 0.00023763 | 0.0012474 | 0.00099372 | -0.00025369
36 | YES | 751.719 | 34.282 | 0.00026 | 0.00024796 | 0.00085012 | 0.00060216
36 | NO | 6.3288 | 31.7194 | 0.00021232 | -0.0026169 | 0.00042198 | 0.0030389
37 | NO | 7.6516 | 31.7816 | 0.00039402 | -0.0021591 | 0.00042274 | 0.0025819
38 | YES | 2001.93 | 34.5986 | 0.00026073 | 0.00054169 | 0.0015682 | 0.0010265
38 | YES | 1503.6 | 34.5183 | 0.00036303 | 0.00098419 | 0.001282 | 0.00029779
38 | NO | 7.3055 | 32.0588 | 0.0002367 | -0.00098038 | 0.00042254 | 0.0014029
39 | YES | 1500.45 | 34.5037 | 0.00028344 | 0.00083542 | 0.0012802 | 0.00044476
39 | NO | 1250.83 | 34.45 | 0.0002589 | 0.23264 | 0.0011368 | -0.23151
39 | NO | 5.8579 | 31.76 | 0.00024936 | -0.0009861 | 0.00042171 | 0.0014078
40 | YES | 752.575 | 34.2687 | 0.00024681 | -0.0019455 | 0.00085061 | 0.0027961
40 | YES | 502.033 | 34.1077 | 0.0006392 | -0.0016594 | 0.0007067 | 0.0023661
40 | NO | 6.1654 | 31.5182 | 0.0010317 | -0.0079022 | 0.00042189 | 0.008324
41 | NO | 76.3358 | 32.4916 | 0.00515 | 0.011707 | 0.00046219 | -0.011245
41 | NO | 50.6984 | 32.2477 | 0.00038604 | -0.0023842 | 0.00044747 | 0.0028317
41 | NO | 5.2834 | 31.8071 | 0.00088408 | 0.0071259 | 0.00042138 | -0.0067045
42 | NO | 0 | 0 | 0 | -32.1178 | 0.00041834 | 32.1182
42 | NO | 30.6527 | 31.7018 | 0.020751 | 0.0038814 | 0.00043595 | -0.0034455
42 | NO | 6.9949 | 31.282 | 0.0019302 | 0.0072002 | 0.00042236 | -0.0067779
43 | YES | 176.465 | 33.8521 | 0.00026813 | -0.00027847 | 0.0005197 | 0.00079818
43 | YES | 150.874 | 33.8035 | 0.00066855 | 0.0011406 | 0.000505 | -0.00063559
43 | NO | 6.1346 | 31.4058 | 0.0011088 | -0.00061226 | 0.00042187 | 0.0010341
44 | NO | 5.3415 | 30.7492 | 0.010093 | 0.010471 | 0.00042141 | -0.01005
45 | NO | 6.0539 | 30.3751 | 0.03089 | 0.034855 | 0.00042182 | -0.034433
46 | NO | 6.0999 | 30.4219 | 0.06701 | 0.37708 | 0.00042185 | -0.37666
47 | NO | 6.4627 | 29.5417 | 0.0024267 | -0.16307 | 0.00042206 | 0.1635
48 | NO | 6.5954 | 29.9546 | 0.043957 | -0.31267 | 0.00042213 | 0.31309
49 | NO | 6.4956 | 29.5558 | 0.002805 | -0.022533 | 0.00042208 | 0.022955
50 | NO | 6.1399 | 31.2758 | 0.00024905 | -4.5776e-05 | 0.00042187 | 0.00046765
51 | NO | 6.2691 | 31.1021 | 0.00026121 | -0.0026417 | 0.00042195 | 0.0030636
52 | NO | 6.6509 | 31.197 | 0.00069771 | -0.0020065 | 0.00042217 | 0.0024287
53 | NO | 6.5395 | 31.7041 | 0.00021574 | -0.0024681 | 0.0004221 | 0.0028902
54 | NO | 6.2682 | 31.6885 | 0.0011196 | -0.0018311 | 0.00042195 | 0.002253
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
YES
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
YES
NO
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal1 - Sal_Bot) (2)
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.043537 |  | X | CALC Y
m= | -3.5615e-05 | 2.435e-05 |  | Std Dev of Resid |  | 0.0017999 |  | 1 | 0.0016249
b= | 0.0016605 | 0.00067044 |  | Resid Sum of Squares |  | 0.00015226 |  | 43 | 0.00012907
# of Values Extracted = |  |  | 133
# of Values Used in Fit = |  |  | 49
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 200.896 | 31.2685 | 0.00019071 | 0.000597 | 0.0016249 | 0.0010279 |  | 1 | 0.000597
1 | YES | 176.086 | 31.2572 | 0.00022002 | -0.0021057 | 0.0016249 | 0.0037306 |  | 1 | -0.0021057
1 | NO | 6.1381 | 29.4314 | 0.012534 | 0.0031796 | 0.0016249 | -0.0015546 |  | 8 | 0.0020142
2 | NO | 21.2537 | 31.9083 | 0.00030931 | 0.00030518 | 0.0015893 | 0.0012841 |  | 9 | 0.00024414
2 | NO | 5.6697 | 31.5926 | 0.003441 | 0.01507 | 0.0015893 | -0.013481 |  | 9 | 0.0013466
3 | NO | 30.3238 | 32.0315 | 0.00091539 | -8.7738e-05 | 0.0015537 | 0.0016414 |  | 10 | 0.0028343
3 | NO | 5.9804 | 31.5608 | 0.00082051 | -0.0066032 | 0.0015537 | 0.0081569 |  | 10 | 0.0022507
4 | NO | 74.8465 | 32.5186 | 0.006668 | 0.011192 | 0.0015181 | -0.0096742 |  | 14 | 0.002842
4 | NO | 6.1404 | 31.5587 | 0.00030457 | -0.00020981 | 0.0015181 | 0.0017279 |  | 14 | 0.00039673
5 | NO | 77.1555 | 32.9755 | 0.0084633 | 0.018879 | 0.0014825 | -0.017396 |  | 15 | 0.0018654
5 | NO | 6.5184 | 31.7429 | 0.0016025 | 0.0065708 | 0.0014825 | -0.0050884 |  | 15 | 0.0020561
6 | NO | 126.97 | 33.5207 | 0.0018135 | 0.025761 | 0.0014468 | -0.024314 |  | 16 | 0.0015373
6 | NO | 101.076 | 33.3987 | 0.00067391 | 0.0087624 | 0.0014468 | -0.0073155 |  | 16 | 0.0027657
6 | NO | 6.3452 | 31.8573 | 0.00026521 | -0.00061226 | 0.0014468 | 0.0020591 |  | 18 | 0.0021286
7 | NO | 101.005 | 33.5109 | 0.0006562 | 0.0076218 | 0.0014112 | -0.0062105 |  | 18 | -0.0022392
7 | NO | 76.0012 | 33.0877 | 0.014548 | 0.068588 | 0.0014112 | -0.067177 |  | 25 | 0.0027084
7 | NO | 5.7881 | 31.9249 | 0.0011202 | 0.00019264 | 0.0014112 | 0.0012186 |  | 26 | 0.00096512
8 | NO | 176.741 | 33.8574 | 0.00059256 | -0.006916 | 0.0013756 | 0.0082917 |  | 26 | -7.2479e-05
8 | YES | 125.859 | 33.6631 | 0.00046028 | 0.0020142 | 0.0013756 | -0.00063855 |  | 27 | 0.001442
8 | NO | 21.4336 | 31.901 | 0.00035395 | -0.0099678 | 0.0013756 | 0.011343 |  | 27 | 0.00069427
9 | YES | 401.027 | 34.0606 | 0.00025356 | 0.00024414 | 0.00134 | 0.0010958 |  | 28 | 0.0011253
9 | YES | 300.392 | 33.9823 | 0.00085293 | 0.0013466 | 0.00134 | -6.6001e-06 |  | 28 | 0.0021095
9 | NO | 6.442 | 31.9741 | 0.00073878 | -0.00021934 | 0.00134 | 0.0015593 |  | 28 | 0.0017242
10 | YES | 1000.96 | 34.3953 | 0.00026701 | 0.0028343 | 0.0013044 | -0.0015299 |  | 28 | 0.0020447
10 | YES | 750.491 | 34.2849 | 0.00033635 | 0.0022507 | 0.0013044 | -0.0009463 |  | 28 | 0.00027084
10 | NO | 5.9333 | 31.9646 | 0.0002237 | -0.00034904 | 0.0013044 | 0.0016534 |  | 28 | 0.0018234
11 | NO | 6.2321 | 32.3143 | 0.00021629 | 0.0010147 | 0.0012688 | 0.00025405 |  | 28 | 0.0032844
12 | NO | 6.2509 | 32.3666 | 0.002317 | 0.0022163 | 0.0012331 | -0.0009832 |  | 28 | 0.0018959
13 | NO | 5.966 | 32.3225 | 0.00062429 | 9.9182e-05 | 0.0011975 | 0.0010983 |  | 28 | 0.0014343
14 | YES | 1000.81 | 34.4059 | 0.00029707 | 0.002842 | 0.0011619 | -0.00168 |  | 28 | 0.0026398
14 | YES | 750.721 | 34.2878 | 0.00033198 | 0.00039673 | 0.0011619 | 0.00076518 |  | 28 | -0.002121
14 | NO | 6.1049 | 32.1149 | 0.0008836 | -0.0012627 | 0.0011619 | 0.0024246 |  | 29 | 0.00020218
15 | YES | 500.646 | 34.124 | 0.00020015 | 0.0018654 | 0.0011263 | -0.00073909 |  | 29 | 0.0013046
15 | YES | 401.081 | 34.0732 | 0.00026341 | 0.0020561 | 0.0011263 | -0.00092982 |  | 29 | 0.0037689
15 | NO | 5.9434 | 31.9971 | 0.00030598 | 0.00045776 | 0.0011263 | 0.00066853 |  | 29 | -0.0043793
16 | YES | 175.825 | 33.8813 | 0.00047635 | 0.0015373 | 0.0010907 | -0.00044664 |  | 31 | -0.00090027
16 | YES | 151.921 | 33.8161 | 0.00051674 | 0.0027657 | 0.0010907 | -0.001675 |  | 31 | -0.001194
16 | NO | 6.1782 | 31.9223 | 0.00043003 | -0.00074005 | 0.0010907 | 0.0018307 |  | 32 | -0.0045967
17 | NO | 76.7712 | 32.9022 | 0.0076515 | 0.011845 | 0.0010551 | -0.01079 |  | 35 | 0.0016136
17 | NO | 51.4401 | 32.5613 | 0.00066571 | 0.0020523 | 0.0010551 | -0.00099724 |  | 35 | -0.00044632
17 | NO | 6.06 | 31.9061 | 0.0011017 | -6.485e-05 | 0.0010551 | 0.0011199 |  | 36 | 0.0012474
18 | YES | 125.848 | 32.8958 | 0.00041762 | 0.0021286 | 0.0010195 | -0.0011091 |  | 36 | 0.00024796
18 | YES | 101.462 | 32.774 | 0.00022162 | -0.0022392 | 0.0010195 | 0.0032587 |  | 38 | 0.00054169
18 | NO | 5.9945 | 31.1543 | 0.0024349 | -0.00038147 | 0.0010195 | 0.0014009 |  | 38 | 0.00098419
19 | NO | 6.5841 | 31.266 | 0.00052424 | -1.3742 | 0.00098384 | 1.3752 |  | 39 | 0.00083542
20 | NO | 51.1952 | 32.2118 | 0.0050313 | 0.0087433 | 0.00094822 | -0.0077951 |  | 40 | -0.0019455
20 | NO | 6.2821 | 31.4458 | 0.00037688 | -0.0012302 | 0.00094822 | 0.0021785 |  | 40 | -0.0016594
21 | NO | 76.5439 | 32.3277 | 0.0071398 | -0.0036735 | 0.00091261 | 0.0045862 |  | 43 | -0.00027847
21 | NO | 51.962 | 32.0465 | 0.00061213 | -0.0029068 | 0.00091261 | 0.0038194 |  | 43 | 0.0011406
21 | NO | 6.5195 | 31.1328 | 0.0013684 | -0.0046921 | 0.00091261 | 0.0056047
22 | NO | 6.8571 | 31.42 | 0.00046083 | -0.0068912 | 0.00087699 | 0.0077682
23 | NO | 5.9884 | 31.4795 | 0.00027412 | -0.0022545 | 0.00084138 | 0.0030959
24 | NO | 5.6775 | 31.5928 | 0.0010654 | -0.0014248 | 0.00080576 | 0.0022306
25 | YES | 102.244 | 33.3581 | 0.002151 | 0.0027084 | 0.00077015 | -0.0019383
25 | NO | 76.2694 | 33.1024 | 0.00039038 | 0.41437 | 0.00077015 | -0.4136
25 | NO | 5.553 | 31.8793 | 0.00020813 | -0.0011578 | 0.00077015 | 0.0019279
26 | YES | 502.011 | 34.1361 | 0.00030548 | 0.00096512 | 0.00073453 | -0.00023059
26 | YES | 402.288 | 34.0564 | 0.00025961 | -7.2479e-05 | 0.00073453 | 0.00080701
26 | NO | 6.3492 | 31.8714 | 0.00021834 | -0.00076866 | 0.00073453 | 0.0015032
27 | YES | 751.968 | 34.2825 | 0.00021833 | 0.001442 | 0.00069892 | -0.00074304
27 | YES | 500.941 | 34.1191 | 0.00040943 | 0.00069427 | 0.00069892 | 4.6406e-06
27 | NO | 6.1653 | 31.9342 | 0.0002148 | -0.001358 | 0.00069892 | 0.0020569
28 | YES | 1499.58 | 34.5285 | 0.00025359 | 0.0011253 | 0.0006633 | -0.00046204
28 | YES | 1250.25 | 34.4711 | 0.00036415 | 0.0021095 | 0.0006633 | -0.0014462
28 | YES | 999.776 | 34.39 | 0.00040117 | 0.0017242 | 0.0006633 | -0.0010609
28 | YES | 749.8 | 34.2697 | 0.00024069 | 0.0020447 | 0.0006633 | -0.0013814
28 | YES | 500.963 | 34.1086 | 0.00057253 | 0.00027084 | 0.0006633 | 0.00039246
28 | YES | 401.139 | 34.0244 | 0.00087598 | 0.0018234 | 0.0006633 | -0.0011601
28 | YES | 301.457 | 33.99 | 0.00060918 | 0.0032844 | 0.0006633 | -0.0026211
28 | YES | 249.801 | 33.956 | 0.00022707 | 0.0018959 | 0.0006633 | -0.0012326
28 | YES | 200.849 | 33.9003 | 0.00030609 | 0.0014343 | 0.0006633 | -0.00077103
28 | YES | 175.561 | 33.864 | 0.00052059 | 0.0026398 | 0.0006633 | -0.0019765
28 | NO | 151.047 | 33.7944 | 0.00061519 | -0.0090027 | 0.0006633 | 0.009666
28 | NO | 125.502 | 33.603 | 0.0038594 | 0.0098534 | 0.0006633 | -0.0091901
28 | YES | 101.056 | 33.3073 | 0.0007446 | -0.002121 | 0.0006633 | 0.0027843
28 | NO | 75.6507 | 32.719 | 0.0089461 | 0.0085258 | 0.0006633 | -0.0078625
28 | NO | 50.3247 | 32.2847 | 0.0019878 | -0.015697 | 0.0006633 | 0.016361
28 | NO | 31.5045 | 32.1437 | 0.0027933 | 0.00085831 | 0.0006633 | -0.00019501
28 | NO | 21.4633 | 32.0468 | 0.0009124 | -0.00072861 | 0.0006633 | 0.0013919
28 | NO | 11.3507 | 32.0437 | 0.00034083 | -0.00068665 | 0.0006633 | 0.0013499
28 | NO | 5.4783 | 32.0437 | 0.00019597 | -0.0011215 | 0.0006633 | 0.0017848
29 | YES | 1751.97 | 34.5616 | 0.00023278 | 0.00020218 | 0.00062769 | 0.00042551
29 | YES | 1501.23 | 34.5153 | 0.00025822 | 0.0013046 | 0.00062769 | -0.00067694
29 | YES | 501.929 | 34.0661 | 0.00030188 | 0.0037689 | 0.00062769 | -0.0031412
29 | YES | 125.721 | 33.6286 | 0.0010708 | -0.0043793 | 0.00062769 | 0.005007
29 | NO | 6.2288 | 31.9769 | 0.000242 | -0.00022507 | 0.00062769 | 0.00085275
30 | NO | 6.0448 | 31.9679 | 0.00025891 | 0.00015449 | 0.00059207 | 0.00043757
31 | YES | 751.975 | 34.2718 | 0.00023102 | -0.00090027 | 0.00055645 | 0.0014567
31 | YES | 500.686 | 34.1156 | 0.00030875 | -0.001194 | 0.00055645 | 0.0017505
31 | NO | 6.1856 | 31.9334 | 0.00086033 | 0.0099754 | 0.00055645 | -0.009419
32 | YES | 101.313 | 33.5703 | 0.00079388 | -0.0045967 | 0.00052084 | 0.0051175
32 | NO | 76.1829 | 32.9463 | 0.0018186 | 0.020668 | 0.00052084 | -0.020147
32 | NO | 5.7854 | 31.5534 | 0.0061522 | -0.011061 | 0.00052084 | 0.011582
33 | NO | 76.4155 | 32.7994 | 0.00067846 | -0.0035744 | 0.00048522 | 0.0040596
33 | NO | 51.4991 | 32.4094 | 0.0010915 | 0.0084305 | 0.00048522 | -0.0079453
33 | NO | 6.6133 | 31.4556 | 0.00066376 | 0.0084305 | 0.00048522 | -0.0079453
34 | NO | 5.915 | 31.4197 | 0.0024731 | -0.00015449 | 0.00044961 | 0.0006041
35 | YES | 152.77 | 33.6721 | 0.00026808 | 0.0016136 | 0.00041399 | -0.0011996
35 | YES | 127.022 | 33.6218 | 0.0010691 | -0.00044632 | 0.00041399 | 0.00086031
35 | NO | 6.6552 | 31.7449 | 0.00041603 | -0.00022125 | 0.00041399 | 0.00063525
36 | YES | 1001.73 | 34.3897 | 0.00023763 | 0.0012474 | 0.00037838 | -0.00086903
36 | YES | 751.719 | 34.282 | 0.00026 | 0.00024796 | 0.00037838 | 0.00013042
36 | NO | 6.3288 | 31.7194 | 0.00021232 | -0.0026169 | 0.00037838 | 0.0029953
37 | NO | 7.6516 | 31.7816 | 0.00039402 | -0.0021591 | 0.00034276 | 0.0025019
38 | YES | 2001.93 | 34.5986 | 0.00026073 | 0.00054169 | 0.00030715 | -0.00023454
38 | YES | 1503.6 | 34.5183 | 0.00036303 | 0.00098419 | 0.00030715 | -0.00067704
38 | NO | 7.3055 | 32.0588 | 0.0002367 | -0.00098038 | 0.00030715 | 0.0012875
39 | YES | 1500.45 | 34.5037 | 0.00028344 | 0.00083542 | 0.00027153 | -0.00056389
39 | NO | 1250.83 | 34.45 | 0.0002589 | 0.23264 | 0.00027153 | -0.23237
39 | NO | 5.8579 | 31.76 | 0.00024936 | -0.0009861 | 0.00027153 | 0.0012576
40 | YES | 752.575 | 34.2687 | 0.00024681 | -0.0019455 | 0.00023592 | 0.0021814
40 | YES | 502.033 | 34.1077 | 0.0006392 | -0.0016594 | 0.00023592 | 0.0018953
40 | NO | 6.1654 | 31.5182 | 0.0010317 | -0.0079022 | 0.00023592 | 0.0081381
41 | NO | 76.3358 | 32.4916 | 0.00515 | 0.011707 | 0.0002003 | -0.011507
41 | NO | 50.6984 | 32.2477 | 0.00038604 | -0.0023842 | 0.0002003 | 0.0025845
41 | NO | 5.2834 | 31.8071 | 0.00088408 | 0.0071259 | 0.0002003 | -0.0069255
42 | NO | 0 | 0 | 0 | -32.1178 | 0.00016469 | 32.118
42 | NO | 30.6527 | 31.7018 | 0.020751 | 0.0038814 | 0.00016469 | -0.0037168
42 | NO | 6.9949 | 31.282 | 0.0019302 | 0.0072002 | 0.00016469 | -0.0070356
43 | YES | 176.465 | 33.8521 | 0.00026813 | -0.00027847 | 0.00012907 | 0.00040754
43 | YES | 150.874 | 33.8035 | 0.00066855 | 0.0011406 | 0.00012907 | -0.0010115
43 | NO | 6.1346 | 31.4058 | 0.0011088 | -0.00061226 | 0.00012907 | 0.00074133
44 | NO | 5.3415 | 30.7492 | 0.010093 | 0.010471 | 9.3456e-05 | -0.010378
45 | NO | 6.0539 | 30.3751 | 0.03089 | 0.034855 | 5.7841e-05 | -0.034797
46 | NO | 6.0999 | 30.4219 | 0.06701 | 0.37708 | 2.2226e-05 | -0.37706
47 | NO | 6.4627 | 29.5417 | 0.0024267 | -0.16307 | -1.3389e-05 | 0.16306
48 | NO | 6.5954 | 29.9546 | 0.043957 | -0.31267 | -4.9004e-05 | 0.31262
49 | NO | 6.4956 | 29.5558 | 0.002805 | -0.022533 | -8.462e-05 | 0.022449
50 | NO | 6.1399 | 31.2758 | 0.00024905 | -4.5776e-05 | -0.00012023 | -7.4459e-05
51 | NO | 6.2691 | 31.1021 | 0.00026121 | -0.0026417 | -0.00015585 | 0.0024858
52 | NO | 6.6509 | 31.197 | 0.00069771 | -0.0020065 | -0.00019147 | 0.0018151
53 | NO | 6.5395 | 31.7041 | 0.00021574 | -0.0024681 | -0.00022708 | 0.002241
54 | NO | 6.2682 | 31.6885 | 0.0011196 | -0.0018311 | -0.0002627 | 0.0015684
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
YES
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
YES
NO
NO
YES
YES
NO
YES
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.3 | Run at 2010/10/12 14:08:40
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2010-12-0001.mrgcor1 | 2010-12-0001.samcor1 | 121 | 200.89 | 200.896 | 0.037495 | 0.0059509 | 31.2679 | 31.2695 | 0.00028717 | 0.0015507 | 31.2679 | 31.2685 | 0.00019071 | 0.000597
2010-12-0001.mrgcor1 | 2010-12-0001.samcor1 | 121 | 176.09 | 176.086 | 0.029772 | -0.0035553 | 31.2593 | 31.2582 | 0.00020694 | -0.0011101 | 31.2593 | 31.2572 | 0.00022002 | -0.0021057
2010-12-0001.mrgcor1 | 2010-12-0001.samcor1 | 121 | 6.14 | 6.1381 | 0.032257 | -0.0019007 | 29.4282 | 29.4337 | 0.01511 | 0.0054913 | 29.4282 | 29.4314 | 0.012534 | 0.0031796
2010-12-0002.mrgcor1 | 2010-12-0002.samcor1 | 132 | 21.32 | 21.2537 | 0.38557 | -0.066288 | 31.908 | 31.909 | 0.00033189 | 0.00095367 | 31.908 | 31.9083 | 0.00030931 | 0.00030518
2010-12-0002.mrgcor1 | 2010-12-0002.samcor1 | 121 | 5.67 | 5.6697 | 0.2454 | -0.00024796 | 31.5775 | 31.5949 | 0.0032116 | 0.017384 | 31.5775 | 31.5926 | 0.003441 | 0.01507
2010-12-0003.mrgcor1 | 2010-12-0003.samcor1 | 112 | 30.24 | 30.3238 | 0.49622 | 0.083839 | 32.0316 | 32.0325 | 0.0019827 | 0.00093842 | 32.0316 | 32.0315 | 0.00091539 | -8.7738e-05
2010-12-0003.mrgcor1 | 2010-12-0003.samcor1 | 121 | 5.98 | 5.9804 | 0.15586 | 0.00041342 | 31.5674 | 31.5616 | 0.00035362 | -0.0058479 | 31.5674 | 31.5608 | 0.00082051 | -0.0066032
2010-12-0005.mrgcor1 | 2010-12-0005.samcor1 | 103 | 75.08 | 74.8465 | 0.50439 | -0.2335 | 32.5074 | 32.5182 | 0.0071206 | 0.010815 | 32.5074 | 32.5186 | 0.006668 | 0.011192
2010-12-0005.mrgcor1 | 2010-12-0005.samcor1 | 121 | 6.14 | 6.1404 | 0.15382 | 0.00041342 | 31.5589 | 31.5581 | 0.0002102 | -0.00078011 | 31.5589 | 31.5587 | 0.00030457 | -0.00020981
2010-12-0009.mrgcor1 | 2010-12-0009.samcor1 | 121 | 77.16 | 77.1555 | 0.28219 | -0.0045471 | 32.9566 | 32.9714 | 0.0081611 | 0.014797 | 32.9566 | 32.9755 | 0.0084633 | 0.018879
2010-12-0009.mrgcor1 | 2010-12-0009.samcor1 | 121 | 6.52 | 6.5184 | 0.11662 | -0.0015702 | 31.7363 | 31.7427 | 0.0020856 | 0.0064278 | 31.7363 | 31.7429 | 0.0016025 | 0.0065708
2010-12-0022.mrgcor1 | 2010-12-0022.samcor1 | 121 | 126.97 | 126.97 | 0.13178 | 0.00032806 | 33.4949 | 33.5212 | 0.0012625 | 0.026321 | 33.4949 | 33.5207 | 0.0018135 | 0.025761
2010-12-0022.mrgcor1 | 2010-12-0022.samcor1 | 121 | 101.08 | 101.076 | 0.074682 | -0.0039673 | 33.3899 | 33.0583 | 0.14551 | -0.33157 | 33.3899 | 33.3987 | 0.00067391 | 0.0087624
2010-12-0022.mrgcor1 | 2010-12-0022.samcor1 | 121 | 6.34 | 6.3452 | 0.12225 | 0.0052066 | 31.8579 | 31.8552 | 0.00029141 | -0.0026913 | 31.8579 | 31.8573 | 0.00026521 | -0.00061226
2010-12-0024.mrgcor1 | 2010-12-0024.samcor1 | 121 | 101.01 | 101.005 | 0.092997 | -0.005043 | 33.5033 | 33.5104 | 0.00048154 | 0.0070839 | 33.5033 | 33.5109 | 0.0006562 | 0.0076218
2010-12-0024.mrgcor1 | 2010-12-0024.samcor1 | 121 | 76 | 76.0012 | 0.4477 | 0.0011597 | 33.0191 | 33.1014 | 0.010022 | 0.082287 | 33.0191 | 33.0877 | 0.014548 | 0.068588
2010-12-0024.mrgcor1 | 2010-12-0024.samcor1 | 121 | 5.79 | 5.7881 | 0.10845 | -0.0019007 | 31.9247 | 31.9225 | 0.0011482 | -0.0022469 | 31.9247 | 31.9249 | 0.0011202 | 0.00019264
2010-12-0026.mrgcor1 | 2010-12-0026.samcor1 | 121 | 176.74 | 176.741 | 0.18405 | 0.0013123 | 33.8643 | 33.8595 | 0.001367 | -0.004818 | 33.8643 | 33.8574 | 0.00059256 | -0.006916
2010-12-0026.mrgcor1 | 2010-12-0026.samcor1 | 121 | 125.86 | 125.859 | 0.22587 | -0.00057983 | 33.6611 | 33.6615 | 0.00043787 | 0.00038528 | 33.6611 | 33.6631 | 0.00046028 | 0.0020142
2010-12-0026.mrgcor1 | 2010-12-0026.samcor1 | 121 | 21.43 | 21.4336 | 0.30367 | 0.0036354 | 31.911 | 31.8986 | 0.00041983 | -0.012421 | 31.911 | 31.901 | 0.00035395 | -0.0099678
2010-12-0028.mrgcor1 | 2010-12-0028.samcor1 | 121 | 401.03 | 401.027 | 0.16044 | -0.0027161 | 34.0604 | 34.06 | 0.00024127 | -0.00043487 | 34.0604 | 34.0606 | 0.00025356 | 0.00024414
2010-12-0028.mrgcor1 | 2010-12-0028.samcor1 | 121 | 300.39 | 300.392 | 0.42279 | 0.0021362 | 33.981 | 33.9815 | 0.0012242 | 0.00050736 | 33.981 | 33.9823 | 0.00085293 | 0.0013466
2010-12-0028.mrgcor1 | 2010-12-0028.samcor1 | 121 | 6.44 | 6.442 | 0.31159 | 0.0019836 | 31.9743 | 31.9708 | 0.001038 | -0.0035381 | 31.9743 | 31.9741 | 0.00073878 | -0.00021934
2010-12-0030.mrgcor1 | 2010-12-0030.samcor1 | 121 | 1000.96 | 1000.96 | 0.4132 | 0.0040894 | 34.3925 | 34.3932 | 0.00032377 | 0.00066376 | 34.3925 | 34.3953 | 0.00026701 | 0.0028343
2010-12-0030.mrgcor1 | 2010-12-0030.samcor1 | 121 | 750.49 | 750.491 | 0.18171 | 0.0010986 | 34.2826 | 34.282 | 0.00035577 | -0.00056076 | 34.2826 | 34.2849 | 0.00033635 | 0.0022507
2010-12-0030.mrgcor1 | 2010-12-0030.samcor1 | 121 | 5.93 | 5.9333 | 0.13422 | 0.0033059 | 31.9649 | 31.9617 | 0.00023692 | -0.0031567 | 31.9649 | 31.9646 | 0.0002237 | -0.00034904
2010-12-0031.mrgcor1 | 2010-12-0031.samcor1 | 121 | 6.23 | 6.2321 | 0.28377 | 0.0020661 | 32.3133 | 32.3097 | 0.00017892 | -0.0036354 | 32.3133 | 32.3143 | 0.00021629 | 0.0010147
2010-12-0032.mrgcor1 | 2010-12-0032.samcor1 | 121 | 6.25 | 6.2509 | 0.10926 | 0.00090933 | 32.3644 | 32.3618 | 0.002684 | -0.0026169 | 32.3644 | 32.3666 | 0.002317 | 0.0022163
2010-12-0033.mrgcor1 | 2010-12-0033.samcor1 | 121 | 5.97 | 5.966 | 0.34791 | -0.0040493 | 32.3224 | 32.3179 | 0.00033287 | -0.0044975 | 32.3224 | 32.3225 | 0.00062429 | 9.9182e-05
2010-12-0037.mrgcor1 | 2010-12-0037.samcor1 | 121 | 1000.82 | 1000.81 | 0.31733 | -0.0053711 | 34.4031 | 34.4034 | 0.00028027 | 0.0002861 | 34.4031 | 34.4059 | 0.00029707 | 0.002842
2010-12-0037.mrgcor1 | 2010-12-0037.samcor1 | 121 | 750.72 | 750.721 | 0.056954 | 0.0011597 | 34.2874 | 34.2852 | 0.00022602 | -0.0022202 | 34.2874 | 34.2878 | 0.00033198 | 0.00039673
2010-12-0037.mrgcor1 | 2010-12-0037.samcor1 | 121 | 6.1 | 6.1049 | 0.31346 | 0.0048761 | 32.1162 | 32.1113 | 0.00030188 | -0.0048561 | 32.1162 | 32.1149 | 0.0008836 | -0.0012627
2010-12-0039.mrgcor1 | 2010-12-0039.samcor1 | 121 | 500.65 | 500.646 | 0.17787 | -0.0039673 | 34.1221 | 34.1222 | 0.00024268 | 6.8665e-05 | 34.1221 | 34.124 | 0.00020015 | 0.0018654
2010-12-0039.mrgcor1 | 2010-12-0039.samcor1 | 121 | 401.08 | 401.081 | 0.22149 | 0.0010986 | 34.0711 | 34.0709 | 0.00021511 | -0.00014877 | 34.0711 | 34.0732 | 0.00026341 | 0.0020561
2010-12-0039.mrgcor1 | 2010-12-0039.samcor1 | 121 | 5.94 | 5.9434 | 0.11133 | 0.0033884 | 31.9966 | 31.9935 | 0.00029328 | -0.0030689 | 31.9966 | 31.9971 | 0.00030598 | 0.00045776
2010-12-0043.mrgcor1 | 2010-12-0043.samcor1 | 121 | 175.83 | 175.825 | 0.36303 | -0.0046234 | 33.8798 | 33.8818 | 0.00098186 | 0.0020027 | 33.8798 | 33.8813 | 0.00047635 | 0.0015373
2010-12-0043.mrgcor1 | 2010-12-0043.samcor1 | 121 | 151.92 | 151.921 | 0.18035 | 0.0014954 | 33.8133 | 33.8154 | 0.00037002 | 0.0020752 | 33.8133 | 33.8161 | 0.00051674 | 0.0027657
2010-12-0043.mrgcor1 | 2010-12-0043.samcor1 | 121 | 6.18 | 6.1782 | 0.07233 | -0.0018182 | 31.923 | 31.9198 | 0.00029824 | -0.0032082 | 31.923 | 31.9223 | 0.00043003 | -0.00074005
2010-12-0046.mrgcor1 | 2010-12-0046.samcor1 | 121 | 76.77 | 76.7712 | 0.13717 | 0.0011597 | 32.8904 | 32.9054 | 0.012237 | 0.014984 | 32.8904 | 32.9022 | 0.0076515 | 0.011845
2010-12-0046.mrgcor1 | 2010-12-0046.samcor1 | 121 | 51.44 | 51.4401 | 0.30042 | 8.3923e-05 | 32.5592 | 32.5596 | 0.00051839 | 0.00041962 | 32.5592 | 32.5613 | 0.00066571 | 0.0020523
2010-12-0046.mrgcor1 | 2010-12-0046.samcor1 | 121 | 6.06 | 6.06 | 0.12925 | 0 | 31.9062 | 31.9036 | 0.0008474 | -0.0025902 | 31.9062 | 31.9061 | 0.0011017 | -6.485e-05
2010-12-0050.mrgcor1 | 2010-12-0050.samcor1 | 103 | 126.03 | 125.848 | 0.64596 | -0.18214 | 32.8937 | 32.8944 | 0.00032298 | 0.00069427 | 32.8937 | 32.8958 | 0.00041762 | 0.0021286
2010-12-0050.mrgcor1 | 2010-12-0050.samcor1 | 121 | 101.46 | 101.462 | 0.073828 | 0.0016556 | 32.7762 | 32.7726 | 0.00034315 | -0.0036278 | 32.7762 | 32.774 | 0.00022162 | -0.0022392
2010-12-0050.mrgcor1 | 2010-12-0050.samcor1 | 121 | 5.99 | 5.9945 | 0.22692 | 0.0044632 | 31.1547 | 31.1519 | 0.0019882 | -0.0028324 | 31.1547 | 31.1543 | 0.0024349 | -0.00038147
2010-12-0051.mrgcor1 | 2010-12-0051.samcor1 | 121 | 6.58 | 6.5841 | 0.15096 | 0.0040498 | 32.6402 | 31.2639 | 0.00039419 | -1.3763 | 32.6402 | 31.266 | 0.00052424 | -1.3742
2010-12-0052.mrgcor1 | 2010-12-0052.samcor1 | 121 | 51.2 | 51.1952 | 0.33694 | -0.0047951 | 32.2031 | 32.2246 | 0.0042871 | 0.021481 | 32.2031 | 32.2118 | 0.0050313 | 0.0087433
2010-12-0052.mrgcor1 | 2010-12-0052.samcor1 | 121 | 6.28 | 6.2821 | 0.18999 | 0.0020661 | 31.447 | 31.4431 | 0.00035543 | -0.003891 | 31.447 | 31.4458 | 0.00037688 | -0.0012302
2010-12-0053.mrgcor1 | 2010-12-0053.samcor1 | 117 | 76.51 | 76.5439 | 0.51279 | 0.033928 | 32.3314 | 32.3246 | 0.0047432 | -0.0067635 | 32.3314 | 32.3277 | 0.0071398 | -0.0036735
2010-12-0053.mrgcor1 | 2010-12-0053.samcor1 | 121 | 51.96 | 51.962 | 0.12499 | 0.0019836 | 32.0494 | 32.0443 | 0.00045136 | -0.0050545 | 32.0494 | 32.0465 | 0.00061213 | -0.0029068
2010-12-0053.mrgcor1 | 2010-12-0053.samcor1 | 121 | 6.52 | 6.5195 | 0.25401 | -0.00049591 | 31.1375 | 31.1307 | 0.0011478 | -0.006752 | 31.1375 | 31.1328 | 0.0013684 | -0.0046921
2010-12-0054.mrgcor1 | 2010-12-0054.samcor1 | 121 | 6.86 | 6.8571 | 0.077669 | -0.0028925 | 31.4269 | 31.4177 | 0.00021193 | -0.0092011 | 31.4269 | 31.42 | 0.00046083 | -0.0068912
2010-12-0055.mrgcor1 | 2010-12-0055.samcor1 | 121 | 5.99 | 5.9884 | 0.087264 | -0.0015697 | 31.4818 | 31.477 | 0.00027725 | -0.004818 | 31.4818 | 31.4795 | 0.00027412 | -0.0022545
2010-12-0058.mrgcor1 | 2010-12-0058.samcor1 | 121 | 5.68 | 5.6775 | 0.41954 | -0.0024791 | 31.5942 | 31.5893 | 0.00077286 | -0.0048714 | 31.5942 | 31.5928 | 0.0010654 | -0.0014248
2010-12-0061.mrgcor1 | 2010-12-0061.samcor1 | 121 | 102.24 | 102.244 | 0.26608 | 0.0038834 | 33.3554 | 33.3548 | 0.0014111 | -0.00061417 | 33.3554 | 33.3581 | 0.002151 | 0.0027084
2010-12-0061.mrgcor1 | 2010-12-0061.samcor1 | 121 | 76.27 | 76.2694 | 0.44667 | -0.00057983 | 32.688 | 33.1003 | 0.00040197 | 0.4123 | 32.688 | 33.1024 | 0.00039038 | 0.41437
2010-12-0061.mrgcor1 | 2010-12-0061.samcor1 | 121 | 5.55 | 5.553 | 0.19433 | 0.002975 | 31.8805 | 31.8762 | 0.00020793 | -0.0043259 | 31.8805 | 31.8793 | 0.00020813 | -0.0011578
2010-12-0063.mrgcor1 | 2010-12-0063.samcor1 | 121 | 502.01 | 502.011 | 0.33037 | 0.0014038 | 34.1351 | 34.1328 | 0.00023199 | -0.0023231 | 34.1351 | 34.1361 | 0.00030548 | 0.00096512
2010-12-0063.mrgcor1 | 2010-12-0063.samcor1 | 121 | 402.29 | 402.288 | 0.49988 | -0.0019836 | 34.0565 | 34.0529 | 0.00028898 | -0.0035782 | 34.0565 | 34.0564 | 0.00025961 | -7.2479e-05
2010-12-0063.mrgcor1 | 2010-12-0063.samcor1 | 121 | 6.35 | 6.3492 | 0.15182 | -0.00082636 | 31.8722 | 31.8677 | 0.00021415 | -0.0044975 | 31.8722 | 31.8714 | 0.00021834 | -0.00076866
2010-12-0064.mrgcor1 | 2010-12-0064.samcor1 | 121 | 751.97 | 751.968 | 0.27351 | -0.0020142 | 34.2811 | 34.2792 | 0.00021984 | -0.0018959 | 34.2811 | 34.2825 | 0.00021833 | 0.001442
2010-12-0064.mrgcor1 | 2010-12-0064.samcor1 | 121 | 500.94 | 500.941 | 0.45095 | 0.00057983 | 34.1184 | 34.1157 | 0.00034261 | -0.0026894 | 34.1184 | 34.1191 | 0.00040943 | 0.00069427
2010-12-0064.mrgcor1 | 2010-12-0064.samcor1 | 121 | 6.17 | 6.1653 | 0.12476 | -0.0047107 | 31.9356 | 31.9303 | 0.00019212 | -0.005291 | 31.9356 | 31.9342 | 0.0002148 | -0.001358
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 1499.58 | 1499.58 | 0.22897 | 0.0032959 | 34.5274 | 34.5253 | 0.00026216 | -0.0021324 | 34.5274 | 34.5285 | 0.00025359 | 0.0011253
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 1250.25 | 1250.25 | 0.25353 | -0.0013428 | 34.469 | 34.4671 | 0.00049083 | -0.0018539 | 34.469 | 34.4711 | 0.00036415 | 0.0021095
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 999.78 | 999.776 | 0.26842 | -0.0036011 | 34.3883 | 34.3862 | 0.00034453 | -0.0020981 | 34.3883 | 34.39 | 0.00040117 | 0.0017242
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 40 | 749.91 | 749.8 | 0.59822 | -0.10998 | 34.2677 | 34.2656 | 0.00023363 | -0.0020676 | 34.2677 | 34.2697 | 0.00024069 | 0.0020447
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 500.96 | 500.963 | 0.30195 | 0.0028076 | 34.1083 | 34.1053 | 0.00065976 | -0.0029869 | 34.1083 | 34.1086 | 0.00057253 | 0.00027084
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 401.14 | 401.139 | 0.30283 | -0.0010986 | 34.0226 | 34.0206 | 0.00071322 | -0.0020409 | 34.0226 | 34.0244 | 0.00087598 | 0.0018234
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 301.46 | 301.457 | 0.48387 | -0.0028076 | 33.9867 | 33.9856 | 0.00058018 | -0.0011139 | 33.9867 | 33.99 | 0.00060918 | 0.0032844
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 249.8 | 249.801 | 0.41777 | 0.00082398 | 33.9541 | 33.9516 | 0.00025913 | -0.0024986 | 33.9541 | 33.956 | 0.00022707 | 0.0018959
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 200.85 | 200.849 | 0.39011 | -0.0011597 | 33.8989 | 33.8959 | 0.00027761 | -0.0029717 | 33.8989 | 33.9003 | 0.00030609 | 0.0014343
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 175.56 | 175.561 | 0.14823 | 0.0011597 | 33.8614 | 33.8597 | 0.0006382 | -0.0016899 | 33.8614 | 33.864 | 0.00052059 | 0.0026398
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 151.05 | 151.047 | 0.4185 | -0.003067 | 33.8034 | 33.7894 | 0.0010651 | -0.014011 | 33.8034 | 33.7944 | 0.00061519 | -0.0090027
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 125.5 | 125.502 | 0.46429 | 0.0023956 | 33.5931 | 33.5982 | 0.002602 | 0.0050888 | 33.5931 | 33.603 | 0.0038594 | 0.0098534
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 101.06 | 101.056 | 0.32555 | -0.0037155 | 33.3094 | 33.303 | 0.00051033 | -0.0064201 | 33.3094 | 33.3073 | 0.0007446 | -0.002121
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 113 | 75.73 | 75.6507 | 0.46572 | -0.079292 | 32.7105 | 32.7187 | 0.0073477 | 0.0082436 | 32.7105 | 32.719 | 0.0089461 | 0.0085258
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 50.32 | 50.3247 | 0.2551 | 0.0047111 | 32.3004 | 32.2821 | 0.001366 | -0.018322 | 32.3004 | 32.2847 | 0.0019878 | -0.015697
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 31.5 | 31.5045 | 0.24702 | 0.0045452 | 32.1428 | 32.1365 | 0.0025385 | -0.006279 | 32.1428 | 32.1437 | 0.0027933 | 0.00085831
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 21.46 | 21.4633 | 0.1544 | 0.0033073 | 32.0475 | 32.0419 | 0.0010819 | -0.0055733 | 32.0475 | 32.0468 | 0.0009124 | -0.00072861
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 11.35 | 11.3507 | 0.083693 | 0.0006609 | 32.0444 | 32.0394 | 0.00026972 | -0.0050468 | 32.0444 | 32.0437 | 0.00034083 | -0.00068665
2010-12-0067.mrgcor1 | 2010-12-0067.samcor1 | 121 | 5.48 | 5.4783 | 0.13182 | -0.0016527 | 32.0448 | 32.0395 | 0.00020145 | -0.0053444 | 32.0448 | 32.0437 | 0.00019597 | -0.0011215
2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 1751.97 | 1751.97 | 0.311 | -0.00036621 | 34.5614 | 34.5579 | 0.00027157 | -0.0034904 | 34.5614 | 34.5616 | 0.00023278 | 0.00020218
2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 1501.23 | 1501.23 | 0.44668 | 0.00036621 | 34.514 | 34.5114 | 0.0002207 | -0.0025864 | 34.514 | 34.5153 | 0.00025822 | 0.0013046
2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 501.93 | 501.929 | 0.26469 | -0.00097656 | 34.0623 | 34.0615 | 0.000408 | -0.00076675 | 34.0623 | 34.0661 | 0.00030188 | 0.0037689
2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 125.72 | 125.721 | 0.27682 | 0.0010757 | 33.633 | 33.6244 | 0.00098502 | -0.0086289 | 33.633 | 33.6286 | 0.0010708 | -0.0043793
2010-12-0069.mrgcor1 | 2010-12-0069.samcor1 | 121 | 6.23 | 6.2288 | 0.23323 | -0.0011568 | 31.9771 | 31.9725 | 0.00030085 | -0.0045834 | 31.9771 | 31.9769 | 0.000242 | -0.00022507
2010-12-0070.mrgcor1 | 2010-12-0070.samcor1 | 121 | 6.04 | 6.0448 | 0.22614 | 0.0047936 | 31.9677 | 31.9615 | 0.00028741 | -0.0062256 | 31.9677 | 31.9679 | 0.00025891 | 0.00015449
2010-12-0074.mrgcor1 | 2010-12-0074.samcor1 | 121 | 751.98 | 751.975 | 0.1486 | -0.0048218 | 34.2727 | 34.2685 | 0.00023923 | -0.0041771 | 34.2727 | 34.2718 | 0.00023102 | -0.00090027
2010-12-0074.mrgcor1 | 2010-12-0074.samcor1 | 121 | 500.69 | 500.686 | 0.15422 | -0.0042114 | 34.1168 | 34.1115 | 0.00038183 | -0.0052605 | 34.1168 | 34.1156 | 0.00030875 | -0.001194
2010-12-0074.mrgcor1 | 2010-12-0074.samcor1 | 121 | 6.19 | 6.1856 | 0.22828 | -0.0043802 | 31.9234 | 31.9286 | 0.00094935 | 0.0052357 | 31.9234 | 31.9334 | 0.00086033 | 0.0099754
2010-12-0076.mrgcor1 | 2010-12-0076.samcor1 | 121 | 101.31 | 101.313 | 0.12736 | 0.0025635 | 33.5749 | 33.5675 | 0.00068147 | -0.0073547 | 33.5749 | 33.5703 | 0.00079388 | -0.0045967
2010-12-0076.mrgcor1 | 2010-12-0076.samcor1 | 121 | 76.18 | 76.1829 | 0.18562 | 0.0028915 | 32.9256 | 32.9426 | 0.0023334 | 0.017044 | 32.9256 | 32.9463 | 0.0018186 | 0.020668
2010-12-0076.mrgcor1 | 2010-12-0076.samcor1 | 121 | 5.79 | 5.7854 | 0.12389 | -0.0045452 | 31.5645 | 31.5425 | 0.0057899 | -0.022043 | 31.5645 | 31.5534 | 0.0061522 | -0.011061
2010-12-0079.mrgcor1 | 2010-12-0079.samcor1 | 121 | 76.42 | 76.4155 | 0.041411 | -0.0044632 | 32.803 | 32.7971 | 0.0010227 | -0.0059242 | 32.803 | 32.7994 | 0.00067846 | -0.0035744
2010-12-0079.mrgcor1 | 2010-12-0079.samcor1 | 121 | 51.5 | 51.4991 | 0.12152 | -0.0009079 | 32.401 | 32.404 | 0.0026555 | 0.0029945 | 32.401 | 32.4094 | 0.0010915 | 0.0084305
2010-12-0079.mrgcor1 | 2010-12-0079.samcor1 | 121 | 6.61 | 6.6133 | 0.10877 | 0.0033054 | 31.4472 | 31.4521 | 0.00068544 | 0.0048714 | 31.4472 | 31.4556 | 0.00066376 | 0.0084305
2010-12-0081.mrgcor1 | 2010-12-0081.samcor1 | 328 | 5.89 | 5.915 | 0.13867 | 0.025 | 31.4199 | 31.4161 | 0.0026361 | -0.0037517 | 31.4199 | 31.4197 | 0.0024731 | -0.00015449
2010-12-0085.mrgcor1 | 2010-12-0085.samcor1 | 79 | 152.89 | 152.77 | 0.60363 | -0.1205 | 33.6705 | 33.6694 | 0.00030662 | -0.0011482 | 33.6705 | 33.6721 | 0.00026808 | 0.0016136
2010-12-0085.mrgcor1 | 2010-12-0085.samcor1 | 78 | 126.68 | 127.022 | 0.6419 | 0.34205 | 33.6222 | 33.6178 | 0.00089118 | -0.0044479 | 33.6222 | 33.6218 | 0.0010691 | -0.00044632
2010-12-0085.mrgcor1 | 2010-12-0085.samcor1 | 121 | 6.65 | 6.6552 | 0.32171 | 0.0052066 | 31.7451 | 31.7402 | 0.0002397 | -0.0048981 | 31.7451 | 31.7449 | 0.00041603 | -0.00022125
2010-12-0088.mrgcor1 | 2010-12-0088.samcor1 | 103 | 1001.52 | 1001.73 | 0.54768 | 0.2063 | 34.3885 | 34.3862 | 0.00029179 | -0.0023155 | 34.3885 | 34.3897 | 0.00023763 | 0.0012474
2010-12-0088.mrgcor1 | 2010-12-0088.samcor1 | 121 | 751.72 | 751.719 | 0.32904 | -0.00097656 | 34.2818 | 34.2783 | 0.00020924 | -0.0035324 | 34.2818 | 34.282 | 0.00026 | 0.00024796
2010-12-0088.mrgcor1 | 2010-12-0088.samcor1 | 121 | 6.33 | 6.3288 | 0.29376 | -0.0012398 | 31.722 | 31.7147 | 0.00031965 | -0.0073013 | 31.722 | 31.7194 | 0.00021232 | -0.0026169
2010-12-0090.mrgcor1 | 2010-12-0090.samcor1 | 171 | 7.47 | 7.6516 | 0.36357 | 0.18158 | 31.7838 | 31.7754 | 0.00033878 | -0.0083923 | 31.7838 | 31.7816 | 0.00039402 | -0.0021591
2010-12-0091.mrgcor1 | 2010-12-0091.samcor1 | 73 | 2001.88 | 2001.93 | 0.60194 | 0.046631 | 34.5981 | 34.5947 | 0.00023125 | -0.003437 | 34.5981 | 34.5986 | 0.00026073 | 0.00054169
2010-12-0091.mrgcor1 | 2010-12-0091.samcor1 | 116 | 1503.65 | 1503.6 | 0.5559 | -0.047852 | 34.5173 | 34.5136 | 0.00028743 | -0.0036964 | 34.5173 | 34.5183 | 0.00036303 | 0.00098419
2010-12-0091.mrgcor1 | 2010-12-0091.samcor1 | 121 | 7.31 | 7.3055 | 0.3336 | -0.0044627 | 32.0598 | 32.0535 | 0.0002377 | -0.0062981 | 32.0598 | 32.0588 | 0.0002367 | -0.00098038
2010-12-0099.mrgcor1 | 2010-12-0099.samcor1 | 121 | 1500.45 | 1500.45 | 0.52516 | 0.003418 | 34.5029 | 34.4984 | 0.00024635 | -0.004509 | 34.5029 | 34.5037 | 0.00028344 | 0.00083542
2010-12-0099.mrgcor1 | 2010-12-0099.samcor1 | 121 | 1250.83 | 1250.83 | 0.36648 | -0.0020752 | 34.2174 | 34.4445 | 0.00025353 | 0.22712 | 34.2174 | 34.45 | 0.0002589 | 0.23264
2010-12-0099.mrgcor1 | 2010-12-0099.samcor1 | 121 | 5.86 | 5.8579 | 0.15368 | -0.0020661 | 31.761 | 31.7546 | 0.00051394 | -0.0063724 | 31.761 | 31.76 | 0.00024936 | -0.0009861
2010-12-0104.mrgcor1 | 2010-12-0104.samcor1 | 121 | 752.57 | 752.575 | 0.21638 | 0.0045166 | 34.2706 | 34.2637 | 0.00023386 | -0.0068665 | 34.2706 | 34.2687 | 0.00024681 | -0.0019455
2010-12-0104.mrgcor1 | 2010-12-0104.samcor1 | 121 | 502.03 | 502.033 | 0.22704 | 0.0032959 | 34.1094 | 34.1026 | 0.00026253 | -0.0068359 | 34.1094 | 34.1077 | 0.0006392 | -0.0016594
2010-12-0104.mrgcor1 | 2010-12-0104.samcor1 | 121 | 6.17 | 6.1654 | 0.31568 | -0.0045457 | 31.5261 | 31.5127 | 0.002083 | -0.013414 | 31.5261 | 31.5182 | 0.0010317 | -0.0079022
2010-12-0108.mrgcor1 | 2010-12-0108.samcor1 | 121 | 76.34 | 76.3358 | 0.48908 | -0.0042114 | 32.4799 | 32.4879 | 0.0038896 | 0.0079613 | 32.4799 | 32.4916 | 0.00515 | 0.011707
2010-12-0108.mrgcor1 | 2010-12-0108.samcor1 | 121 | 50.7 | 50.6984 | 0.25123 | -0.0015717 | 32.2501 | 32.2441 | 0.00060399 | -0.0060082 | 32.2501 | 32.2477 | 0.00038604 | -0.0023842
2010-12-0108.mrgcor1 | 2010-12-0108.samcor1 | 121 | 5.28 | 5.2834 | 0.04488 | 0.0033884 | 31.8 | 31.8014 | 0.0011236 | 0.0014057 | 31.8 | 31.8071 | 0.00088408 | 0.0071259
2010-12-0116.mrgcor1 | 2010-12-0116.samcor1 | 0 | 51.53 | 0 | 0 | -51.53 | 32.1178 | 0 | 0 | -32.1178 | 32.1178 | 0 | 0 | -32.1178
2010-12-0116.mrgcor1 | 2010-12-0116.samcor1 | 193 | 30.68 | 30.6527 | 0.17319 | -0.027254 | 31.6979 | 31.7005 | 0.0024313 | 0.002615 | 31.6979 | 31.7018 | 0.020751 | 0.0038814
2010-12-0116.mrgcor1 | 2010-12-0116.samcor1 | 121 | 6.99 | 6.9949 | 0.15441 | 0.0048761 | 31.2748 | 31.2758 | 0.0026485 | 0.00099373 | 31.2748 | 31.282 | 0.0019302 | 0.0072002
2010-12-0122.mrgcor1 | 2010-12-0122.samcor1 | 121 | 176.46 | 176.465 | 0.061469 | 0.0047913 | 33.8524 | 33.8487 | 0.00025578 | -0.0036888 | 33.8524 | 33.8521 | 0.00026813 | -0.00027847
2010-12-0122.mrgcor1 | 2010-12-0122.samcor1 | 121 | 150.87 | 150.874 | 0.1402 | 0.0040588 | 33.8024 | 33.7995 | 0.0016723 | -0.0029411 | 33.8024 | 33.8035 | 0.00066855 | 0.0011406
2010-12-0122.mrgcor1 | 2010-12-0122.samcor1 | 121 | 6.13 | 6.1346 | 0.15689 | 0.0046277 | 31.4064 | 31.4007 | 0.001433 | -0.0057316 | 31.4064 | 31.4058 | 0.0011088 | -0.00061226
2010-12-0126.mrgcor1 | 2010-12-0126.samcor1 | 166 | 5.34 | 5.3415 | 0.10852 | 0.0014458 | 30.7387 | 30.7332 | 0.012531 | -0.0054588 | 30.7387 | 30.7492 | 0.010093 | 0.010471
2010-12-0127.mrgcor1 | 2010-12-0127.samcor1 | 242 | 6.04 | 6.0539 | 0.07756 | 0.013884 | 30.3402 | 30.3548 | 0.03423 | 0.014631 | 30.3402 | 30.3751 | 0.03089 | 0.034855
2010-12-0128.mrgcor1 | 2010-12-0128.samcor1 | 156 | 6.06 | 6.0999 | 0.057271 | 0.039872 | 30.0448 | 30.3925 | 0.044283 | 0.34769 | 30.0448 | 30.4219 | 0.06701 | 0.37708
2010-12-0129.mrgcor1 | 2010-12-0129.samcor1 | 155 | 6.57 | 6.4627 | 0.077922 | -0.10729 | 29.7048 | 29.5396 | 0.0030516 | -0.16522 | 29.7048 | 29.5417 | 0.0024267 | -0.16307
2010-12-0130.mrgcor1 | 2010-12-0130.samcor1 | 121 | 6.6 | 6.5954 | 0.034976 | -0.0045452 | 30.2673 | 29.9766 | 0.054357 | -0.29074 | 30.2673 | 29.9546 | 0.043957 | -0.31267
2010-12-0148.mrgcor1 | 2010-12-0148.samcor1 | 121 | 6.5 | 6.4956 | 0.061372 | -0.0043802 | 29.5783 | 29.551 | 0.0018187 | -0.02726 | 29.5783 | 29.5558 | 0.002805 | -0.022533
2010-12-0149.mrgcor1 | 2010-12-0149.samcor1 | 121 | 6.14 | 6.1399 | 0.063672 | -8.2493e-05 | 31.2758 | 31.2707 | 0.00026005 | -0.0050716 | 31.2758 | 31.2758 | 0.00024905 | -4.5776e-05
2010-12-0151.mrgcor1 | 2010-12-0151.samcor1 | 204 | 6.05 | 6.2691 | 0.27157 | 0.21912 | 31.1047 | 31.096 | 0.00022174 | -0.0086765 | 31.1047 | 31.1021 | 0.00026121 | -0.0026417
2010-12-0152.mrgcor1 | 2010-12-0152.samcor1 | 188 | 6.56 | 6.6509 | 0.18852 | 0.090904 | 31.199 | 31.1915 | 0.00084291 | -0.007494 | 31.199 | 31.197 | 0.00069771 | -0.0020065
2010-12-0154.mrgcor1 | 2010-12-0154.samcor1 | 121 | 6.54 | 6.5395 | 0.21264 | -0.00049591 | 31.7066 | 31.6975 | 0.00023122 | -0.0091076 | 31.7066 | 31.7041 | 0.00021574 | -0.0024681
2010-12-0155.mrgcor1 | 2010-12-0155.samcor1 | 170 | 6.02 | 6.2682 | 0.29292 | 0.24824 | 31.6903 | 31.6836 | 0.001568 | -0.0067272 | 31.6903 | 31.6885 | 0.0011196 | -0.0018311
Create Fit Sheet