#### 2010-13

CCGS JP Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### June 5 to June 21, 2010

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from all major stations (P2, P4, P12, P16, P20 & P26) for DMS, DMSPD (dissolved) & DMSPT (total).  Incubation samples were also run for Laval University.

  - DMS

Fourteen water samples from various depths (300m, 200m, 175m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at each of the stations P4, P12, P16, P20 & P26.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface).  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in the dark and removed one at a time before analysis.  Laval University incubation samples were collected in 50mL, crimp top, serum vials.

  - DMSP

Six samples for both DMSPD and DMSPT were collected at each station; two at the surface (0m, 5m), one in the mixed layer (100m), one in the deep chlorophyll max (20m) and two in the salinity mix layer (175m, 200m).  The only exception to this was P2 where there were no 175m or 200m samples, hence, only 4 samples were collected.

- Analysis

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Helium for 4 minutes at ~50mL/min.  The DMS was extracted from the water and absorbed onto a 100cm Teflon trap kept at -180oC (liquid nitrogen).  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasorb 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run immediately after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

3.5mL of seawater was collected directly from the niskin into a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve is valid for 12 hours.  If analysis exceeds 12 hours, a continuing calibration standard is run to ensure the calibration curve is still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.  Stripping efficiency was evaluated at the beginning of the cruise and was proven to be acceptable at over 95%.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by simply treating MQ water as an actual sample.  For example, in the case of DMSPD it was put through a separate funnel and for DMSPT it was added directly to the polypropylene tube.

- Data & Results

  - DMS

This was the first cruise where the DMS System #2 Purge and Trap was used along with the newer Agilent 6890 GC with the more sensitive Flame Photometric Detector.  This system was different than the previous system in that it allowed for a faster analysis time (10 minutes versus 20) and greater FPD sensitivity.

  - DMSP

Samples were run at IOS.

- Conclusions

6.1 DMS

Instrument and analysis performed relatively well on this cruise.  There were issues however with obtaining a blank that was completely free of DMS.  The new detector is in some ways too sensitive and I was detecting DMS in the water I needed to use to make the standards.  I suspect these low levels have always been present but were previously non-detectable.  For the most part the problem was solved by using MQ water directly from the system in the ships lab, however, it was difficult to rid the system of minute residual carry-over DMS from previous samples.  This “carry-over” was less than 1% but with the more sensitive detector still presented a problem when one was trying to achieve a blank 100% free of DMS.  In some cases the lower level standards were blank subtracted in order to obtain a more accurate calibration curve.  These problems may be solved in a couple ways.  First, it may require a better cleaning procedure for the amber, glass, bottles used to store the MQ water used for standards.  Second, it may be possible to adjust the detector so as to reduce its sensitivity.

Another potential problem is that with the increased sensitivity of the detector there is a reduced linear range.  That is, levels of DMS that ran under the old system (i.e. > 8nM) were able to be resolved without problem but with the new system any level of DMS over 10nM causes over-saturation of the detector.  In other words the peak would not be resolved and was “flat topped.”  Reducing the sensitivity (as discussed above) of the detector may also help with this problem or an alternate solution would be to dilute a suspected high sample before it is run.

  - DMSP

No problems to report.