## File List
Number of Files= |  | 19
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2010-13-0001.mrg | 2010-13-0001.sam | 11 | 21 | 2010-13-0001.mrg & 2010-13-0001.sam
2 | 2010-13-0005.mrg | 2010-13-0005.sam | 23 | 29 | 2010-13-0005.mrg & 2010-13-0005.sam
3 | 2010-13-0012.mrg | 2010-13-0012.sam | 31 | 47 | 2010-13-0012.mrg & 2010-13-0012.sam
4 | 2010-13-0014.mrg | 2010-13-0014.sam | 49 | 49 | 2010-13-0014.mrg & 2010-13-0014.sam
5 | 2010-13-0015.mrg | 2010-13-0015.sam | 51 | 51 | 2010-13-0015.mrg & 2010-13-0015.sam
6 | 2010-13-0016.mrg | 2010-13-0016.sam | 53 | 53 | 2010-13-0016.mrg & 2010-13-0016.sam
7 | 2010-13-0018.mrg | 2010-13-0018.sam | 55 | 72 | 2010-13-0018.mrg & 2010-13-0018.sam
8 | 2010-13-0019.mrg | 2010-13-0019.sam | 74 | 74 | 2010-13-0019.mrg & 2010-13-0019.sam
9 | 2010-13-0020.mrg | 2010-13-0020.sam | 76 | 76 | 2010-13-0020.mrg & 2010-13-0020.sam
10 | 2010-13-0021.mrg | 2010-13-0021.sam | 78 | 78 | 2010-13-0021.mrg & 2010-13-0021.sam
11 | 2010-13-0025.mrg | 2010-13-0025.sam | 80 | 102 | 2010-13-0025.mrg & 2010-13-0025.sam
12 | 2010-13-0028.mrg | 2010-13-0028.sam | 104 | 114 | 2010-13-0028.mrg & 2010-13-0028.sam
13 | 2010-13-0030.mrg | 2010-13-0030.sam | 116 | 116 | 2010-13-0030.mrg & 2010-13-0030.sam
14 | 2010-13-0031.mrg | 2010-13-0031.sam | 118 | 118 | 2010-13-0031.mrg & 2010-13-0031.sam
15 | 2010-13-0035.mrg | 2010-13-0035.sam | 120 | 143 | 2010-13-0035.mrg & 2010-13-0035.sam
16 | 2010-13-0040.mrg | 2010-13-0040.sam | 145 | 145 | 2010-13-0040.mrg & 2010-13-0040.sam
17 | 2010-13-0041.mrg | 2010-13-0041.sam | 147 | 168 | 2010-13-0041.mrg & 2010-13-0041.sam
18 | 2010-13-0046.mrg | 2010-13-0046.sam | 170 | 170 | 2010-13-0046.mrg & 2010-13-0046.sam
19 | 2010-13-0049.mrg | 2010-13-0049.sam | 172 | 195 | 2010-13-0049.mrg & 2010-13-0049.sam

## (DOX_CTD - DOX_BOT)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.94728 |  | X | CALC Y
m= | -0.041367 | 0.00087635 |  | Std Dev of Resid |  | 0.02335 |  | 0.072215 | -0.0086693
b= | -0.005682 | 0.0035422 |  | Resid Sum of Squares |  | 0.067607 |  | 7.0009 | -0.29529
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 126
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 218.403 | 0.10197 | 0.0017817 | 0.10197 | -0.0099004 | -0.11188 |  | 0.072215 | 0.010215 | 0
1 | NO | 200.531 | 0.088653 | 0.0016917 | 0.088653 | -0.0093493 | -0.098002 |  | 4.7946 | -0.19345 | 0
1 | NO | 175.42 | 0.08076 | 0.0010249 | 0.08076 | -0.0090228 | -0.089783 |  | 3.1153 | -0.11473 | 0
1 | NO | 150.458 | 0.07595 | 0.0018967 | 0.07595 | -0.0088239 | -0.084774 |  | 3.6672 | -0.12076 | 0
1 | YES | 125.596 | 0.072215 | 0.0016442 | 0.010215 | -0.0086693 | -0.018884 |  | 3.6672 | -0.13076 | 0.062
1 | NO | 100.522 | 1.4962 | 0.008499 | 0.17716 | -0.067574 | -0.24473 |  | 5.8906 | -0.21443 | 1.319
1 | NO | 75.0778 | 3.6331 | 0.0030488 | -0.0049338 | -0.15597 | -0.15104 |  | 6.8152 | -0.29277 | 3.638
1 | NO | 50.4725 | 4.5052 | 0.0047135 | -0.13576 | -0.19205 | -0.056291 |  | 7.0009 | -0.33307 | 4.641
1 | YES | 25.1291 | 4.7946 | 0.0042427 | -0.19345 | -0.20402 | -0.010565 |  | 0.42322 | -0.055781 | 4.988
1 | NO | 10.0099 | 5.5954 | 0.010131 | 0.079356 | -0.23715 | -0.3165 |  | 0.26691 | -0.034091 | 5.516
1 | NO | 5.1566 | 6.2031 | 0.21295 | -1.5569 | -0.26229 | 1.2946 |  | 0.22779 | -0.029207 | 7.76
2 | YES | 101.269 | 3.1153 | 0.0048921 | -0.11473 | -0.13455 | -0.019825 |  | 0.33769 | -0.071306
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.12076 | -0.15739 | -0.036622 |  | 0.83171 | -0.057289
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.13076 | -0.15739 | -0.026622 |  | 1.3764 | -0.085612
2 | NO | 50.9797 | 4.5985 | 0.0056717 | -0.14654 | -0.19591 | -0.049371 |  | 1.799 | -0.109
2 | YES | 24.8055 | 5.8906 | 0.0044437 | -0.21443 | -0.24936 | -0.034929 |  | 2.5945 | -0.12946
2 | YES | 10.171 | 6.8152 | 0.009922 | -0.29277 | -0.28761 | 0.0051585 |  | 2.8032 | -0.12675
2 | YES | 5.2564 | 7.0009 | 0.017248 | -0.33307 | -0.29529 | 0.037775 |  | 3.361 | -0.145
3 | YES | 1250.61 | 0.42322 | 0.0035786 | -0.055781 | -0.023189 | 0.032591 |  | 4.2772 | -0.14783
3 | YES | 1001.1 | 0.26691 | 0.0026646 | -0.034091 | -0.016723 | 0.017368 |  | 5.5548 | -0.22117
3 | YES | 800.671 | 0.22779 | 0.0026706 | -0.029207 | -0.015105 | 0.014101 |  | 6.0704 | -0.22559
3 | YES | 600.701 | 0.33769 | 0.0025028 | -0.071306 | -0.019651 | 0.051654 |  | 6.3624 | -0.24163
3 | YES | 400.813 | 0.83171 | 0.0019682 | -0.057289 | -0.040088 | 0.017202 |  | 6.2546 | -0.26943
3 | YES | 300.322 | 1.3764 | 0.0026626 | -0.085612 | -0.062619 | 0.022992 |  | 6.4424 | -0.27061
3 | YES | 250.571 | 1.799 | 0.0017464 | -0.109 | -0.080102 | 0.028898 |  | 6.3588 | -0.32422
3 | NO | 200.192 | 2.1873 | 0.0020919 | -0.15166 | -0.096166 | 0.055495 |  | 6.4118 | -0.25124
3 | YES | 175.145 | 2.5945 | 0.0028461 | -0.12946 | -0.11301 | 0.016452 |  | 6.4089 | -0.31712
3 | YES | 150.231 | 2.8032 | 0.0020788 | -0.12675 | -0.12164 | 0.0051072 |  | 2.2131 | -0.09786
3 | YES | 125.157 | 3.361 | 0.0032146 | -0.145 | -0.14472 | 0.00028261 |  | 2.1816 | -0.088427
3 | YES | 100.168 | 4.2772 | 0.0066675 | -0.14783 | -0.18262 | -0.034782 |  | 2.1816 | -0.088427
3 | YES | 75.274 | 5.5548 | 0.010038 | -0.22117 | -0.23547 | -0.014296 |  | 1.787 | -0.066
3 | YES | 50.0366 | 6.0704 | 0.0098299 | -0.22559 | -0.2568 | -0.031211 |  | 1.2673 | -0.036661
3 | NO | 25.2465 | 6.4055 | 0.0068873 | -0.22652 | -0.27066 | -0.044138 |  | 0.66384 | -0.004157
3 | YES | 10.5579 | 6.3624 | 0.0088479 | -0.24163 | -0.26888 | -0.027248 |  | 0.3805 | -0.0034959
3 | NO | 5.6538 | 6.3019 | 0.011329 | -0.21912 | -0.26637 | -0.047258 |  | 0.25207 | 0.010074
4 | YES | 5.3088 | 6.2546 | 0.0071529 | -0.26943 | -0.26442 | 0.0050135 |  | 0.24929 | 0.00028926
5 | YES | 5.4542 | 6.4424 | 0.0023322 | -0.27061 | -0.27219 | -0.0015751 |  | 0.44213 | -0.0028678
6 | NO | 5.8932 | 6.4976 | 0.016952 | -0.22142 | -0.27447 | -0.053048 |  | 0.97554 | -0.047463
7 | NO | 1499.59 | 0.52926 | 0.010928 | -0.19874 | -0.027576 | 0.17116 |  | 1.2546 | -0.068421
7 | NO | 1249.7 | 0.29893 | 0.0049936 | -0.15507 | -0.018048 | 0.13703 |  | 1.4783 | -0.061711
7 | NO | 1000.11 | 0.14787 | 0.0095638 | -0.14113 | -0.011799 | 0.12933 |  | 1.6469 | -0.066132
7 | NO | 799.938 | 0.12507 | 0.003645 | -0.13193 | -0.010856 | 0.12107 |  | 1.8105 | -0.073529
7 | NO | 600.069 | 0.26533 | 0.0057277 | -0.14267 | -0.016658 | 0.12601 |  | 2.0813 | -0.079719
7 | NO | 401.436 | 0.76881 | 0.0048275 | -0.14219 | -0.037486 | 0.1047 |  | 2.2659 | -0.092107
7 | NO | 300.308 | 1.3372 | 0.0098088 | -0.17777 | -0.061 | 0.11677 |  | 2.7438 | -0.12924
7 | NO | 250.267 | 1.7647 | 0.006877 | -0.17034 | -0.078681 | 0.091658 |  | 4.629 | -0.16099
7 | NO | 200.949 | 2.2732 | 0.0036794 | -0.19575 | -0.09972 | 0.096032 |  | 6.135 | -0.22701
7 | NO | 175.622 | 2.4642 | 0.0047875 | -0.19279 | -0.10762 | 0.085165 |  | 6.5646 | -0.29942
7 | NO | 150.306 | 2.9429 | 0.0089897 | -0.31513 | -0.12742 | 0.18771 |  | 6.4439 | -0.27907
7 | NO | 124.541 | 3.1468 | 0.0061658 | -1.0352 | -0.13586 | 0.89932 |  | 6.4652 | -0.26282
7 | NO | 100.118 | 4.5387 | 0.017814 | 0.036661 | -0.19343 | -0.2301 |  | 1.2571 | -0.067942
7 | NO | 75.4503 | 5.8218 | 0.013968 | -0.3332 | -0.24651 | 0.086685 |  | 0.58515 | -0.030851
7 | NO | 49.9647 | 6.4887 | 0.0076196 | -0.50426 | -0.2741 | 0.23015 |  | 0.25176 | -0.03124
7 | NO | 25.5491 | 6.3986 | 0.0049993 | -0.34341 | -0.27037 | 0.073031 |  | 0.32921 | -0.028785
7 | NO | 10.936 | 6.3303 | 0.0068274 | -0.34669 | -0.26755 | 0.079144 |  | 0.5568 | -0.032198
7 | YES | 5.573 | 6.3588 | 0.015454 | -0.32422 | -0.26873 | 0.055496 |  | 0.99347 | -0.097529
8 | NO | 5.2317 | 6.2715 | 0.006222 | -0.35153 | -0.26512 | 0.086419 |  | 6.1798 | -0.28821
9 | YES | 4.7448 | 6.4118 | 0.0098755 | -0.25124 | -0.27092 | -0.019681 |  | 6.4295 | -0.26547
10 | YES | 5.7805 | 6.4089 | 0.0076651 | -0.31712 | -0.2708 | 0.046316 |  | 6.415 | -0.26302
11 | YES | 3275.5 | 2.2131 | 0.0049637 | -0.09786 | -0.097234 | 0.00062604 |  | 6.4109 | -0.27111
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.095928 | -0.0075003 |  | 6.4124 | -0.26159
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.095928 | -0.0075003 |  | 2.5804 | -0.15655
11 | YES | 2500.36 | 1.787 | 0.0055976 | -0.066 | -0.079605 | -0.013605 |  | 2.5804 | -0.15955
11 | YES | 2000.62 | 1.2673 | 0.0039445 | -0.036661 | -0.058108 | -0.021447 |  | 2.375 | -0.11004
11 | YES | 1500.02 | 0.66384 | 0.0030056 | -0.004157 | -0.033143 | -0.028986 |  | 1.8804 | -0.10855
11 | YES | 1249.97 | 0.3805 | 0.0031441 | -0.0034959 | -0.021422 | -0.017927 |  | 1.2312 | -0.060777
11 | YES | 1000.45 | 0.25207 | 0.003006 | 0.010074 | -0.01611 | -0.026184 |  | 0.5678 | -0.016198
11 | YES | 800.106 | 0.24929 | 0.0026939 | 0.00028926 | -0.015994 | -0.016284 |  | 0.33898 | 0.0019752
11 | YES | 600.603 | 0.44213 | 0.0029352 | -0.0028678 | -0.023972 | -0.021104 |  | 0.29317 | -0.0088264
11 | YES | 400.012 | 0.97554 | 0.0024834 | -0.047463 | -0.046037 | 0.0014255 |  | 0.37046 | -0.012539
11 | YES | 300.556 | 1.2546 | 0.0017404 | -0.068421 | -0.057581 | 0.010841 |  | 0.63053 | -0.041471
11 | YES | 250.721 | 1.4783 | 0.0017051 | -0.061711 | -0.066835 | -0.0051241 |  | 1.6482 | -0.080843
11 | YES | 200.4 | 1.6469 | 0.0017839 | -0.066132 | -0.073809 | -0.0076762 |  | 2.4146 | -0.12442
11 | YES | 175.636 | 1.8105 | 0.002266 | -0.073529 | -0.080576 | -0.0070473 |  | 3.1558 | -0.14422
11 | YES | 150.415 | 2.0813 | 0.0032177 | -0.079719 | -0.091779 | -0.01206 |  | 3.9694 | -0.20264
11 | YES | 125.849 | 2.2659 | 0.0027683 | -0.092107 | -0.099416 | -0.0073082 |  | 4.2895 | -0.19352
11 | YES | 100.422 | 2.7438 | 0.0021603 | -0.12924 | -0.11918 | 0.010056 |  | 4.6717 | -0.19227
11 | YES | 74.8778 | 4.629 | 0.014433 | -0.16099 | -0.19717 | -0.03618 |  | 5.1518 | -0.21019
11 | YES | 50.3333 | 6.135 | 0.0067311 | -0.22701 | -0.25947 | -0.032462 |  | 5.8719 | -0.25209
11 | YES | 25.0889 | 6.5646 | 0.011212 | -0.29942 | -0.27724 | 0.02218 |  | 6.6217 | -0.27326
11 | YES | 10.3182 | 6.4439 | 0.010426 | -0.27907 | -0.27225 | 0.0068162 |  | 6.7685 | -0.26449
11 | YES | 5.6855 | 6.4652 | 0.0029411 | -0.26282 | -0.27313 | -0.010311 |  | 6.4155 | -0.25953
12 | YES | 2000.19 | 1.2571 | 0.0037998 | -0.067942 | -0.057683 | 0.010259 |  | 6.4084 | -0.2476
12 | YES | 1499.41 | 0.58515 | 0.0034026 | -0.030851 | -0.029888 | 0.00096318 |  | 6.4214 | -0.23364
12 | YES | 998.622 | 0.25176 | 0.0021796 | -0.03124 | -0.016097 | 0.015143 |  | 6.4384 | -0.28461
12 | YES | 800.39 | 0.32921 | 0.0026558 | -0.028785 | -0.019301 | 0.0094844 |  | 2.7524 | -0.16961
12 | YES | 600.518 | 0.5568 | 0.0023224 | -0.032198 | -0.028715 | 0.0034829 |  | 2.4194 | -0.12455
12 | YES | 400.245 | 0.99347 | 0.0021451 | -0.097529 | -0.046779 | 0.05075 |  | 1.8721 | -0.062901
12 | NO | 201.031 | 2.9307 | 0.006522 | -0.19628 | -0.12692 | 0.069363 |  | 1.2756 | -0.044397
12 | NO | 100.738 | 5.574 | 0.0090563 | -0.30702 | -0.23626 | 0.070754 |  | 0.584 | 0.007
12 | YES | 75.1238 | 6.1798 | 0.025755 | -0.28821 | -0.26132 | 0.026884 |  | 0.37144 | 0.023438
12 | YES | 30.1212 | 6.4295 | 0.0045019 | -0.26547 | -0.27165 | -0.0061831 |  | 0.52383 | -0.018173
12 | YES | 5.1268 | 6.415 | 0.0083676 | -0.26302 | -0.27105 | -0.0080352 |  | 1.0774 | -0.058636
13 | YES | 5.3761 | 6.4109 | 0.016095 | -0.27111 | -0.27088 | 0.00022499 |  | 1.9071 | -0.13192
14 | YES | 5.4002 | 6.4124 | 0.010385 | -0.26159 | -0.27095 | -0.0093589 |  | 2.5519 | -0.12007
15 | NO | 3681.03 | 2.6256 | 0.0054848 | -0.18145 | -0.11429 | 0.06716 |  | 3.9304 | -0.15965
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15655 | -0.11243 | 0.044126 |  | 4.3963 | -0.19268
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15955 | -0.11243 | 0.047126 |  | 4.8779 | -0.22512
15 | YES | 2999.99 | 2.375 | 0.0039569 | -0.11004 | -0.10393 | 0.0061134 |  | 5.8287 | -0.29727
15 | YES | 2500.88 | 1.8804 | 0.0068361 | -0.10855 | -0.083471 | 0.025083 |  | 6.6846 | -0.26644
15 | YES | 2000.65 | 1.2312 | 0.0025215 | -0.060777 | -0.056614 | 0.0041627 |  | 6.6043 | -0.27967
15 | YES | 1499.86 | 0.5678 | 0.0028683 | -0.016198 | -0.02917 | -0.012972 |  | 6.5625 | -0.2655
15 | YES | 1250.04 | 0.33898 | 0.0028707 | 0.0019752 | -0.019704 | -0.02168 |  | 6.5223 | -0.24274
15 | YES | 1000.74 | 0.29317 | 0.0034992 | -0.0088264 | -0.01781 | -0.0089834 |  | 6.5365 | -0.28447
15 | YES | 799.17 | 0.37046 | 0.0021492 | -0.012539 | -0.021007 | -0.0084679 |  | 3.0625 | -0.13845
15 | YES | 599.933 | 0.63053 | 0.002021 | -0.041471 | -0.031765 | 0.0097058 |  | 3.0161 | -0.14688
15 | YES | 399.65 | 1.6482 | 0.0024393 | -0.080843 | -0.073862 | 0.0069811 |  | 2.8164 | -0.14059
15 | YES | 300 | 2.4146 | 0.0036073 | -0.12442 | -0.10557 | 0.018855 |  | 2.4589 | -0.10606
15 | YES | 249.213 | 3.1558 | 0.0036844 | -0.14422 | -0.13623 | 0.007995 |  | 1.9097 | -0.059289
15 | YES | 201.385 | 3.9694 | 0.0087082 | -0.20264 | -0.16988 | 0.032752 |  | 1.3096 | -0.030418
15 | YES | 176.241 | 4.2895 | 0.0066959 | -0.19352 | -0.18313 | 0.010395 |  | 0.4281 | 0.0090992
15 | YES | 150.259 | 4.6717 | 0.0052742 | -0.19227 | -0.19894 | -0.0066657 |  | 0.38405 | -0.0059504
15 | YES | 126.022 | 5.1518 | 0.0054625 | -0.21019 | -0.2188 | -0.0086083 |  | 0.43269 | 0.00069419
15 | YES | 100.338 | 5.8719 | 0.0099315 | -0.25209 | -0.24859 | 0.0035041 |  | 0.49294 | -0.0010578
15 | YES | 74.8615 | 6.6217 | 0.0059463 | -0.27326 | -0.27961 | -0.0063493 |  | 0.80394 | -0.039058
15 | YES | 51.62 | 6.7685 | 0.0088243 | -0.26449 | -0.28568 | -0.021189 |  | 1.2659 | -0.085058
15 | YES | 25.8385 | 6.4155 | 0.012656 | -0.25953 | -0.27107 | -0.011543 |  | 1.6588 | -0.076231
15 | YES | 10.6251 | 6.4084 | 0.012711 | -0.2476 | -0.27078 | -0.023185 |  | 2.4929 | -0.12306
15 | YES | 5.959 | 6.4214 | 0.016474 | -0.23364 | -0.27132 | -0.03768 |  | 3.0154 | -0.14162
16 | YES | 5.6584 | 6.4384 | 0.0094793 | -0.28461 | -0.27202 | 0.012591 |  | 3.5806 | -0.12239
17 | NO | 4023 | 2.8556 | 0.00702 | -0.25743 | -0.12381 | 0.13362 |  | 4.2171 | -0.22392
17 | YES | 3500.03 | 2.7524 | 0.0040812 | -0.16961 | -0.11954 | 0.05007 |  | 6.6079 | -0.31913
17 | YES | 3000.73 | 2.4194 | 0.0057594 | -0.12455 | -0.10577 | 0.018777 |  | 6.671 | -0.31003
17 | YES | 2498.77 | 1.8721 | 0.0036797 | -0.062901 | -0.083126 | -0.020225 |  | 6.7 | -0.27302
17 | YES | 2001.65 | 1.2756 | 0.0036934 | -0.044397 | -0.05845 | -0.014053 |  | 6.7037 | -0.25326
17 | YES | 1499.84 | 0.584 | 0.0024461 | 0.007 | -0.02984 | -0.03684 |  | 6.7063 | -0.26866
17 | YES | 1250.74 | 0.37144 | 0.0029718 | 0.023438 | -0.021047 | -0.044485
17 | NO | 600.318 | 0.52383 | 0.003292 | 0.10783 | -0.027351 | -0.13518
17 | YES | 600.318 | 0.52383 | 0.003292 | -0.018173 | -0.027351 | -0.0091787
17 | YES | 401.395 | 1.0774 | 0.0023309 | -0.058636 | -0.050249 | 0.008387
17 | YES | 300.778 | 1.9071 | 0.0024103 | -0.13192 | -0.084573 | 0.047344
17 | YES | 250.418 | 2.5519 | 0.0033954 | -0.12007 | -0.11125 | 0.0088176
17 | NO | 200.911 | 3.4487 | 0.0024619 | -0.2133 | -0.14835 | 0.064952
17 | YES | 175.849 | 3.9304 | 0.0051943 | -0.15965 | -0.16827 | -0.0086253
17 | YES | 149.91 | 4.3963 | 0.0051059 | -0.19268 | -0.18755 | 0.0051322
17 | YES | 124.964 | 4.8779 | 0.0035967 | -0.22512 | -0.20747 | 0.017649
17 | YES | 99.8412 | 5.8287 | 0.01541 | -0.29727 | -0.2468 | 0.050464
17 | YES | 75.0487 | 6.6846 | 0.0044119 | -0.26644 | -0.2822 | -0.015766
17 | YES | 49.7955 | 6.6043 | 0.015588 | -0.27967 | -0.27889 | 0.00078391
17 | YES | 25.8528 | 6.5625 | 0.010572 | -0.2655 | -0.27715 | -0.011659
17 | NO | 10.4646 | 6.5178 | 0.0081332 | -0.37018 | -0.27531 | 0.094876
17 | YES | 5.3326 | 6.5223 | 0.0059579 | -0.24274 | -0.27549 | -0.032755
18 | YES | 4.8354 | 6.5365 | 0.0033914 | -0.28447 | -0.27608 | 0.0083906
19 | YES | 4316.66 | 3.0625 | 0.0073575 | -0.13845 | -0.13237 | 0.0060747
19 | YES | 4000.97 | 3.0161 | 0.0086493 | -0.14688 | -0.13045 | 0.016425
19 | YES | 3500.19 | 2.8164 | 0.0036599 | -0.14059 | -0.12219 | 0.018398
19 | YES | 2999.35 | 2.4589 | 0.0047492 | -0.10606 | -0.1074 | -0.0013436
19 | YES | 2498.97 | 1.9097 | 0.004458 | -0.059289 | -0.084681 | -0.025392
19 | YES | 1999.87 | 1.3096 | 0.0039175 | -0.030418 | -0.059856 | -0.029437
19 | NO | 1501.11 | 0.64938 | 0.0032666 | 0.01738 | -0.032545 | -0.049925
19 | YES | 1250.88 | 0.4281 | 0.0028472 | 0.0090992 | -0.023391 | -0.03249
19 | YES | 1001.36 | 0.38405 | 0.0022835 | -0.0059504 | -0.021569 | -0.015619
19 | YES | 799.943 | 0.43269 | 0.0025028 | 0.00069419 | -0.023581 | -0.024276
19 | YES | 600.641 | 0.49294 | 0.0033821 | -0.0010578 | -0.026074 | -0.025016
19 | YES | 400.367 | 0.80394 | 0.0031919 | -0.039058 | -0.038939 | 0.00011901
19 | YES | 301.185 | 1.2659 | 0.0036224 | -0.085058 | -0.05805 | 0.027007
19 | YES | 250.879 | 1.6588 | 0.0042284 | -0.076231 | -0.074301 | 0.0019307
19 | YES | 200.507 | 2.4929 | 0.004011 | -0.12306 | -0.10881 | 0.01425
19 | YES | 174.492 | 3.0154 | 0.0056939 | -0.14162 | -0.13042 | 0.0112
19 | YES | 151.221 | 3.5806 | 0.026746 | -0.12239 | -0.1538 | -0.031413
19 | YES | 125.937 | 4.2171 | 0.0090026 | -0.22392 | -0.18013 | 0.043786
19 | NO | 125.937 | 4.2171 | 0.0090026 | -0.24592 | -0.18013 | 0.065786
19 | YES | 75.471 | 6.6079 | 0.0043759 | -0.31913 | -0.27903 | 0.0401
19 | YES | 51.0093 | 6.671 | 0.01042 | -0.31003 | -0.28164 | 0.028391
19 | YES | 26.1199 | 6.7 | 0.010029 | -0.27302 | -0.28284 | -0.0098259
19 | YES | 10.9632 | 6.7037 | 0.0038741 | -0.25326 | -0.283 | -0.029741
19 | YES | 5.7636 | 6.7063 | 0.0022641 | -0.26866 | -0.2831 | -0.014444
NO
NO
NO
NO
YES
NO
NO
NO
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.074291 |  | X | CALC Y
m= | 3.0092e-05 | 8.3457e-06 |  | Std Dev of Resid |  | 0.11002 |  | 4.7448 | -0.16232
b= | -0.16246 | 0.010362 |  | Resid Sum of Squares |  | 1.9609 |  | 4316.66 | -0.032568
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 164
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 218.403 | 0.10197 | 0.0017817 | 0.10197 | -0.15589 | -0.25787 |  | 218.403 | 0.10197
1 | YES | 200.531 | 0.088653 | 0.0016917 | 0.088653 | -0.15643 | -0.24508 |  | 200.531 | 0.088653
1 | YES | 175.42 | 0.08076 | 0.0010249 | 0.08076 | -0.15719 | -0.23795 |  | 175.42 | 0.08076
1 | YES | 150.458 | 0.07595 | 0.0018967 | 0.07595 | -0.15794 | -0.23389 |  | 150.458 | 0.07595
1 | YES | 125.596 | 0.072215 | 0.0016442 | 0.010215 | -0.15868 | -0.1689 |  | 125.596 | 0.010215
1 | YES | 100.522 | 1.4962 | 0.008499 | 0.17716 | -0.15944 | -0.3366 |  | 100.522 | 0.17716
1 | YES | 75.0778 | 3.6331 | 0.0030488 | -0.0049338 | -0.1602 | -0.15527 |  | 75.0778 | -0.0049338
1 | YES | 50.4725 | 4.5052 | 0.0047135 | -0.13576 | -0.16095 | -0.025185 |  | 50.4725 | -0.13576
1 | YES | 25.1291 | 4.7946 | 0.0042427 | -0.19345 | -0.16171 | 0.031746 |  | 25.1291 | -0.19345
1 | YES | 10.0099 | 5.5954 | 0.010131 | 0.079356 | -0.16216 | -0.24152 |  | 10.0099 | 0.079356
1 | NO | 5.1566 | 6.2031 | 0.21295 | -1.5569 | -0.16231 | 1.3946 |  | 101.269 | -0.11473
2 | YES | 101.269 | 3.1153 | 0.0048921 | -0.11473 | -0.15942 | -0.04469 |  | 74.8171 | -0.12076
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.12076 | -0.16021 | -0.039449 |  | 74.8171 | -0.13076
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.13076 | -0.16021 | -0.029449 |  | 50.9797 | -0.14654
2 | YES | 50.9797 | 4.5985 | 0.0056717 | -0.14654 | -0.16093 | -0.014393 |  | 24.8055 | -0.21443
2 | YES | 24.8055 | 5.8906 | 0.0044437 | -0.21443 | -0.16172 | 0.052712 |  | 10.171 | -0.29277
2 | YES | 10.171 | 6.8152 | 0.009922 | -0.29277 | -0.16216 | 0.13061 |  | 5.2564 | -0.33307
2 | YES | 5.2564 | 7.0009 | 0.017248 | -0.33307 | -0.16231 | 0.17076 |  | 1250.61 | -0.055781
3 | YES | 1250.61 | 0.42322 | 0.0035786 | -0.055781 | -0.12483 | -0.06905 |  | 1001.1 | -0.034091
3 | YES | 1001.1 | 0.26691 | 0.0026646 | -0.034091 | -0.13234 | -0.098248 |  | 800.671 | -0.029207
3 | YES | 800.671 | 0.22779 | 0.0026706 | -0.029207 | -0.13837 | -0.10916 |  | 600.701 | -0.071306
3 | YES | 600.701 | 0.33769 | 0.0025028 | -0.071306 | -0.14439 | -0.073082 |  | 400.813 | -0.057289
3 | YES | 400.813 | 0.83171 | 0.0019682 | -0.057289 | -0.1504 | -0.093114 |  | 300.322 | -0.085612
3 | YES | 300.322 | 1.3764 | 0.0026626 | -0.085612 | -0.15343 | -0.067815 |  | 250.571 | -0.109
3 | YES | 250.571 | 1.799 | 0.0017464 | -0.109 | -0.15492 | -0.045924 |  | 200.192 | -0.15166
3 | YES | 200.192 | 2.1873 | 0.0020919 | -0.15166 | -0.15644 | -0.004779 |  | 175.145 | -0.12946
3 | YES | 175.145 | 2.5945 | 0.0028461 | -0.12946 | -0.15719 | -0.027731 |  | 150.231 | -0.12675
3 | YES | 150.231 | 2.8032 | 0.0020788 | -0.12675 | -0.15794 | -0.031191 |  | 125.157 | -0.145
3 | YES | 125.157 | 3.361 | 0.0032146 | -0.145 | -0.1587 | -0.013698 |  | 100.168 | -0.14783
3 | YES | 100.168 | 4.2772 | 0.0066675 | -0.14783 | -0.15945 | -0.011615 |  | 75.274 | -0.22117
3 | YES | 75.274 | 5.5548 | 0.010038 | -0.22117 | -0.1602 | 0.060975 |  | 50.0366 | -0.22559
3 | YES | 50.0366 | 6.0704 | 0.0098299 | -0.22559 | -0.16096 | 0.064629 |  | 25.2465 | -0.22652
3 | YES | 25.2465 | 6.4055 | 0.0068873 | -0.22652 | -0.1617 | 0.064817 |  | 10.5579 | -0.24163
3 | YES | 10.5579 | 6.3624 | 0.0088479 | -0.24163 | -0.16215 | 0.079482 |  | 5.6538 | -0.21912
3 | YES | 5.6538 | 6.3019 | 0.011329 | -0.21912 | -0.16229 | 0.056822 |  | 5.3088 | -0.26943
4 | YES | 5.3088 | 6.2546 | 0.0071529 | -0.26943 | -0.1623 | 0.10713 |  | 5.4542 | -0.27061
5 | YES | 5.4542 | 6.4424 | 0.0023322 | -0.27061 | -0.1623 | 0.10831 |  | 5.8932 | -0.22142
6 | YES | 5.8932 | 6.4976 | 0.016952 | -0.22142 | -0.16229 | 0.059134 |  | 1499.59 | -0.19874
7 | YES | 1499.59 | 0.52926 | 0.010928 | -0.19874 | -0.11734 | 0.081401 |  | 1249.7 | -0.15507
7 | YES | 1249.7 | 0.29893 | 0.0049936 | -0.15507 | -0.12486 | 0.030216 |  | 1000.11 | -0.14113
7 | YES | 1000.11 | 0.14787 | 0.0095638 | -0.14113 | -0.13237 | 0.008763 |  | 799.938 | -0.13193
7 | YES | 799.938 | 0.12507 | 0.003645 | -0.13193 | -0.13839 | -0.0064665 |  | 600.069 | -0.14267
7 | YES | 600.069 | 0.26533 | 0.0057277 | -0.14267 | -0.14441 | -0.001738 |  | 401.436 | -0.14219
7 | YES | 401.436 | 0.76881 | 0.0048275 | -0.14219 | -0.15038 | -0.0081942 |  | 300.308 | -0.17777
7 | YES | 300.308 | 1.3372 | 0.0098088 | -0.17777 | -0.15343 | 0.024342 |  | 250.267 | -0.17034
7 | YES | 250.267 | 1.7647 | 0.006877 | -0.17034 | -0.15493 | 0.015406 |  | 200.949 | -0.19575
7 | YES | 200.949 | 2.2732 | 0.0036794 | -0.19575 | -0.15642 | 0.039335 |  | 175.622 | -0.19279
7 | YES | 175.622 | 2.4642 | 0.0047875 | -0.19279 | -0.15718 | 0.035606 |  | 150.306 | -0.31513
7 | YES | 150.306 | 2.9429 | 0.0089897 | -0.31513 | -0.15794 | 0.15719 |  | 100.118 | 0.036661
7 | NO | 124.541 | 3.1468 | 0.0061658 | -1.0352 | -0.15872 | 0.87646 |  | 75.4503 | -0.3332
7 | YES | 100.118 | 4.5387 | 0.017814 | 0.036661 | -0.15945 | -0.19611 |  | 25.5491 | -0.34341
7 | YES | 75.4503 | 5.8218 | 0.013968 | -0.3332 | -0.16019 | 0.17301 |  | 10.936 | -0.34669
7 | NO | 49.9647 | 6.4887 | 0.0076196 | -0.50426 | -0.16096 | 0.3433 |  | 5.573 | -0.32422
7 | YES | 25.5491 | 6.3986 | 0.0049993 | -0.34341 | -0.1617 | 0.18171 |  | 5.2317 | -0.35153
7 | YES | 10.936 | 6.3303 | 0.0068274 | -0.34669 | -0.16214 | 0.18456 |  | 4.7448 | -0.25124
7 | YES | 5.573 | 6.3588 | 0.015454 | -0.32422 | -0.1623 | 0.16193 |  | 5.7805 | -0.31712
8 | YES | 5.2317 | 6.2715 | 0.006222 | -0.35153 | -0.16231 | 0.18923 |  | 3275.5 | -0.09786
9 | YES | 4.7448 | 6.4118 | 0.0098755 | -0.25124 | -0.16232 | 0.088917 |  | 3000.46 | -0.088427
10 | YES | 5.7805 | 6.4089 | 0.0076651 | -0.31712 | -0.16229 | 0.15483 |  | 3000.46 | -0.088427
11 | YES | 3275.5 | 2.2131 | 0.0049637 | -0.09786 | -0.063898 | 0.033961 |  | 2500.36 | -0.066
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.072175 | 0.016253 |  | 2000.62 | -0.036661
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.072175 | 0.016253 |  | 1500.02 | -0.004157
11 | YES | 2500.36 | 1.787 | 0.0055976 | -0.066 | -0.087224 | -0.021224 |  | 1249.97 | -0.0034959
11 | YES | 2000.62 | 1.2673 | 0.0039445 | -0.036661 | -0.10226 | -0.065601 |  | 1000.45 | 0.010074
11 | YES | 1500.02 | 0.66384 | 0.0030056 | -0.004157 | -0.11733 | -0.11317 |  | 800.106 | 0.00028926
11 | YES | 1249.97 | 0.3805 | 0.0031441 | -0.0034959 | -0.12485 | -0.12135 |  | 600.603 | -0.0028678
11 | YES | 1000.45 | 0.25207 | 0.003006 | 0.010074 | -0.13236 | -0.14243 |  | 400.012 | -0.047463
11 | YES | 800.106 | 0.24929 | 0.0026939 | 0.00028926 | -0.13839 | -0.13868 |  | 300.556 | -0.068421
11 | YES | 600.603 | 0.44213 | 0.0029352 | -0.0028678 | -0.14439 | -0.14152 |  | 250.721 | -0.061711
11 | YES | 400.012 | 0.97554 | 0.0024834 | -0.047463 | -0.15043 | -0.10296 |  | 200.4 | -0.066132
11 | YES | 300.556 | 1.2546 | 0.0017404 | -0.068421 | -0.15342 | -0.084998 |  | 175.636 | -0.073529
11 | YES | 250.721 | 1.4783 | 0.0017051 | -0.061711 | -0.15492 | -0.093209 |  | 150.415 | -0.079719
11 | YES | 200.4 | 1.6469 | 0.0017839 | -0.066132 | -0.15643 | -0.090301 |  | 125.849 | -0.092107
11 | YES | 175.636 | 1.8105 | 0.002266 | -0.073529 | -0.15718 | -0.08365 |  | 100.422 | -0.12924
11 | YES | 150.415 | 2.0813 | 0.0032177 | -0.079719 | -0.15794 | -0.078219 |  | 74.8778 | -0.16099
11 | YES | 125.849 | 2.2659 | 0.0027683 | -0.092107 | -0.15868 | -0.06657 |  | 50.3333 | -0.22701
11 | YES | 100.422 | 2.7438 | 0.0021603 | -0.12924 | -0.15944 | -0.030202 |  | 25.0889 | -0.29942
11 | YES | 74.8778 | 4.629 | 0.014433 | -0.16099 | -0.16021 | 0.00078107 |  | 10.3182 | -0.27907
11 | YES | 50.3333 | 6.135 | 0.0067311 | -0.22701 | -0.16095 | 0.066058 |  | 5.6855 | -0.26282
11 | YES | 25.0889 | 6.5646 | 0.011212 | -0.29942 | -0.16171 | 0.13771 |  | 2000.19 | -0.067942
11 | YES | 10.3182 | 6.4439 | 0.010426 | -0.27907 | -0.16215 | 0.11691 |  | 1499.41 | -0.030851
11 | YES | 5.6855 | 6.4652 | 0.0029411 | -0.26282 | -0.16229 | 0.10052 |  | 998.622 | -0.03124
12 | YES | 2000.19 | 1.2571 | 0.0037998 | -0.067942 | -0.10227 | -0.034332 |  | 800.39 | -0.028785
12 | YES | 1499.41 | 0.58515 | 0.0034026 | -0.030851 | -0.11734 | -0.086493 |  | 600.518 | -0.032198
12 | YES | 998.622 | 0.25176 | 0.0021796 | -0.03124 | -0.13241 | -0.10117 |  | 400.245 | -0.097529
12 | YES | 800.39 | 0.32921 | 0.0026558 | -0.028785 | -0.13838 | -0.10959 |  | 201.031 | -0.19628
12 | YES | 600.518 | 0.5568 | 0.0023224 | -0.032198 | -0.14439 | -0.1122 |  | 100.738 | -0.30702
12 | YES | 400.245 | 0.99347 | 0.0021451 | -0.097529 | -0.15042 | -0.052891 |  | 75.1238 | -0.28821
12 | YES | 201.031 | 2.9307 | 0.006522 | -0.19628 | -0.15641 | 0.039866 |  | 30.1212 | -0.26547
12 | YES | 100.738 | 5.574 | 0.0090563 | -0.30702 | -0.15943 | 0.14758 |  | 5.1268 | -0.26302
12 | YES | 75.1238 | 6.1798 | 0.025755 | -0.28821 | -0.1602 | 0.128 |  | 5.3761 | -0.27111
12 | YES | 30.1212 | 6.4295 | 0.0045019 | -0.26547 | -0.16156 | 0.10391 |  | 5.4002 | -0.26159
12 | YES | 5.1268 | 6.415 | 0.0083676 | -0.26302 | -0.16231 | 0.10071 |  | 3681.03 | -0.18145
13 | YES | 5.3761 | 6.4109 | 0.016095 | -0.27111 | -0.1623 | 0.10881 |  | 3500.68 | -0.15655
14 | YES | 5.4002 | 6.4124 | 0.010385 | -0.26159 | -0.1623 | 0.099285 |  | 3500.68 | -0.15955
15 | YES | 3681.03 | 2.6256 | 0.0054848 | -0.18145 | -0.051695 | 0.12976 |  | 2999.99 | -0.11004
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15655 | -0.057122 | 0.099432 |  | 2500.88 | -0.10855
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15955 | -0.057122 | 0.10243 |  | 2000.65 | -0.060777
15 | YES | 2999.99 | 2.375 | 0.0039569 | -0.11004 | -0.072189 | 0.037852 |  | 1499.86 | -0.016198
15 | YES | 2500.88 | 1.8804 | 0.0068361 | -0.10855 | -0.087208 | 0.021346 |  | 1250.04 | 0.0019752
15 | YES | 2000.65 | 1.2312 | 0.0025215 | -0.060777 | -0.10226 | -0.041484 |  | 1000.74 | -0.0088264
15 | YES | 1499.86 | 0.5678 | 0.0028683 | -0.016198 | -0.11733 | -0.10113 |  | 799.17 | -0.012539
15 | YES | 1250.04 | 0.33898 | 0.0028707 | 0.0019752 | -0.12485 | -0.12682 |  | 599.933 | -0.041471
15 | YES | 1000.74 | 0.29317 | 0.0034992 | -0.0088264 | -0.13235 | -0.12352 |  | 399.65 | -0.080843
15 | YES | 799.17 | 0.37046 | 0.0021492 | -0.012539 | -0.13842 | -0.12588 |  | 300 | -0.12442
15 | YES | 599.933 | 0.63053 | 0.002021 | -0.041471 | -0.14441 | -0.10294 |  | 249.213 | -0.14422
15 | YES | 399.65 | 1.6482 | 0.0024393 | -0.080843 | -0.15044 | -0.069595 |  | 201.385 | -0.20264
15 | YES | 300 | 2.4146 | 0.0036073 | -0.12442 | -0.15344 | -0.029015 |  | 176.241 | -0.19352
15 | YES | 249.213 | 3.1558 | 0.0036844 | -0.14422 | -0.15496 | -0.010742 |  | 150.259 | -0.19227
15 | YES | 201.385 | 3.9694 | 0.0087082 | -0.20264 | -0.1564 | 0.046232 |  | 126.022 | -0.21019
15 | YES | 176.241 | 4.2895 | 0.0066959 | -0.19352 | -0.15716 | 0.03636 |  | 100.338 | -0.25209
15 | YES | 150.259 | 4.6717 | 0.0052742 | -0.19227 | -0.15794 | 0.03433 |  | 74.8615 | -0.27326
15 | YES | 126.022 | 5.1518 | 0.0054625 | -0.21019 | -0.15867 | 0.051518 |  | 51.62 | -0.26449
15 | YES | 100.338 | 5.8719 | 0.0099315 | -0.25209 | -0.15944 | 0.092646 |  | 25.8385 | -0.25953
15 | YES | 74.8615 | 6.6217 | 0.0059463 | -0.27326 | -0.16021 | 0.11304 |  | 10.6251 | -0.2476
15 | YES | 51.62 | 6.7685 | 0.0088243 | -0.26449 | -0.16091 | 0.10358 |  | 5.959 | -0.23364
15 | YES | 25.8385 | 6.4155 | 0.012656 | -0.25953 | -0.16169 | 0.097842 |  | 5.6584 | -0.28461
15 | YES | 10.6251 | 6.4084 | 0.012711 | -0.2476 | -0.16214 | 0.085451 |  | 4023 | -0.25743
15 | YES | 5.959 | 6.4214 | 0.016474 | -0.23364 | -0.16228 | 0.071351 |  | 3500.03 | -0.16961
16 | YES | 5.6584 | 6.4384 | 0.0094793 | -0.28461 | -0.16229 | 0.12232 |  | 3000.73 | -0.12455
17 | YES | 4023 | 2.8556 | 0.00702 | -0.25743 | -0.041405 | 0.21603 |  | 2498.77 | -0.062901
17 | YES | 3500.03 | 2.7524 | 0.0040812 | -0.16961 | -0.057142 | 0.11247 |  | 2001.65 | -0.044397
17 | YES | 3000.73 | 2.4194 | 0.0057594 | -0.12455 | -0.072167 | 0.052378 |  | 1499.84 | 0.007
17 | YES | 2498.77 | 1.8721 | 0.0036797 | -0.062901 | -0.087272 | -0.024371 |  | 1250.74 | 0.023438
17 | YES | 2001.65 | 1.2756 | 0.0036934 | -0.044397 | -0.10223 | -0.057834 |  | 600.318 | 0.10783
17 | YES | 1499.84 | 0.584 | 0.0024461 | 0.007 | -0.11733 | -0.12433 |  | 600.318 | -0.018173
17 | YES | 1250.74 | 0.37144 | 0.0029718 | 0.023438 | -0.12483 | -0.14827 |  | 401.395 | -0.058636
17 | YES | 600.318 | 0.52383 | 0.003292 | 0.10783 | -0.1444 | -0.25223 |  | 300.778 | -0.13192
17 | YES | 600.318 | 0.52383 | 0.003292 | -0.018173 | -0.1444 | -0.12623 |  | 250.418 | -0.12007
17 | YES | 401.395 | 1.0774 | 0.0023309 | -0.058636 | -0.15039 | -0.091749 |  | 200.911 | -0.2133
17 | YES | 300.778 | 1.9071 | 0.0024103 | -0.13192 | -0.15341 | -0.021496 |  | 175.849 | -0.15965
17 | YES | 250.418 | 2.5519 | 0.0033954 | -0.12007 | -0.15493 | -0.034863 |  | 149.91 | -0.19268
17 | YES | 200.911 | 3.4487 | 0.0024619 | -0.2133 | -0.15642 | 0.056879 |  | 124.964 | -0.22512
17 | YES | 175.849 | 3.9304 | 0.0051943 | -0.15965 | -0.15717 | 0.0024725 |  | 99.8412 | -0.29727
17 | YES | 149.91 | 4.3963 | 0.0051059 | -0.19268 | -0.15795 | 0.034725 |  | 75.0487 | -0.26644
17 | YES | 124.964 | 4.8779 | 0.0035967 | -0.22512 | -0.1587 | 0.066412 |  | 49.7955 | -0.27967
17 | YES | 99.8412 | 5.8287 | 0.01541 | -0.29727 | -0.15946 | 0.13781 |  | 25.8528 | -0.2655
17 | YES | 75.0487 | 6.6846 | 0.0044119 | -0.26644 | -0.16021 | 0.10623 |  | 10.4646 | -0.37018
17 | YES | 49.7955 | 6.6043 | 0.015588 | -0.27967 | -0.16097 | 0.1187 |  | 5.3326 | -0.24274
17 | YES | 25.8528 | 6.5625 | 0.010572 | -0.2655 | -0.16169 | 0.10381 |  | 4.8354 | -0.28447
17 | YES | 10.4646 | 6.5178 | 0.0081332 | -0.37018 | -0.16215 | 0.20803 |  | 4316.66 | -0.13845
17 | YES | 5.3326 | 6.5223 | 0.0059579 | -0.24274 | -0.1623 | 0.080431 |  | 4000.97 | -0.14688
18 | YES | 4.8354 | 6.5365 | 0.0033914 | -0.28447 | -0.16232 | 0.12215 |  | 3500.19 | -0.14059
19 | YES | 4316.66 | 3.0625 | 0.0073575 | -0.13845 | -0.032568 | 0.10588 |  | 2999.35 | -0.10606
19 | YES | 4000.97 | 3.0161 | 0.0086493 | -0.14688 | -0.042068 | 0.10481 |  | 2498.97 | -0.059289
19 | YES | 3500.19 | 2.8164 | 0.0036599 | -0.14059 | -0.057137 | 0.08345 |  | 1999.87 | -0.030418
19 | YES | 2999.35 | 2.4589 | 0.0047492 | -0.10606 | -0.072208 | 0.03385 |  | 1501.11 | 0.01738
19 | YES | 2498.97 | 1.9097 | 0.004458 | -0.059289 | -0.087266 | -0.027976 |  | 1250.88 | 0.0090992
19 | YES | 1999.87 | 1.3096 | 0.0039175 | -0.030418 | -0.10228 | -0.071866 |  | 1001.36 | -0.0059504
19 | YES | 1501.11 | 0.64938 | 0.0032666 | 0.01738 | -0.11729 | -0.13467 |  | 799.943 | 0.00069419
19 | YES | 1250.88 | 0.4281 | 0.0028472 | 0.0090992 | -0.12482 | -0.13392 |  | 600.641 | -0.0010578
19 | YES | 1001.36 | 0.38405 | 0.0022835 | -0.0059504 | -0.13233 | -0.12638 |  | 400.367 | -0.039058
19 | YES | 799.943 | 0.43269 | 0.0025028 | 0.00069419 | -0.13839 | -0.13909 |  | 301.185 | -0.085058
19 | YES | 600.641 | 0.49294 | 0.0033821 | -0.0010578 | -0.14439 | -0.14333 |  | 250.879 | -0.076231
19 | YES | 400.367 | 0.80394 | 0.0031919 | -0.039058 | -0.15042 | -0.11136 |  | 200.507 | -0.12306
19 | YES | 301.185 | 1.2659 | 0.0036224 | -0.085058 | -0.1534 | -0.068343 |  | 174.492 | -0.14162
19 | YES | 250.879 | 1.6588 | 0.0042284 | -0.076231 | -0.15491 | -0.078683 |  | 151.221 | -0.12239
19 | YES | 200.507 | 2.4929 | 0.004011 | -0.12306 | -0.15643 | -0.033373 |  | 125.937 | -0.22392
19 | YES | 174.492 | 3.0154 | 0.0056939 | -0.14162 | -0.15721 | -0.015593 |  | 125.937 | -0.24592
19 | YES | 151.221 | 3.5806 | 0.026746 | -0.12239 | -0.15791 | -0.035525 |  | 75.471 | -0.31913
19 | YES | 125.937 | 4.2171 | 0.0090026 | -0.22392 | -0.15867 | 0.065243 |  | 51.0093 | -0.31003
19 | YES | 125.937 | 4.2171 | 0.0090026 | -0.24592 | -0.15867 | 0.087243 |  | 26.1199 | -0.27302
19 | YES | 75.471 | 6.6079 | 0.0043759 | -0.31913 | -0.16019 | 0.15894 |  | 10.9632 | -0.25326
19 | YES | 51.0093 | 6.671 | 0.01042 | -0.31003 | -0.16093 | 0.1491 |  | 5.7636 | -0.26866
19 | YES | 26.1199 | 6.7 | 0.010029 | -0.27302 | -0.16168 | 0.11134
19 | YES | 10.9632 | 6.7037 | 0.0038741 | -0.25326 | -0.16213 | 0.091122
19 | YES | 5.7636 | 6.7063 | 0.0022641 | -0.26866 | -0.16229 | 0.10637
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (3)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0035989 |  | X | CALC Y
m= | 0.0017382 | 0.0022516 |  | Std Dev of Resid |  | 0.17392 |  | 1 | -0.17541
b= | -0.17714 | 0.028699 |  | Resid Sum of Squares |  | 4.9907 |  | 19 | -0.14412
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 167
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 218.403 | 0.10197 | 0.0017817 | 0.10197 | -0.17541 | -0.27738 |  | 1 | 0.10197
1 | NO | 200.531 | 0.088653 | 0.0016917 | 0.088653 | -0.17541 | -0.26406 |  | 1 | 0.088653
1 | NO | 175.42 | 0.08076 | 0.0010249 | 0.08076 | -0.17541 | -0.25617 |  | 1 | 0.08076
1 | NO | 150.458 | 0.07595 | 0.0018967 | 0.07595 | -0.17541 | -0.25136 |  | 1 | 0.07595
1 | NO | 125.596 | 0.072215 | 0.0016442 | 0.010215 | -0.17541 | -0.18562 |  | 1 | 0.010215
1 | NO | 100.522 | 1.4962 | 0.008499 | 0.17716 | -0.17541 | -0.35256 |  | 1 | 0.17716
1 | YES | 75.0778 | 3.6331 | 0.0030488 | -0.0049338 | -0.17541 | -0.17047 |  | 1 | -0.0049338
1 | YES | 50.4725 | 4.5052 | 0.0047135 | -0.13576 | -0.17541 | -0.039646 |  | 1 | -0.13576
1 | YES | 25.1291 | 4.7946 | 0.0042427 | -0.19345 | -0.17541 | 0.018048 |  | 1 | -0.19345
1 | YES | 10.0099 | 5.5954 | 0.010131 | 0.079356 | -0.17541 | -0.25476 |  | 1 | 0.079356
1 | YES | 5.1566 | 6.2031 | 0.21295 | -1.5569 | -0.17541 | 1.3815 |  | 1 | -1.5569
2 | YES | 101.269 | 3.1153 | 0.0048921 | -0.11473 | -0.17367 | -0.05894 |  | 2 | -0.11473
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.12076 | -0.17367 | -0.052903 |  | 2 | -0.12076
2 | YES | 74.8171 | 3.6672 | 0.0068631 | -0.13076 | -0.17367 | -0.042903 |  | 2 | -0.13076
2 | YES | 50.9797 | 4.5985 | 0.0056717 | -0.14654 | -0.17367 | -0.02713 |  | 2 | -0.14654
2 | YES | 24.8055 | 5.8906 | 0.0044437 | -0.21443 | -0.17367 | 0.040763 |  | 2 | -0.21443
2 | YES | 10.171 | 6.8152 | 0.009922 | -0.29277 | -0.17367 | 0.1191 |  | 2 | -0.29277
2 | YES | 5.2564 | 7.0009 | 0.017248 | -0.33307 | -0.17367 | 0.1594 |  | 2 | -0.33307
3 | YES | 1250.61 | 0.42322 | 0.0035786 | -0.055781 | -0.17193 | -0.11615 |  | 3 | -0.055781
3 | YES | 1001.1 | 0.26691 | 0.0026646 | -0.034091 | -0.17193 | -0.13784 |  | 3 | -0.034091
3 | YES | 800.671 | 0.22779 | 0.0026706 | -0.029207 | -0.17193 | -0.14272 |  | 3 | -0.029207
3 | YES | 600.701 | 0.33769 | 0.0025028 | -0.071306 | -0.17193 | -0.10062 |  | 3 | -0.071306
3 | YES | 400.813 | 0.83171 | 0.0019682 | -0.057289 | -0.17193 | -0.11464 |  | 3 | -0.057289
3 | YES | 300.322 | 1.3764 | 0.0026626 | -0.085612 | -0.17193 | -0.086318 |  | 3 | -0.085612
3 | YES | 250.571 | 1.799 | 0.0017464 | -0.109 | -0.17193 | -0.062929 |  | 3 | -0.109
3 | YES | 200.192 | 2.1873 | 0.0020919 | -0.15166 | -0.17193 | -0.020268 |  | 3 | -0.15166
3 | YES | 175.145 | 2.5945 | 0.0028461 | -0.12946 | -0.17193 | -0.042466 |  | 3 | -0.12946
3 | YES | 150.231 | 2.8032 | 0.0020788 | -0.12675 | -0.17193 | -0.045177 |  | 3 | -0.12675
3 | YES | 125.157 | 3.361 | 0.0032146 | -0.145 | -0.17193 | -0.026929 |  | 3 | -0.145
3 | YES | 100.168 | 4.2772 | 0.0066675 | -0.14783 | -0.17193 | -0.024094 |  | 3 | -0.14783
3 | YES | 75.274 | 5.5548 | 0.010038 | -0.22117 | -0.17193 | 0.049245 |  | 3 | -0.22117
3 | YES | 50.0366 | 6.0704 | 0.0098299 | -0.22559 | -0.17193 | 0.053658 |  | 3 | -0.22559
3 | YES | 25.2465 | 6.4055 | 0.0068873 | -0.22652 | -0.17193 | 0.054592 |  | 3 | -0.22652
3 | YES | 10.5579 | 6.3624 | 0.0088479 | -0.24163 | -0.17193 | 0.069699 |  | 3 | -0.24163
3 | YES | 5.6538 | 6.3019 | 0.011329 | -0.21912 | -0.17193 | 0.047187 |  | 3 | -0.21912
4 | YES | 5.3088 | 6.2546 | 0.0071529 | -0.26943 | -0.17019 | 0.099239 |  | 4 | -0.26943
5 | YES | 5.4542 | 6.4424 | 0.0023322 | -0.27061 | -0.16845 | 0.10216 |  | 5 | -0.27061
6 | YES | 5.8932 | 6.4976 | 0.016952 | -0.22142 | -0.16671 | 0.054707 |  | 6 | -0.22142
7 | NO | 1499.59 | 0.52926 | 0.010928 | -0.19874 | -0.16498 | 0.033764 |  | 7 | -0.19874
7 | NO | 1249.7 | 0.29893 | 0.0049936 | -0.15507 | -0.16498 | -0.0099023 |  | 7 | -0.15507
7 | NO | 1000.11 | 0.14787 | 0.0095638 | -0.14113 | -0.16498 | -0.023844 |  | 7 | -0.14113
7 | NO | 799.938 | 0.12507 | 0.003645 | -0.13193 | -0.16498 | -0.03305 |  | 7 | -0.13193
7 | NO | 600.069 | 0.26533 | 0.0057277 | -0.14267 | -0.16498 | -0.022307 |  | 7 | -0.14267
7 | NO | 401.436 | 0.76881 | 0.0048275 | -0.14219 | -0.16498 | -0.022786 |  | 7 | -0.14219
7 | NO | 300.308 | 1.3372 | 0.0098088 | -0.17777 | -0.16498 | 0.012793 |  | 7 | -0.17777
7 | NO | 250.267 | 1.7647 | 0.006877 | -0.17034 | -0.16498 | 0.0053627 |  | 7 | -0.17034
7 | NO | 200.949 | 2.2732 | 0.0036794 | -0.19575 | -0.16498 | 0.030776 |  | 7 | -0.19575
7 | NO | 175.622 | 2.4642 | 0.0047875 | -0.19279 | -0.16498 | 0.027809 |  | 7 | -0.19279
7 | NO | 150.306 | 2.9429 | 0.0089897 | -0.31513 | -0.16498 | 0.15016 |  | 7 | -0.31513
7 | NO | 124.541 | 3.1468 | 0.0061658 | -1.0352 | -0.16498 | 0.8702 |  | 7 | -1.0352
7 | NO | 100.118 | 4.5387 | 0.017814 | 0.036661 | -0.16498 | -0.20164 |  | 7 | 0.036661
7 | NO | 75.4503 | 5.8218 | 0.013968 | -0.3332 | -0.16498 | 0.16822 |  | 7 | -0.3332
7 | NO | 49.9647 | 6.4887 | 0.0076196 | -0.50426 | -0.16498 | 0.33928 |  | 7 | -0.50426
7 | NO | 25.5491 | 6.3986 | 0.0049993 | -0.34341 | -0.16498 | 0.17843 |  | 7 | -0.34341
7 | NO | 10.936 | 6.3303 | 0.0068274 | -0.34669 | -0.16498 | 0.18172 |  | 7 | -0.34669
7 | NO | 5.573 | 6.3588 | 0.015454 | -0.32422 | -0.16498 | 0.15925 |  | 7 | -0.32422
8 | YES | 5.2317 | 6.2715 | 0.006222 | -0.35153 | -0.16324 | 0.1883 |  | 8 | -0.35153
9 | YES | 4.7448 | 6.4118 | 0.0098755 | -0.25124 | -0.1615 | 0.089738 |  | 9 | -0.25124
10 | YES | 5.7805 | 6.4089 | 0.0076651 | -0.31712 | -0.15976 | 0.15735 |  | 10 | -0.31712
11 | YES | 3275.5 | 2.2131 | 0.0049637 | -0.09786 | -0.15802 | -0.060164 |  | 11 | -0.09786
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.15802 | -0.069596 |  | 11 | -0.088427
11 | YES | 3000.46 | 2.1816 | 0.0042614 | -0.088427 | -0.15802 | -0.069596 |  | 11 | -0.088427
11 | YES | 2500.36 | 1.787 | 0.0055976 | -0.066 | -0.15802 | -0.092023 |  | 11 | -0.066
11 | YES | 2000.62 | 1.2673 | 0.0039445 | -0.036661 | -0.15802 | -0.12136 |  | 11 | -0.036661
11 | YES | 1500.02 | 0.66384 | 0.0030056 | -0.004157 | -0.15802 | -0.15387 |  | 11 | -0.004157
11 | YES | 1249.97 | 0.3805 | 0.0031441 | -0.0034959 | -0.15802 | -0.15453 |  | 11 | -0.0034959
11 | YES | 1000.45 | 0.25207 | 0.003006 | 0.010074 | -0.15802 | -0.1681 |  | 11 | 0.010074
11 | YES | 800.106 | 0.24929 | 0.0026939 | 0.00028926 | -0.15802 | -0.15831 |  | 11 | 0.00028926
11 | YES | 600.603 | 0.44213 | 0.0029352 | -0.0028678 | -0.15802 | -0.15516 |  | 11 | -0.0028678
11 | YES | 400.012 | 0.97554 | 0.0024834 | -0.047463 | -0.15802 | -0.11056 |  | 11 | -0.047463
11 | YES | 300.556 | 1.2546 | 0.0017404 | -0.068421 | -0.15802 | -0.089602 |  | 11 | -0.068421
11 | YES | 250.721 | 1.4783 | 0.0017051 | -0.061711 | -0.15802 | -0.096313 |  | 11 | -0.061711
11 | YES | 200.4 | 1.6469 | 0.0017839 | -0.066132 | -0.15802 | -0.091891 |  | 11 | -0.066132
11 | YES | 175.636 | 1.8105 | 0.002266 | -0.073529 | -0.15802 | -0.084494 |  | 11 | -0.073529
11 | YES | 150.415 | 2.0813 | 0.0032177 | -0.079719 | -0.15802 | -0.078304 |  | 11 | -0.079719
11 | YES | 125.849 | 2.2659 | 0.0027683 | -0.092107 | -0.15802 | -0.065916 |  | 11 | -0.092107
11 | YES | 100.422 | 2.7438 | 0.0021603 | -0.12924 | -0.15802 | -0.028783 |  | 11 | -0.12924
11 | YES | 74.8778 | 4.629 | 0.014433 | -0.16099 | -0.15802 | 0.0029687 |  | 11 | -0.16099
11 | YES | 50.3333 | 6.135 | 0.0067311 | -0.22701 | -0.15802 | 0.068985 |  | 11 | -0.22701
11 | YES | 25.0889 | 6.5646 | 0.011212 | -0.29942 | -0.15802 | 0.1414 |  | 11 | -0.29942
11 | YES | 10.3182 | 6.4439 | 0.010426 | -0.27907 | -0.15802 | 0.12104 |  | 11 | -0.27907
11 | YES | 5.6855 | 6.4652 | 0.0029411 | -0.26282 | -0.15802 | 0.10479 |  | 11 | -0.26282
12 | YES | 2000.19 | 1.2571 | 0.0037998 | -0.067942 | -0.15629 | -0.088343 |  | 12 | -0.067942
12 | YES | 1499.41 | 0.58515 | 0.0034026 | -0.030851 | -0.15629 | -0.12543 |  | 12 | -0.030851
12 | YES | 998.622 | 0.25176 | 0.0021796 | -0.03124 | -0.15629 | -0.12505 |  | 12 | -0.03124
12 | YES | 800.39 | 0.32921 | 0.0026558 | -0.028785 | -0.15629 | -0.1275 |  | 12 | -0.028785
12 | YES | 600.518 | 0.5568 | 0.0023224 | -0.032198 | -0.15629 | -0.12409 |  | 12 | -0.032198
12 | YES | 400.245 | 0.99347 | 0.0021451 | -0.097529 | -0.15629 | -0.058756 |  | 12 | -0.097529
12 | YES | 201.031 | 2.9307 | 0.006522 | -0.19628 | -0.15629 | 0.039996 |  | 12 | -0.19628
12 | YES | 100.738 | 5.574 | 0.0090563 | -0.30702 | -0.15629 | 0.15073 |  | 12 | -0.30702
12 | YES | 75.1238 | 6.1798 | 0.025755 | -0.28821 | -0.15629 | 0.13192 |  | 12 | -0.28821
12 | YES | 30.1212 | 6.4295 | 0.0045019 | -0.26547 | -0.15629 | 0.10919 |  | 12 | -0.26547
12 | YES | 5.1268 | 6.415 | 0.0083676 | -0.26302 | -0.15629 | 0.10673 |  | 12 | -0.26302
13 | YES | 5.3761 | 6.4109 | 0.016095 | -0.27111 | -0.15455 | 0.11656 |  | 13 | -0.27111
14 | YES | 5.4002 | 6.4124 | 0.010385 | -0.26159 | -0.15281 | 0.10878 |  | 14 | -0.26159
15 | YES | 3681.03 | 2.6256 | 0.0054848 | -0.18145 | -0.15107 | 0.030384 |  | 15 | -0.18145
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15655 | -0.15107 | 0.0054836 |  | 15 | -0.15655
15 | YES | 3500.68 | 2.5804 | 0.0037371 | -0.15955 | -0.15107 | 0.0084836 |  | 15 | -0.15955
15 | YES | 2999.99 | 2.375 | 0.0039569 | -0.11004 | -0.15107 | -0.041029 |  | 15 | -0.11004
15 | YES | 2500.88 | 1.8804 | 0.0068361 | -0.10855 | -0.15107 | -0.042516 |  | 15 | -0.10855
15 | YES | 2000.65 | 1.2312 | 0.0025215 | -0.060777 | -0.15107 | -0.090294 |  | 15 | -0.060777
15 | YES | 1499.86 | 0.5678 | 0.0028683 | -0.016198 | -0.15107 | -0.13487 |  | 15 | -0.016198
15 | YES | 1250.04 | 0.33898 | 0.0028707 | 0.0019752 | -0.15107 | -0.15305 |  | 15 | 0.0019752
15 | YES | 1000.74 | 0.29317 | 0.0034992 | -0.0088264 | -0.15107 | -0.14224 |  | 15 | -0.0088264
15 | YES | 799.17 | 0.37046 | 0.0021492 | -0.012539 | -0.15107 | -0.13853 |  | 15 | -0.012539
15 | YES | 599.933 | 0.63053 | 0.002021 | -0.041471 | -0.15107 | -0.1096 |  | 15 | -0.041471
15 | YES | 399.65 | 1.6482 | 0.0024393 | -0.080843 | -0.15107 | -0.070227 |  | 15 | -0.080843
15 | YES | 300 | 2.4146 | 0.0036073 | -0.12442 | -0.15107 | -0.026648 |  | 15 | -0.12442
15 | YES | 249.213 | 3.1558 | 0.0036844 | -0.14422 | -0.15107 | -0.0068474 |  | 15 | -0.14422
15 | YES | 201.385 | 3.9694 | 0.0087082 | -0.20264 | -0.15107 | 0.051566 |  | 15 | -0.20264
15 | YES | 176.241 | 4.2895 | 0.0066959 | -0.19352 | -0.15107 | 0.042451 |  | 15 | -0.19352
15 | YES | 150.259 | 4.6717 | 0.0052742 | -0.19227 | -0.15107 | 0.041203 |  | 15 | -0.19227
15 | YES | 126.022 | 5.1518 | 0.0054625 | -0.21019 | -0.15107 | 0.05912 |  | 15 | -0.21019
15 | YES | 100.338 | 5.8719 | 0.0099315 | -0.25209 | -0.15107 | 0.10102 |  | 15 | -0.25209
15 | YES | 74.8615 | 6.6217 | 0.0059463 | -0.27326 | -0.15107 | 0.12219 |  | 15 | -0.27326
15 | YES | 51.62 | 6.7685 | 0.0088243 | -0.26449 | -0.15107 | 0.11342 |  | 15 | -0.26449
15 | YES | 25.8385 | 6.4155 | 0.012656 | -0.25953 | -0.15107 | 0.10846 |  | 15 | -0.25953
15 | YES | 10.6251 | 6.4084 | 0.012711 | -0.2476 | -0.15107 | 0.096525 |  | 15 | -0.2476
15 | YES | 5.959 | 6.4214 | 0.016474 | -0.23364 | -0.15107 | 0.082566 |  | 15 | -0.23364
16 | YES | 5.6584 | 6.4384 | 0.0094793 | -0.28461 | -0.14933 | 0.13528 |  | 16 | -0.28461
17 | YES | 4023 | 2.8556 | 0.00702 | -0.25743 | -0.14759 | 0.10984 |  | 17 | -0.25743
17 | YES | 3500.03 | 2.7524 | 0.0040812 | -0.16961 | -0.14759 | 0.022017 |  | 17 | -0.16961
17 | YES | 3000.73 | 2.4194 | 0.0057594 | -0.12455 | -0.14759 | -0.023049 |  | 17 | -0.12455
17 | YES | 2498.77 | 1.8721 | 0.0036797 | -0.062901 | -0.14759 | -0.084693 |  | 17 | -0.062901
17 | YES | 2001.65 | 1.2756 | 0.0036934 | -0.044397 | -0.14759 | -0.1032 |  | 17 | -0.044397
17 | YES | 1499.84 | 0.584 | 0.0024461 | 0.007 | -0.14759 | -0.15459 |  | 17 | 0.007
17 | YES | 1250.74 | 0.37144 | 0.0029718 | 0.023438 | -0.14759 | -0.17103 |  | 17 | 0.023438
17 | YES | 600.318 | 0.52383 | 0.003292 | 0.10783 | -0.14759 | -0.25542 |  | 17 | 0.10783
17 | YES | 600.318 | 0.52383 | 0.003292 | -0.018173 | -0.14759 | -0.12942 |  | 17 | -0.018173
17 | YES | 401.395 | 1.0774 | 0.0023309 | -0.058636 | -0.14759 | -0.088958 |  | 17 | -0.058636
17 | YES | 300.778 | 1.9071 | 0.0024103 | -0.13192 | -0.14759 | -0.015677 |  | 17 | -0.13192
17 | YES | 250.418 | 2.5519 | 0.0033954 | -0.12007 | -0.14759 | -0.027528 |  | 17 | -0.12007
17 | YES | 200.911 | 3.4487 | 0.0024619 | -0.2133 | -0.14759 | 0.065703 |  | 17 | -0.2133
17 | YES | 175.849 | 3.9304 | 0.0051943 | -0.15965 | -0.14759 | 0.012051 |  | 17 | -0.15965
17 | YES | 149.91 | 4.3963 | 0.0051059 | -0.19268 | -0.14759 | 0.045084 |  | 17 | -0.19268
17 | YES | 124.964 | 4.8779 | 0.0035967 | -0.22512 | -0.14759 | 0.077522 |  | 17 | -0.22512
17 | YES | 99.8412 | 5.8287 | 0.01541 | -0.29727 | -0.14759 | 0.14967 |  | 17 | -0.29727
17 | YES | 75.0487 | 6.6846 | 0.0044119 | -0.26644 | -0.14759 | 0.11884 |  | 17 | -0.26644
17 | YES | 49.7955 | 6.6043 | 0.015588 | -0.27967 | -0.14759 | 0.13208 |  | 17 | -0.27967
17 | YES | 25.8528 | 6.5625 | 0.010572 | -0.2655 | -0.14759 | 0.1179 |  | 17 | -0.2655
17 | YES | 10.4646 | 6.5178 | 0.0081332 | -0.37018 | -0.14759 | 0.22259 |  | 17 | -0.37018
17 | YES | 5.3326 | 6.5223 | 0.0059579 | -0.24274 | -0.14759 | 0.095141 |  | 17 | -0.24274
18 | YES | 4.8354 | 6.5365 | 0.0033914 | -0.28447 | -0.14586 | 0.13862 |  | 18 | -0.28447
19 | YES | 4316.66 | 3.0625 | 0.0073575 | -0.13845 | -0.14412 | -0.0056715 |  | 19 | -0.13845
19 | YES | 4000.97 | 3.0161 | 0.0086493 | -0.14688 | -0.14412 | 0.0027585 |  | 19 | -0.14688
19 | YES | 3500.19 | 2.8164 | 0.0036599 | -0.14059 | -0.14412 | -0.0035305 |  | 19 | -0.14059
19 | YES | 2999.35 | 2.4589 | 0.0047492 | -0.10606 | -0.14412 | -0.038059 |  | 19 | -0.10606
19 | YES | 2498.97 | 1.9097 | 0.004458 | -0.059289 | -0.14412 | -0.084828 |  | 19 | -0.059289
19 | YES | 1999.87 | 1.3096 | 0.0039175 | -0.030418 | -0.14412 | -0.1137 |  | 19 | -0.030418
19 | YES | 1501.11 | 0.64938 | 0.0032666 | 0.01738 | -0.14412 | -0.1615 |  | 19 | 0.01738
19 | YES | 1250.88 | 0.4281 | 0.0028472 | 0.0090992 | -0.14412 | -0.15322 |  | 19 | 0.0090992
19 | YES | 1001.36 | 0.38405 | 0.0022835 | -0.0059504 | -0.14412 | -0.13817 |  | 19 | -0.0059504
19 | YES | 799.943 | 0.43269 | 0.0025028 | 0.00069419 | -0.14412 | -0.14481 |  | 19 | 0.00069419
19 | YES | 600.641 | 0.49294 | 0.0033821 | -0.0010578 | -0.14412 | -0.14306 |  | 19 | -0.0010578
19 | YES | 400.367 | 0.80394 | 0.0031919 | -0.039058 | -0.14412 | -0.10506 |  | 19 | -0.039058
19 | YES | 301.185 | 1.2659 | 0.0036224 | -0.085058 | -0.14412 | -0.05906 |  | 19 | -0.085058
19 | YES | 250.879 | 1.6588 | 0.0042284 | -0.076231 | -0.14412 | -0.067886 |  | 19 | -0.076231
19 | YES | 200.507 | 2.4929 | 0.004011 | -0.12306 | -0.14412 | -0.021059 |  | 19 | -0.12306
19 | YES | 174.492 | 3.0154 | 0.0056939 | -0.14162 | -0.14412 | -0.0024975 |  | 19 | -0.14162
19 | YES | 151.221 | 3.5806 | 0.026746 | -0.12239 | -0.14412 | -0.021728 |  | 19 | -0.12239
19 | YES | 125.937 | 4.2171 | 0.0090026 | -0.22392 | -0.14412 | 0.0798 |  | 19 | -0.22392
19 | YES | 125.937 | 4.2171 | 0.0090026 | -0.24592 | -0.14412 | 0.1018 |  | 19 | -0.24592
19 | YES | 75.471 | 6.6079 | 0.0043759 | -0.31913 | -0.14412 | 0.17501 |  | 19 | -0.31913
19 | YES | 51.0093 | 6.671 | 0.01042 | -0.31003 | -0.14412 | 0.16592 |  | 19 | -0.31003
19 | YES | 26.1199 | 6.7 | 0.010029 | -0.27302 | -0.14412 | 0.1289 |  | 19 | -0.27302
19 | YES | 10.9632 | 6.7037 | 0.0038741 | -0.25326 | -0.14412 | 0.10914 |  | 19 | -0.25326
19 | YES | 5.7636 | 6.7063 | 0.0022641 | -0.26866 | -0.14412 | 0.12454 |  | 19 | -0.26866
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.3 | Run at 2010/08/28 17:02:33
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX_BOT | DOX_CTD |  | DOX_CTD - DOX_BOT
Average Size: | 2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX_BOT | DOX_CTD | DOX_CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 218.403 | 218.403 | 0.028684 | 0.00021362 | 0 | 0.10197 | 0.0017817 | 0.10197
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 200.531 | 200.531 | 0.02978 | 0.00048828 | 0 | 0.088653 | 0.0016917 | 0.088653
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 175.42 | 175.42 | 0.027039 | 0 | 0 | 0.08076 | 0.0010249 | 0.08076
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 150.458 | 150.458 | 0.032989 | 0.00010681 | 0 | 0.07595 | 0.0018967 | 0.07595
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 125.596 | 125.596 | 0.031595 | 0.00046539 | 0.062 | 0.072215 | 0.0016442 | 0.010215
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 100.522 | 100.522 | 0.034545 | -0.00041962 | 1.319 | 1.4962 | 0.008499 | 0.17716
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 75.078 | 75.0778 | 0.034015 | -0.00022888 | 3.638 | 3.6331 | 0.0030488 | -0.0049338
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 50.472 | 50.4725 | 0.032148 | 0.00048065 | 4.641 | 4.5052 | 0.0047135 | -0.13576
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 25.129 | 25.1291 | 0.028881 | 8.3923e-05 | 4.988 | 4.7946 | 0.0042427 | -0.19345
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 10.01 | 10.0099 | 0.041351 | -5.8174e-05 | 5.516 | 5.5954 | 0.010131 | 0.079356
2010-13-0001.mrg | 2010-13-0001.sam | 121 | 5.157 | 5.1566 | 0.030569 | -0.00040483 | 7.76 | 6.2031 | 0.21295 | -1.5569
2010-13-0005.mrg | 2010-13-0005.sam | 121 | 101.269 | 101.269 | 0.22782 | 3.0518e-05 | 3.23 | 3.1153 | 0.0048921 | -0.11473
2010-13-0005.mrg | 2010-13-0005.sam | 242 | 74.702 | 74.8171 | 0.18509 | 0.11507 | 3.788 | 3.6672 | 0.0068631 | -0.12076
2010-13-0005.mrg | 2010-13-0005.sam | 242 | 74.932 | 74.8171 | 0.18509 | -0.11493 | 3.798 | 3.6672 | 0.0068631 | -0.13076
2010-13-0005.mrg | 2010-13-0005.sam | 121 | 50.98 | 50.9797 | 0.25062 | -0.00033951 | 4.745 | 4.5985 | 0.0056717 | -0.14654
2010-13-0005.mrg | 2010-13-0005.sam | 121 | 24.806 | 24.8055 | 0.19568 | -0.00047874 | 6.105 | 5.8906 | 0.0044437 | -0.21443
2010-13-0005.mrg | 2010-13-0005.sam | 121 | 10.171 | 10.171 | 0.062367 | 0 | 7.108 | 6.8152 | 0.009922 | -0.29277
2010-13-0005.mrg | 2010-13-0005.sam | 121 | 5.256 | 5.2564 | 0.11315 | 0.00039673 | 7.334 | 7.0009 | 0.017248 | -0.33307
2010-13-0012.mrg | 2010-13-0012.sam | 187 | 1250.56 | 1250.61 | 0.29304 | 0.053955 | 0.479 | 0.42322 | 0.0035786 | -0.055781
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 1001.1 | 1001.1 | 0.52506 | -0.00012207 | 0.301 | 0.26691 | 0.0026646 | -0.034091
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 800.671 | 800.671 | 0.36334 | -0.00024414 | 0.257 | 0.22779 | 0.0026706 | -0.029207
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 600.701 | 600.701 | 0.086834 | -0.00042725 | 0.409 | 0.33769 | 0.0025028 | -0.071306
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 400.813 | 400.813 | 0.14586 | -0.00036621 | 0.889 | 0.83171 | 0.0019682 | -0.057289
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 300.322 | 300.322 | 0.13772 | 0.00012207 | 1.462 | 1.3764 | 0.0026626 | -0.085612
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 250.571 | 250.571 | 0.056098 | -0.00032043 | 1.908 | 1.799 | 0.0017464 | -0.109
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 200.192 | 200.192 | 0.13648 | 0.00038147 | 2.339 | 2.1873 | 0.0020919 | -0.15166
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 175.145 | 175.145 | 0.068834 | 0.00015259 | 2.724 | 2.5945 | 0.0028461 | -0.12946
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 150.231 | 150.231 | 0.26176 | -0.00010681 | 2.93 | 2.8032 | 0.0020788 | -0.12675
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 125.157 | 125.157 | 0.24598 | 0.00048065 | 3.506 | 3.361 | 0.0032146 | -0.145
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 100.168 | 100.168 | 0.06612 | 1.5259e-05 | 4.425 | 4.2772 | 0.0066675 | -0.14783
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 75.274 | 75.274 | 0.26833 | -1.5259e-05 | 5.776 | 5.5548 | 0.010038 | -0.22117
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 50.037 | 50.0366 | 0.063782 | -0.00040436 | 6.296 | 6.0704 | 0.0098299 | -0.22559
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 25.247 | 25.2465 | 0.13408 | -0.00048828 | 6.632 | 6.4055 | 0.0068873 | -0.22652
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 10.558 | 10.5579 | 0.07077 | -9.9182e-05 | 6.604 | 6.3624 | 0.0088479 | -0.24163
2010-13-0012.mrg | 2010-13-0012.sam | 121 | 5.654 | 5.6538 | 0.18114 | -0.00017357 | 6.521 | 6.3019 | 0.011329 | -0.21912
2010-13-0014.mrg | 2010-13-0014.sam | 121 | 5.309 | 5.3088 | 0.13399 | -0.00023985 | 6.524 | 6.2546 | 0.0071529 | -0.26943
2010-13-0015.mrg | 2010-13-0015.sam | 121 | 5.454 | 5.4542 | 0.10765 | 0.00014877 | 6.713 | 6.4424 | 0.0023322 | -0.27061
2010-13-0016.mrg | 2010-13-0016.sam | 121 | 5.893 | 5.8932 | 0.21967 | 0.00023127 | 6.719 | 6.4976 | 0.016952 | -0.22142
2010-13-0018.mrg | 2010-13-0018.sam | 242 | 1499.63 | 1499.59 | 0.2204 | -0.035645 | 0.728 | 0.52926 | 0.010928 | -0.19874
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 1249.7 | 1249.7 | 0.22928 | -0.00024414 | 0.454 | 0.29893 | 0.0049936 | -0.15507
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 1000.11 | 1000.11 | 0.16072 | 0.00036621 | 0.289 | 0.14787 | 0.0095638 | -0.14113
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 799.938 | 799.938 | 0.27455 | 0 | 0.257 | 0.12507 | 0.003645 | -0.13193
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 600.069 | 600.069 | 0.18892 | 0.00048828 | 0.408 | 0.26533 | 0.0057277 | -0.14267
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 401.436 | 401.436 | 0.089219 | -0.00021362 | 0.911 | 0.76881 | 0.0048275 | -0.14219
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 300.308 | 300.308 | 0.15703 | -0.00018311 | 1.515 | 1.3372 | 0.0098088 | -0.17777
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 250.267 | 250.267 | 0.08014 | 9.1553e-05 | 1.935 | 1.7647 | 0.006877 | -0.17034
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 200.949 | 200.949 | 0.22979 | -0.00022888 | 2.469 | 2.2732 | 0.0036794 | -0.19575
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 175.622 | 175.622 | 0.15996 | 0.00041199 | 2.657 | 2.4642 | 0.0047875 | -0.19279
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 150.306 | 150.306 | 0.13349 | -0.00047302 | 3.258 | 2.9429 | 0.0089897 | -0.31513
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 124.541 | 124.541 | 0.093553 | 0.0003891 | 4.182 | 3.1468 | 0.0061658 | -1.0352
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 100.118 | 100.118 | 0.14693 | 2.2888e-05 | 4.502 | 4.5387 | 0.017814 | 0.036661
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 75.45 | 75.4503 | 0.085401 | 0.00028229 | 6.155 | 5.8218 | 0.013968 | -0.3332
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 49.965 | 49.9647 | 0.40376 | -0.00029755 | 6.993 | 6.4887 | 0.0076196 | -0.50426
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 25.549 | 25.5491 | 0.10177 | 0.00013161 | 6.742 | 6.3986 | 0.0049993 | -0.34341
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 10.936 | 10.936 | 0.14435 | -2.4796e-05 | 6.677 | 6.3303 | 0.0068274 | -0.34669
2010-13-0018.mrg | 2010-13-0018.sam | 121 | 5.573 | 5.573 | 0.054331 | 3.2902e-05 | 6.683 | 6.3588 | 0.015454 | -0.32422
2010-13-0019.mrg | 2010-13-0019.sam | 217 | 5.323 | 5.2317 | 0.193 | -0.091272 | 6.623 | 6.2715 | 0.006222 | -0.35153
2010-13-0020.mrg | 2010-13-0020.sam | 202 | 4.546 | 4.7448 | 0.20884 | 0.19879 | 6.663 | 6.4118 | 0.0098755 | -0.25124
2010-13-0021.mrg | 2010-13-0021.sam | 121 | 5.781 | 5.7805 | 0.12206 | -0.00044632 | 6.726 | 6.4089 | 0.0076651 | -0.31712
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 3275.5 | 3275.5 | 0.3063 | 0 | 2.311 | 2.2131 | 0.0049637 | -0.09786
2010-13-0025.mrg | 2010-13-0025.sam | 234 | 3000.49 | 3000.46 | 0.12851 | -0.030762 | 2.27 | 2.1816 | 0.0042614 | -0.088427
2010-13-0025.mrg | 2010-13-0025.sam | 234 | 3000.43 | 3000.46 | 0.12851 | 0.028076 | 2.27 | 2.1816 | 0.0042614 | -0.088427
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 2500.36 | 2500.36 | 0.34391 | -0.00024414 | 1.853 | 1.787 | 0.0055976 | -0.066
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 2000.62 | 2000.62 | 0.17096 | -0.00012207 | 1.304 | 1.2673 | 0.0039445 | -0.036661
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 1500.02 | 1500.02 | 0.17036 | 0.00036621 | 0.668 | 0.66384 | 0.0030056 | -0.004157
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 1249.97 | 1249.97 | 0.16059 | -0.00048828 | 0.384 | 0.3805 | 0.0031441 | -0.0034959
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 1000.45 | 1000.45 | 0.0654 | 0.00024414 | 0.242 | 0.25207 | 0.003006 | 0.010074
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 800.106 | 800.106 | 0.039062 | -0.00036621 | 0.249 | 0.24929 | 0.0026939 | 0.00028926
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 600.603 | 600.603 | 0.1325 | -0.00030518 | 0.445 | 0.44213 | 0.0029352 | -0.0028678
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 400.012 | 400.012 | 0.17319 | 0.00030518 | 1.023 | 0.97554 | 0.0024834 | -0.047463
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 300.556 | 300.556 | 0.15691 | 9.1553e-05 | 1.323 | 1.2546 | 0.0017404 | -0.068421
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 250.721 | 250.721 | 0.11416 | 9.1553e-05 | 1.54 | 1.4783 | 0.0017051 | -0.061711
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 200.4 | 200.4 | 0.043027 | 0 | 1.713 | 1.6469 | 0.0017839 | -0.066132
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 175.636 | 175.636 | 0.19148 | -0.0004425 | 1.884 | 1.8105 | 0.002266 | -0.073529
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 150.415 | 150.415 | 0.1375 | 6.1035e-05 | 2.161 | 2.0813 | 0.0032177 | -0.079719
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 125.849 | 125.849 | 0.21278 | 0.00032043 | 2.358 | 2.2659 | 0.0027683 | -0.092107
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 100.422 | 100.422 | 0.073248 | -0.00045776 | 2.873 | 2.7438 | 0.0021603 | -0.12924
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 74.878 | 74.8778 | 0.088442 | -0.00023651 | 4.79 | 4.629 | 0.014433 | -0.16099
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 50.333 | 50.3333 | 0.18156 | 0.0003128 | 6.362 | 6.135 | 0.0067311 | -0.22701
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 25.089 | 25.0889 | 0.1216 | -0.00014114 | 6.864 | 6.5646 | 0.011212 | -0.29942
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 10.318 | 10.3182 | 0.071946 | 0.00024796 | 6.723 | 6.4439 | 0.010426 | -0.27907
2010-13-0025.mrg | 2010-13-0025.sam | 121 | 5.685 | 5.6855 | 0.15547 | 0.00046301 | 6.728 | 6.4652 | 0.0029411 | -0.26282
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 2000.19 | 2000.19 | 0.21731 | 0.00036621 | 1.325 | 1.2571 | 0.0037998 | -0.067942
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 1499.41 | 1499.41 | 0.26251 | 0.00024414 | 0.616 | 0.58515 | 0.0034026 | -0.030851
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 998.622 | 998.622 | 0.1412 | 6.1035e-05 | 0.283 | 0.25176 | 0.0021796 | -0.03124
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 800.39 | 800.39 | 0.1069 | -0.00018311 | 0.358 | 0.32921 | 0.0026558 | -0.028785
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 600.518 | 600.518 | 0.1773 | 0.00012207 | 0.589 | 0.5568 | 0.0023224 | -0.032198
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 400.245 | 400.245 | 0.24251 | 0.00027466 | 1.091 | 0.99347 | 0.0021451 | -0.097529
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 201.031 | 201.031 | 0.12102 | 0.00045776 | 3.127 | 2.9307 | 0.006522 | -0.19628
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 100.738 | 100.738 | 0.21452 | -0.0003891 | 5.881 | 5.574 | 0.0090563 | -0.30702
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 75.124 | 75.1238 | 0.069661 | -0.00022125 | 6.468 | 6.1798 | 0.025755 | -0.28821
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 30.121 | 30.1212 | 0.088549 | 0.00019836 | 6.695 | 6.4295 | 0.0045019 | -0.26547
2010-13-0028.mrg | 2010-13-0028.sam | 121 | 5.127 | 5.1268 | 0.19344 | -0.00018168 | 6.678 | 6.415 | 0.0083676 | -0.26302
2010-13-0030.mrg | 2010-13-0030.sam | 121 | 5.376 | 5.3761 | 0.59046 | 0.00014067 | 6.682 | 6.4109 | 0.016095 | -0.27111
2010-13-0031.mrg | 2010-13-0031.sam | 121 | 5.4 | 5.4002 | 0.43798 | 0.00020647 | 6.674 | 6.4124 | 0.010385 | -0.26159
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 3681.02 | 3681.03 | 0.27907 | 0.00048828 | 2.807 | 2.6256 | 0.0054848 | -0.18145
2010-13-0035.mrg | 2010-13-0035.sam | 242 | 3500.52 | 3500.68 | 0.33907 | 0.15454 | 2.737 | 2.5804 | 0.0037371 | -0.15655
2010-13-0035.mrg | 2010-13-0035.sam | 242 | 3500.83 | 3500.68 | 0.33907 | -0.15454 | 2.74 | 2.5804 | 0.0037371 | -0.15955
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 2999.99 | 2999.99 | 0.80952 | 0.00048828 | 2.485 | 2.375 | 0.0039569 | -0.11004
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 2500.88 | 2500.88 | 0.5557 | 0.00024414 | 1.989 | 1.8804 | 0.0068361 | -0.10855
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 2000.65 | 2000.65 | 0.57936 | 0.00036621 | 1.292 | 1.2312 | 0.0025215 | -0.060777
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 1499.86 | 1499.86 | 0.34505 | -0.00048828 | 0.584 | 0.5678 | 0.0028683 | -0.016198
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 1250.04 | 1250.04 | 0.22702 | 0 | 0.337 | 0.33898 | 0.0028707 | 0.0019752
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 1000.74 | 1000.74 | 0.22819 | -6.1035e-05 | 0.302 | 0.29317 | 0.0034992 | -0.0088264
2010-13-0035.mrg | 2010-13-0035.sam | 115 | 799.277 | 799.17 | 1.1049 | -0.10681 | 0.383 | 0.37046 | 0.0021492 | -0.012539
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 599.933 | 599.933 | 0.24701 | 0.00030518 | 0.672 | 0.63053 | 0.002021 | -0.041471
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 399.65 | 399.65 | 0.57486 | -0.00048828 | 1.729 | 1.6482 | 0.0024393 | -0.080843
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 300 | 300 | 0.59163 | -0.00033569 | 2.539 | 2.4146 | 0.0036073 | -0.12442
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 249.213 | 249.213 | 0.52919 | 0.00039673 | 3.3 | 3.1558 | 0.0036844 | -0.14422
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 201.385 | 201.385 | 0.61036 | -0.00015259 | 4.172 | 3.9694 | 0.0087082 | -0.20264
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 176.241 | 176.241 | 0.37202 | -0.00024414 | 4.483 | 4.2895 | 0.0066959 | -0.19352
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 150.259 | 150.259 | 0.28359 | -0.00021362 | 4.864 | 4.6717 | 0.0052742 | -0.19227
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 126.022 | 126.022 | 0.47285 | 0.00035858 | 5.362 | 5.1518 | 0.0054625 | -0.21019
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 100.338 | 100.338 | 0.62822 | 0.00035095 | 6.124 | 5.8719 | 0.0099315 | -0.25209
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 74.862 | 74.8615 | 0.19787 | -0.00048065 | 6.895 | 6.6217 | 0.0059463 | -0.27326
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 51.62 | 51.62 | 0.072191 | 2.6703e-05 | 7.033 | 6.7685 | 0.0088243 | -0.26449
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 25.838 | 25.8385 | 0.19803 | 0.00048828 | 6.675 | 6.4155 | 0.012656 | -0.25953
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 10.625 | 10.6251 | 0.11384 | 7.4387e-05 | 6.656 | 6.4084 | 0.012711 | -0.2476
2010-13-0035.mrg | 2010-13-0035.sam | 121 | 5.959 | 5.959 | 0.26721 | -2.4796e-05 | 6.655 | 6.4214 | 0.016474 | -0.23364
2010-13-0040.mrg | 2010-13-0040.sam | 121 | 5.658 | 5.6584 | 0.10828 | 0.00036383 | 6.723 | 6.4384 | 0.0094793 | -0.28461
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 4023 | 4023 | 0.5095 | -0.00048828 | 3.113 | 2.8556 | 0.00702 | -0.25743
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 3500.03 | 3500.03 | 0.40278 | -0.00024414 | 2.922 | 2.7524 | 0.0040812 | -0.16961
2010-13-0041.mrg | 2010-13-0041.sam | 242 | 3000.84 | 3000.73 | 0.56008 | -0.10742 | 2.544 | 2.4194 | 0.0057594 | -0.12455
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 2498.77 | 2498.77 | 0.49948 | 0.00024414 | 1.935 | 1.8721 | 0.0036797 | -0.062901
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 2001.65 | 2001.65 | 0.58031 | 0.00036621 | 1.32 | 1.2756 | 0.0036934 | -0.044397
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 1499.84 | 1499.84 | 0.34864 | -0.00012207 | 0.577 | 0.584 | 0.0024461 | 0.007
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 1250.74 | 1250.74 | 0.60693 | 0 | 0.348 | 0.37144 | 0.0029718 | 0.023438
2010-13-0041.mrg | 2010-13-0041.sam | 197 | 600.568 | 600.318 | 0.3936 | -0.25024 | 0.416 | 0.52383 | 0.003292 | 0.10783
2010-13-0041.mrg | 2010-13-0041.sam | 197 | 600.16 | 600.318 | 0.3936 | 0.15778 | 0.542 | 0.52383 | 0.003292 | -0.018173
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 401.395 | 401.395 | 0.6307 | -0.00033569 | 1.136 | 1.0774 | 0.0023309 | -0.058636
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 300.778 | 300.778 | 0.78834 | -0.00030518 | 2.039 | 1.9071 | 0.0024103 | -0.13192
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 250.418 | 250.418 | 0.24313 | -0.00024414 | 2.672 | 2.5519 | 0.0033954 | -0.12007
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 200.911 | 200.911 | 0.65838 | 0.00047302 | 3.662 | 3.4487 | 0.0024619 | -0.2133
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 175.849 | 175.849 | 0.52605 | 0.00032043 | 4.09 | 3.9304 | 0.0051943 | -0.15965
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 149.91 | 149.91 | 0.35224 | 0.00033569 | 4.589 | 4.3963 | 0.0051059 | -0.19268
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 124.964 | 124.964 | 0.28266 | 0.0003891 | 5.103 | 4.8779 | 0.0035967 | -0.22512
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 99.841 | 99.8412 | 0.30197 | 0.00016785 | 6.126 | 5.8287 | 0.01541 | -0.29727
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 75.049 | 75.0487 | 0.18252 | -0.00032806 | 6.951 | 6.6846 | 0.0044119 | -0.26644
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 49.795 | 49.7955 | 0.32581 | 0.00049973 | 6.884 | 6.6043 | 0.015588 | -0.27967
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 25.853 | 25.8528 | 0.48603 | -0.00015831 | 6.828 | 6.5625 | 0.010572 | -0.2655
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 10.465 | 10.4646 | 0.16334 | -0.00043011 | 6.888 | 6.5178 | 0.0081332 | -0.37018
2010-13-0041.mrg | 2010-13-0041.sam | 121 | 5.333 | 5.3326 | 0.14139 | -0.00044632 | 6.765 | 6.5223 | 0.0059579 | -0.24274
2010-13-0046.mrg | 2010-13-0046.sam | 121 | 4.835 | 4.8354 | 0.57492 | 0.00042963 | 6.821 | 6.5365 | 0.0033914 | -0.28447
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 4316.65 | 4316.66 | 0.17615 | 0.00048828 | 3.201 | 3.0625 | 0.0073575 | -0.13845
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 4000.97 | 4000.97 | 0.81374 | -0.00024414 | 3.163 | 3.0161 | 0.0086493 | -0.14688
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 3500.2 | 3500.19 | 0.97168 | -0.00048828 | 2.957 | 2.8164 | 0.0036599 | -0.14059
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 2999.35 | 2999.35 | 0.79636 | 0 | 2.565 | 2.4589 | 0.0047492 | -0.10606
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 2498.97 | 2498.97 | 1.039 | 0 | 1.969 | 1.9097 | 0.004458 | -0.059289
2010-13-0049.mrg | 2010-13-0049.sam | 110 | 2000.1 | 1999.87 | 0.92064 | -0.23169 | 1.34 | 1.3096 | 0.0039175 | -0.030418
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 1501.11 | 1501.11 | 0.37913 | -0.00024414 | 0.632 | 0.64938 | 0.0032666 | 0.01738
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 1250.88 | 1250.88 | 0.77736 | -0.00012207 | 0.419 | 0.4281 | 0.0028472 | 0.0090992
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 1001.36 | 1001.36 | 1.033 | -0.00042725 | 0.39 | 0.38405 | 0.0022835 | -0.0059504
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 799.943 | 799.943 | 0.63384 | 0.00030518 | 0.432 | 0.43269 | 0.0025028 | 0.00069419
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 600.641 | 600.641 | 0.11832 | 0.00024414 | 0.494 | 0.49294 | 0.0033821 | -0.0010578
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 400.367 | 400.367 | 0.10069 | 9.1553e-05 | 0.843 | 0.80394 | 0.0031919 | -0.039058
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 301.185 | 301.185 | 0.255 | 3.0518e-05 | 1.351 | 1.2659 | 0.0036224 | -0.085058
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 250.879 | 250.879 | 0.52419 | 0.00045776 | 1.735 | 1.6588 | 0.0042284 | -0.076231
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 200.507 | 200.507 | 0.76893 | 0.00048828 | 2.616 | 2.4929 | 0.004011 | -0.12306
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 174.492 | 174.492 | 0.66126 | -9.1553e-05 | 3.157 | 3.0154 | 0.0056939 | -0.14162
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 151.221 | 151.221 | 0.68354 | -0.00042725 | 3.703 | 3.5806 | 0.026746 | -0.12239
2010-13-0049.mrg | 2010-13-0049.sam | 242 | 125.934 | 125.937 | 0.35384 | 0.003273 | 4.441 | 4.2171 | 0.0090026 | -0.22392
2010-13-0049.mrg | 2010-13-0049.sam | 242 | 125.941 | 125.937 | 0.35384 | -0.0037308 | 4.463 | 4.2171 | 0.0090026 | -0.24592
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 75.471 | 75.471 | 0.59418 | 4.5776e-05 | 6.927 | 6.6079 | 0.0043759 | -0.31913
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 51.009 | 51.0093 | 0.65645 | 0.00034714 | 6.981 | 6.671 | 0.01042 | -0.31003
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 26.12 | 26.1199 | 0.48584 | -0.00014114 | 6.973 | 6.7 | 0.010029 | -0.27302
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 10.963 | 10.9632 | 0.33102 | 0.00018978 | 6.957 | 6.7037 | 0.0038741 | -0.25326
2010-13-0049.mrg | 2010-13-0049.sam | 121 | 5.764 | 5.7636 | 0.53558 | -0.00037193 | 6.975 | 6.7063 | 0.0022641 | -0.26866
Create Fit Sheet