#### 2010-14

CCGS JP Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### August 17 to September 2, 2010

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from all major stations (P2, P4, P12, P16, P20, P26) for DMS, DMSPD (dissolved) & DMSPT (total).  Incubation samples were also run for Laval University.  Finally, a Langrangian-type DMS diurnal cycle experiment was conducted at P26 over a 28 hour period.

  - DMS

Fourteen water samples from various depths (300m, 200m, 175m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at each of the stations P4, P12, P16, P20 & P26.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface).  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in the dark and removed one at a time before analysis.  Laval University incubation samples were collected in 50mL, crimp top, serum vials and ran within an hour of being collected.  The diurnal experiment consisted of collecting 2 samples (as per regular DMS samples) in duplicate, at 5m and 20m, every two hours, for 28 hours.  A total of 60 samples that were also run within an hour of being collected.

  - DMSP

Six samples for both DMSPD and DMSPT were collected at each station; two at the surface (0m, 5m), one in the mixed layer (100m), one in the deep chlorophyll max (20m) and two in the salinity mix layer (175m, 200m).  The only exception to this was P2 where there were no 175m or 200m samples, hence, only 4 samples were collected.

- Analysis

2.1 DMS

For this cruise there were two DMS systems used.  The older DMS system running on the GC5890 and the newer system running on the GC6890.  The only differences between the two systems is the GC6890 is more sensitive than the GC5890 and the purge gas for the GC6890 is UHP Helium whereas it is UHP Nitrogen for the GC5890.  Nevertheless, in both cases, a sample was loaded onto the stripper and purged with the above specified purge gas for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasorb 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

3.5mL of seawater was collected directly from the niskin into a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was then stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.  Stripping efficiency was evaluated at the beginning of the cruise and was proven to be acceptable at over 93%.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by simply treating MQ water as an actual sample.  For example, in the case of DMSPD it was put through a separate funnel and for DMSPT it was added directly to the polypropylene tube.

- Data & Results

  - DMS

Both systems worked well on this cruise but the newer DMS system (GC6890) was too sensitive to run many of the samples on this cruise due to the higher levels of DMS detected at some of the stations.  Whereas the saturation point on the GC5890 Flame Photometric Detector (FPD) is about 60nM it is about 12 nM on the GC6890.  This added sensitivity is good for low level DMS but poses a problem when one encounters levels above 12nM.  For this reason, when it became apparent that DMS levels for this cruise had the potential to be higher than normal, the system used primarily for analysis was the older, less sensitive one.

  - DMSP

Samples are to be run here at IOS within the next few months.

- Conclusions

6.1 DMS

It will become necessary to try and find a way to alter and manipulate the sensitivity for the new DMS system (GC6890).  Ultimately the goal was to replace the aging system with the new one but it was not anticipated that increased sensitivity would be a stumbling block.  There are a few ways to adjust sensitivity such as adjusting hydrogen flow, doing a split injection and/or reducing the sample volume (i.e. stripping say 5mL instead of 20mL).  All these options will have to be investigated and a suitable solution implemented.  Until such time as the sensitivity issue is rectified it would not be advisable to phase out the old DMS system.

  - DMSP

No problems to report.