## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 October 2011 | Fixed error in Oxygen:Dissolved:SBE channel in CTD files #37-129. Added PAR and fixed channel names in CHE files #48-129 Fixed file 2010-14-0084.CHE which had samples associated with wrong pressures & fixed the file that caused that error, “ADDSAMP.csv”. Fixed SBE DO values in loop file 2010-14-surface.loop. |
| 2 August 2011 | Reprocessed Oxygen:Dissolved:SBE channel for CTD #0443. See notes at end of report. Fixed comments in bottle files. Added Loop file. |

## PROCESSING NOTES

Cruise: 2010-14

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: August 17, 2010 – 2 September 2010

Processed by: Germaine Gatien

Date of Processing: 27 January 2011 – 10 May 2011

Number of original HEX files: 88 CTD and 6 TSG

Number of CTD files:   76     		Number of CTD casts processed:  73

Number of bottle casts:	69			Number of bottle casts processed: 69

Number of original TSG files:  6	         	Number of TSG files processed:  6

# INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs were used for this cruise.

- CTD #0506 was used for casts 1-33. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1176) on the primary pump), a Seapoint Fluorometer (#2345) with a 3X cable (on the primary pump), a Biospherical QSP-400 PAR sensor (#4694) and an altimeter (#1024). All casts were run with deck unit #0424.

- CTD #0443 was used for casts 36-129. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1483) on the primary pump), a Seapoint Fluorometer (#2356) with a 3X cable (on the primary pump) and an altimeter (#1252). Two Biospherical QSP-400 PAR sensors were used, #4694 for casts 36-74 and 116-129 and #4615 for casts 75-115.

The deck unit was #0424 for casts 1-59 and #0506 for casts 36-129.

All casts were run with the LARS mid-ship station.

Seasave version 7.16 was used for casts 1-69 and version 7.2g for casts 73-129. The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

After processing had been completed it was discovered that the wrong serial number had been entered for the SBE Dissolved Oxygen Sensor on CTD #0443, so channel Oxygen:Dissolved:SBE was reprocessed for casts 37-129 and the results substituted in the CHE and CTD files. The details on what was done are entered at the end of the report. The comments on DO quality remain valid.

The CTD and rosette logs were generally in good order though the equipment list had the wrong serial number for the dissolved oxygen sensors as mentioned above and there was no information entered about the thermosalinograph. The sampling notes were extremely helpful in dealing with problems noted at sea.

This was a very complex cruise with frequent computer crashes leading to many split casts and corrupted data. This was further complicated by the use of user-driven bottle firing where decisions were made on the fly. Recent changes to IOS SHELL will make it easier in future to process data using this firing method, but good record keeping will still be important. The rosette sheets have columns for firing order and Niskin #, but during this cruise only 1 column was completed for most casts, and while it was labelled “Firing #”, it was generally the Niskin # that was entered in that column. The BL files do provide a good record of the order of Niskin firing, but it is extremely helpful to have the information recorded on the rosette sheets with the sample numbers.

In cases where data acquisition was interrupted so that there are 2 files for each deployment, data were combined as required and event numbers were set to match the first of each of the set. Where bottles were collected from a separate upcast file, the event number and file names for the CHE files were changed to match the corresponding CTD file. There are gaps in the CTD files created by merging.

For CTD #0506 the secondary salinity is closest to bottles and showed no obvious pressure or time dependence.

For CTD #0443 the secondary salinity is also closest to bottles. There is some evidence of time dependence in the secondary channel, but no obvious pressure dependence, while the primary does vary with pressure and is noisier. The December 2010 factory calibration shows a higher than usual drift in the secondary temperature sensor used for this CTD, but data from this cruise show most of that drift must have occurred after August 2010. The salinity from CTD #0443 was noisy, with 2-sided spikes diagnostic of flow rate irregularity. Some editing was done to these spikes, but many remain. Bin-averaging will remove much of the problem, but the overall quality of the salinity is likely lower than usual for this CTD.

Inter-comparison of the two CTDs was possible since some stations were occupied on both the outward and inward journey. The salinity comparisons between the two systems were in agreement with the results of the CTD-bottle comparisons.

The primary channels were selected for archiving for file 2010-14-0040 because the secondary sensors performed badly. The secondary channels were selected for all other files.

This cruise was the first in 2010 since the salinity bottle liners were replaced. There is still more scatter than we might hope for in the comparison of CTD and bottles, but it is better than in the comparison for the Line P cruise in June, in that there are fewer unexplained outliers. This improvement is despite the fact that the samples spent longer between collection and analysis than for the June cruise, suggesting that there has been an improvement in the sealing of bottles.

A concern arises about salinity analysis from the study of 22 bottles fired at the bottom whose values drifted fairly steadily downwards through a range of 0.005. There is no obvious explanation for this other than analysis problems. It would be easier to determine the source of such variability if salinity samples were analyzed in a more random way. Analysis in sample # order makes it impossible to distinguish salinometer drift from other sources of error.

There were some errors in salinity sample labels and for one cast no bottle liners were used and the results were all outliers in the CTD-bottle comparison.

Two different SBE Dissolved Oxygen sensors were used for this cruise and both were problematic. The first (#1176) behaved much as it had during the previous use on 2010-36. There were 2 stable states for the sensor, one having higher noise level than the other in the raw voltage channel. The casts with higher-noise levels produce DO data that is approximately 0.24mL/L lower than the lower-noise casts (based on comparisons with bottles.) For this cruise the early casts were fine, but casts #20-32 had the noisier data and are considered of lower quality than the earlier casts. Unlike 2010-36 there were no cases of noise level changing mid-cast. There are only 3 casts with titrated oxygen samples from the “quiet” group and 1 from the “noisy” group. The results of 2010-36 were used for recalibration of the “noisy” casts.

The second DO sensor (#1438) showed a large change since it was last used and the bottle comparison was very difficult to interpret. There were only 4 bottle casts with titrated DO and for one of them the CTD DO was not connnected. One of the others had some spikes and looked odd in COMPARE, so recalibration was based on the last two. The recalibration of these data is based on very little information, so the quality of the calibration must be considered lower than usual and no estimate is made of precision for either DO sensor. Sensor #1438 shows evidence of hysteresis so SBE Dissolved Oxygen data from 2200db down were removed for casts using that sensor.

Acquisition of data for cast #61 did not begin until the CTD was at about 770db on the downcast. A file was prepared for the partial downcast. The upcast data were processed to provide information above 770db, but the data quality was very low due, so the file will not be archived. The data are available from the Chief Scientist.

The Thermosalinograph data had some sections with zero flow rate; all data in those sections were replaced with pad values except for date, time and positions. The first two TSG files have some low salinity values that are suspicious, but given the location in Haro Strait, they are not clearly bad.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as a report from the Chief Scientist summarizing problems and points of interest with reference to processing. There were many problems with spikes, interrupted casts and negative dissolved oxygen values. A number of steps were taken to try to reduce the spiking, including disabling individual sensors on casts as tests. So special care will be needed to determine which channels should be removed before archiving. The spiking problem did not occur after the SeaSave version was switched before cast #73.

Extracted chlorophyll, nutrients, dissolved oxygen, salinity and DMS data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. It was determined that there were 3 different configurations, one for CTD #506 and 2 for CTD #0443. There were a number of errors in the con files:

For CTD #0506:

- The pressure slope and offset were wrong.

- The data entered for the dissolved oxygen sensor were from an earlier calibration; a more recent one in November 2009 should have been used. There were also fine-tuned values for E, H1 and H3 available from tests run during 2010-13, so those were entered.

- The transmissometer calibration was not on file and the date was not entered, but the technician confirmed that a check was done on 15 August 2010 and the values in the con file are correct.

For CTD #0443 (for both configurations – only difference between them is change of PAR sensor)

- The pressure offset has been increased to +7db based on other 2010 cruises.

- The dissolved oxygen parameters for this sensor were not on file, but they were confirmed by the technician when they were used during 2010-01. There were fine-tuning parameters available E, H1 and H3 for the dissolved oxygen sensor from 2010-01 so those were entered.

After those changes were made, the configuration files were saved as 2010-14-ctd1, 2010-13-ctd2.con and 2010-14-ctd3.con.

##### Conversion of Raw Data

Data were converted using configuration files:

2010-14-ctd1.con for casts #1-32

2010-14-ctd2.con for casts #36- 73 & 116-129

2010-14-ctd3.con for casts #76-115

After conversion some corrections were made to headers based on the sampling notes.

Two casts have odd names and were each checked to determine whether they were just misnamed or were test casts that should not be processed.

File 2010-13-0105hex.cnv does contain reasonable data. so was renamed 2010-13-0105.cnv.

File 2010-13-0062test.cnv contains only a little surface data, so will not be processed further.

A few casts were examined and all expected channels are present.

For CTD #0506 the two temperature channels are reasonably close during the downcasts but the primary has some small spikes. The upcast traces differ more at times and differ from the downcasts significantly. Where they differ the primary tends to be smoother than the secondary. Examination of one cast shows the primary to be slightly lower than the secondary during downcasts and bottle stops, but higher just before the stop. These look like they could be alignment issues, due to either physical arrangement or problems in the plumbing.

The pair of conductivity channels is much like the temperature pair.

The fluorescence and DO voltage look as expected with a dark value of about 0.06 for the former.

PAR looks fine. The altimetry looks ok where the CTD got close to the bottom. The transmissivity data look ok.

For CTD #0443 the differences between temperature channels are similar to those for #0506, but the downcasts do not contain spikes. The upcasts and downcasts are quite different, but the descent rate is mostly so noisy as to make this a hard judgment. The conductivity channels are much like the temperature pair. The DO voltage looks normal. The fluorescence dark value is sometimes ~0 and at other times varies from 0.02 to 0.07 looking rather odd, but it was not always mounted, so some of the odd results might be noise or zero signal. PAR looks fine. The altimetry and transmissivity look ok.

During 2010-36 which preceded this cast it was found that the DO sensor used on CTD #0506 flipped between two states with different fits of SBE versus bottles and with a different basic noise level in the DO voltage. Examination of casts #7 and #20 shows that this problem continued during this cruise with cast #20 having a noise level twice as high as #7 through both downcast and upcast. The sampling notes mentioned negative DO values in the #20 upcast. The DO sensor was later changed and the noise level looks low from then on, so the problem is likely with the first sensor. There were many communication problems during this cruise and through 2010-36 that might have been thought to contribute to the DO problem. However, changing the DO sensor put an end to noisy DO signals even though the spikes and crashes continued, so it looks like DO sensor #1176 needs repairs.

Some casts were interrupted by equipment problems. In most cases the downcast sections of casts are all in one file, but the following two files will need to be combined later:

- File #52 contains surface-22m and File #53 contains 48m to the bottom. These can be combined but will have a large gap from 22-48m. The only alternative is to use the upcast from #53 but given that there were many bottle stops the quality of such data would be poor. There are other casts at this site, so this profile is not critical. So the two downcast files will be combined and named 2010-14-0052. There is no IOS sampling from this file, but a bottle file will be needed.

- Files #55 and #56 are from the same cast and should be combined as #55.

The downcast files will be joined after DELETE is run.

##### Rosette File Conversion and Study of Split Casts

Rosette files were converted using a start time of -5s and duration of 10s. The TAU and hysteresis corrections were chosen. Because Niskin bottles were fired out of order on many casts, the Bottle Position in Carousel was converted. There were rosette files for every CTD cast, but 13 files were empty, so those were deleted.

Corrections were made to headers based on the sampling notes from the Chief Scientist.

As was done for the CNV files casts with odd names were examined but only 1 actually contained any data. File 2010-13-0105hex.ros contains reasonable data so was renamed 2010-13-0105.ros.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers, with output named *.BOT. CLEAN was also used to remove off-scale bad fluorescence values based on plots to determine what are spikes and what are likely real. Values >21ug/L were removed from casts 1-128 and values >44.5 for cast #129.

Temperature and salinity were plotted for all BOT files and a few casts appeared to have outliers. Those casts were examined in CTDEDIT.

- For cast #40 1 record was removed from the 100m bottle and primary salinity was cleaned very lightly.

- There were a few spikes in cast #43 – records were removed using CTDEDIT in most cases, but a text editor was used to replace a few individual values with pad values for a few others.

Editing details were added to the header comments of the edited files. The output files were then copied to *.BOT.

The upcasts were interrupted by computer crashes or restarts for some casts, so the bottle files will require careful merging and the first file # will be the one selected for the final files. There were problems during 2010-36 when bottles were sometimes fired during each of the fragments making the merge process more complex, so the BL files were checked for that.

In the following description of the problem casts [#-# & #-#] will be used to indicate [Firing # range & Niskin # range] in each file.

- 58/59 [1-8 & 1-8] / [1-6 & 9-14]. Straight-forward merge possible.

- 63/64 – [1 & 1] / [1-3 & 2-4]. Straight-forward merge possible.

- 68/69 - [1-11 & 1-11 / [1-6 & 1-3 and 12-14] – This is the most confusing case. It looks like Niskins 1-3 were fired twice, but in fact they were not sampled, so it is not important. But those records should be removed from file #69 just to make the merge possible.

- 76/77 - [1-4& 1-4] / [1-14 & bottles 5-18 but user-driven bottle firing was used] Checks will have to made that the merge is done right. This process would have been much easier if the rosette sheet had been filled in with firing # and Niskin #. There was only 1 column completed and it appears to be the Niskin # though it is listed under Firing #. The BL files do contain the required information.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using 2 casts from each CTD. The goal is to make upcasts look closer to downcasts on a T-S surface. Most casts had stops for bottles and few casts without those stops had a steady-enough descent rate to make it easy to judge which settings are best. That was especially true for CTD #0443. Near-surface data often differed significantly, and those changes may well be real, so the results were extremely difficult to judge. It was clear that some adjustment helped, but most choices looked best in some part of the profile.

The choice (α = 0.02, β=9) looked best overall for the both sensors on CTD #0506.

The choice of (α = 0.03, β=9) was best for the primary and (α = 0.02, β=7) was best for the secondary conductivity for CTD #0443.

CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The first three rows are from 2010-36 which preceded this cruise and had the same sensors as were used on CTD #1 for this cruise. The next two rows are from CTD#1 for this cruise. The last 3 rows are from CTD #2 for this cruise (#0443). The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2010-36-0020 #0506 | 500 1000 1900 | Too noisy -0.0002 -0.0003 | Too noisy 0.0001 0.0001 | Too noisy 0.0015 VN 0.0014 | High, Very Noisy |
| 2010-36-0065 #506 | 500 1000 1900 | -0.0008 -0.0007 -0.0007 | 0.00005 0.0001 0.0001 | 0.0015 0.0017 0.0018 | High, noisy |
| 2010-36-0162 #506 | 500 1000 1900 | ~0 -0.0007 -0.0007 | 0.0002 0.0002 0.0002 | 0.0025 0.0033 0.0032 | High, very noisy |
| 2010-14-0016 #506 | 500 1000 1900 | -0.0005 -0.001 -0.0005 | +0.00025 +0.0002 +0.00025 | +0.0035 +0.004 VN +0.003 | Noisy, High |
| 2010-14-0022 #506 | 500 1000 1900 | -0.002 VN -0.0002 -0.0005 | +0.0001 +0.0002 +0.0002 | +0.003 N +0.003 +0.003 | Noisy, High |
| 2010-14-0045 #0443 | 500 1000 1900 | -0.0008 VN ~0 VN -0.0001 | +0.0001 +0.0003 +0.00035 | +0.002 +0.0035 +0.0045 | Noisy, High |
| 2010-14-0073 #0443 | 500 1000 1900 3800 | +0.0002 0.0000 -0.0002 -0.0002 | +0.0001 +0.0002 +0.0003 +0.00035 | +0.001 +0.003 +0.0035 +0.0045 | V noisy, High |
| 2010-14-0090 #0443 | 500 1000 1900 3800 | 0.0000 0.0000 -0.0001 -0.00015 | +0.0001 +0.00025 +0.0003 +0.00035 | +0.0015 +0.003 +0.0035 +0.005 | V noisy, High |
| 2010-14-0127 | 500 1000 1900 | ~0 VN +0.0003 +0.0003XN | +0.00004 +0.0002 +0.00025 | +0.0006VN +0.002 +0.003 | X Noisy High |

For the first CTD system, the temperature differences are small and similar to the earlier cruise. The conductivity and salinity differences continued the temporal drift noted in 2010-36. During the earlier cruise the primary salinity was found to have been drifting with time more than the secondary.

For CTD #0443, the differences in temperature are small and show no time or pressure dependence, except for the last cast checked, where the difference changes sign but is still small. The conductivity and salinity show significant pressure dependence, but little temporal change.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

Some files could not be converted, so the CNV files for those casts were examined and were possible were edited to enable conversion.

- 32: Spiky records were removed from the end of the file and then conversion worked.

- 33: This file was aborted and is full of spikes, so will not be processed further

- 37: Spiky records were removed from the end of the file and then conversion worked.

- 44: Spiky records were removed from the end of the file and then conversion worked

- 52: Spiky records were removed from the end of the file and then conversion worked

- 76: Spiky records were removed from the end of the file and then conversion worked – all removed data were from the upcast.

While some casts with spikes had to be edited to enable conversion, others could be converted but still were heavily corrupted. Fixing them at this stage seemed preferable to doing it later so that alignment and DELETE have the best chance of working correctly. So plots were produced for all casts to see if removal of more bad records was required, and the following cases were examined:

- Cast #17 – Records were removed from the end of the file due to heavy corruption by spikes.

- Cast #23 - Records were removed from the end of the file due to heavy corruption by spikes.

- Cast #26 – The pressure is ok, but other channels have spikes around 1000db, i.e. in the middle of the downcast. Some spikes were replaced with pad values in the Wildedit stage, but many others were not. It may be possible to remove fluorescence spikes using the CLEAN program by determining the highest “real” values in the cruise and replace all that are higher than that. But the temperature channels are heavily corrupted and salinity is worst with large sections of all data being bad. Ultraedit was used to remove many of these spikes; those left can be addressed later with CTDEDIT. While it is expected that the dissolved oxygen channel would be affected as well, there is no evidence of this in the oxygen voltage trace. In the DO concentration channel there are spikes which are followed by large sections of negative values. This does not look like the sort of shifts seen in 2010-36 where there were 2 different fits between bottles and CTD. What is seen here is a change in values on the order 1ml/l rather than 0.24 as seen in the earlier cruise. These data were clearly bad and were removed.

- Cast #29 – The problems are all in the upcast, so no editing is required for the profile file.

- Cast #40 - The problems are all in the upcast, so no editing is required for the profile file.

- Cast #55 – Spiky records near the end of the file were replaced with pad values.

- Cast #65 – There were many spiky records near the end of the file. Fixing this with a text editor was not practical. The spikes in pressure line up with spikes in primary conductivity and salinity, though there are many more spikes in those variables than in pressure. The spikes don’t line up with secondary channel spikes, but the secondary salinity is very bad overall. When some clearly bad data were removed from the salinity data, it was found that what remained contained smaller spikes which were more difficult to deal with because it is not clear which values are right, which wrong. There is a patch of data where the pressure suddenly drops and this looks like a case of a block of data being dumped twice. In previous cruises when this occurred there was no need to edit it because DELETE selected the first block and that works well. Primary temperature does seem ok.

##### Checking Headers

The header check was run. One cast was found to have the wrong CTD serial number so that was corrected. There are many negative values which is not surprising given the known spikes in many channels.

The fluorescence maximum was checked in the header check to see if there are likely to be any off-scale values. There were a few casts for which it wasn’t connected, but the signal does contain a few spikes, so those casts were removed from the list for the header check. In the course of checking which sensors were connected on which casts it was found that while the log says only a DO sensor was logged for cast #37, there are believable signals in fluorescence and PAR channels.

Casts #26, 29, 45 and 58 appeared to have high fluorescence values, but they proved to be spikes only. The highest value when those casts were excluded (and others for which the sensor was not connected but spikes were logged) was ~44.3 during cast #129. The peak was well below the maximum and showed no sign of being off-scale. When that cast was excluded the next highest value was ~20.5ug/l from cast #1.

Surface check was run and shows an average surface pressure for the cruise was 2.9db, 2.84 for CTD #1 and 2.91 for CTD #2, with no significant difference between the two systems. During cast #46 the pressure was very close to 0 at the end of the upcast, with values as low as -0.3db. The pumps were on and associated conductivity and salinity show that the CTD was indeed right at the surface, and could have been dipping in and out of the water. There were no external sensors on for that cast, so we can’t check fluorescence or transmissivity. Cast #61 also has some very low values in the upcast, and in that case fluorescence and transmissivity read 0 for some of the negative values; the pumps were on. So CTD#0443 pressure looks fine. No negative pressures were found for CTD #0506 except for very large negative spikes in some casts.

The cross-reference check was compared with the log book and the only inconsistency was a station name which did not match the log book and a few station names that were not in standard format; those were fixed.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. Most casts did not have a header entry because they did not get within 15m of the bottom. There were 12 cases where there was a header entry, but the CTD did not come close to the bottom according to the log; these entries are caused by spikes in the altimetry, so the header entries were removed. Only 9 casts with header entries remained and the values appear to be correct.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. As noted above few casts got within 15m of the bottom. Of the 9 casts with header entries, 4 were from shallow casts in deep water so the header entries were removed.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number. This step was also used to replace SBE:Fluorescence values >44.5ug/l with pad values in cast #129 and >21ug/l for all other casts.

##### BOTTLE FILE PREPARATION

A BOT file exists for event #32 but is empty, so will not be processed further.

As discussed in section 4, there are 4 cases where 2 files need to be merged to create complete bottle files.

A text editor was used to remove data from firing #s 1-3 from file #69. The edited file was saved as BOT1, and put through clean to produce a new 69.BOT with adjusted header limits.

The file names for all the split casts (multiple files) were then adjusted so that the partial casts all have the name of the 1st file in the group, but the extensions were named BOTx and BOTy. So for the files that were named 58.BOT and 59.BOT, we now have 58.BOTx and 58.BOTy.

The BOTy files were opened in Ultraedit to adjust the Bottle number to match the rosette sheet number.  Then JOIN was used to produce a single BOT file for each of the split casts: 58, 63, 68 & 76.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Because sample numbers were assigned in order of Niskin # rather than Firing #, a column with the Niskin #s was first added, then BL files were used to complete that information. The data were reordered on Niskin # and sample #s were then added.

Cast #29 was confusing because the bottle file contains only 4 bottles, 3 at2000m and 1 at 1000m. The rosette sheet indicates there were 12 bottles closed and samples taken from 4 bottles at 2000, 1000, 500 and 10. The salinity samples from those bottles do indicate that the samples come from the levels shown on the rosette sheet. So where is the CTD data to match them? The full file does contain data from those levels. The BL file shows 12 bottles fired with scan numbers available, so we can create data to go with these bottles. An examination of the ROS file provided the answer. There are a lot of spikes in the file, but the critical problem is that the DO values went off-scale. When those were replaced with pad values conversion to IOS SHELL format was repeated and all bottle data are in the output. The file was still full of spikes, but all bottles were included. The BOT file was renamed as BOT1 and then several passes through CTDEDIT were made to remove records with bad pressure data and individual spiky values for other channels. Even the bottle position number had spikes. Even after repeated passes through CTDEDIT spikes remained and were gradually removed using a text editor until plots and header check looked ok. (Later, more problems were found by checking that the bin-averaged SAMAVG files had similar values for 2 sensor pairs and that data from multiple bottles at a single depth looked similar.)

Cast #40 caused many problems – some were not discovered until late in the processing job. There were many spikes that led to truncated bottle files.

Cast #101 is another confusing case, as the rosette log shows 24 bottles fired, but the BL file has 25. Niskin #16 was fired twice according to the BL file, it is listed for the 16th and 19th bottles fired. It will be assumed at this point that the 19th firing should be dropped from the list. (This cast was checked when the bottle files were assembled; as there were no IOS samples to be added the evidence is limited to confirming that the Niskin bottle # entries match the rosette log.)

Cast #109 has a bottle file, but the bottles were leaking so the cast was rerun. The bottle data will not be processed further.

The addsamp.csv file was ordered on sample number and then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files.

The SAM files were then bin-averaged.

For cast #1 bottles #12-24 were removed from the SAMAVG file with output SAMAVG1; that file was then cleaned to fix the headers and saved as SAMAVG. Similarly, bottles 4 and 5 were removed from cast #16 since they were for test purposes only.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. Comments and were included but no quality flags. The sheet from the analyst was saved as 2010-14-sal.csv. Loop samples were copied to a separate spreadsheet (2010-14-loop-sal.csv) and removed from the main spreadsheet. Also removed were some samples taken for iron studies and other non-CTD purposes. Some columns were removed and a column was added for the event number – the latter information was found in the log book.

The spreadsheet included duplicates which were removed to a separate sheet, averaged and the differences analysed (See 2010-14-sal-duplicates.xls). Of the 8 pairs 6 are within 0.003 of each other and 3 are closer than 0.001. There are two larger differences. In one case the difference is 0.28, but when the rosette sheet was checked there was no indication that duplicate samples had been taken, so it is assumed that one of these samples is mislabelled. One is close to the CTD salinity while the other is not, so the one that looks wrong will not be included in the average. The other pair differ from each other by 0.01 but when compared with the CTD they both seem off. It is notable that there was another bottle at the same level and that salinity is differs significantly from the pair from the other bottle, though the CTD salinity varies little. These 3 samples had been noted by the analyst as having a “single cap screw only no liner”. The samples were analyzed 2 months after the cruise ended allowing time for evaporation which would explain why they are significantly higher than the CTD. After COMPARE is run it should be clearer if the samples with no liners are all affected in this way.  When those 2 outliers are excluded the value of Sp is 0.0013 indicating excellent reliability of the analysis and sampling.

All salinity values were checked against rosette sheets to see if the mislabeled samples belong somewhere else. In the course of checking a few other problems were noted:

- Cast #3, Sample #22 – the value ~34 is too high for 20m and there is no indication that a sample was taken. The Autosal analysis sheet indicates it is from P20 – this looks like it should be labeled as an Iron Sample. The SAL file was deleted and the merges rerun.

- Cast #52, Samples 268-273. There is no note of intended salinity sampling and the values look too high. The Autosal analysis sheet indicates it is from P26 – again this should be labeled as an Iron Sample. The SAL file was deleted. The merge process was rerun.

- Cast #126 – Sample #631 – There is no note of intended salinity sampling on the rosette sheet but it looks ok.

- Cast #127/128 - Given that there were 5m samples from casts 125, 126 and 128, there is a reasonable chance that the extra sample labelled 635 (cast 128) should really be #633 (cast #127). Looking at the Autosal analysis sheet, it is clear that there were samples #633 and #635, and the higher value was the first analyzed. The first one analyzed is close to the CTD data for cast #127. This looks just an error in entering sample # on the computer. The sample # was fixed in the spreadsheet from the analyst, and corrected in the simplified spreadsheet.

File 2010-14-sal.csv was converted to individual SAL files.

The salinity spreadsheet was updated with the changes to sample #633 and iron samples.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2010-14oxy.xls which includes a precision study. The spreadsheet was simplified, ordered on sample numbers, flags changed to numeric format and headings changed to standard format. Event #55 was changed to #56 to match what was done with the nutrient samples from this split cast. Loop samples were removed and placed in file 2010-14-loop-oxy.csv. The spreadsheet was then saved as 2010-14-oxy.csv and that file was converted into individual *.ADD files.  There were some samples flagged “2” which should be re-examined after COMPARE is run and either be removed or changed to 3, 4 or 5.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2010-14nuts.xls which included a report on precisions. Loop data were copied to a separate spreadsheet (2010-14-loop-nuts.csv) and removed from the rosette spreadsheet which was saved as 2010-14-nuts.csv. Headers were changed to standard format, event #s were adjusted as needed for the split casts and quality flags were changed to numeric values. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. Colour coding was used in the spreadsheet to indicate that a few samples were stored cool rather than frozen, so a comment was entered for those since the colour will not show in the bottle files. The same was done for the coding for “PO4 rerun”. Pad values were added to a few blank entries with “9” flags. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2010-14chl.xls. The name of the file was changed by adding the date of creation. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which headers were changed to standard format, data were sorted on sample number, pad values were entered for missing samples, unnecessary columns were removed and the file was saved as 2010-14-chl.csv. Loop samples were placed in a new file, 2010-14-loop-chl.csv, and removed from the main spreadsheet which was then converted to individual CHL files.

DMS

DMS data were obtained in file DMS 2010-14 summary.xls. The file was saved as 2010-14-dms.csv and edited. There were flags and comments but not in column format so the details were transferred to columns. One sample number and depth was corrected in the P2 data after consultation with the analyst. All entries “<” were replaced with “0”; a note in the header will explain that the minimum detectable level is 0.1. Duplicates were averaged and “6” flags attached. Headers were changed to standard format and unnecessary columns were removed; an event number column was added and filled in with information from the log book. The file was then converted to individual DMS files.

The SAL, CHL, ADD, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

Because this process was complex, all bottle data from the MRG files were exported to a spreadsheet and reordered on Bottle Position (Niskin #) order to check that all expected samples are present and in the right spot. Many errors were found and corrected including small errors in the analyst’s spreadsheets, wrong event numbers added to the salinity spreadsheet, errors in the ADDSAMP file and files that had not been renamed properly from the split casts. File #84 was the hardest to deal with because sample numbers had not been assigned in Niskin # order like all other casts.

##### Compare

Salinity

Compare was run for each CTD system with pressure as reference channel.

For CTD #0506 all points from cast #12 were major outliers – they had been flagged by the analyst due to the lack of liners and in one case salt on the lid. A bottle from cast #3 was also an outlier and there is no note on the rosette log of any sample being taken at that site. There is a lot of variability in the results even after all values above 200db are excluded; this left only 18 bottles. The primary salinity was low by an average of 0.0032 and the secondary salinity was high by an average of 0.0005; the standard deviation in both channels was ~0.002. While there is a hint of time dependence in the primary, the variability is too high to put much confidence in that observation. The results are consistent with the findings of section 8. There is no suggestion of time-dependence in either channel.

For CTD #0443 all points from cast #52 were removed – once again they appear bad and there is no record in the rosette log of any sampling at this cast. The only other notable outliers are from very close to the surface where we do not expect good results. Two other outliers are from cast #40 which was corrupted by spikes during the upcast, so it is possible a few problems remain in the CTD data in the bottle file. The average of differences below 200db shows the primary CTD salinity to be lower than the bottles by 0.0060 and the secondary lower by 0.0016 with standard deviations of ~0.0015. The differences are consistent with the results of section 8 where the secondary CTD salinity was found to be higher than the primary by ~0.0045 at 1900db towards the end of the cruise. There is pressure dependence in the primary salinity but not in the secondary; there is time dependence in both having opposite signs and being slightly higher in the secondary salinity.

There were 22 bottles fired at 2001db during cast #118 (CTD#0443). The average of those bottles indicates that the primary salinity was low by 0.006 and the secondary by 0.0016 with a standard deviation of ~0.0012 and 0.0016, respectively. A plot shows a drift (with a range of ~0.005) towards lower differences in both series so that by the end of the stop the differences are close to those found as the average for the whole cruise. The CTD data are identical because this was a pressure-based comparison so all the deep bottles get combined. Looking at the full file shows there was some pressure variation during the stop, but it is small and would only explain variations on the order of 0.001 in salinity. This cast stopped far above bottom, so the variations are not caused by bottom reflections. The bottle values become lower with increasing Niskin # which rules out slow flushing. So it seems most likely that the variations are due to analysis variations since the samples were analyzed in sample # order. This is more variation than we would expect to see, so is worrying. Randomizing analysis order would help assess whether salinometer drift is a significant problem – these samples were all analyzed in sample # order.

Another run of COMPARE was done using Niskin bottle numbers as reference channel to see if there is an indication that any Niskin bottle was malfunctioning. This comparison included both CTD systems, but separate plots were produced for each. There is a trend of slightly higher differences for the higher Niskin #s but it is small (especially in the secondary) and probably due to the higher numbers only being used for shallow sampling where local gradients are more significant. There are 2 clear outliers for CTD #0506 but none for CTD #0443 (other than those from casts #3, 12 and 52 discussed earlier). The 22 bottles fired at 2000db showed the opposite trend with the smallest differences for the higher Niskin #s. There is no evidence of bottle malfunctions.

This cruise was the first for which new liners were used on the bottles. The new bottle washing system was not yet in use.

The outliers from cast #12 were flagged “5” and replaced with pad values. There is already a comment about the reason, but it was amended by adding that they were severe outliers in COMPARE.

The outliers from casts 3 and 52 are severe and given there was no salinity sampling shown on the rosette log sheets for those casts, it is assumed that these are mislabelled. It is likely that they are iron samples given where they are written in the Autosal analysis log sheets; the sample numbers do not match the station they are said to be from on the rosette sheets. The SAL files were removed for those 2 casts and the merge process rerun.

The other 3 outliers that stand out are surface samples from casts 2, 5 and 15; the vertical distance between the CTD and bottles may be significant at these levels and profiles show a lot of variability in the CTD data. No flags will be attached.

For details see 2010-14-sal-comp1-506.xls and 2010-14-sal-comp1-443.

Dissolved Oxygen

The notes about the sensor on CTD #0443 are from the original processing. See end of report for reprocessing details. The fits were different but variability and outlier comments are valid.

As a first quality check plots were made of DO from both CTD and Bottles versus CTD Salinity. No major outliers stood out in these plots.

COMPARE was run for Dissolved Oxygen. There were 8 files with titrated samples, but for 1 of those the SBE sensor was not connected. There were 2 DO sensors used, so the initial fits were done in 2 groups using 4 casts for CTD #0506 and 3 for CTD #0443.

For sensor #1176 (CTD #0506) the data from casts 1, 5, 12 and 22 were plotted together. Clearly, cast #22 does not fit with the others. This is reminiscent of 2010-36 when data from this sensor fell into 2 groups that differed by a roughly constant offset.

The fit for casts 1, 5 and 12 when the hypoxic values from Saanich Inlet plus a few outliers are excluded is: 	Bottle DO = 1.0497 * CTD DO - 0.0115

For 2010-36 casts with low noise level in the DO channel the results were:

Bottle DO = 1.0492 * CTD DO - 0.0237 (2010-36 Fit #1)

For cast #22

Bottle DO = 1.0378 * CTD DO + 0.2588

For 2010-36 casts with higher noise level in the DO channel the results were:

Bottle DO = 1.0347 * CTD DO + 0.2383 (2010-36 Fit #2)

There is 1 point in each of casts 1 and 12 that are outliers from the group and could conceivably be better represented in the cast #22 group, but 1 is from 10db in Saanich Inlet where the fits are often poor due to the extreme gradient and the other is above a DO maximum, so not well represented in data which is chiefly from regions where DO is increasing as the CTD rises There are other outliers, but they look bad in both fits.

Plots were made of the casts to see if there is a distinctive noise level in the DO voltage that can guide us on which fits will suit which casts – we do not have bottles to help with most of the casts. For casts 1, 5 and 12 the DO voltage has a noise level of approximately ±0.0006 volts. For cast #22 it is higher by a factor of 2 or more. It is a little harder to judge the noisy casts, but the quiet ones are easy to pick out. From 2010-36 we know that this can change in the middle of a cast and is sometimes marked by a sudden shift in DO values, so the checks should involve looking for such shifts and then an examination of sections of each cast. Examination of full files turned up no cases of the noise level changing mid-cast and all casts up to #17 fell into the Fit #1 group and those from cast #20 to 32 into Fit #2 group.

It is noteworthy that the shifts in noise level are seen in the Oxygen Voltage channel and not in the temperature channel, so this is not a “derived” problem. No other channel shows the same behaviour.

The next step was to look at DO #1438 (CTD #0443). There were only 3 casts available and the results are even more puzzling than those for the other DO sensor. The plot with all data is extremely noisy. When each cast is analyzed separately there is noise but it becomes obvious that it is from the top and bottom of each cast. If data are restricted to 27db to 2200db there are reasonable fits for each cast. But while casts 73 and 86 have similar slopes, cast #55 is quite different. Cast #55 was interrupted by spikes and restarted, but all of the upcast data came from one file which contained no large spikes. It does look very noisy, but the errors that would result from such noise are too small to explain the odd fit. There is some evidence that the deep values are getting more out of line as time goes along – this is seen at 3500db and maybe at 3000db, but not at 2500db. In each case there only 3 or 4 bottles on which to base a conclusion. The picture is unclear as to whether differences are growing with time at the surface. The odd surface behaviour may be related to a subsurface DO maximum at ~40db, so that a fit that suits data below 40 would not do as good a job above that level.

This sensor was used on 2009-43 and 2010-01, though not on many casts. No problems of this sort were noted. The deep data suggest hysteresis, so the parameters chosen should be reviewed. The noise level in the DO voltage is of the same order as noted in the “quiet” casts for the other DO sensor.

A plot was made of casts #29 and #128 which were both at P12 using the two different DO sensors. The earlier cast was full of spikes. The two profiles look very different, with the DO minimum values 0.02 and 0.2ml/l. But this is before recalibration and cast #29 falls into the Fit #2 group, so it will come to about 0.28. The #128 value will come up to 0.23 if the same calibration is used as found for cast #86 and 0.255 if we use the fit for cast #55. The two sensors are producing minimum values at P12 that are in reasonable agreement. DO minima in June and August of 2008 and 2009 were 0.2 to 0.25ml/l.

Further consideration of how to handle the DO data from sensor #1438 will be suspended until the full profiles have been processed further. After the data have been aligned, we can do a comparison of downcast data to bottle data and see if we can learn anything from that.

A few bottles were checked to see if the flags already assigned are appropriate:
Cast #40, Sample #186 – This sample had been flagged “2” as it was not clear if the value was affected by the fact that the Niskin bottle was leaking. Unfortunately the DO sensor was not connected for this cast, so a “4” flag was added since there is some doubt.

Cast #40, Sample #187 – This sample had been flagged “4” with the note “very stepped curve, endpoint likely overshot”.  The absence of SBE DO data means there is no check on this.

Cast #55, Samples #298 and 300 – Both were outliers in COMPARE but most near-surface samples were. This is no evidence of a problem, but does not prove there is no problem, so flag was changed to “4” since the analyst had noted problems with seal for one and vent for the other.

Cast #73, Sample #360 – Like the two from #55, this is an outlier, but as a surface sample this is no proof of a problem. However, the vent was open and this does not prove it was ok, so flag was changed to “4”.

For CTD #0506 there were a few outliers that did not fit either Fit #1 or Fit #2 but none suggest that the bottles require flags:

- Cast #1 – Saanich Inlet samples are usually problematic due to the very high DO gradient. The 5m sample appears to be lower than the CTD which is not expected, but there is a lot of variability in the CTD data that is large enough to account for the difference. The 25db bottle can be explained by the 1.5m between the bottle and the CTD. The 11m sample is only a little off, but again is easily explained by the local gradient. The bottom bottles were removed from the comparison because the CTD sensor is known to not reach equilibrium in anoxic waters unless there is a very long wait at the bottom.

- Cast #5 has a minor outlier at 11db, but the CTD never reached equilibrium at that level so that is no evidence of a bottle problem.

- Cast #12 had an outlier at 26db, but the CTD trace shows a lot of variability with a complex structure obvious in the downcast. The standard deviation in the CTD data in the few seconds around firing is not particularly high, but a larger window would have shown a different picture.

The only outliers for CTD #0443 are the shallow and deep values noted before; these are almost certainly problems with the CTD not the bottles.

(See 2010-14-dox-comp1*.xls.).

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. For 2 casts the fluorometer was not connected so those values were excluded from the fits. There remains a lot of scatter. The outliers all come from near-shore casts (P1 to P7) between 0 and 22db. When those are removed the fit leads to the relationship:

Extracted Chlorophyll = 0.097 * SBE Fluorescence +0.256.

The most severe outliers are from P1 to P4. Chlorophyll is much higher than fluorescence when CHL>1 and Fluorescence is often much higher than CHL when CHL<1ug/L. (See 2010-14-chl-fluor-comp.xls.)

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

A few MRG files were examined because of queries from the nutrient analyst, wondering if there might have been some misfires:

- Cast 12 – Sample 61 at 400m – the nitrate is considered a little low. The DO versus CTD salinity shows no problem. All salinity samples were flagged due to poor seals and values way out of line, so this does not help. There is no evidence of a misfire but also no evidence that all is well.

- Cast 12 – Sample 56 at 1000m – the nitrate and phosphate were considered out of line and the DO looks a little off. All the bottle samples had poor seals and stand out in COMPARE, but the one at 1000m does look most “off” and could be a sign that the bottle closed above the level at which it was fired. All samples were flagged “3” except the salinity which was already flagged “4”.

- Cast 12 – Sample 55 at 1250m – the nitrate and phosphate were flagged as outliers. The DO versus CTD salinity plot looks fine. The 2 salinity samples do look quite different but those were flagged due to poor seals so that is poor evidence. DO was varying significantly through the stop.

- Cast #22 – Samples 127 and 128 – Plots of DO against salinity look fine and the salinity bottles are not significant outliers in COMPARE. So no sign of a misfire.

- Cast #40 – Sample 173 – The salinity bottle looks fine. There was no DO sensor connected so that is no help, but titrated DO versus salinity looks ok. Looks like only Phosphate was off.

- Cast #55 – Sample 287 – The salinity bottle looks fine. The plot of DO (titrated and CTD) against Sal looks fine. No evidence of a misfire.

Once again data were exported to spreadsheet 2010-14-bottles.xls and compared to the rosette sheets to ensure all expected data are present. Errors were found in cast #78 in the association of bottle #s and sample #s. That was fixed and the files were put through the Merge process again.

The MRG file data were repeatedly exported to a spreadsheet and checked until all the above problems had been resolved.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise there were few casts without stops and with a steady descent rate, so it is hard to make a judgment about the best setting. A few casts were examined for each fluorometer and the setting seems appropriate.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 or 4 casts from each CTD system and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing.

(a) CTD #0506 - The results looked best overall when a shift of -0.5s was applied to the primary and to the secondary conductivity.

(b) CTD #0443 – 4 casts were tested and the best overall result was with a setting of -0.5 records for both conductivity channels, though for the primary -0.7 looked better for a few features and for the secondary both -0.2 or -0.7 looked best in some places. There are problems in some casts, especially at P26, with salinity spikes in high T and C gradients around 20 to 40db. Fine-tuning of the shift settings reduced the secondary spikes greatly. While it reduced the very large primary spikes there remain some spikes and it is noted that the primary salinity seems lower relative to the secondary salinity from the level of the spikes until about 100db. Fine-tuning of the shift parameter did not reduce this offset noticeably.

SHIFT was run on conductivity for casts 1-33 using -0.5 records on both channels and for casts 36-129 using -0.5 records for the primary and +0.2 records for the secondary.

Dissolved Oxygen

The usual tests for Dissolved Oxygen alignment did not work well for sensor #1176 because so many casts were interrupted by crashes and because there were so many bottles fired. A check was made of one cast to see that the setting of +90records looked reasonable; that has been used for other recent cruises for this sensor. It did look ok, so an advance of +90 records was applied to casts 1-33.

For sensor #1438 there were some casts with only a few bottles fired, so 2 were checked and the result suggests an advance of between +60 and +70 records is appropriate, so all data were advanced by +65 records for casts #36-129. For sensor #1483 a shift of +60 records was applied.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were many warnings:

- Cast #26 – The file contains a repeated section as has been noted before in cruises with the sort of communication problems experienced during this cruise. Fortunately DELETE deals very well with such “double dumps” and made an appropriate choice for data between 919db and 924db.

- Cast #29 – There were many warnings for this cast. They all derive from the upcast where there were pressure spikes. Examination of the DELETE file shows the last 3 records come from upcast spikes and the record just before those 3 has bad salinity and DO values. The last 3 records were removed and pad values were entered for the salinity and DO concentration for the record just before those. The edited file was named DEL1 and that was put through CLEAN to fix the headers, with output DEL.

- Cast #40 – There were many warnings for this cast, but they all come from the upcast and the DEL file looks fine, though there are some odd small spikes in the salinity. These are best dealt with in CTDEDIT.

- Cast #42 – There were many warnings for this cast, but they come from the upcast section and the only problem in the DEL file is 2 bad records at the end. Those were removed and CLEAN was used to fix the file headers.

- Cast #45 – There were many warnings for this cast, but all come from the upcast section. There were 33 bad records at the end of the DEL file which were removed and the file headers were cleaned.

- Cast #55 – There were many warnings for this cast. Some are due to a problem like that in cast #26 where a block of data was dumped twice; again Delete worked well for that. There are also some large spikes right at the bottom of the cast that led to 15 bad records at the end of the DEL file. Those were removed and the headers were cleaned.

- Cast #58 – There were many warnings for this cast that derived from the upcast but left some bad values at the end of the DEL file. 34 records were removed from the end of the DEL file and the headers were cleaned.

- Cast #65 – There were many warnings for this cast and they all come from pressure spikes towards the end of the cast. DELETE got rid of many of the bad records, but so many records are removed that what is left is suspect, with large gaps and bad salinity values. There are few records below 1624db so that it is hard to judge the quality of what remains. All data below 1624db were removed and the file was put through CLEAN to fix the headers.

- Cast #68 – There were many warnings for this cast. They derive from the upcast and led to 18 bad records at the end of the DEL file. Those were removed and the headers cleaned.

Plots were made after these adjustments and no serious problems remain, though clearly some spikes will need cleaning in CTDEDIT.

Because cast #61 did not start until 700db and there was only the single cast at that station, the data were put through REVERSE and DELETE to produce a DELREV file. These will be considered as a source of a full profile, though the data look very noisy and upcast data are difficult to edit.

Files 52.DEL and 53.DEL were renamed 52.DEL1 and 52.DEL2 and put through JOIN to produce 52.DEL. The output file was examined and a plot looks fine other than the expected 28db gap below 19.7db.

The same process was applied to 55.DEL and 56.DEL and it looks fine other than a 7db gap below 2278.7db.

##### Dissolved Oxygen Calibration Study – sensor #1438

These notes are from the original processing. See end of report for reprocessing details. The fits were different but variability comments are valid.

Study of 2010-14

The problems with CTD #1176 have been seen before and a scheme for recalibration has been found. The problems for CTD #1438 have not been seen before and it is not obvious how to handle this data.

There are only 3 bottle casts using this sensor. The COMPARE results show reasonable fits between 25 and 2000db, but near the surface and bottom the bottles from all 3 casts are outliers. For the top 25db this is not terribly unusual as the CTD sensor has some difficulty adjusting to subsurface maxima and sometimes the local gradient may make the 1.5m offset of bottles to CTD DO sensor more significant than elsewhere. The slope of the fit of differences changes at subsurface maxima, making fits less reliable there. The same happens at the DO minimum, so we do expect a little misfit at that point. But we do not expect the major outliers seen below 2400db and getting worse as pressure increases. This looks like a hysteresis problem. More confusing is that the slope for cast #55 is notably higher than for the other two and has an offset somewhat higher than the other two. For all three the slopes are much higher than in previous uses of this sensor.

The slope/offsets are 1.143/+0.027 for cast #55, 1.126/-0.009 for #73 and 1.130/+0.004 for #86.

The larger offset of cast #55 looks a little like the phenomenon noted from DO sensor #1176, but it is much smaller and a similar offset was found when it was used during 2010-01. A check was made of DO voltage profiles, but there does not seem to be any increased noise level for cast #55.

To see if there might be a problem with the upcast CTD DO data, a comparison of downcast CTD data with upcast bottles was done. The 3 bottle casts were opened in EXCEL and downcast CTD values were added, based on matching pressure in the DEL files. That is the easiest approach, but if a more rigorous comparison is justified then density should be derived and used for the match. The plot of differences between downcast CTD data and bottles differed from the COMPARE results in that the most outliers were at 25db and 125-300db, in the highest gradients. That is to be expected from samples taken “on the fly”. The comparison with upcast data is better overall, as expected, so there is no evidence of upcast problems. (See files 2010-14-****-down_vs_Bottle_DO.xls)

If the problem is hysteresis, then do we need to reconvert the CTD data with different parameters? Tests were run on cast #86, the deepest of the bottle casts for this DO sensor and the ones with the most striking outliers. Rosette files were converted using a variety of settings to see if there was any possibility of improving these data. The following runs were included:
1. Hysteresis correction applied and same H1 & H3 parameters as used for 2010-01 (-3.3e-2 & 1.45e3).

2. No hysteresis correction.

3 / 4. Hysteresis correction applied with same H1 as for 2010-01 and H3 settings -3.6e-2/-3.0e-2.

5 / 6. Hysteresis correction applied with same H3 as for 2010-01 and H3 settings.1.35e3/1.55e3.

The SeaBird routine Bottle Summary was used to average the data. A table was produced with average values for bottles from 1500db to 4300db from each of these runs. Plots showed little difference in all runs except for #2 which looked notably worse than all the others. So the hysteresis correction is important and has improved the data. Test 4 produced the best results at depth with test 5 showing some improvement as well, so more depths were plotted for those to ensure that matters had not gotten worse in shallower waters and they had not. So while a higher value of H1 and lower H3 improved things slightly, the improvement is slight, so more tests were run with more extreme settings of H1 and H3.

7. Hysteresis correction applied with H1 set to -2.5e-002 and H3 set to 1.1e+003. This case went too far, did not help deep data and made shallower data worse.

While fine-tuning might come up with a slight improvement that does not make shallow data worse, the results do not indicate that any setting is going to make this data acceptable. The best we can do to address the hysteresis problem is to replace deep data with pad values, and do fits based on only data above that level. The bottles look fine at 2000db and just a little off at 2500db, so a cut-off of 2200db looks reasonable.

This still leaves the mystery of cast #55 to be considered. This cast included some very low titrated DO values, ~0.32mL/L, and from the distribution it seems likely there were lower values around 1150db. We might imagine that the problem involves failure of the SBE sensor to measure low values well, but the fit looks much the same whether or not we include the lowest DO values and even allowing all the shallow comparisons makes little difference to the fit. Furthermore, the sensor values are lower than expected, not higher, whereas in cases where the sensor fails in low oxygen environments, it is because it fails to get low enough. Forcing the fit through the origin leads to a fit that only seems to suit the middle part of the range, not the low or high DO areas. This is a puzzle, and unlike the case with the other sensor, we have no way of predicting which casts will fit this pattern.

Study of 2010-15

We do however, have the possibility that cruise 2010-15 may shed light on the problem since the same DO sensor was used and the cruise immediately followed 2010-14. The results of COMPARE for the later cruise also looked strange with many outliers. When many outliers were removed the fit was:

Bottle DO = 1.1931 * CTD DO - 0.0097

This slope is much higher than seen in 2010-14 and there is no hint of drift with time during 2010-15. There was however, a notable difference seen between casts inshore of the 100m contour and those offshore of that. The offshore outliers were seen mostly in the top 300db where there were some very low T, S and DO gradients with minimal stability in the profiles. The DO signal often failed to equilibrate during 30s+ bottle stops and even the temperature was still changing in some cases.

When separated into two groups the fit for on-shelf casts when outliers were removed based on residuals was:

Bottle DO = 1.1942 * CTD DO - 0.0129

The fit for the off-shelf casts with outliers removed based on residuals was:

Bottle DO = 1.1717 * CTD DO - 0.027

Do these two groups also explain the change in slope between this cruise and 2010-15, and do they explain the variations in slope during this cruise?

Conclusions

Profiles from the 3 2010-14 casts were re-examined but there is no indication of the sort of mixing seen in the later cruise, though there are sections of very low gradient, where the DO sensor might “catch up” better and thus produce a lower slope. For cast #55 the temperature trace has a local maximum around 125db but the salinity gradient is high so this profile is stable. The dissolved oxygen appears to have some little blips and primary salinity is noisy, so there may have been a problem with the flow rate. Casts #73 and 86 do not have that problem. While there are differences in the profiles, the cast at P20 looks more like the P16 cast than the P26 cast, so there is nothing in that to explain the different fit of oxygen bottles against SBE DO. Most likely the spikiness of cast #55 (also seen in the primary salinity) is related to the different fit, so the fits from cast #73 and 86 are a better choice overall.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

(a) The primary sensors for CTD #0506 have been used for only one cruise that has been processed while the secondary have been used for 3. During 2010-36 the primary salinity was found to be low by 0.002 and the secondary very close to bottles. The primary salinity drifted downwards through the cruise and there was only a slight change in the secondary salinity. The differences between the sensors were about 0.0014 at the beginning and 0.004 at the end with the primary lower. The secondary sensors were found to be high by 0.003 and 0.002 in May and June 2010.

(b) The primary conductivity sensor for CTD #0443 had been used many times with a variety of temperature sensors. Bottle comparisons were often unavailable or limited to surface samples or showing a lot of scatter that led to concerns about sampling protocols. Poor seals and long waits for analysis likely led to bottle salinity being too high, thus making CTD salinity look low. The secondary conductivity sensor was used for only 2 cruises that have been processed; one had only surface bottles and the other looked poor with a lot of scatter. A comparison with an Argo float showed the primary to be low by ~0.006 and the secondary to be within 0.001.

2. Dissolved Oxygen

(a) Since being calibrated in Nov. 2009, DO sensor 1176 has been used for 3 cruises that have been processed. The corrections found for the first 2 were:

Bottle DO = 1.0339 * CTD DO - 0.0043 (2010-12)

Bottle DO = 1.0414 * CTD DO + 0.0057 (2010-13)

For 2010-36 serious problems were found with the sensor and the bottle comparison suggested 2 possible fits that were randomly spread through the cruise. The fits were associated with a different level of noise in the DO voltage.

Bottle DO = 1.0491 * CTD DO - 0.0237 (Fit for quiet DO voltage)

Bottle DO = 1.0347 * CTD DO + 0.2383 (Fir for noisy DO voltage)

(b) Sensor #1438 was used for cruises in June 2009 and February 2010 when the slope/offset for the fits were 1.0551/+.0204 and 1.0784/.0218. (NOTE 1438 was not used for this cruise.)

3. Pressure

Since its latest factory recalibration in May 2006 this sensor has been used mainly on the Ricker. The factory offset was used for the Ricker cruises (~-0.8db). It is likely since it is an older sensor that there is some drift, but it is harder to judge from Ricker data since the CTD usually is started fairly deep. During 3 recent Tully cruises the pressure looked low by at least 0.7db so a net offset of 0db was used (an addition of 0.8db to the factory offset.)

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were many excursions from the historic ranges. Some involved spikes that will be removed later. Others appear to be from the seasonal mixed layer being a little deeper than usual. There are excursions towards both ends of the ranges, so these do not look like systemic instrumental problems. However, they may indicate sporadic effects of the spiking and will be examined individually:

Cast 16 –Salinity high around 50-65db – temperature within the 3std dev lines, but near the minima. Below 800m it looks like nearby casts, but stands out as different from both P4 and P6 above that. It is assumed this is an incursion from the south as there is no sign of instrumental problems.

Cast 17 – Salinity is at low end of range from 34-39db. This P4 cast is similar to the P6 cast but very different from the P5 cast. It looks ok.

Cast 26 – Only excursions were from a very spiky section in salinity, 1000-1200db, both channels, the bad section will be removed in editing.

Cast 40 – Salinity odd around 1600db, will need editing.

Cast 41 –Higher T and lower S than usual around 40db.

Cast 42 –Salinity on the low end near surface.

Cast 76 – Salinity high 100-125db but does not look like spiking.

Cast 78 – As for cast 76.

Repeat Casts –

Two casts at station P15 separated by about 8 days and using CTD #0443 showed significant differences above 900db with the later cruise looking cooler and fresher, but below that they looked very close, with temperature differences along lines of constant sigma-t being ~0.002C°  to 0.004C° and salinity differences being <0.001. However, the later cast appeared warmer and saltier in the primary channels, but cooler and fresher in the secondary. So one or both of these #0443 T-C pairs must have drifted a little; this fits the observation in COMPARE that both salinity channels seemed to be drifting in opposite directions. The differences over the 8 days are much smaller than the differences between the 2 channels along sigma-t lines, so it appears that neither T-C pair were drifting seriously.

Station P12 is very useful in that we have 2 casts using CTD #0506 that are close in time (#29 and 32) and then a CTD #0443 cast (#40) about 8 hours after those. Then on the return trip cast #128 using CTD #0443 also sampled P12 so we can study the drift in those sensors over 9 days. This is an area subject to rapid change but there are some deep casts and we at least hope for smaller changes below 1500db.

The two early casts  which were run 5 hours apart, show considerable variability above 1100db with much interleaving obvious on a T-S plot, but the two casts are virtually indistinguishable below that for both primary and secondary channels. The secondary pair are warmer (by 0.01C°) and saltier (by 0.003) than the primary. The bottle comparison found the secondary salinity from CTD #0506 was higher than the primary by about 0.0035, so roughly the same as seen in these plots.

Next T-S plots of casts #32 and 40 were examined; they occurred about 8 hours apart and with different CTDs. The secondary data from #0443 looks almost exactly the same as the secondary from #0506. Since the #0506 primary salinity was found to be higher than bottles by about 0.0005 and for #0443 lower by 0.0016, we would expect them to differ by ~0.002. The primary from #0506 is cooler and fresher than the primary from #0506 and fresher than both secondary channels.

Finally comparing casts #40 and 129 shows temporal drift in CD #0443. The primary looks a little warmer and saltier while the secondary looks cooler and fresher. Again, this fits the results from COMPARE where the primary salinity of #0443 started low and drifted upwards, while the secondary started slightly high and drifted lower.

Post-Cruise Calibration

CTD #0506 – not available.

CTD #0443 – A post-cruise calibration on Dec. 28, 2010 indicates that both the primary and secondary conductivity sensors would lead to salinity values low by 0.003. A further error would arise from the secondary temperature channel which was reading high by about 0.003°C which would lead to salinity being low by a further 0.003 for a cumulative error of -0.006 in the secondary salinity. The primary temperature drift was negligible. These tests are done at 3S/m conductivity, which is lower than all values seen on this cruise.

If we assume that the drift for the primary conductivity occurred mostly before August, then the finding in COMPARE that it is low by 0.006 while the drift estimate is that it is low by only 0.003 implies either the samples were slightly contaminated by salt and/or evaporation or the Autosal is reading low. And since the same conditions would apply to the secondary which appears to be low by ~0.002 that would imply it is actually high by 0.001, which is way out of line with the drift estimate of -0.006.

If we assume that the secondary conductivity has not drifted much since August, then while it appears to be low by 0.002 and is really low by 0.006, then the Autosal must be reading high or somehow fresh water got in the samples. And it would imply that the primary which appears to be low by 0.006 is really low by 0.01 which is out of line with the drift estimate for that sensor.

So it seems likely that there was conductivity drift between August and December. COMPARE suggests that the primary salinity was increasing while the secondary was decreasing. So it not unreasonable that the conductivity effect in August of -0.006 salinity units for the primary could move upwards to -0.003 by December, while the secondary being low by <-0.002 salinity units in August could move down to -0.003 in December. The secondary temperature drift, however, seems out of line and will be investigated in section 17.

##### Temperature and Conductivity Sensor Choice for Archive

For CTD #0506 the secondary temperature and salinity look best based on history and the bottle comparison. No recalibration of salinity appears justified.

For CTD #0443 the post-cruise calibration makes it more difficult to decide which sensor to choose for archiving. Based on history and the bottle comparison, the secondary channels look like the better choice. However, the factory report shows secondary salinity low by 0.003 due to conductivity sensor drift and low by a further 0.003 due to temperature sensor drift, while the bottle comparison indicated it was low by <0.002. We have had many problems in the past 2 years that appear to be due to poor sampling protocols and delays in analysis, but they all lead to the result that CTD salinity looks lower than it really is, not higher. So how do we reconcile these conclusions?

One thing that stands out in the history of the temperature sensors is that the difference between primary and secondary sensors during 5 recent cruises has never been >0.0005C°. At the time of the post-cruise calibration the primary temperature was low by about 0.0002°C and the secondary high by 0.0034°C, so the difference at that time was 0.0036C°. That suggests that the secondary temperature sensor has sustained a major shift in calibration after August 2010. The sign of the temperature differences changed towards the end of this cruise; the differences were still small, but that might have marked the beginning of the drift, though the only evidence for this is from some very noisy casts and the differences are still very small. There has been at least one cruise that used these sensors since 2010-14, so we can assume some of the drift occurred during that cruise. If we assume most of the secondary temperature drift came after 2010-14, as well as a little of each of the other sensor drifts, then we would expect to find both primary and secondary salinity to be low by ~0.003. There is some evidence of pressure dependence in the primary conductivity and the factory calibration data might not catch that because of the low conductivity value used for the test. The salinity differences at 500db are only 0.001 which is more in line with the factory observations (if we assume no drift due to temperature).

Another possible issue is the difference between downcast and upcast data. The bottles come from the upcast where we generally get poorer data quality because there is more wake effect and alignment seems to vary perhaps due to vibrations or other motion of the CTD. However, we don’t usually see that during stops for bottles and we find consistent differences between sensors based on the bottle comparison and on the downcast differences like those noted in section 8. Just to be sure that the differences are consistent, the SAM file from cast #118 was examined in EXCEL and the differences and standard deviations found for 2000db data. The differences in temperature, conductivity and salinity (Secondary – Primary) were found to be -0.00011, +0.00026 and +0.00335 with standard deviations of 0.0002, 0.00025 and 0.00033. The 1900db downcast results for cast #90 were very similar at -0.0001, +0.0003 and +0.0035. So there does not appear to be any significant difference between data from downcasts and upcasts during stops.

Finally, a look at profiles argues in favour of using the secondary sensors. Many casts using this CTD were affected by spiking around 20 to 40db where there was a very sharp temperature change, especially at station P26. The alignment step greatly reduced these spikes in the secondary salinity, but while the primary was slightly improved, spikes remain. Furthermore, in the region of fairly high T and C gradients that occur between about 20db and 150db, the primary salinity data are lower relative to the secondary salinity than they are in the rest of the profile. The temperature and conductivity are not spiky in that region, so the problem is presumably a matter of alignment. The pumps were on through that period. It looks as though there was some plumbing problem that prevented good alignment, yet fine-tuning of the SHIFT settings did not help.

One cast with many bottles was examined in detail. The secondary salinity data are reasonably close to the bottles both in both down and upcasts and the primary salinity data are quite close for the upcasts, but for the downcast they are low by 0.01 to 0.02 at 25 and 50db. So when the CTD is stopped the primary data generally look ok, but not while in motion downwards.

So there is some problem with either plumbing or alignment causing spiky and offset primary salinity in the high temperature gradient zone, and there is pressure dependence in the primary salinity. The secondary salinity looks like the best choice, especially given the comparison of CTDs #0506 and #0443 showing secondary channels from each in excellent agreement. The bottle comparison suggests the salinity is low by 0.0016 while the post-cruise calibration report indicates it was low by 0.003 in December. No temperature recalibration will be done since the evidence suggests that most of the drift in that sensor occurred after August 2010. There is some evidence that the secondary salinity was drifting downwards but likely by only 0.002 over the period during which it was used during this cruise, being very close to the bottles early and low by 0.003 at the end. However, there is so much noise that a time-dependent correction is not justified.

All salinity channels should be recalibrated since in at least one case the primary channels will be needed.

##### DETAILED EDITING

The secondary T and S sensors were chosen for archiving. For CTD #0443, the salinity was very noisy in the top 200db for many casts; this noise is not associated with shed wakes or large spikes, and is mostly 2-sided which suggests it is due to flow-rate problems, not alignment of T and C sensors. Given that the spikes are 2-sided and fairly small, they should not have a large effect on metre-averaged data. This sort of noise stands out more in moderate gradients; the noise is probably masked in stronger gradients and does not occur in very well-mixed waters. Editing was limited to larger spikes likely to survive metre-averaging.

A few problems were encountered:
26. As noted earlier the salinity data are bad between 978 and 1239db for both primary and secondary channels. Those points were removed.

40. The secondary salinity has a large section with bad data, and there is a short section of bad secondary temperature. Primary channels will be selected for this cast.

61. The downcast started at 770db. That file was edited and then an attempt was made to edit the reversed file 61.delrev. The data were very noisy and distinguishing good from bad was mostly impossible since the CTD was measuring in the wake of the CTD which had a noisy ascent rate. A file was prepared for the use of the chief scientist in case there are requests for this data, but it is not considered suitable for archiving.

110. This was a repeat of #109 and was run only to obtain bottle samples from the top 20db because the bottles leaked during #109. The data look very poor from both the downcast and upcast. The bottle files named #110 were changed to #109 and the cast list adjusted. The profile file #110 will not be processed further and that cast list was also adjusted.

All EDU files were copied to EDT.

T-S plots were examined and cast #79 was found to need a 2nd run through CTDEDIT. The output file, *.EDV, was then copied to EDT.

The EDT files were bin-averaged.

Profile plots were made of T, S, PAR, Fluorescence, DO, transmissivity and a few problems were identified:
Cast #55 – There are spikes in DO around 2350db, so CTDEDIT was used to remove them, producing file EDV. PAR spikes were edited in that file using a text editor. The file was then copied to EDT.

Cast #65 had a few spikes at the bottom, so a text editor was used to fix those in file EDU which was then copied to EDT.

##### Initial Recalibration

The SBE Dissolved Oxygen data were recalibrated in three groups depending on which fit worked best in COMPARE.

Fit #1 for casts 1-19 is based on 3 bottle casts. It is very close to Fit #1 from 2010-36.

Fit #2 for casts 20-32 is taken from 2010-26 since there was only 1 bottle cast available from 2010-14.

Fit #3 was applied to all casts using DO sensor 1438, casts 36 to the end, and is based on 2 bottle casts only.

Fit 1:    Corrected CTD DO = 1.0497 * CTD DO – 0.0115

Fit 2:    Corrected CTD DO = 1.0347 * CTD DO + 0.2383

Fit 3:    Corrected CTD DO = 1.1273 * CTD DO – 0.0016

For casts #1-32 a correction of +0.0032 was applied to the primary salinity channel and -0.0005 to the secondary salinity.

For casts #36 – 129 a correction of +0.006 was applied to the primary salinity channel and +0.0015 was applied to the secondary channel.

The SAM files were recalibrated using file 2010-14-recal1.ccf.

COMPARE was rerun and shows that the recalibrations were applied correctly. There is some evidence of time-dependence in the salinity for CTD #0443, but the errors appear to be ~0.002.

(See 2010-14-sal-comp2.xls and 2010-14-dox-comp2.xls.)

The MRGCLN2 files were then recalibrated using the same calibration control file.

The same calibrations were applied to the EDT files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The scatter in the differences against DO concentration is much higher than usual, with many of the outliers coming from the top 25db.

- Sensor #1176: The differences against file pair number shows that the first 3 casts using sensor #1176 are high by about 0.06mL/L. Cast #22 had an average difference of 0.001mL/L.  When this sensor was used in July 2010 it was found that data above the DO minimum would be improved by applying an offset of -0.4mL/L to -0.7mL/L, so that is a similar result to this. But the offset was not applied because it made data below the oxygen minimum worse. For 2010-14 we have no bottle data near or below the oxygen minimum except for cast #22 which is the one with ~0 average difference. The deepest bottles for that cast do show an offset of opposite sign to that in most of the profile. Since many of the casts to which this correction would be applied do to to 2000db, it is not wise to apply a correction.

- For sensor #1438 the picture is quite different with a lot of noise and the sensor being low by an average of 0.1mL/L when some outliers are excluded including the points thought to display hysteresis. A plot of differences against pressure shows larger differences at the surface and smaller ones from 500m to 2000db. This sort of study usually leads to a clear picture in plots against pressure and/or DO concentration. There is no clear picture here from either plot, so no further correction is justified. However, the plots do at least offer some information about the limitations of the DO data for this sensor, suggestion potential errors of ±0.4mL/L from 0 to 300db, ±0.2mL/L from 300 to 500db and ±0.1mL/L from 500 to 2000db. Below 2000db

No further corrections will be applied to the SBE Dissolved Oxygen channel.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db, except for files for which the fluorometer was not connected – those were skipped. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

Lists were produced of files with and without PAR sensor channel.

Plots were made of those casts to ensure they were accurate; a few of the No Par casts did have a few spikes, but nothing that looks reasonable.

The AVG files were copied to CLN2.

The AVG files for casts #36 to the end were put through CLEAN to replace the DO values below 2200db with pad values, with output CLN2.

REMOVE was on all casts except #40 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on cast #40 only to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

As well the PAR channel was removed from casts for which it was not mounted.

SBE DO, Transmissivity and SBE Fluorescence were removed from casts 41, 42 and 45.

Transmissivity was also removed from 36 and 37 and DO from 36, 38 and 40.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the CTD serial number, to add “Mid-ship” to the instrument location section and to add the following comments for CTD #0506:

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

As has been noted from several other cruises using SBE Dissolved Oxygen

sensor #1176 in summer 2010, the data are odd for some casts (casts #22 to 32)

with concentration values offset by approximately 0.24mL/L from the normal

looking casts and a higher noise level in the raw oxygen voltage.

Data were recalibrated to correct for this offset, but there were only 3 casts

with titrated DO for the normal casts and 1 from the odd casts. Establishing

which casts require the special recalibration was based on the noise level

in the DO voltage.

The quality of SBE Dissolved Oxygen calibration is considered lower than

usual and there are insufficient data to make precision estimates.

For details on the processing see processing report: 2010-14-proc.doc.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments for CTD #0443:

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

There are only 3 casts with titrated dissolved oxygen samples that could be

compared to the SBE Dissolved Oxygen data for sensor #1438, and one

of them looked significantly different in the comparison from the other 2.

Recalibration of DO was based on the 2 casts that looked similar.

There are insufficient data to make a precision estimate.

For SBE Dissolved Oxygen sensor #1438 values below 2200db were replaced

by pad values because there is evidence of significant hysteresis.

The salinity from CTD #0443 showed evidence of flow rate problems in the top 200m;

most of that effect will be removed by bin-averaging.

For details on the processing see processing report: 2010-14-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 85% at P1 to 140% in Saanich Inlet with all values to the west of P3 being between 90% and 110%. At P2 and P3 the values were from 110% to 125%. P2 also had the highest SBE Fluorescence values.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

They were copied to MRGCLN3 and then CLEAN was run on casts #36 to 129 to replace DO values with pad values when pressure is >2200db.

REMOVE was run using the same settings as for the CTD files.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data..

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no discrepancies were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no outliers were identified.

Standards check was run on all files and no errors were found.

##### 25. Thermosalinograph Data

Data were provided in 6 hex files. There was loop sampling for Salinity, Chlorophyll and Nutrient samples at 4 locations in Juan de Fuca Strait; those were combined in file 2010-14-tsg-loop-comp.xls. There were 4 Dissolved Oxygen samples taken during stops (one is flagged as the values look unbelievable), but there was no DO sampling on the TSG, so these data were not added to the comparison spreadsheet. Times were added to the spreadsheet based on log records.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong date and scale factor. After that correction the CON file was saved as 2010-14-tsg.con.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were produced. Overall the records look good, but a few problems were noted:

- Flow interruption – The flow rate is mostly steady at ~1 but it was higher during file #1 which was very short and dropped to zero for about 10 hours during file #4.

- The salinity look oddly low at the end of file #1 and beginning of file #2, but these sections are from Haro Strait, so possibly they are ok.

There are a few significant shifts in temperature, salinity or fluorescence, often while the ship was on station or just starting or stopping, so they may be associated with ship movements, thrusters etc.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2010-14-ctd-tsg-comp.xls. All the data came from ~4.5db.

File #1 does not overlap with any CTD casts. There are some odd jumps in salinity and in temperature differences between the intake and lab, but given that the ship was travelling through Haro Strait, these may well be real. The temperature does not show the same variation but judging by CTD casts in the same area in September, the temperature is better mixed near the surface. With a ship in motion water in the loop is likely to come from above 4.5m. There is a drop in flow rate near the end of the file with some very low salinity values. These look suspicious and should be edited. An estimate of loop transit time was made by finding the minimum standard deviation in the differences calculated as Temp Intake scan (X) – Temp Lab scan (X + N). The standard deviations were lowest when N=7, indicating a transit time of 7 scans or 3.5min. The average flow rate at that time was 1.2. In later files, the flow rate was lower so the transit time is likely to be a little higher.

File #2 has a lot of variability in the early section which is in Haro Strait, so much of what looks like noise would be real. A study of the time in the loop for this file suggests between 8 and 9 scans, or ~4.25min. The average flow rate was 1.03.

File #3 has an average flow rate of 1.07. There is a suspicious drop in salinity late in the file with no associated changes in flow rate, fluorescence or temperature. There are other jumps in values that are likely due to ship manoeuvring.

File #4 has a large section with no flow – all records should be replaced with pad values except the times and positions for that section.

Files #5 and 6 look ok.

The 5 TSG files that overlap CTD casts were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. Those data were added to 2010-14-ctd-tsg-comp.xls. TSG values were found in the same way for times of underway loop sampling and added to file 2010-14-tsg-loop-comp.xls.

File 2010-14-ctd-tsg-comp.xls has 70 matches, but the flow rate was ~0 for 8 of them. In some recent uses of this equipment latitude and longitude have become stuck, or times have been wrong due to missing scans. In this file the times from the CTD and TSG are matched, so comparing positions will turn up problems of that sort. The differences in latitude were all <0.0004° and in latitude <0.0003°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

File 2010-14-ROS-LOOP-comp.xls was prepared combining with all loop samples and corresponding data from rosette bottles at 6db.

- T1 vs T2 The intake thermistor was connected throughout the cruise. The first two files were short and noisy, the first so much so that it will not be analyzed further. The average differences for files #2, 3, 4, 5, and 6 were 0.24, 0.19, 0.15, 0.13 and 0.15Cº and the average intake temperatures were 11.0, 11.8, 15.0, 14.1 and 14.7ºC, respectively. It is expected that as the average intake increases the differences will go down, so file #5 stands out as odd.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. As noted earlier the flow rate was ~0 during 8 CTD casts. Plots show that the lab temperature and salinity eventually became quite flat during that time. Then when the flow rate once again went to values ~1 variability returns, so clearly the flow really was off. The differences between the TSG and CTD also became larger during the time with no flow. While the differences between the intake temperature and CTD should be the same when the flow is off, those data were not included in comparisons that follow in case the intake thermistor is affected.

When all data with non-zero flow were included the TSG intake temperature was higher than the CTD by 0.057Cº while the median difference was 0.005Cº. Many of the largest outliers occur when the standard deviation in the TSG temperatures are low. When the differences and standard deviations are plotted against CTD cast number it is clear that the highest differences occur early and late in the cruise when the CTD was sampling waters with higher variability such as at station P2. A section that is harder to understand occurs from casts #114 to 121 where the temperature differences are ~0.08Cº, very consistently. These differences are higher than expected, especially since the standard deviations are low and the salinity differences do not stand out from other values. The differences return to ~0.01Cº for cast #124 onwards. The differences between the CTD and surface rosette samples show the CTD to be quite consistently lower than the bottle samples by ~0.002 through that period. All of the unusual data come from a single TSG file, #5, which was noted above to have smaller differences between intake and lab temperatures than expected given the average intake temperature. There are no notes in the log to indicate there was a problem with this file, but it is relatively short, so might have been stopped because of a problem noted at sea.

The TSG salinity is lower than the CTD by an average of 0.416 and a median of 0.108.  A plot made of differences versus standard deviation in TSG salinity shows values approaching a value close to the median. There was one extreme outlier at station P4 with TSG salinity ~15 and CTD salinity ~32. The TSG data are obviously bad in this area – towards the end of file #3.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.8 to 21.2 and a median of 2.4. When plotted together the patterns are similar. The largest ratios are associated with the lowest CTD fluorescence. (See 2010-14-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons  A comparison was done in spreadsheet 2010-14-loop-TSG-comp.xls of loop salinity and chlorophyll samples with TSG salinity and fluorescence from the same time. There are only 4 points of comparison and all are from Juan de Fuca Strait. The TSG salinity was lower than bottle salinity by between 0.65 and 0.73. The TSG fluorescence ranged from 16 to 44 times the extracted chlorophyll. (See file 2010-14-loop-TSG-comp.xls.).

- Loop Bottle - Rosette Comparisons

There were no cases of loop bottles being taken during a CTD cast.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and were used for 2009-10 (with a different intake thermistor), 2009-11, 2010-01 2010-13 and 2010-22 (and 2 other cruises with clock problems so that data were not useful for recalibration). The salinity was low by about 0.02, 0.02, 0.06/0.16, 0.06 and 0.06 for those cruises with the change for 2010-01 being associated with a change of flow rate. During 2010-36 the differences started at 0.06 but were at ~0.4 by the end of the cruise, likely due to a fouling problem.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was steady except for interruptions in files #1 and #4. The transit time from intake to lab varies from about 3.5 to 4.25 minutes.

3. The temperature in the increases by an average of 0.14Cº between intake and lab, but warming varies with intake temperatures.

4. The TSG intake temperature appears to be within 0.005Cº of the CTD, based on median values, but there is a lot of variability. For file #4 the difference is 0.08Cº.

5. Salinity is low by a median value of ~0.108 with little variation between casts #40 and 124. The differences are very noisy up to cast #12.

6. There is a lot of variability in the ratio of TSG fluorescence to either CTD fluorescence or loop CHL. The TSG fluorometer had values higher than the CTD fluorometer by a median factor of 2.4 (range of ~1 to ~21) and was higher than the 4 loop CHL samples by a factor of 16 to 44.

7. The difference between the TSG salinity and the loop salinity is much higher than the difference between the TSG and CTD salinity. There were only 4 loops samples and none during a CTD stop. A possible explanation might be bubbles in samples taken while underway.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. Because only 3 channels can be displayed, the files were opened several times to remove data where the flow rate was zero – position and time channels need to be kept, so simple removal of records won’t work. A few salinity points were also removed where values were extremely low, but many low values were left in the first 2 files, because they are from Haro Strait where it is not clear what is real and what not.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

File 2010-14-tsg-recal1.ccf was prepared to adjust salinity by adding 0.108. A few values were checked to ensure it was applied correctly and after an error was fixed, the results looked right.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

The final loop file 2010-14-che-loop.csv was prepared by the chief scientist including data from the final CTD files and samples from the loop or from 5m bottles. That spreadsheet was simplified, header names and formats were adjusted and unneeded channels were removed and saved as a CSV file and the 6-line header template inserted and adjusted as needed. The resulting file (2010-14-loop-6linehdr.csv) was converted to IOS format, put through CLEAN and HEADEDIT to get start and stop times and positions, and to add general comments and specific comments for flagged values. The final file was named 2010-14-surface.loop.

Particulars

3 – Pumps off for downcast

17/18/19 – 1st ends with spikes at bottom, 18 no file, l9 upcast from 25db only

23/24 – 1st down to 461 then all spikes – 2nd upcast

26. Spikes

32/33 - 1st down to 3200 then all spikes – 2nd all spikes

37 – Down to 296 then spiky, aborted – possibly of some use, but rerun done

38 – Rerun of 37.

41. This cast was aborted at 100db due to error message, but data to that point are ok, so was processed.

45/46 – 1st has spikes at end of upcast – 46 finishes upcast

52/53 – 1st part of downcast, 2nd rest of down plus up

55/56 - 1st part of downcast, 2nd rest of down plus up

58/59 – 1st downcast plus part of upcast, 2nd rest of upcast

61. Downcast archiving did not start until 766db; upcast data not suitable for archive.

63/64 - 1st downcast plus part of upcast, 2nd rest of upcast

65/66 – 1st very messy – downcast, 2nd upcast

68/69 - 1st downcast plus part of upcast, 2nd rest of upcast

76/77 -1st downcast plus part of upcast, 2nd rest of upcast

110 – Just down to 20db for bottles because of leak on 109

121/122 - Separate up and down done deliberately, no spikes.

Reprocessing of SBE DO channel. July 2011

In June 2011 it was discovered that there was an error in the original processing of data from this cruise. The serial number of the SBE Dissolved Oxygen sensor was entered incorrectly in the log and in the 2 configuration files used for CTD #0443 during casts 36-129. Initial tests showed the results made sufficient difference to the results to justify reprocessing, though the hysteresis noted in the original processing remains and all the comments on the questions about the DO data are still valid. There is little on which to base the DO calibration.

Since the error only affects DO data and the SBE DO sensor was not mounted for casts 36-46, only casts 47 to 129 were reprocessed. The same configuration file, 2010-14-ctd2-1483.con was used for all casts since the only difference between the 2 configuration files used in the original processing concerned PAR sensors. The only channels that need be processed again are pressure, temperature and dissolved oxygen, and for the bottle files Bottle Number, Bottle Position and sample number.

Because there were problems with split casts, care will be needed for casts 58/59, 63/64, 68/69 and 76/77.

- Bottle File Reprocessing

The task is to replace the 2 SBE DO channels in the original CHE files. This is complicated by the split casts. First the original files were sorted on bottle number and saved as CHE1

Then using the appropriate configuration file for the DO sensor #1483, all steps were followed in the way described in the original processing for the SBE DO channel except as mentioned. The ROS files were converted with all channels, they were converted to IOS header files (*.BOT). Those files were joined as needed, after adjusting file names with Ultraedit. Sample Numbers were added and they were bin-averaged. The original CST files were sorted on bottle number and merged with SAMAVG files. CLEAN was run again and included replacing data from below 2200db with pad values. COMPARE was rerun to determine how to recalibrate the data. As in the original processing it was found that of the 3 casts with bottle sampling, two were similar and the other quite different. This sensor showed similar variability when it was used during 2010-01, but the slopes were higher for that cruise. The results from casts #73 and 86 look most reliable, but the quality from this sensor must be considered lower than usual.

The files were then recalibrated using the COMPARE result:

SBE DO corrected = 1.0335 *SBD DO – 0.0115

The 2nd DO channel was added, and HEADEDIT was used to fix formats and change the bottle number and bottle position channels to usual names (MRGHDR); then the data were sorted on bottle number as that channel name now refers to Niskin # and that is how the original data is ordered. (There was a problem using firing sequence as merge channel due to name length.). These files are called *.CHE2.

Merge was run choosing all channels except SBE:DO from the CHE1 files and the 2 SBE:Oxygen:Dissolved channels from CHE2, with bottle number as reference channel (CHE3). Those files were reordered to get the 2 SBE:DO channels together and sorted on pressure (CHE5) and finally HEADEDIT was used to add the following comment (CHE):

***********************************************************************

In June 2011 it was discovered that the wrong information was used for

conversion of the SBE Dissolved Oxygen sensor data for CTD #0443.

Sensor #1483 was actually used for this cruise.

The SBE Dissolved Oxygen channel was reprocessed. A new recalibration

scheme was found by rerunning the bottle comparison. The files were

then merged with all the other channels from the original processing.

The hysteresis noted in the original processing remains, and the quality

of the data is still considered of lower quality than usual. For details

see the end of processing report 2010-14-proc.doc.

************************************************************************

The CHE files for casts #36-45 were not reprocessed since the DO sensor was not mounted on the

CTD; however, the calibration summary does contain wrong information so that line was replaced in those files a few irrelevant comments about the DO sensor were removed.

This was a very complex process, so many checks were run to ensure that no data had been lost.

1. A header check found several errors – a channel had been missed in the merge process for all casts, and the upcast from 2 of the split casts hadn’t been renamed to match the downcast, so they were missed in the final merge process.

2. The standard check was run and no problems were found.

3. Data from the new CHE files were exported to a spreadsheet which was added to a similar spreadsheet from the original files. Each entry was compared to be sure no data were missing. Two files had missed a step so did not get sorted on pressure. An error was found because one file had a non-standard name which was fixed in the original processing, but not on the corrections, and extracted chlorophyll samples had not been included in the merges. The only differences found after those corrections were in a few cases where the pressures of 2 bottles were the same, the SORT function had produced a different order. This does not matter, so no further work is needed on the files.

The new DO values were almost all lower by 0.01mL/L or 0.02mL/L except for a few outliers where the differences were very large, up to 1.2 mL/L in either direction. Those were all cases of high temperatures and are all from the top 50db. Only casts #78 or 84 had large differences. These were not casts that were highly affected by spiking, so why are the differences so high there? The profiles do not stand out from other nearby plots, but the DO derivation is complex and may be influenced by combinations of DO and temperature gradients and pressure. The spikes that occurred through this cruise be the source of the few larger differences. See file “2010-14-1438vs1483 files.xls.”

- CTD file reprocessing

# The raw hex files were converted to CNV files with just scan #, pressure, temperature, conductivity and dissolved oxygen voltage.

# DERIVE was used to compute Dissolved Oxygen in ml/L.

# The files were then converted to IOS HEADER format. This required use of a text editor to remove spiky records from the end of files 52, 63 and 76.

# CLEAN was run to replace pad values in the pressure channel.

SHIFT was run on DO only with a setting of +60 records.

DELETE was run with the same parameters as noted in the original processing, with output named *.DELNEW.

JOIN was used to combine 52/53 and 55/56.

BIN AVERAGE was run using 1db bins. Output: AVGNEW

CALIBRATE was run using 2010-15-1483-recal1.ccf.

The 2nd DO channel was derived with output files and CLEAN was then used to remove the SeaBird headers. OXCLN

MERGE was run merging the original CTD files with OXCLN taking all but DO from the CTD files and only the 2 DO channels from OXCLN. OUTPUT MRG

HEADEDIT was used to add a comment about the reprocessing and replace the calibration information for the DO channels and fix the DO formats.

Plots were made to check for spikes and ensure the merges were done properly. Header Check was rerun and no differences were found form the original other than slight

The parameters listed in the Seabird header in the CTD files includes calibration parameters

for the wrong dissolved oxygen sensor. The following are the correct parameters:

#   <sensor Channel="7" >

#     <!-- A/D voltage 1, Oxygen, SBE 43 -->

#     <OxygenSensor SensorID="38" >

#       <SerialNumber>1483</SerialNumber>

#       <CalibrationDate>03-feb-09</CalibrationDate>

#         <!-- Coefficients for Sea-Bird equation - SBE calibration in 2007 and later. -->

#         <Soc>4.1000e-001</Soc>

#         <offset>-0.5055</offset>

#         <A>-1.9487e-003</A>

#         <B> 1.7439e-004</B>

#         <C>-2.8684e-006</C>

#         <D0> 2.5826e+000</D0>

#         <D1> 1.92630e-004</D1>

#         <D2>-4.64800e-002</D2>

#         <E> 3.6000e-002</E>

#         <Tau20> 1.1000</Tau20>

#         <H1>-3.3000e-002</H1>

#         <H2> 5.0000e+003</H2>

#         <H3> 1.4500e+003</H3>

#       </CalibrationCoefficients>

7 October 2011: Corrected error in how merge was done in CTD files.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 06May08 | Factory “ |  |  |
| Conductivity | 1764 | 10Feb09 | “ |  |  |
| Secondary Temp. | 2038 | 06May08 | “ |  |  |
| Secondary Cond. | 3394 | 06Mar09 | “ |  |  |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 10Nov2009 | Factory |  |  |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Fluorometer | 2345 |  | IOS |  |  |
| Pressure Sensor | 69698 | 26May2006 | Factory |  |  |
| Altimeter | 1024 |  |  |  |  |
| Calibration Information CTD #443 |  |  |  |  |  |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 06Mar07 | Factory “ | 5Jan11 | factory |
| Conductivity | 2173 | 7May08 | “ | 28Dec10 | factory |
| Secondary Temp. | 2968 | 22Aug07 | “ | 30Dec10 | factory |
| Secondary Cond. | 2399 | 13Jun08 | “ | 28Dec10 | factory |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 03Feb2009 | Factory |  |  |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 63507 | 11Dec2009 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2010-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 2416 | 23Dec06 |  |  |  |