## File List
Number of Files= |  | 3
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2010-14-0055.che | 2010-14-0055.sam | 11 | 34 | 2010-14-0055.che & 2010-14-0055.sam
2 | 2010-14-0073.che | 2010-14-0073.sam | 36 | 59 | 2010-14-0073.che & 2010-14-0073.sam
3 | 2010-14-0086.che | 2010-14-0086.sam | 61 | 84 | 2010-14-0086.che & 2010-14-0086.sam

## (DOX_CTD - DOX_BOT)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.67933 |  | X | CALC Y
m= | -0.045629 | 0.0043898 |  | Std Dev of Resid |  | 0.072901 |  | 0.28192 | -0.0062345
b= | 0.0066292 | 0.018295 |  | Resid Sum of Squares |  | 0.27104 |  | 6.6959 | -0.2989
# of Values Extracted = |  |  | 72
# of Values Used in Fit = |  |  | 53
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 5.5131 | 5.6947 | 0.0068385 | -0.36627 | -0.25322 | 0.11305 |  | 5.6947 | -0.36627
1 | YES | 11.4237 | 5.7052 | 0.012238 | -0.35982 | -0.25369 | 0.10612 |  | 5.7052 | -0.35982
1 | YES | 25.9505 | 5.7399 | 0.013592 | -0.39509 | -0.25528 | 0.13981 |  | 5.7399 | -0.39509
1 | YES | 50.8061 | 6.4258 | 0.008567 | -0.33921 | -0.28658 | 0.052635 |  | 6.4258 | -0.33921
1 | YES | 75.8634 | 6.3624 | 0.012954 | -0.31664 | -0.28368 | 0.032957 |  | 6.3624 | -0.31664
1 | YES | 101.357 | 5.9388 | 0.0041329 | -0.28917 | -0.26436 | 0.02481 |  | 5.9388 | -0.28917
1 | YES | 126.286 | 5.0785 | 0.010088 | -0.25351 | -0.2251 | 0.028415 |  | 5.0785 | -0.25351
1 | YES | 151.066 | 4.5396 | 0.0068377 | -0.22943 | -0.20051 | 0.028923 |  | 4.5396 | -0.22943
1 | YES | 175.828 | 4.206 | 0.0068607 | -0.23404 | -0.18529 | 0.048751 |  | 4.206 | -0.23404
1 | YES | 201.975 | 3.9226 | 0.0086621 | -0.21845 | -0.17235 | 0.046093 |  | 3.9226 | -0.21845
1 | YES | 250.858 | 3.2153 | 0.0031427 | -0.18366 | -0.14008 | 0.043571 |  | 3.2153 | -0.18366
1 | YES | 301.724 | 2.4638 | 0.0030017 | -0.1352 | -0.10579 | 0.029406 |  | 2.4638 | -0.1352
1 | YES | 401.369 | 1.4749 | 0.0029703 | -0.10105 | -0.060672 | 0.040378 |  | 1.4749 | -0.10105
1 | YES | 601.075 | 0.64261 | 0.0028425 | -0.049386 | -0.022693 | 0.026693 |  | 0.64261 | -0.049386
1 | YES | 801.373 | 0.37141 | 0.0014291 | -0.028593 | -0.010318 | 0.018276 |  | 0.37141 | -0.028593
1 | YES | 1001.47 | 0.28192 | 0.0028725 | -0.040083 | -0.0062345 | 0.033849 |  | 0.28192 | -0.040083
1 | YES | 1251.68 | 0.2886 | 0.0017713 | -0.028399 | -0.0065395 | 0.02186 |  | 0.2886 | -0.028399
1 | YES | 1501.92 | 0.46542 | 0.0039448 | -0.024581 | -0.014608 | 0.0099733 |  | 0.46542 | -0.024581
1 | YES | 2001.36 | 1.1462 | 0.0044275 | -0.063851 | -0.045669 | 0.018182 |  | 1.1462 | -0.063851
1 | NO | 2501.01 | 1.8312 | 0.0067979 | -0.11283 | -0.076926 | 0.0359 |  | 5.5459 | -0.39809
1 | NO | 3000.17 | 2.3001 | 0.0038255 | -0.15688 | -0.098324 | 0.058552 |  | 5.5671 | -0.37586
1 | NO | 3500.91 | 2.5186 | 0.007738 | -0.17739 | -0.10829 | 0.069094 |  | 6.6959 | -0.35907
1 | NO | 3500.91 | 2.5186 | 0.007738 | -0.18139 | -0.10829 | 0.073094 |  | 6.6093 | -0.22866
1 | NO | 3685.45 | 2.5436 | 0.0056329 | -0.19643 | -0.10943 | 0.087 |  | 6.4442 | -0.19181
2 | YES | 5.4688 | 5.5459 | 0.0067324 | -0.39809 | -0.24643 | 0.15166 |  | 5.6728 | -0.14122
2 | YES | 10.5053 | 5.5671 | 0.0030017 | -0.37586 | -0.2474 | 0.12847 |  | 4.7795 | -0.13848
2 | YES | 26.0889 | 6.6959 | 0.012576 | -0.35907 | -0.2989 | 0.060164 |  | 4.4324 | -0.11458
2 | YES | 50.6567 | 6.6093 | 0.0041654 | -0.22866 | -0.29495 | -0.066291 |  | 4.0931 | -0.11985
2 | YES | 75.5444 | 6.4442 | 0.0060387 | -0.19181 | -0.28742 | -0.095611 |  | 3.7668 | -0.08122
2 | YES | 101.001 | 5.6728 | 0.0046242 | -0.14122 | -0.25222 | -0.11099 |  | 2.9673 | -0.075701
2 | YES | 126.114 | 4.7795 | 0.0073161 | -0.13848 | -0.21146 | -0.072981 |  | 2.1054 | -0.04361
2 | YES | 150.772 | 4.4324 | 0.0054393 | -0.11458 | -0.19562 | -0.081043 |  | 1.1989 | -0.024058
2 | YES | 175.66 | 4.0931 | 0.0029804 | -0.11985 | -0.18014 | -0.060285 |  | 0.57274 | -0.0072572
2 | YES | 200.852 | 3.7668 | 0.0047685 | -0.08122 | -0.16525 | -0.084027 |  | 1.2154 | -0.031556
2 | YES | 251.018 | 2.9673 | 0.0032625 | -0.075701 | -0.12877 | -0.053066 |  | 5.8483 | -0.35368
2 | YES | 300.847 | 2.1054 | 0.0027305 | -0.04361 | -0.089439 | -0.045829 |  | 5.8435 | -0.36253
2 | YES | 400.897 | 1.1989 | 0.0023248 | -0.024058 | -0.048078 | -0.02402 |  | 6.5895 | -0.025477
2 | YES | 598.714 | 0.57274 | 0.0017176 | -0.0072572 | -0.019505 | -0.012248 |  | 6.5724 | -0.23163
2 | YES | 801.728 | 0.4122 | 0.0027877 | 0.0012033 | -0.012179 | -0.013383 |  | 6.4829 | -0.19712
2 | YES | 1000.71 | 0.29106 | 0.0026593 | 0.0080581 | -0.0066516 | -0.01471 |  | 5.9724 | -0.22058
2 | YES | 1251.27 | 0.3221 | 0.0021618 | 0.0060996 | -0.008068 | -0.014168 |  | 4.1488 | -0.15422
2 | YES | 1499.86 | 0.51848 | 0.0029241 | 0.010477 | -0.017029 | -0.027506 |  | 3.5112 | -0.1038
2 | YES | 2001.13 | 1.2154 | 0.0034925 | -0.031556 | -0.048831 | -0.017275 |  | 3.1758 | -0.10616
2 | NO | 2501.51 | 1.8231 | 0.0048254 | -0.097946 | -0.076556 | 0.02139 |  | 2.8848 | -0.10025
2 | NO | 3000.2 | 2.278 | 0.0032979 | -0.18595 | -0.097317 | 0.088629 |  | 2.1244 | -0.089556
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23726 | -0.11209 | 0.12517 |  | 1.6672 | -0.047751
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23426 | -0.11209 | 0.12217 |  | 0.84169 | -0.030307
2 | NO | 4024.87 | 2.7297 | 0.0079746 | -0.23932 | -0.11792 | 0.1214 |  | 0.4576 | -0.0064025
3 | YES | 6.2646 | 5.8483 | 0.0060404 | -0.35368 | -0.26023 | 0.09345 |  | 0.41795 | -0.0060539
3 | YES | 11.1444 | 5.8435 | 0.0072509 | -0.36253 | -0.26001 | 0.10253 |  | 0.34633 | -0.005668
3 | YES | 25.6832 | 6.5895 | 0.033892 | -0.025477 | -0.29405 | -0.26857 |  | 0.41278 | -0.014224
3 | YES | 51.18 | 6.5724 | 0.0051882 | -0.23163 | -0.29326 | -0.061633 |  | 0.63808 | -0.001917
3 | YES | 76.1665 | 6.4829 | 0.0062408 | -0.19712 | -0.28918 | -0.092061 |  | 1.2864 | -0.038614
3 | YES | 101.248 | 5.9724 | 0.0039774 | -0.22058 | -0.26589 | -0.045308
3 | YES | 126.171 | 4.1488 | 0.0062533 | -0.15422 | -0.18268 | -0.028461
3 | YES | 151.535 | 3.5112 | 0.0050697 | -0.1038 | -0.15358 | -0.049784
3 | YES | 176.02 | 3.1758 | 0.0039257 | -0.10616 | -0.13828 | -0.032121
3 | YES | 202.384 | 2.8848 | 0.0035617 | -0.10025 | -0.125 | -0.024756
3 | YES | 251.625 | 2.1244 | 0.0043058 | -0.089556 | -0.090308 | -0.00075179
3 | YES | 300.174 | 1.6672 | 0.0034863 | -0.047751 | -0.069446 | -0.021695
3 | YES | 399.816 | 0.84169 | 0.0026889 | -0.030307 | -0.031777 | -0.0014697
3 | YES | 600.79 | 0.4576 | 0.0041109 | -0.0064025 | -0.014251 | -0.0078483
3 | YES | 801.411 | 0.41795 | 0.0011767 | -0.0060539 | -0.012441 | -0.0063875
3 | YES | 1000.23 | 0.34633 | 0.003397 | -0.005668 | -0.0091737 | -0.0035057
3 | YES | 1251.18 | 0.41278 | 0.0039412 | -0.014224 | -0.012206 | 0.0020186
3 | YES | 1500.09 | 0.63808 | 0.0046932 | -0.001917 | -0.022486 | -0.020569
3 | YES | 2001.07 | 1.2864 | 0.0054716 | -0.038614 | -0.052068 | -0.013454
3 | NO | 2498.79 | 1.8684 | 0.0061695 | -0.096614 | -0.078624 | 0.01799
3 | NO | 3000.14 | 2.3878 | 0.0036375 | -0.16625 | -0.10232 | 0.063931
3 | NO | 3499.95 | 2.7079 | 0.0038866 | -0.24405 | -0.11693 | 0.12712
3 | NO | 4001.02 | 2.8882 | 0.0045882 | -0.2838 | -0.12516 | 0.15864
3 | NO | 4318.28 | 2.924 | 0.012767 | -0.28397 | -0.12679 | 0.15718
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.99154 |  | X | CALC Y
m= | -0.04833 | 0.0011932 |  | Std Dev of Resid |  | 0.010693 |  | 0.28192 | -0.030642
b= | -0.017017 | 0.0043976 |  | Resid Sum of Squares |  | 0.0016008 |  | 6.4258 | -0.32757
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 16
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 5.5131 | 5.6947 | 0.0068385 | -0.36627 | -0.29224 | 0.074029 |  | 6.4258 | -0.33921
1 | NO | 11.4237 | 5.7052 | 0.012238 | -0.35982 | -0.29275 | 0.067072 |  | 6.3624 | -0.31664
1 | NO | 25.9505 | 5.7399 | 0.013592 | -0.39509 | -0.29442 | 0.10066 |  | 5.9388 | -0.28917
1 | YES | 50.8061 | 6.4258 | 0.008567 | -0.33921 | -0.32757 | 0.011638 |  | 5.0785 | -0.25351
1 | YES | 75.8634 | 6.3624 | 0.012954 | -0.31664 | -0.32451 | -0.0078686 |  | 4.5396 | -0.22943
1 | YES | 101.357 | 5.9388 | 0.0041329 | -0.28917 | -0.30404 | -0.014873 |  | 4.206 | -0.23404
1 | YES | 126.286 | 5.0785 | 0.010088 | -0.25351 | -0.26246 | -0.0089446 |  | 3.9226 | -0.21845
1 | YES | 151.066 | 4.5396 | 0.0068377 | -0.22943 | -0.23641 | -0.0069807 |  | 3.2153 | -0.18366
1 | YES | 175.828 | 4.206 | 0.0068607 | -0.23404 | -0.22029 | 0.013748 |  | 2.4638 | -0.1352
1 | YES | 201.975 | 3.9226 | 0.0086621 | -0.21845 | -0.20659 | 0.011856 |  | 1.4749 | -0.10105
1 | YES | 250.858 | 3.2153 | 0.0031427 | -0.18366 | -0.17241 | 0.011243 |  | 0.64261 | -0.049386
1 | YES | 301.724 | 2.4638 | 0.0030017 | -0.1352 | -0.13609 | -0.00089242 |  | 0.37141 | -0.028593
1 | YES | 401.369 | 1.4749 | 0.0029703 | -0.10105 | -0.088301 | 0.012749 |  | 0.28192 | -0.040083
1 | YES | 601.075 | 0.64261 | 0.0028425 | -0.049386 | -0.048074 | 0.0013117 |  | 0.2886 | -0.028399
1 | YES | 801.373 | 0.37141 | 0.0014291 | -0.028593 | -0.034967 | -0.0063733 |  | 0.46542 | -0.024581
1 | YES | 1001.47 | 0.28192 | 0.0028725 | -0.040083 | -0.030642 | 0.0094413 |  | 1.1462 | -0.063851
1 | YES | 1251.68 | 0.2886 | 0.0017713 | -0.028399 | -0.030965 | -0.0025656
1 | YES | 1501.92 | 0.46542 | 0.0039448 | -0.024581 | -0.03951 | -0.014929
1 | YES | 2001.36 | 1.1462 | 0.0044275 | -0.063851 | -0.07241 | -0.0085592
1 | NO | 2501.01 | 1.8312 | 0.0067979 | -0.11283 | -0.10552 | 0.0073094
1 | NO | 3000.17 | 2.3001 | 0.0038255 | -0.15688 | -0.12818 | 0.028695
1 | NO | 3500.91 | 2.5186 | 0.007738 | -0.17739 | -0.13874 | 0.038648
1 | NO | 3500.91 | 2.5186 | 0.007738 | -0.18139 | -0.13874 | 0.042648
1 | NO | 3685.45 | 2.5436 | 0.0056329 | -0.19643 | -0.13995 | 0.056485
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (3)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.96931 |  | X | CALC Y
m= | -0.032442 | 0.0015429 |  | Std Dev of Resid |  | 0.013722 |  | 0.29106 | 0.0093546
b= | 0.018797 | 0.005562 |  | Resid Sum of Squares |  | 0.0026361 |  | 6.6093 | -0.19562
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 16
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
2 | NO | 5.4688 | 5.5459 | 0.0067324 | -0.39809 | -0.16112 | 0.23697 |  | 6.6093 | -0.22866
2 | NO | 10.5053 | 5.5671 | 0.0030017 | -0.37586 | -0.16181 | 0.21405 |  | 6.4442 | -0.19181
2 | NO | 26.0889 | 6.6959 | 0.012576 | -0.35907 | -0.19843 | 0.16064 |  | 5.6728 | -0.14122
2 | YES | 50.6567 | 6.6093 | 0.0041654 | -0.22866 | -0.19562 | 0.033039 |  | 4.7795 | -0.13848
2 | YES | 75.5444 | 6.4442 | 0.0060387 | -0.19181 | -0.19026 | 0.0015414 |  | 4.4324 | -0.11458
2 | YES | 101.001 | 5.6728 | 0.0046242 | -0.14122 | -0.16524 | -0.024014 |  | 4.0931 | -0.11985
2 | YES | 126.114 | 4.7795 | 0.0073161 | -0.13848 | -0.13626 | 0.0022182 |  | 3.7668 | -0.08122
2 | YES | 150.772 | 4.4324 | 0.0054393 | -0.11458 | -0.125 | -0.010421 |  | 2.9673 | -0.075701
2 | YES | 175.66 | 4.0931 | 0.0029804 | -0.11985 | -0.11399 | 0.0058622 |  | 2.1054 | -0.04361
2 | YES | 200.852 | 3.7668 | 0.0047685 | -0.08122 | -0.1034 | -0.022184 |  | 1.1989 | -0.024058
2 | YES | 251.018 | 2.9673 | 0.0032625 | -0.075701 | -0.077467 | -0.0017661 |  | 0.57274 | -0.0072572
2 | YES | 300.847 | 2.1054 | 0.0027305 | -0.04361 | -0.049505 | -0.0058956 |  | 0.4122 | 0.0012033
2 | YES | 400.897 | 1.1989 | 0.0023248 | -0.024058 | -0.020099 | 0.0039594 |  | 0.29106 | 0.0080581
2 | YES | 598.714 | 0.57274 | 0.0017176 | -0.0072572 | 0.00021623 | 0.0074735 |  | 0.3221 | 0.0060996
2 | YES | 801.728 | 0.4122 | 0.0027877 | 0.0012033 | 0.0054244 | 0.0042211 |  | 0.51848 | 0.010477
2 | YES | 1000.71 | 0.29106 | 0.0026593 | 0.0080581 | 0.0093546 | 0.0012965 |  | 1.2154 | -0.031556
2 | YES | 1251.27 | 0.3221 | 0.0021618 | 0.0060996 | 0.0083475 | 0.0022479
2 | YES | 1499.86 | 0.51848 | 0.0029241 | 0.010477 | 0.0019767 | -0.0085005
2 | YES | 2001.13 | 1.2154 | 0.0034925 | -0.031556 | -0.020634 | 0.010922
2 | NO | 2501.51 | 1.8231 | 0.0048254 | -0.097946 | -0.040346 | 0.0576
2 | NO | 3000.2 | 2.278 | 0.0032979 | -0.18595 | -0.055107 | 0.13084
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23726 | -0.065608 | 0.17165
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23426 | -0.065608 | 0.16865
2 | NO | 4024.87 | 2.7297 | 0.0079746 | -0.23932 | -0.069758 | 0.16957
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (4)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.9779 |  | X | CALC Y
m= | -0.035319 | 0.001419 |  | Std Dev of Resid |  | 0.012274 |  | 0.34633 | -0.0064986
b= | 0.0057335 | 0.0047538 |  | Resid Sum of Squares |  | 0.0021091 |  | 6.5724 | -0.2264
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 16
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
3 | NO | 6.2646 | 5.8483 | 0.0060404 | -0.35368 | -0.20082 | 0.15285 |  | 6.5724 | -0.23163
3 | NO | 11.1444 | 5.8435 | 0.0072509 | -0.36253 | -0.20065 | 0.16188 |  | 6.4829 | -0.19712
3 | NO | 25.6832 | 6.5895 | 0.033892 | -0.025477 | -0.227 | -0.20152 |  | 5.9724 | -0.22058
3 | YES | 51.18 | 6.5724 | 0.0051882 | -0.23163 | -0.2264 | 0.005235 |  | 4.1488 | -0.15422
3 | YES | 76.1665 | 6.4829 | 0.0062408 | -0.19712 | -0.22324 | -0.026115 |  | 3.5112 | -0.1038
3 | YES | 101.248 | 5.9724 | 0.0039774 | -0.22058 | -0.20521 | 0.015375 |  | 3.1758 | -0.10616
3 | YES | 126.171 | 4.1488 | 0.0062533 | -0.15422 | -0.1408 | 0.013419 |  | 2.8848 | -0.10025
3 | YES | 151.535 | 3.5112 | 0.0050697 | -0.1038 | -0.11828 | -0.014478 |  | 2.1244 | -0.089556
3 | YES | 176.02 | 3.1758 | 0.0039257 | -0.10616 | -0.10643 | -0.00027195 |  | 1.6672 | -0.047751
3 | YES | 202.384 | 2.8848 | 0.0035617 | -0.10025 | -0.096153 | 0.0040917 |  | 0.84169 | -0.030307
3 | YES | 251.625 | 2.1244 | 0.0043058 | -0.089556 | -0.0693 | 0.020256 |  | 0.4576 | -0.0064025
3 | YES | 300.174 | 1.6672 | 0.0034863 | -0.047751 | -0.053152 | -0.005401 |  | 0.41795 | -0.0060539
3 | YES | 399.816 | 0.84169 | 0.0026889 | -0.030307 | -0.023994 | 0.0063129 |  | 0.34633 | -0.005668
3 | YES | 600.79 | 0.4576 | 0.0041109 | -0.0064025 | -0.010428 | -0.0040259 |  | 0.41278 | -0.014224
3 | YES | 801.411 | 0.41795 | 0.0011767 | -0.0060539 | -0.0090279 | -0.002974 |  | 0.63808 | -0.001917
3 | YES | 1000.23 | 0.34633 | 0.003397 | -0.005668 | -0.0064986 | -0.00083052 |  | 1.2864 | -0.038614
3 | YES | 1251.18 | 0.41278 | 0.0039412 | -0.014224 | -0.0088453 | 0.0053788
3 | YES | 1500.09 | 0.63808 | 0.0046932 | -0.001917 | -0.016803 | -0.014886
3 | YES | 2001.07 | 1.2864 | 0.0054716 | -0.038614 | -0.0397 | -0.0010863
3 | NO | 2498.79 | 1.8684 | 0.0061695 | -0.096614 | -0.060256 | 0.036358
3 | NO | 3000.14 | 2.3878 | 0.0036375 | -0.16625 | -0.078599 | 0.087654
3 | NO | 3499.95 | 2.7079 | 0.0038866 | -0.24405 | -0.089909 | 0.15414
3 | NO | 4001.02 | 2.8882 | 0.0045882 | -0.2838 | -0.096275 | 0.18753
3 | NO | 4318.28 | 2.924 | 0.012767 | -0.28397 | -0.09754 | 0.18643
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (5)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.9533 |  | X | CALC Y
m= | -0.03354 | 0.0013553 |  | Std Dev of Resid |  | 0.016848 |  | 0.29106 | 0.0017852
b= | 0.011547 | 0.0047163 |  | Resid Sum of Squares |  | 0.0085159 |  | 6.6093 | -0.21013
# of Values Extracted = |  |  | 48
# of Values Used in Fit = |  |  | 32
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
2 | NO | 5.4688 | 5.5459 | 0.0067324 | -0.39809 | -0.17446 | 0.22363 |  | 6.6093 | -0.22866
2 | NO | 10.5053 | 5.5671 | 0.0030017 | -0.37586 | -0.17517 | 0.20069 |  | 6.4442 | -0.19181
2 | NO | 26.0889 | 6.6959 | 0.012576 | -0.35907 | -0.21303 | 0.14603 |  | 5.6728 | -0.14122
2 | YES | 50.6567 | 6.6093 | 0.0041654 | -0.22866 | -0.21013 | 0.01853 |  | 4.7795 | -0.13848
2 | YES | 75.5444 | 6.4442 | 0.0060387 | -0.19181 | -0.20459 | -0.012785 |  | 4.4324 | -0.11458
2 | YES | 101.001 | 5.6728 | 0.0046242 | -0.14122 | -0.17872 | -0.037493 |  | 4.0931 | -0.11985
2 | YES | 126.114 | 4.7795 | 0.0073161 | -0.13848 | -0.14876 | -0.01028 |  | 3.7668 | -0.08122
2 | YES | 150.772 | 4.4324 | 0.0054393 | -0.11458 | -0.13712 | -0.022539 |  | 2.9673 | -0.075701
2 | YES | 175.66 | 4.0931 | 0.0029804 | -0.11985 | -0.12574 | -0.0058827 |  | 2.1054 | -0.04361
2 | YES | 200.852 | 3.7668 | 0.0047685 | -0.08122 | -0.11479 | -0.03357 |  | 1.1989 | -0.024058
2 | YES | 251.018 | 2.9673 | 0.0032625 | -0.075701 | -0.087976 | -0.012275 |  | 0.57274 | -0.0072572
2 | YES | 300.847 | 2.1054 | 0.0027305 | -0.04361 | -0.059067 | -0.015457 |  | 0.4122 | 0.0012033
2 | YES | 400.897 | 1.1989 | 0.0023248 | -0.024058 | -0.028665 | -0.004607 |  | 0.29106 | 0.0080581
2 | YES | 598.714 | 0.57274 | 0.0017176 | -0.0072572 | -0.0076625 | -0.00040524 |  | 0.3221 | 0.0060996
2 | YES | 801.728 | 0.4122 | 0.0027877 | 0.0012033 | -0.002278 | -0.0034813 |  | 0.51848 | 0.010477
2 | YES | 1000.71 | 0.29106 | 0.0026593 | 0.0080581 | 0.0017852 | -0.0062729 |  | 1.2154 | -0.031556
2 | YES | 1251.27 | 0.3221 | 0.0021618 | 0.0060996 | 0.00074409 | -0.0053555 |  | 6.5724 | -0.23163
2 | YES | 1499.86 | 0.51848 | 0.0029241 | 0.010477 | -0.0058424 | -0.01632 |  | 6.4829 | -0.19712
2 | YES | 2001.13 | 1.2154 | 0.0034925 | -0.031556 | -0.029218 | 0.0023375 |  | 5.9724 | -0.22058
2 | NO | 2501.51 | 1.8231 | 0.0048254 | -0.097946 | -0.049598 | 0.048348 |  | 4.1488 | -0.15422
2 | NO | 3000.2 | 2.278 | 0.0032979 | -0.18595 | -0.064858 | 0.12109 |  | 3.5112 | -0.1038
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23726 | -0.075715 | 0.16154 |  | 3.1758 | -0.10616
2 | NO | 3501.65 | 2.6017 | 0.0069916 | -0.23426 | -0.075715 | 0.15854 |  | 2.8848 | -0.10025
2 | NO | 4024.87 | 2.7297 | 0.0079746 | -0.23932 | -0.080006 | 0.15932 |  | 2.1244 | -0.089556
3 | NO | 6.2646 | 5.8483 | 0.0060404 | -0.35368 | -0.1846 | 0.16907 |  | 1.6672 | -0.047751
3 | NO | 11.1444 | 5.8435 | 0.0072509 | -0.36253 | -0.18444 | 0.17809 |  | 0.84169 | -0.030307
3 | NO | 25.6832 | 6.5895 | 0.033892 | -0.025477 | -0.20946 | -0.18399 |  | 0.4576 | -0.0064025
3 | YES | 51.18 | 6.5724 | 0.0051882 | -0.23163 | -0.20889 | 0.022741 |  | 0.41795 | -0.0060539
3 | YES | 76.1665 | 6.4829 | 0.0062408 | -0.19712 | -0.20589 | -0.0087681 |  | 0.34633 | -0.005668
3 | YES | 101.248 | 5.9724 | 0.0039774 | -0.22058 | -0.18877 | 0.031814 |  | 0.41278 | -0.014224
3 | YES | 126.171 | 4.1488 | 0.0062533 | -0.15422 | -0.1276 | 0.026613 |  | 0.63808 | -0.001917
3 | YES | 151.535 | 3.5112 | 0.0050697 | -0.1038 | -0.10622 | -0.0024171 |  | 1.2864 | -0.038614
3 | YES | 176.02 | 3.1758 | 0.0039257 | -0.10616 | -0.09497 | 0.011192
3 | YES | 202.384 | 2.8848 | 0.0035617 | -0.10025 | -0.085207 | 0.015038
3 | YES | 251.625 | 2.1244 | 0.0043058 | -0.089556 | -0.059706 | 0.02985
3 | YES | 300.174 | 1.6672 | 0.0034863 | -0.047751 | -0.044372 | 0.0033789
3 | YES | 399.816 | 0.84169 | 0.0026889 | -0.030307 | -0.016683 | 0.013624
3 | YES | 600.79 | 0.4576 | 0.0041109 | -0.0064025 | -0.0038005 | 0.002602
3 | YES | 801.411 | 0.41795 | 0.0011767 | -0.0060539 | -0.0024706 | 0.0035833
3 | YES | 1000.23 | 0.34633 | 0.003397 | -0.005668 | -6.8654e-05 | 0.0055994
3 | YES | 1251.18 | 0.41278 | 0.0039412 | -0.014224 | -0.0022972 | 0.011927
3 | YES | 1500.09 | 0.63808 | 0.0046932 | -0.001917 | -0.009854 | -0.007937
3 | YES | 2001.07 | 1.2864 | 0.0054716 | -0.038614 | -0.031598 | 0.0070161
3 | NO | 2498.79 | 1.8684 | 0.0061695 | -0.096614 | -0.051118 | 0.045496
3 | NO | 3000.14 | 2.3878 | 0.0036375 | -0.16625 | -0.068538 | 0.097715
3 | NO | 3499.95 | 2.7079 | 0.0038866 | -0.24405 | -0.079277 | 0.16477
3 | NO | 4001.02 | 2.8882 | 0.0045882 | -0.2838 | -0.085323 | 0.19848
3 | NO | 4318.28 | 2.924 | 0.012767 | -0.28397 | -0.086524 | 0.19745
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.3 | Run at 2011/07/11 09:52:37
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX_BOT | DOX_CTD |  | DOX_CTD - DOX_BOT
Average Size: | 2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX_BOT | DOX_CTD | DOX_CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2010-14-0055.che | 2010-14-0055.sam | 241 | 5.5 | 5.5131 | 0.23134 | 0.0131 | 6.061 | 5.6947 | 0.0068385 | -0.36627
2010-14-0055.che | 2010-14-0055.sam | 241 | 11.4 | 11.4237 | 0.37966 | 0.023681 | 6.065 | 5.7052 | 0.012238 | -0.35982
2010-14-0055.che | 2010-14-0055.sam | 241 | 26 | 25.9505 | 0.31188 | -0.049515 | 6.135 | 5.7399 | 0.013592 | -0.39509
2010-14-0055.che | 2010-14-0055.sam | 241 | 50.8 | 50.8061 | 0.15998 | 0.0060921 | 6.765 | 6.4258 | 0.008567 | -0.33921
2010-14-0055.che | 2010-14-0055.sam | 241 | 75.9 | 75.8634 | 0.64475 | -0.036583 | 6.679 | 6.3624 | 0.012954 | -0.31664
2010-14-0055.che | 2010-14-0055.sam | 241 | 101.4 | 101.357 | 0.38783 | -0.04316 | 6.228 | 5.9388 | 0.0041329 | -0.28917
2010-14-0055.che | 2010-14-0055.sam | 241 | 126.3 | 126.286 | 0.13883 | -0.01429 | 5.332 | 5.0785 | 0.010088 | -0.25351
2010-14-0055.che | 2010-14-0055.sam | 241 | 151.1 | 151.066 | 0.45748 | -0.034012 | 4.769 | 4.5396 | 0.0068377 | -0.22943
2010-14-0055.che | 2010-14-0055.sam | 241 | 175.8 | 175.828 | 0.24602 | 0.027527 | 4.44 | 4.206 | 0.0068607 | -0.23404
2010-14-0055.che | 2010-14-0055.sam | 241 | 202 | 201.975 | 0.54819 | -0.024689 | 4.141 | 3.9226 | 0.0086621 | -0.21845
2010-14-0055.che | 2010-14-0055.sam | 241 | 250.9 | 250.858 | 0.41026 | -0.041901 | 3.399 | 3.2153 | 0.0031427 | -0.18366
2010-14-0055.che | 2010-14-0055.sam | 241 | 301.7 | 301.724 | 0.37491 | 0.024078 | 2.599 | 2.4638 | 0.0030017 | -0.1352
2010-14-0055.che | 2010-14-0055.sam | 241 | 401.4 | 401.369 | 0.71218 | -0.031067 | 1.576 | 1.4749 | 0.0029703 | -0.10105
2010-14-0055.che | 2010-14-0055.sam | 241 | 601.1 | 601.075 | 0.39379 | -0.025208 | 0.692 | 0.64261 | 0.0028425 | -0.049386
2010-14-0055.che | 2010-14-0055.sam | 241 | 801.4 | 801.373 | 0.61079 | -0.027283 | 0.4 | 0.37141 | 0.0014291 | -0.028593
2010-14-0055.che | 2010-14-0055.sam | 241 | 1001.5 | 1001.47 | 0.57477 | -0.027283 | 0.322 | 0.28192 | 0.0028725 | -0.040083
2010-14-0055.che | 2010-14-0055.sam | 233 | 1251.6 | 1251.68 | 1.1578 | 0.082031 | 0.317 | 0.2886 | 0.0017713 | -0.028399
2010-14-0055.che | 2010-14-0055.sam | 241 | 1501.9 | 1501.92 | 0.73152 | 0.016724 | 0.49 | 0.46542 | 0.0039448 | -0.024581
2010-14-0055.che | 2010-14-0055.sam | 241 | 2001.4 | 2001.36 | 0.53882 | -0.035034 | 1.21 | 1.1462 | 0.0044275 | -0.063851
2010-14-0055.che | 2010-14-0055.sam | 241 | 2501 | 2501.01 | 0.5942 | 0.014893 | 1.944 | 1.8312 | 0.0067979 | -0.11283
2010-14-0055.che | 2010-14-0055.sam | 241 | 3000.2 | 3000.17 | 0.76 | -0.032227 | 2.457 | 2.3001 | 0.0038255 | -0.15688
2010-14-0055.che | 2010-14-0055.sam | 327 | 3500.7 | 3500.91 | 0.58309 | 0.20581 | 2.696 | 2.5186 | 0.007738 | -0.17739
2010-14-0055.che | 2010-14-0055.sam | 327 | 3501 | 3500.91 | 0.58309 | -0.094238 | 2.7 | 2.5186 | 0.007738 | -0.18139
2010-14-0055.che | 2010-14-0055.sam | 241 | 3685.4 | 3685.45 | 0.8291 | 0.045898 | 2.74 | 2.5436 | 0.0056329 | -0.19643
2010-14-0073.che | 2010-14-0073.sam | 241 | 5.5 | 5.4688 | 0.31622 | -0.031253 | 5.944 | 5.5459 | 0.0067324 | -0.39809
2010-14-0073.che | 2010-14-0073.sam | 241 | 10.5 | 10.5053 | 0.16259 | 0.0053282 | 5.943 | 5.5671 | 0.0030017 | -0.37586
2010-14-0073.che | 2010-14-0073.sam | 241 | 26.1 | 26.0889 | 0.0947 | -0.011084 | 7.055 | 6.6959 | 0.012576 | -0.35907
2010-14-0073.che | 2010-14-0073.sam | 241 | 50.7 | 50.6567 | 0.067913 | -0.04327 | 6.838 | 6.6093 | 0.0041654 | -0.22866
2010-14-0073.che | 2010-14-0073.sam | 241 | 75.5 | 75.5444 | 0.1829 | 0.044441 | 6.636 | 6.4442 | 0.0060387 | -0.19181
2010-14-0073.che | 2010-14-0073.sam | 241 | 101 | 101.001 | 0.35623 | 0.0014343 | 5.814 | 5.6728 | 0.0046242 | -0.14122
2010-14-0073.che | 2010-14-0073.sam | 241 | 126.1 | 126.114 | 0.2421 | 0.014175 | 4.918 | 4.7795 | 0.0073161 | -0.13848
2010-14-0073.che | 2010-14-0073.sam | 241 | 150.8 | 150.772 | 0.36292 | -0.028153 | 4.547 | 4.4324 | 0.0054393 | -0.11458
2010-14-0073.che | 2010-14-0073.sam | 241 | 175.7 | 175.66 | 0.47182 | -0.040085 | 4.213 | 4.0931 | 0.0029804 | -0.11985
2010-14-0073.che | 2010-14-0073.sam | 241 | 200.9 | 200.852 | 0.23492 | -0.048111 | 3.848 | 3.7668 | 0.0047685 | -0.08122
2010-14-0073.che | 2010-14-0073.sam | 241 | 251 | 251.018 | 0.16903 | 0.017914 | 3.043 | 2.9673 | 0.0032625 | -0.075701
2010-14-0073.che | 2010-14-0073.sam | 241 | 300.8 | 300.847 | 0.212 | 0.046783 | 2.149 | 2.1054 | 0.0027305 | -0.04361
2010-14-0073.che | 2010-14-0073.sam | 241 | 400.9 | 400.897 | 0.43366 | -0.0026245 | 1.223 | 1.1989 | 0.0023248 | -0.024058
2010-14-0073.che | 2010-14-0073.sam | 241 | 598.7 | 598.714 | 0.4674 | 0.014099 | 0.58 | 0.57274 | 0.0017176 | -0.0072572
2010-14-0073.che | 2010-14-0073.sam | 241 | 801.7 | 801.728 | 0.30123 | 0.028015 | 0.411 | 0.4122 | 0.0027877 | 0.0012033
2010-14-0073.che | 2010-14-0073.sam | 241 | 1000.7 | 1000.71 | 0.61876 | 0.0083618 | 0.283 | 0.29106 | 0.0026593 | 0.0080581
2010-14-0073.che | 2010-14-0073.sam | 241 | 1251.3 | 1251.27 | 0.46304 | -0.025757 | 0.316 | 0.3221 | 0.0021618 | 0.0060996
2010-14-0073.che | 2010-14-0073.sam | 241 | 1499.9 | 1499.86 | 0.65522 | -0.042358 | 0.508 | 0.51848 | 0.0029241 | 0.010477
2010-14-0073.che | 2010-14-0073.sam | 241 | 2001.1 | 2001.13 | 0.62122 | 0.030273 | 1.247 | 1.2154 | 0.0034925 | -0.031556
2010-14-0073.che | 2010-14-0073.sam | 241 | 2501.5 | 2501.51 | 0.22308 | 0.012695 | 1.921 | 1.8231 | 0.0048254 | -0.097946
2010-14-0073.che | 2010-14-0073.sam | 241 | 3000.2 | 3000.2 | 0.74808 | -0.0031738 | 2.464 | 2.278 | 0.0032979 | -0.18595
2010-14-0073.che | 2010-14-0073.sam | 482 | 3501.5 | 3501.65 | 0.74717 | 0.14551 | 2.839 | 2.6017 | 0.0069916 | -0.23726
2010-14-0073.che | 2010-14-0073.sam | 482 | 3501.7 | 3501.65 | 0.74717 | -0.054443 | 2.836 | 2.6017 | 0.0069916 | -0.23426
2010-14-0073.che | 2010-14-0073.sam | 241 | 4024.9 | 4024.87 | 0.5202 | -0.026123 | 2.969 | 2.7297 | 0.0079746 | -0.23932
2010-14-0086.che | 2010-14-0086.sam | 241 | 6.3 | 6.2646 | 0.28349 | -0.035432 | 6.202 | 5.8483 | 0.0060404 | -0.35368
2010-14-0086.che | 2010-14-0086.sam | 241 | 11.1 | 11.1444 | 0.28054 | 0.044386 | 6.206 | 5.8435 | 0.0072509 | -0.36253
2010-14-0086.che | 2010-14-0086.sam | 241 | 25.7 | 25.6832 | 0.2248 | -0.016785 | 6.615 | 6.5895 | 0.033892 | -0.025477
2010-14-0086.che | 2010-14-0086.sam | 241 | 51.2 | 51.18 | 0.14153 | -0.020012 | 6.804 | 6.5724 | 0.0051882 | -0.23163
2010-14-0086.che | 2010-14-0086.sam | 241 | 76.2 | 76.1665 | 0.26936 | -0.033501 | 6.68 | 6.4829 | 0.0062408 | -0.19712
2010-14-0086.che | 2010-14-0086.sam | 241 | 101.2 | 101.248 | 0.26384 | 0.047653 | 6.193 | 5.9724 | 0.0039774 | -0.22058
2010-14-0086.che | 2010-14-0086.sam | 241 | 126.2 | 126.171 | 0.29101 | -0.028854 | 4.303 | 4.1488 | 0.0062533 | -0.15422
2010-14-0086.che | 2010-14-0086.sam | 241 | 151.5 | 151.535 | 0.22007 | 0.034622 | 3.615 | 3.5112 | 0.0050697 | -0.1038
2010-14-0086.che | 2010-14-0086.sam | 241 | 176 | 176.02 | 0.073205 | 0.020447 | 3.282 | 3.1758 | 0.0039257 | -0.10616
2010-14-0086.che | 2010-14-0086.sam | 241 | 202.4 | 202.384 | 0.31793 | -0.015915 | 2.985 | 2.8848 | 0.0035617 | -0.10025
2010-14-0086.che | 2010-14-0086.sam | 241 | 251.6 | 251.625 | 0.27082 | 0.025284 | 2.214 | 2.1244 | 0.0043058 | -0.089556
2010-14-0086.che | 2010-14-0086.sam | 241 | 300.2 | 300.174 | 0.41279 | -0.026489 | 1.715 | 1.6672 | 0.0034863 | -0.047751
2010-14-0086.che | 2010-14-0086.sam | 241 | 399.8 | 399.816 | 0.26572 | 0.015625 | 0.872 | 0.84169 | 0.0026889 | -0.030307
2010-14-0086.che | 2010-14-0086.sam | 241 | 600.8 | 600.79 | 0.13617 | -0.0098267 | 0.464 | 0.4576 | 0.0041109 | -0.0064025
2010-14-0086.che | 2010-14-0086.sam | 241 | 801.4 | 801.411 | 0.2437 | 0.01062 | 0.424 | 0.41795 | 0.0011767 | -0.0060539
2010-14-0086.che | 2010-14-0086.sam | 241 | 1000.2 | 1000.23 | 0.35159 | 0.028564 | 0.352 | 0.34633 | 0.003397 | -0.005668
2010-14-0086.che | 2010-14-0086.sam | 241 | 1251.2 | 1251.18 | 0.43655 | -0.024048 | 0.427 | 0.41278 | 0.0039412 | -0.014224
2010-14-0086.che | 2010-14-0086.sam | 241 | 1500.1 | 1500.09 | 0.38164 | -0.010376 | 0.64 | 0.63808 | 0.0046932 | -0.001917
2010-14-0086.che | 2010-14-0086.sam | 241 | 2001.1 | 2001.07 | 0.15358 | -0.03064 | 1.325 | 1.2864 | 0.0054716 | -0.038614
2010-14-0086.che | 2010-14-0086.sam | 241 | 2498.8 | 2498.79 | 0.24967 | -0.013184 | 1.965 | 1.8684 | 0.0061695 | -0.096614
2010-14-0086.che | 2010-14-0086.sam | 241 | 3000.1 | 3000.14 | 0.49349 | 0.044189 | 2.554 | 2.3878 | 0.0036375 | -0.16625
2010-14-0086.che | 2010-14-0086.sam | 241 | 3500 | 3499.95 | 0.40324 | -0.047363 | 2.952 | 2.7079 | 0.0038866 | -0.24405
2010-14-0086.che | 2010-14-0086.sam | 241 | 4001 | 4001.02 | 0.24736 | 0.019531 | 3.172 | 2.8882 | 0.0045882 | -0.2838
2010-14-0086.che | 2010-14-0086.sam | 241 | 4318.3 | 4318.28 | 0.41052 | -0.015625 | 3.208 | 2.924 | 0.012767 | -0.28397
Create Fit Sheet