## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2025 | Updated channel names & formats in TOB files. GG |
| 25 Feb 2025 | Removed original Silicate & Flag:Silicate channels and renamed corrected versions of those channels. G.G. |
| 25 July 2011 | Reprocessed Oxygen:Dissolved:SBE channel. See notes at end of report. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2010-15

Agency: OSD

Location: WCVI

Project: La Perouse / WCVI

Party Chief: Yelland D.

Platform: John P. Tully

Date: 3 September 2010 – 14 September 2010

Processed by: Germaine Gatien

Date of Processing: 2 May 2011 – 25 July 2011

Number of CTD files:   97	Number of CTD casts processed: 96 (1 file with upcast only)

Number of bottle casts:	37 		Number of bottle casts processed: 36 (1 file had no sampling)

Number of original TSG files:	3  	Number of TSG files processed:  3

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (serial number #0443) was used for all casts. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#1483) on the primary pump), a Seapoint Fluorometer (#2845) with a 3X cable (on the primary pump), an altimeter (#1252) and a Biospherical QSP-400 PAR sensor #4694.

The deck unit was #0508.

All casts were run with the LARS mid-ship station.

Seasave version 7.3 was used.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

NOTE: After processing had been completed it was discovered that the wrong serial number had been entered for the SBE Dissolved Oxygen Sensor, so channel Oxygen:Dissolved:SBE was reprocessed and the results substituted in the CHE and CTD files. The details on what was done are entered at the end of the report. The comments on DO quality remain valid.

The CTD and rosette logs were in good order overall but there was no information entered about the thermosalinograph. The equipment list proved to be faulty so it is assumed no visual check was made of serial numbers.

The comparison of salinity bottles with CTD salinity indicated that the CTD values are too high for both sensor pairs, whereas other recent uses of this equipment indicated they were reading low. The analyst noted odd jumps in values between the readings being taken, and those being recorded, and thought that it might just be an adjustment based on the standardization reading. However, the results are out of step with both the history of the sensor and a post-cruise calibration. So it appears that there was a problem with the salinometer. The post-cruise calibration of the secondary temperature sensor showed a significant drift, but there is evidence that most of the drift must have occurred after this cruise. The salinity calibration was primarily based on the results of 2010-14.

This cruise was blessed with plentiful dissolved oxygen sampling. This is not only valuable for calibration purposes, but also provides useful evidence of whether bottles closed at the intended depths. This is particularly important when there is not a lot of salinity sampling.

The DO sensor showed a large change since it was last used and the bottle comparison was very difficult to interpret. Many bottles from 0 to 300db from casts offshore of the 200m contour stood out in the fit. The bottle comparison showed a higher slope (larger correction) for the onshore and deep bottles, but less noise. Profiles from the shallow offshore region showed sections of marginal stability and high variability in DO, but lower DO gradients. In the shallow offshore area the lower gradients would allow the sensor to do a better job of “catching up”, but the noise level would be higher due to mixing. The shallow offshore bottles were not included in the fit used for recalibration and a warning was put in the header files to indicate that the SBE DO data is of lower quality than usual; a note was also entered to indicate if a cast was from the inshore or offshore group.

The precision of the SBE dissolved oxygen channel is difficult to estimate because the comparison with bottles was complex and noisy but, roughly, the DO for off-shelf casts (westward of the 200m contour) should be considered:

•	±0.4ml/l from 0-300 db

•	±0.25ml/l from 300-500 db

•	±0.1ml/l below 500db

For on-shelf casts (including some casts deeper than 200m from Rivers Inlet and the Strait of Georgia) the DO should be considered:

•	±0.4 ml/l from 0-50 db

•	±0.2 ml/l from 50-125 db

•	±0.02 ml/l below 125db

The Thermosalinograph position download failed for part of the record, with latitude and longitude values getting “stuck”. Data were obtained from the ship position records and patched into the TSG file by matching times. This produced a track plot that looks right, so any errors in this patching are thought to be small.

The TSG temperature and salinity data look reliable with no large spikes, but the fluorescence data are unusual, with a sudden drop in values between the LG and LBP lines. While it is possible that there is a biological or instrumental explanation for this, the absence of loop chlorophyll samples and the fact that the TSG fluorometer has not been calibrated in the past 10 years make it impossible to provide a convincing explanation of the problem. The fluorescence data are included but must, as always, be considered nominal.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Extracted chlorophyll, nutrients, dissolved oxygen, salinity and NH4 data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked and two errors were fixed:

- The pressure offset has been increased to +7db based on other 2010 cruises.

- For the Dissolved Oxygen sensor there were fine-tuning parameters available (E, H1 and H3) for the dissolved oxygen sensor based on 2010-01 studies.

After those changes were made, the configuration file was saved as 2010-15-ctd.con.

##### Conversion of Raw Data

Data were converted using the configuration file 2010-15-ctd.con. SPAR was converted, though the log book does not indicate that it was in use. It will probably be removed later.

A few casts were examined. The temperature and conductivity channel pairs track reasonably well with a vertical offset between downcast and upcast, and more noise in the upcast. The DO voltage looks odd, with upcasts not as close to downcasts as usual. PAR looks ok and SPAR has no signal in the casts examined.  The fluorescence dark value is about 0.065. The altimetry and transmissivity look ok.

Two casts were interrupted, one by a SARS call and one by an oil leak. Casts #21 and 22 are at the same site – the second cast is deeper. Cast #66 is a downcast only and #67 is an upcast. Bottles were closed during #67, but there was no sampling, so no CHE file will be produced. These problems will not affect processing, since there is no need to patch bottle files.

##### Rosette File Conversion and Study of Split Casts

Rosette files were converted using a start time of -5s and duration of 10s. The TAU and hysteresis corrections were chosen. There were rosette files for every CTD cast, but many were empty, so those were deleted. A few others will not be processed further because of comments in the log book:

- Cast #1 – bottles were closed at the surface as a test, no samples

- Cast #67 – upcast only, bottles were closed to prevent oil getting in them, no sampling

- Cast #152 - bottles were closed on the fly and there was no sampling

- Cast #154 - bottles were closed on the fly and there was no sampling

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers, with output named *.BOT.

Header Check was run on the files to see if there were off-scale fluorescence values and there were no values close to the maximum. There is no evidence of spikes in any channels.

Temperature and salinity were plotted for all BOT files and a few casts appeared to have outliers. Those casts were examined in CTDEDIT.

- For cast #60 secondary salinity was cleaned for a few points from the surface bottle.

- For cast #85 a few points were cleaned in the primary salinity for the 125db bottle.

- For cast #135 some records were removed from the beginning of the 6db bottle.

Editing details were added to the header comments of the edited files. The output files were then copied to *.BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity, temperature and descent rate channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using 3 casts. The goal is to make upcasts look closer to downcasts on a T-S surface. Most casts were either very shallow or had stops for bottles or very noisy descent rates, which make the plots too noisy to judge, but 3 were found that with as steady a descent rate as possible. The choice of (α = 0.0245, β=9.5) looked best for the primary and (α = 0.02, β=7) was best for the secondary conductivity for CTD #0443.

CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The first 4 rows are from 2010-14 which preceded this cruise and had the same sensors as were used for this cruise. The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2010-14-0045 #0443 | 500 1000 1900 | -0.0008 VN ~0 VN -0.0001 | +0.0001 +0.0003 +0.00035 | +0.002 +0.0035 +0.0045 | Noisy, High |
| 2010-14-0073 #0443 | 500 1000 1900 3800 | +0.0002 0.0000 -0.0002 -0.0002 | +0.0001 +0.0002 +0.0003 +0.00035 | +0.001 +0.003 +0.0035 +0.0045 | V noisy, High |
| 2010-14-0090 #0443 | 500 1000 1900 3800 | 0.0000 0.0000 -0.0001 -0.00015 | +0.0001 +0.00025 +0.0003 +0.00035 | +0.0015 +0.003 +0.0035 +0.005 | V noisy, High |
| 2010-14-0127 | 500 1000 1900 | ~0 VN +0.0003 +0.0003XN | +0.00004 +0.0002 +0.00025 | +0.0006VN +0.002 +0.003 | X Noisy High |
| 2010-15-0030 | 500 850 | -0.0005 +0.0005 | -0.00008 +0.0002 | +0.0016 +0.0015 | Mod, High |
| 2010-15-0034 | 500 1000 1500 | +0.0003 +0.0005 +0.0006 | +0.0008VN +0.00018 +0.00024VN | +0.0007 +0.0017 +0.0024 | V Noisy, High |
| 2010-15-0035 | 500 1000 1500 1900 | +0.0005 +0.0006VN +0.0004VN +0.0003VN | +0.00005 +0.00018 +0.00022VN +0.00025VN | ~0 +0.0016 +0.0024 +0.0028 | V Noisy, High |
| 2010-15-0112 | 500 1000 1500 1900 | ~0 XN +0.0005 +0.0003 +0.0003 | +0.0001XN +0.0002VN +0.00023 +0.00025 | +0.001XN +0.0022VN +0.0025 +0.0028 | Moderate, HIgh |
| 2010-15-0132 | 400 | +0.0006 | +0.00013 | +0.0009VN | Steady, F.High |
| 2010-15-0154 | 350 | +0.0003 | +0.00002 | ~0 | V.Steady, F. High |

The differences in temperature are small and show no time or pressure dependence, but there is a lot of noise with significant spikes that do not seem related to descent rate. The conductivity and salinity show significant pressure dependence, similar to what was observed during 2010-14, but there is little temporal change.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on scan number.

##### Checking Headers

The header check was run and no obvious problems were found. The fluorescence maximum value was 10.2mg/m3, so there is no problem with off-scale values since the maximum with the 30X cable is 50mg/m3. There was no evidence of the spiking problems seen on other Tully cruises from the summer of 2010.

Surface check was run and shows an average surface pressure for the cruise was 2.7db, with values between 0.9 and 4.2db. These are reasonable values. Examination of a few casts shows salinity values of ~30.8 at 0.9db for one cast. For another there are pressure values as low as 0.4-0.5db during the soak period (with pumps on); the transmissivity drops to 0 and for a few casts the conductivity goes off scale and salinity goes as low as 29, so this must be very close to the surface. It is possible that the CTD was very briefly out of the water. So the pressures look reasonably close to correct.

The cross-reference check was compared with the log book and the only problems found were in station name formats – they were corrected in both the CLN files and MRG files.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings and bottom depths from the headers of the CLN files were exported to a spreadsheet. The water depths were checked against the log book. One was missing so that was added to the header and 2 were incorrect – changes were entered in the log book, but not the file. Those were fixed in the CLN files (and in the MRGCLN2 files where appropriate). There were other small differences likely due to real changes between the time of recording in the log and the headers. Some casts did not have an altimeter header entry because they did not get within 15m of the bottom. A plot of random casts turned up one case where there is no obvious signal in the altimetry at the bottom, yet there is a header entry. That entry was removed and then all casts were checked to see if this was a more general problem, but no other problems were found. The MRG files were also checked and no problems were found.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Sample numbers were assigned in order of Niskin # rather than Firing # so the file was reordered on bottle position numbers. There were some bottles fired but no sample numbers assigned, so those lines were removed from the addsamp file. Cast #122 was the most confusing as there were 2 firings of Niskin #4. The 2nd was removed from the addsamp list.

A problem was found in the ADDSAMP list from cast #29 onwards. Looking carefully at cast #29 it can be seen that while 18 bottles closed (as judged by the BOT file), only 17 were given sample numbers. Because the bottles were closed out of order and because actual pressures were very different from the target pressures, it was hard to sort out what happened. Niskin bottle #5 was closed twice. After Niskin bottle #5 was closed around 250m, Niskins #17 and 18 were also fired at that level. The CTD came up to 200m and a bottle was fired, but it was #5 again which was already closed. So it appears there is no sample from 200db. The salinity sample from Niskin #14 looks right for 10m so the near-surface data looks ok. And the nutrient analyst also concluded that there was no 200m sample. To fix the addsamp file the 8th bottle fired was removed from the list.

The addsamp.csv file was ordered on sample number and then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files.

The SAM files were then bin-averaged.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. There was a comment column, though no comments were entered. There was no quality flag channel, so that was added. The Autosal analysis log sheets were available and no comments were entered on those. The spreadsheet was saved as 2010-15-sal.csv. Some columns were removed and a column was added for the event number based on log book entries. The sample #s and station names were entered in a single column, so that was divided into separate columns. The headers were changed to standard format.

There was one entry with no sample number or station name; a check against the salinity analysis log sheets found nothing missing, so this is likely from some other cruise. The rosette log sheets does have a salinity sample indicated from cast #126 that is not in the spreadsheet, but the CTD salinity for that bottle is very different from the unidentified salinity value in the spreadsheet, so that is not likely to be the source.

There were 2 duplicate samples taken (samples #41 and 304). They were copied to sheet 2010-15-sal-duplicates.csv where they were averaged and differences were found. The duplicate values in 2010-15-sal.csv were replaced with averages and flag “6” was added. The differences between the two pairs were 0.03 and 0.0015. Flag “3” was added to sample #41. (Note: COMPARE showed both of the duplicates were outliers and the salinity analysis is likely the problem.)

File 2010-15-sal.csv was converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2010-15oxy.xls which includes a precision study. The date of creation was added to the name of the file in case changes occur in the future. The spreadsheet was simplified, ordered on sample numbers, flags changed to numeric format and headings changed to standard format.

Note that samples 53, 77 and 378 had “b” flags indicating that the analyst was not sure how significant the problem was. These were changed to “4’ but that will be reconsidered after COMPARE is run.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2010-15nuts.xls which included a report on precision. Headers were changed to standard format and quality flags were changed to numeric values, and the file was saved as 2010-15-nuts.csv. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. Colour coding was used in the spreadsheet to indicate that a few samples were stored cool rather than frozen, so a comment was entered for those since the colour will not show in the bottle files. Pad values were added to a few blank entries with “9” flags. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2010_15chl.xls. The name of the file was changed by adding the date of creation. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which headers were changed to standard format, CHL was entered before comments, unnecessary columns were removed and the file was saved as 2010-15-chl.csv.

NH4

NH4 data were obtained in file QF2010-15NH4.xls. The file was saved as 2010-15-NH4.csv and edited. There were flags and comments. The event number and bottle # were missing for 3 samples – those were added and the comment saying why they were missing was removed.  Headers were changed to standard format and unnecessary columns were removed. The file was then converted to individual NH4 files.

The SAL, CHL, ADD, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number. So one or the other set needs to be reordered in order to merge them. It was decided to reorder the MRGCLN1 files on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

Because this process was complex, all bottle data from the MRG files were exported to a spreadsheet and reordered on Bottle Position (Niskin #) order to check that all expected samples are present and in the right spot. The CHL were missing from casts 135 at UBC7, 138 at Ri1, 143 at Ri3, 145 at Ri4, 147 at Ri5 and 150 at Ri6. No such sampling is indicated in the Hydro file, and Marie Robert confirms that these data were collected by university researchers and analyzed elsewhere.

Plots were made of dissolved oxygen (bottles and SBE) versus CTD salinity. All data looked ok.

##### 12) Compare

Salinity

Compare was run for each CTD system with pressure as reference channel.

There were 2 severe outliers from very close to the surface. One was in Rivers Inlet where the surface gradient was fairly high right at the surface, but not at the depth of the sample. The other is associated with a high standard deviation in the CTD data. Most of the bottles are from the surface. The average of differences below 25db shows the primary CTD salinity to be higher than the bottles by ~0.003 and the secondary higher by ~0.007. There is a lot of scatter in the comparison with a few bottles rejected as outliers. The CTD standard deviations were not high for the outliers. These results are very different from the results for 2010-14 and do not fit the post-cruise calibration. Restricting the comparison to bottles below 200db reduces the average differences a little, but does not make them look like the earlier cruise. There is no evidence of pressure dependence or time dependence, though that is hard to judge from so little non-surface data.  A look at cast #104 that had 20 bottles shows considerable variation in the differences.

It has been reported that the analyst noted odd jumps in values between the readings being taken, and those being recorded, though there is no note of that on the salinity sheets or in the spreadsheet. This is known only by personal communication; Hugh Maclean thought this might have been an automatic adjustment based on the standardization reading or that there was a problem with the system. The latter would appear to have been the case.

The duplicates for sample #41 were examined and 1 of them is definitely out of line with the other bottles, so the average could be replaced with the lower salinity bottle, but given the overall unreliability of the analysis this is not clear. On the other hand it looks very close to the sample #32 value – could there be some mislabelled samples? Given all the problems with the analysis, it seems most likely the problem lies with analysis rather than with sampling.

Flag “3” was assigned to all salinity bottles.

This cruise was the second for which new liners were used on the bottles. The new bottle washing system was not yet in use.

Dissolved Oxygen These notes are from the original processing. See end of report for reprocessing details. The fits were different but variability and outlier comments are valid.

COMPARE was run for Dissolved Oxygen.

The plot with all data included is noisy. Outliers were gradually removed leading to the fit:

Bottle DO = 1.1931 * CTD DO - 0.0097

The slope is much higher than noted when the sensor was used in August during 2010-14. When only casts #2-37 were used and outliers excluded, the fit was very similar:

Bottle DO = 1.1936 * CTD DO - 0.0109

There is no evidence of rapid drift to explain the difference from 2010-14. During the earlier cruise there was evidence of hysteresis with apparently bad values below 2200db but for this cruise there are only 11 bottles below 1225db; of those, the only notable outliers (2) had been flagged by the analyst. None of the bottles was below 2010db, so based on the observations of 2010-14 these data are likely not affected by hysteresis. The noise level in the dissolved oxygen voltage looks the same for stops for which the differences are outliers and those which are not outliers.

A list was prepared of the data points that look out of line in COMPARE in the hope that a pattern would emerge. As usual there are many outliers near the surface and for these data quite a few between 50 and 200db, but there are many bottles in that range that do not stand out. There is no obvious time-dependence with casts 29, 39, 69, 74 and 91 standing out. Checking against the cruise track a pattern does emerge, since the outliers (mostly to one side of the fit) come from casts beyond the 200m contour. When the data are divided into on-shelf and off-shelf groups there remain some outliers, but for the on-shelf they are almost all from 10, 20 and 50db. For the on-shelf group, the outliers are mostly from cast #91, the most northerly of these casts with useful DO samples (all bottles from cast #104 were flagged by the analyst.) A look at a few CTD profiles shows there are odd features in T, S and DO with reversals that probably allow the DO sensor to “catch up” more than usual. (These DO reversals are confirmed by titrated DO samples.) For example there is a local DO maximum and temperature and salinity minima at about 175db for cast #74. The fit for on-shelf casts when outliers are removed based on residuals is:

Bottle DO = 1.1942 * CTD DO - 0.0129 (Fit #1)

The fit for the off-shelf casts with outliers removed based on residuals is:

Bottle DO = 1.1717 * CTD DO - 0.027 (Fit #2)

There is still a lot of noise in Fit #2 even after outliers are removed. Fit #1 would reasonably represent the deeper data from the off-shelf casts, and neither fit does a good job of the shallower data.

What accounts for the difference between on and off-shelf when bottle stops were at least 30s long? A look at a few plots shows that the DO sensor did not equilibrate in many cases. For example, for cast #74 the DO at 150db and 175db was varying significantly when those bottles were closed. In some cases temperature was also varying significantly showing that changes were still happening, while for a few others the temperature has settled down, but presumably the DO didn’t get enough time to do that.

This explains the difference between the two groups, and indicates that Fit #1 is best for this cruise, though it will not do a great job in sections of the top 250db of off-shore casts with active mixing. During 2010-14 there were also some odd results, but too few bottle casts to determine what happened. Perhaps this sensor could not keep up with rapid changes as well it should.

Plots were examined from the 3 casts of 2010-14 with bottle sampling and DO sensor #1478 to see if there were any hints. Profiles from those 3 casts do not indicate the sort of variability seen in this cruise. There are sections of very low gradient, where the DO sensor might “catch up” more easily and thus produce a lower slope. For cast #55 at P16 the temperature trace has a local maximum around 125db but salinity gradient is high so this profile is stable. The dissolved oxygen appears to have some little blips and primary salinity is noisy, so there may have been a problem with the flow rate. Casts #73 and 86 do not have that problem. While there are differences in the profiles, the cast at P20 looks more like the P16 cast than the P26 cast, so there is nothing in that to explain the different fit of oxygen bottles against SBE DO. Most likely the spikiness of cast #55 had an affect on the fit, so the fits from cast #73 and 86 were the better choice overall for 2010-14.

As to why the fit changed so much between the two cruises, it is possible that something happened to the sensor near the end of 2010-14 or beginning of 2010-15. There was no DO sampling after cast #86 of the earlier cruise.

A few bottles were checked to see if the flags already assigned are appropriate, in particular, cases where the analyst noted something but considered it not worth a flag unless COMPARE or other bottle observations indicated a problem:

- Cast #9, sample #18 average looks ok in COMPARE; flag left as “3” as duplicates differ significantly.

- Cast #29. sample #53 – looks ok in COMPARE so flag “2” removed

- Cast #31, sample #77  - looks ok in COMPARE so flag “2” removed

- Cast #39, sample #120 – outlier but that looks like it is due to the CTD not equilibrating but flag “2” changed to “4” since there is doubt.

- Cast #104, samples #294 and 295 are definite outliers so flag “4” is appropriate

- Cast #122, sample #325 – average looks ok in COMPARE; flag “3” kept as duplicates differ significantly.

- Cast #145, sample #375 – looks ok in COMPARE, so flag “2” removed.

(See 2010-15-dox-comp1-5May2011.xls.).

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. As usual there is a lot of scatter. Fluorescence is about 80% of the extracted chlorophyll, on average, but it is about 50% for CHL>2.5 ug/L and varies from 50% to 200% at low CHL values. There is no obvious difference between on-shore and off-shore casts, except that there are more low CHL values off-shore, so more instances of high FL/CHL ratio.

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

Data were exported to spreadsheet 2010-15-bottles.xls and compared to the rosette sheets to ensure all expected data are present. A few problems were found and corrected:

- There was no entry for DO sample #107 – the sample had been spilled and not analyzed, but it was added to the ADD file with a pad value and flag “1”.

- For cast #126 the rosette log sheet indicates there should be a salinity sample #342 but no such sample is listed in the analysis sheet. It is assumed it was either lost or never drawn. A file was created with a pad value and flag “9” and comment entered.

- There were no CHL samples from casts #135 to 150 though some are noted on the rosette sheet. They are not recorded in the file 2010-15-hydro.xlsx. As noted earlier these samples were collected by university researchers and are to be analyzed elsewhere.

The merge process was repeated after these corrections.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no problems were detected.

##### 13. Other Comparisons

Previous experience with these sensors –

1. Salinity:

The primary conductivity sensor has been used many times with 2 different temperature sensors. Bottle comparisons were often unavailable or limited to surface samples or showing a lot of scatter that led to concerns about sampling protocols. Poor seals and long waits for analysis likely led to bottle salinity being too high, thus making CTD salinity look low. During 2010-14 in August 2010, the primary sensors were archived for only 1 cast; the salinity appeared low by 0.006 during that cruise. The secondary conductivity sensor was used for 5 cruises that have been processed; only 2010-14 had reliable calibration sampling which showed secondary salinity to be low by 0.0015. During 2010-01 comparison with an Argo float showed the primary to be low by ~0.006 and the secondary to be within 0.001.

2. Dissolved Oxygen

Sensor #1438 was used for cruises in June 2009, February 2010 and Aug. 2010 when the slope/offset for the fits were 1.0551/+.0204 and 1.0784/.0218 and 1.1273/-0.0016. During the last of those cruises there were only 3 bottle casts with DO calibration sampling and they showed more variability than expected.

3. Pressure

Since its latest factory recalibration in December 2009 the pressure sensor has been used on 3 other cruises with offsets of 0, +0.5 and +0.5db.

Historic ranges – Profile plots were made with historic ranges (3 standard deviations) of T and S superimposed. There were some excursions with temperature a little low at LD9 and LG6 around 400db and temperature high and salinity low near the bottom in the Hakai Passage / Rivers Inlet area. These outliers do not look like evidence of systemic instrumental problems.

Repeat Casts –

The only repeat casts were two very shallow casts at B7 where small differences in position and time are likely to be significant. On a T-S surface the two casts are reasonably close.

Post-Cruise Calibration

Conductivity Sensors– A post-cruise calibration on Dec. 28, 2010 indicates that the primary and secondary conductivity sensors would lead to salinity values low by ~0.003 and.~0.002. A further error would arise from the secondary temperature calibration drift since it was reading high by ~0.0025°C which would lead to secondary salinity being low by a further ~0.0025 for a cumulative error of ~ -0.005 in the secondary salinity. The primary temperature drift was negligible.

Dissolved Oxygen Sensor – The membrane was found to be torn in Dec. 2010 and there was drift of about 10%.

Pressures Sensor – The slope and offset were changed slightly.

##### 14. Temperature and Conductivity Sensor Choice for Archive

For this cruise the salinity analysis results indicate the primary was high by 0.003 and the secondary high by 0.007. This is out of line with both the previous cruise, 2010-14, and the post-cruise calibration. Taken together with comments from the salinity analyst that some odd jumps in values were seen between the readings being taken and those being recorded, it looks as though the salinity comparison is not reliable. Poor duplicates and a noisy fit all strengthen that conclusion.

A study of 2010-14 data concluded that the large drift in secondary temperature calibration noted at the factory in December 2010 must have occurred after that cruise and that some drift in conductivity calibrations occurred between August and December 2010. The 2010-14 salinity comparison showed some time-dependence, with the primary moving closer to bottles and the secondary further away. The primary was low by 0.006 in August and low by 0.003 in December. The secondary was low by 0.0015 in August and by December it was low by ~0.0025 due to conductivity drift and a further 0.0025 due to temperature drift. So for this cruise we might expect the primary to be low by ~0.0045. We might expect the secondary to be low by ~0.002 if the bulk of the temperature drift had not yet occurred, or even lower if had. Those estimates would lead to a difference of 0.0035 between salinity channels if we assume the temperature drift came later. Those differences are similar to the section 8 results at depth.

The only evidence we have as to when the temperature drift occurred is to look at the differences between channels given in section 8. The differences are very slightly higher for 2010-15 (~0.0005 C°) than 2010-14 (~0.0002 C°), but the drift noted in December 2010 was an order of magnitude higher at ~0.026C°. So it does not look like this shift occurred during 2010-15.

The secondary channels were selected for archiving for 2010-14. Examination of a few profiles from 2010-15 shows that the primary temperature and salinity are noisier – mostly on upcasts – but also somewhat in downcasts, so they are the better choice.

##### 15. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise tests on a few casts showed that the usual setting is appropriate.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 4 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing.

The best overall result was with a setting of -0.7 records for the primary conductivity channel and -0.5 for the secondary. SHIFT was used to advance primary and secondary conductivity channels by -0.7 and -0.5 records, respectively.

Dissolved Oxygen

The usual tests were run on 4 casts using shifts of from 55 to 75 records. The data look odd and it was sometimes hard to tell which setting was better, but they all looked better than no advancement. Overall +60 looks best and that has been used for this sensor in the past. All DO data were advanced by +60.

##### 16. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #67 which is an upcast only file. It will not be processed further.

##### 17. DETAILED EDITING

The secondary T and S sensors were chosen for archiving. CTDEDIT was used to clean salinity and remove bad records.

A few problems were encountered:

- Cast #29. As noted in the log the pumps came on late, so the downcast data were deleted to 135db. The upcast data contain many large unstable features due to many stops for bottles and a noisy ascent rate between stops. An attempt was made to use the reversed upcast data to produce a full profile, but it was abandoned due to the very high noise level and the impossibility of distinguishing good data from bad.

- Cast #35. The secondary salinity is bad from ~500-600db, so primary sensors were selected for this cast.

In a few cases the editor crashed so the processing history record shows several runs of CTDEDIT. This is because a saved partially-edited file was opened to continue the job.

All EDU files were copied to EDT.

T-S plots were examined and no further editing was found necessary.

##### 18. Initial Recalibration

Pressure does not require recalibration.

As discussed in section 14 salinity calibration is problematic. The post-cruise calibration shows that by December the primary salinity was reading low by ~0.003 and the secondary by ~0.005. Half of the secondary error is due to drift in the secondary temperature sensor, but there was no sign of such drift at the time of this cruise. The differences between the two temperature sensors were <0.001 (see section 8) and close to those of other cruises in summer 2010. The differences between salinity values seen in section 8 suggests that the primary is reading higher than the secondary by ~0.0035 at depth. The salinity bottles are not trusted for absolute values, but COMPARE does support the conclusion that the secondary is higher than the primary by ~0.004, which is close to the difference found for 2010-14. The sensors were used for cruise 2010-69 immediately after this cruise, but only for the first group of casts for which only 3 salinity samples were taken, so waiting for that cruise to be processed is not justified.

There are sources of error other than calibration drift. The differences between salinity channels in the downcasts show pressure dependence, whereas the bottle comparison does not. The same was true for 2010-14. The temperature differences do not show much pressure dependence; the problem is in conductivity. We could apply a pressure-dependent correction, but this is not justified by bottles. Is there something inherent in the in-motion data that does not affect the CTD while stopped? Examination of salinity differences shows that the upcast does not show the same pressure dependence and that is true whether or not there are stops for bottles. The upcast differences at 400db are close to those at 2000db. The most likely cause for this behaviour would seem to be variations in alignment of the two sensors, either in space or in response time or transit time. Where the gradient is higher the differences are lower, presumably because the sensor that is reading higher is sensing change later, so that in high gradients the difference is reduced. The low gradient zones give the best measure of the calibration and that is what we see at depth. Whether the cause is in the pumps, sensors or physical alignment is not known. We don’t see it in temperature, but those differences were so low that such a pattern might be hard to see.

The best we can do is accept that the two sensors differ by ~0.004 and use corrections in line with the last cruise with a good comparison, 2010-14. So +0.006 will be added to the primary salinity channel and +0.002 to the secondary. The primary channels are used only for 1 cast.

Dissolved oxygen will be recalibrated using the best overall fit as discussed in section 12:

Bottle DO = 1.1942 * CTD DO - 0.0129

File 2010-15-recal1.ccf was prepared to recalibrate salinity and dissolved oxygen. It was applied to the SAM and MRGCLN2 files. COMPARE was rerun to see how well the DO recalibration worked. The results are good when points are excluded that were not used for the fit upon which the recalibration was based, showing it was applied correctly. When outliers are excluded based on residuals the results are also good. The outliers chosen by residuals include the suspect bottles from between 0db and 300db of the offshore casts. COMPARE was not rerun for salinity since the results were unreliable, but a file was examined to ensure the corrections were applied correctly. (See 2010-15-dox-comp2.xls.)

The EDT files were then recalibrated using the same calibration control file.

##### 19. Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction can be applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. There is a lot of scatter in the plot of differences against DO concentration, with a tendency towards values that are high by ~0.02 at the low end of the range and by 0.045 at the high end. However, the outliers are mostly on the high side and it is quite possible that more should have been removed as outliers, so this is not a robust fit. The plot against pressure shows values too high above 500db, being high by ~0.04mL/L near the surface, but as noted in the fit against DO there are many values that could easily be rejected as outliers. The CTD values are close to bottles below 600db when 3 obvious outliers are excluded, but that leaves only 11 bottles at those depths. There is a lot of noise in the comparison and the way the fits were done does not lead us to expect perfect results. What we have is probably as good as we can expect from these data, so no further recalibration will be applied. (See 2010-15-dox-comp3.xls.)

Finally, the data were separated into offshore and onshore casts, to see if the comparisons lead to different quality estimates. There was a clear difference in the scatter between the two groups, so separate cast lists were prepared based on the 200m contour on track plots. Casts westward of that contour were put in the offshore group. The onshore group includes a few casts deeper than 200m from Rivers Inlet and the Strait of Georgia. They do not stand out from the shallow casts in the fit.

##### 20. Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 125db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### 21. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### 22. Final CTD File Steps (REMOVE and HEADEDIT)

Lists were produced of files with and without PAR sensor channel.

REMOVE was run 2 times using the 2 lists to removing the following channels:

For all casts except #35: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag (plus PAR for the group without that sensor)

For cast #35 only: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference,Altimeter, Status:Pump, Descent_Rate and Flag and PAR

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments for the offshore casts:

Transmissivity, fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

The precision of the SBE dissolved oxygen channel is difficult to estimate

because the comparison with bottles was complex and noisy, but roughly,

the DO should be considered:

•	±0.4ml/l from 0-300 db

•	±0.2ml/l from 300-500 db

•	±0.1ml/l below 500db

For details on the processing see processing report: 2010-15-proc.doc.

For the onshore casts the same comments were made except that the DO quality estimates were:

•	±0.4ml/l from 0-50 db

•	±0.2ml/l from 50-125 db

•	±0.02ml/l below 125db

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

##### 23. Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 55% to 110% with most values between 90% and 110%. The lowest values are from near-shore casts, with those in Rivers Inlet and near the northern tip of Vancouver Island lowest of all. The only cast with values less than 80% that could be checked against bottles was cast #145 and that was one with very different downcast and upcast values for temperature, salinity and dissolved oxygen; while surface values look different, when salinity values are matched the CTD DO is close to the bottle DO.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run in the same 2 groups as the CTD files as described in section 22.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods. Separate comments were added for the inshore and offshore casts.

The CHE bottle data were exported to a spreadsheet, rearranged in similar order to rosette sheets and compared with the rosette log sheets to check that all data are present and in the correct place. A few discrepancies were noted:

- Cast #81 – Niskin bottles 6 to 24 were all fired at the surface. There is no note in the log or rosette sheet to indicate that this happened and no evidence of any samples taken, so the lines were removed from the CHE file. DO values compare well between bottle and CTD.

- Cast #122 – as noted earlier Niskin #4 was fired twice, first at 26db and then at 20db. The samples from that bottle are entered for the first firing as that is presumed to be when the bottle closed. Dissolved oxygen from CTD and bottle agree well, so this assumption appears correct. A noted was added to the header of the CHE file explaining what happened.

- Cast #126 – there were no nutrient data for sample #341. This was because it was labelled as #342 in the nutrient spreadsheet. Values are exactly the same. Correct data were obtained and substituted in the CHE file for sample #341.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO. There are some outliers but the bottle confirms the shape of the CTD DO. The offshore casts show up with CTD DO being generally higher than bottles at a given salinity above 300db. This fits the earlier observations.

Standards check was run on all files and no errors were found.

##### 25. Thermosalinograph Data

Data were provided in 3 hex files. There was no loop sampling.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong date and scale factor. After that correction the CON file was saved as 2010-15-tsg.con.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were produced. The traces look good:

- Flow interruption – The flow rate is fairly steady at slightly above 1.

- The salinity and temperature data are a little noisy early in the record, but that came from Haro Strait, so the noise is likely real.

- There are a few small shifts in temperature, salinity or fluorescence, some of which are associated with the end of station stops, though the changes are not consistent in direction. There are also features associated with the ship slowing or positioning before a cast. Very low and noisy salinity data are seen in the 2nd file, but this corresponds to the time in Rivers Inlet, so is likely accurate.

- The 3rd file contains only 94 records.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2010-15-ctd-tsg-comp.xls. All the data came from ~4.5db. For one CTD cast there was no data at that level.

A first attempt at checking time turned up a major problem in file #1: the position data became stuck for ~75 hours. Ship positions were obtained in a spreadsheet from the Chief Scientist for that period. The data were for every 10s whereas the TSG file has data every 30s, so they were reduced to records nearest the 18 and 48 second marks as that is what is in file #1. That file was saved as “File 1 position repair –vers2.xls”. File 2010-15-0001.atc was converted to spreadsheet format as 2010-15-positions-orig.csv so the good data could be used and the bad data be replaced. The repaired data were used to replace the bad data, scan numbers were added and the file was saved as 2010-15-positiions-fixed.csv. That file was then converted to 2010-15-0001.ios2. MERGE was then used to combine all columns in 2010-15-0001.atc except latitude and longitude with positions from file 2010-15-0002.ios2 and once again data were reduced to the times of CTD casts.

The TSG files (either MRG or ATC) were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence.

The files were then reduced to the times when CTDs were run to provide data for the CTD-TSG comparison. File #3 is very short and has no overlap with CTD casts.

These data were added to the CTD data in file 2010-15-ctd-tsg-comp.xls. There were 92 matches. In this file the times from the CTD and TSG are matched to find problems with time or positions. The median differences in latitude and longitude are 0.0000° and 0.0003° with all differences <0.0008° in latitude and in latitude <0.0009°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from TSG and CTD data

T1 vs T2 The average and median differences were 0.211C° and 0.206C° for file #1 and 0.212C° and 0.200C° for file #2. The heating in the loop is affected by flow rate, intake temperature and the ambient temperature of the ship. A plot was made of temperature differences against intake temperature, and from the first 2 files there were trends showing heating in the loop to be ~0.32C° & -0.40C° when the intake temperature was about 8°C and ~0.14C° & ~0.08C° when the intake temperature was 15°C. This fits past observations and the trendline allows a very rough estimate of the average ambient temperature around the loop at ~20°C for file #1 and ~16.5°C for file #2. The intake thermistor was connected throughout the cruise.

TSG vs CTD

- The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. When all data were included the TSG intake temperature was higher than the CTD by 0.12Cº while the median was 0.0094Cº. When outliers were excluded based on a standard deviation in the TSG being >0.008 or TSG-CTD differences >0.2Cº, the average difference was 0.018Cº and the median 0.0065Cº. This is a reasonable match.

- The TSG salinity is lower than the CTD by an average of 0.87 and a median of 0.116 using all the data, but there are some very large differences especially in the Rivers Inlet section. When only casts up to #120 are used the average is 0.117 and the median 0.111. On an earlier cruise there was some evidence of calibration drift in the TSG, but for these data while there may be a slight increase with time, there is so much noise late in the cruise that no comment can be made about temporal variations.

- The ratio of TSG fluorescence to CTD fluorescence ranges from 1.5 to 21.2 and an average of 6.5 and median of 5.7. A plot of the ratio against CTD fluorescence reveals a strange pattern with data falling into 2 groups; one shows overall low ratios with a slight drop as CTD fluorescence rises while the other has higher ratios and a steep drop as the CTD fluorescence rises. When the ratio is plotted against event #, there is great variability up to about cast #81. Between LG01 and LBP1 there was a sudden drop in the TSG fluorescence with no corresponding change in other TSG channels including temperature, salinity and flow rate. Even the standard deviation in the TSG fluorescence does not change significantly at that stage, though late in the cruise variability does increase notably as the ship moved into near-shore parts of Queen Charlotte Sound and Rivers Inlet. The CTD fluorescence shows a lot of variability throughout the record, but there is no suggestion of a difference before and after cast #81. The median ratio for casts up to #81 is 7.7, and for cast #82 to 154 it is 2.4. One obvious difference between these two groups is geographic in that the later casts were all to the north. The differences between the TSG salinity and the CTD salinity also get much noisier at about cast #81, which is probably due to the generally higher variability of the waters being sampled after that. (See 2010-15-ctd-tsg-comp.xls.)

e) Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and were used for 2009-10 (with a different intake thermistor), 2009-11, 2010-01, 2010-12, 2010-13 and 2010-14. The salinity was low by about 0.02, 0.02, 0.06 / 0.16, 0.07, 0.06 and 0.108 for those cruises. The 2 values for 2010-01 reflect a change in flow rate; there are problems with the 2010-12 data due to clock problems.

The TSG intake temperature was within 0.00 of the CTD for the two 2009 cruises and 0.001 during 2010-01, but differed by 0.04Cº for 2010-12 and 0.02Cº for 2010-13. For 2010-14 the median difference was 0.005Cº but there was a lot of variability. The TSG fluorometer was high by a factor of 2.2 to 6 for 2009-10, by about 8 for 2009-11, ~2 for 2010-01, 2.5 for 2010-12 and ~3 for 2010-13 and 2.4 for 2010-14 with a lot of variability. For 2010-36 in July the ratio started at about 4 and rose rapidly and went off scale; this was believed to be due to biological fouling.

f) Conclusions

1. The TSG clock appears to have worked well; matching times in the TSG file with those of CTD casts, the positions were mostly very close (where positions were available).

2. The flow rate was steady.

3. The temperature in the loop warms by a median value of ~0.21Cº but values range from 0.09Cº to 0.4Cº, with the highest values associated with the cooler waters sampled early in the cruise.

4. Salinity is low by ~0.11 when outliers are excluded, mostly from Rivers Inlet and other near-shore casts. This is higher than for 2010-14 by 0.01. While there is a suggestion of slight drift in the calibration, a time-dependent recalibration is not justified in the absence of loop samples and reliable data from the later casts.

5. There is always a lot of variability in the ratio of TSG fluorescence to either CTD fluorescence or loop CHL, with the TSG fluorometer reading higher. But this cruise was exceptional in having the ratio change markedly between the LG and LBP lines. This could be due to different conditions in the northern part of the cruise and during some other recent cruises there has been evidence of biological fouling of the loop. The later results are closer to observations on 2010-14. Since there was no loop CHL sampling and the TSG fluorometer has not been calibrated since 2001, further investigation of the variability is likely to be unproductive.

6. The intake temperature is within 0.01Cº of the CTD temperature from 4.5m based on median value, but the standard deviation is high.

g.) Editing

The MRG file for file #1 and the ATC files for file #2 & 3 were copied to *.FIX and *.EDT.

The FIX files were opened in CTDEDIT. Files #1 and 2 required no editing. The flow had not been turned on for the first 2 records for File #3 so rather than remove the records, a text editor was used to replace temperature, salinity and fluorescence with pad values. This maintains the position data to enable ship tracking.

The FIX files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

h.) Recalibration

File 2010-15-tsg-recal1.ccf was prepared to adjust salinity by adding 0.11.

i.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

1-30. PAR on.

1. 24 Niskins closed at the surface as a test. No sampling.

# 20/21 Interrupted by SAR – CTD rerun.

29. Pumps not on until 130m on downcast.

31-40. PAR off.

35. Secondary salinity looks bad around 500-600db. Use primary.

41-63. PAR on.

65-79. PAR off

66/67. The cast was stopped due to an oil leak. A new cast was started and bottles were closed to protect them from oil. There was no sampling.

81-90. PAR on.

91-112. PAR off.

112. Weird layer at 300-350db. Salinity profile down and up different ~160-280db.

114-154. PAR on.

122. Computer froze on start up.

130. Syringes not on since CS10. Not on for all M stations.

132. Long time at surface.

133. Stop at 265db down.

134. Little ways from station.

135. Rosette for UBC.

152. All Niskins closed at surface, no sampling.

154. All Niskins closed on upcast just to close Niskins, no sampling.

Reprocessing of SBE DO channel. July 2011

In June 2011 it was discovered that there was an error in the original processing of data from this cruise. The serial number of the SBE Dissolved Oxygen sensor was entered incorrectly in the log and in the configuration file used for CTD #0443. Initial tests showed only small difference in the results. The most difficult part of these corrections is bottle file reprocessing because those are very complex.

- Bottle File Reprocessing

The task was to replace the 2 SBE DO channels in the original CHE files. This is complicated by the split casts. First the original files were sorted on Bottle:Firing_Sequence and saved as CHE1. That file was put through CLEAN to remove the Seabird headers which had inadvertently not been removed in the original processing. (CHE1CL)

Then using the appropriate configuration file for the DO sensor #1483, all steps were followed in the way described in the original processing for the SBE DO channel except as mentioned. The ROS files were converted (some channels were skipped as they will not change and will not be needed to analyse the results), they were converted to IOS header files (*.BOT). Sample Numbers were added and they were bin-averaged. Now Bottle_Number has a different meaning in these two files, so the header names needed to be changed in the SAMAVG file. HEADER EDIT was used to do that and to fix the formats for the two SBE DO channels. That file was called SAMAVGHDR. It was then merged with the CHE1 file to create CHE2.

COMPARE was rerun using CHE2 and SAMAVG to determine how to recalibrate the data. As found in the original processing, many of the data from offshore between 0 and 200db are outliers in a general fit. Separate fits were made of onshore and offshore casts using the 200m contour as the dividing line. As found in the original processing, the onshore fit of differences vs SBE DO is fairly tight while the offshore has some data that look like the inshore group and others that show no particular trend. The onshore fit will be used for recalibration:

SBE DO corrected = 1.0954 *SBD DO – 0.0049

That correction is much larger than the one found for 2010-14:

SBE DO corrected = 1.0335 *SBD DO – 0.0115

However, it is much closer to the fit from 2010-01, though the latter cruise also had different fits for different casts. Overall, these data must be considered to be less reliable than usual. The SAM files were recalibrated and COMPARE was run again and showed the correction was made appropriately.

The CHE2 files were then recalibrated using the COMPARE result:

SBE DO corrected = 1.0954 *SBD DO – 0.0049

The 2nd DO channel was added, REORDER was used to get the 2 SBE:DO channels together, and SORT was used to reorder on Pressure.

HEADEDIT was used to correct the calibration summary and to add the following comment:

In June 2011 it was discovered that the wrong information was used for

conversion of the SBE Dissolved Oxygen sensor data for CTD #0443.

Sensor #1483 was actually used for this cruise.

The SBE Dissolved Oxygen channel was reprocessed. A new recalibration

scheme was found by rerunning the bottle comparison. The files were

then merged with all the other channels from the original processing.

Previous warnings about the quality of the SBE Dissolved Oxygen remain valid.

For details see the end of processing report 2010-15-proc.doc.

Note that the original comments on accuracy of the DO remain valid as problems remain with these data

Checks were run to ensure that no data had been lost.

1. A header check found no errors and a track plot looks fine.

2. The standard check was run and no problems were found.

3. Data from the new CHE files were exported to a spreadsheet which was added to a similar spreadsheet from the original files. Pressures and bottle contents were all the same. The reprocessed SBE DO values were the same or lower than found in the original processing, with most differences in the 0.00-0.02mL/L range with just a few differences between 0.03 and 0.06mL/L. The larger differences occurred at the higher range of DO values.

- CTD file reprocessing

# The raw hex files were converted to CNV files with just scan #, pressure, temperature, conductivity and dissolved oxygen voltage.

# DERIVE was used to compute Dissolved Oxygen in ml/L.

# The files were then converted to IOS HEADER format.

# CLEAN was run to replace pad values in the pressure channel and remove SeaBird comments.

SHIFT was run on DO only with a setting of +60 records.

DELETE was run with the same parameters as noted in the original processing. DELNEW

BIN AVERAGE on pressure with 1db bins. Output: AVGNEW

CALIBRATE using 2010-14-1483-recal1.ccf.

DERIVE the 2nd DO channel. Call those files OXNEW

MERGE the original CTD files with OXNEW taking all but DO from the CTD files and only the 2 DO channels from OXNEW. OUTPUT MRG

HEADEDIT was used to add a comment about the reprocessing and replace the calibration information for the DO channels and fix the DO formats.

Plots were made to check for spikes and ensure the merges were done properly. Header Check was rerun and no differences were found form the original other than slight

The parameters listed in the Seabird header in the CTD files includes calibration parameters

for the wrong dissolved oxygen sensor. The following are the correct parameters:

#   <sensor Channel="7" >

#     <!-- A/D voltage 1, Oxygen, SBE 43 -->

#     <OxygenSensor SensorID="38" >

#       <SerialNumber>1483</SerialNumber>

#       <CalibrationDate>03-feb-09</CalibrationDate>

#         <!-- Coefficients for Sea-Bird equation - SBE calibration in 2007 and later. -->

#         <Soc>4.1000e-001</Soc>

#         <offset>-0.5055</offset>

#         <A>-1.9487e-003</A>

#         <B> 1.7439e-004</B>

#         <C>-2.8684e-006</C>

#         <D0> 2.5826e+000</D0>

#         <D1> 1.92630e-004</D1>

#         <D2>-4.64800e-002</D2>

#         <E> 3.6000e-002</E>

#         <Tau20> 1.1000</Tau20>

#         <H1>-3.3000e-002</H1>

#         <H2> 5.0000e+003</H2>

#         <H3> 1.4500e+003</H3>

#       </CalibrationCoefficients>

#     </OxygenSensor>

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 06Mar07 | Factory “ | 5Jan11 | factory |
| Conductivity | 2173 | 7May08 | “ | 28Dec10 | factory |
| Secondary Temp. | 2968 | 22Aug07 | “ | 30Dec10 | factory |
| Secondary Cond. | 2399 | 13Jun08 | “ | 28Dec10 | factory |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 03Feb2009 | Factory |  |  |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 63507 | 11Dec2009 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2010-15

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 2416 | 23Dec06 |  |  |  |