## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 24 Nov 2021 | ### Correcting Salinity:Bottle precision lost during HPLC addition. S.H. |
| 14 Dec 2020 | Added HPLC Data. S.H. |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |

## PROCESSING NOTES

Cruise: 2010-16

Agency: OSD

Location: Strait of Georgia

Project: SoG Moorings / AMP ERI

Party Chief: Johannessen S.

Platform: Vector

Date: 3 April 2010 – 6 April 2010

Processed by: Germaine Gatien

Date of Processing: 30 November 2010 – 25 January 2011

Number of original CTD casts: 3		Number of CTD casts processed: 3

Number of bottle casts:	13			Number of bottle casts processed: 12

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2356) with a 30X cable (and later a 3X cable), a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1024).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sheets were in good order with useful notes about problems encountered during the cruise.

An unusual sampling identification scheme was used for this cruise. Because so much sampling was needed at every level, two bottles were fired at each stop, but only 1 sample number was assigned. So the data in each line of the CHE files derive from two different bottle firings. This is very awkward to work with and leads to some loss of information, such as evidence of premature bottle closings. If, in future, a scientist wants output with all samples in a single row coming from more than one bottle, that could be achieved by assigning sample numbers in the usual way, and then bin-averaging based on pressure at the end of the processing.

More duplicate samples would ensure better checks on sampling and analysis quality, especially for dissolved oxygen. For salinity calibration checks it would be good to have more samples from the region of lowest salinity gradient, which is usually in deep water but not right at the bottom.

There were also many spikes in data in the CTD files that were not removed by WILDEDIT. Graphical editing removed many of these spikes, but the overall quality of the data is likely to be lower than usual.

CTD data collected during bottle firing for cast #12 were heavily corrupted. Editing removed the obvious bad values, but the quality of the remaining data may be lower than usual. This impacts the comparisons with rosette samples.

Very light editing (to remove some very large spikes) was applied to the transmissivity of cast #12

With that exception Fluorescence and Transmissivity data are nominal and unedited, except where records were removed in editing temperature and salinity. SBE Fluorescence data were extremely noisy in the top 40db in both upcasts and downcasts and the comparison with bottle CHL is unusual.

The PAR channel was removed from all casts because there was either no signal or unbelievable values.

The pumped channels were removed from file 2010-16-0004.CHE because the pumps were not turned on.

The pH:SBE channel was removed from the files because sensor performance and recalibration are being studied and are not yet considered ready for the archive. Separate files were prepared for the Chief Scientist with pH:SBE included.

Based on the results of cruise 2010-19 the SBE dissolved oxygen data in the CTD files are considered, roughly:

•	±0.5  ml/l from  0– 25db

•	±0.2  ml/l from 25– 200db

•	±0.06  ml/l below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

This cruise had only a few casts and limited sampling. Since it occurred shortly before 2010-19 and used mostly the same equipment, some of the comparisons and methods used for that cruise were applied to this one.

The Log Book and rosette log sheets were obtained. There were obviously many problems during the cruise including pressure spikes as were noted on 2010-19.

Extracted chlorophyll, dissolved oxygen, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The configuration file used for this cruise had a few errors which were corrected and the resulting file was saved as 2010-16-ctd.con:

- The most recent pressure offset had not been entered, so that was changed from 0.1 to 0.8db based on the 2010-19 results.

- The transmissivity date and parameters were wrong as they have been for a few recent cruises. For those other cruises checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It was assumed that this really was 1005DR and the appropriate numbers were entered in the con files. Note that this is the transmissometer that has drifted very significantly and is producing odd deep profiles; a correction factor (2.382) will have to be applied later.

- The dissolved oxygen sensor entries were for the Owens-Millard algorithm. We should be using the SeaBird algorithm, so those parameters were entered.

##### Conversion of Raw Data

Rosette files were converted using a start time of -5s and duration of 10s and the three configuration files listed in the previous section. The Tau correction was chosen, but not the hysteresis correction for DO calculation.

An error was found in the PAR calibration and it was discovered that there was no signal in the Surface PAR channel. The SPAR had not been listed in the log book, so must have been left in the configuration file accidentally. Conversion was redone without SPAR and with the right PAR parameters. PAR values still look odd.

The ROS files were then converted to IOS SHELL format, but cast #12 could not be converted. Bottles #19 to 24 were full of spikes. Similar problems occurred during 2010-19 and it was found that repairs could be done by replacing the spikes with pad values using a text editor; this was done to the pressure, temperature, conductivity, salinity and dissolved oxygen concentration channels before conversion could be done. This process just removes the very large spikes.

CLEAN was then run to add event numbers to the ROS headers and those files were named *.BOT.

All BOT files were then plotted to check for outliers. Many were found in cast #12. Further test editing gradually removed the outliers. For some bottles about 20% of the data were replaced. This task is usually done using CTDEDIT but in this case there were sometimes spikes in only one column, sometimes in all and pressure, temperature, salinity and oxygen are all corrupted, so text editing seemed more appropriate, though it is tedious.

CNV files were converted using the Tau correction. The PAR:Reference channel was converted – this will be removed later since there is no signal.

The first 10 casts sampled only the surface and the data look ok.

Cast #11 was supposed to be a bottle cast but due to excessive spiking the CTD was shut down and returned to the surface. The downcast data are probably ok.

Cast #12 was run at the same site as #11. There was spiking on the upcast, but the downcast looks reasonable.

Cast #13 looks ok.

A few casts were examined with most attention on cast #13. All expected channels are present, but there is no signal in the Surface PAR. That is not listed in the log book, so probably was not mounted, just left in the configuration file from an earlier cruise. Other problems noted are:

- The PAR trace looks very odd with no signal (but high values) near the surface and more variability and higher values at depth. For the casts with only surface data the values are all around 300-350, but these casts were mostly at night. Cast #1 at 7:45pm has lower values than cast #10 at 11pm. For casts #12 and 13 the highest values are at the bottom. A test run using the parameters in the configuration file used at sea, produced smaller values but exactly the same shape and that shape looks somewhat like the temperature trace suggesting there was no PAR signal, just a temperature correction factor.

- The upcast temperature and conductivity are much noisier than the downcast.

- The fluorescence values are noisy but have a reasonable overall shape. There are a few cases that may be off-scale, but likely these are instrumental spikes, not natural ones. This will be examined more closely later.

- Dissolved oxygen voltage looks as usual with an offset.

- The descent rate is hard to assess because of pressure spikes but it appears to have been fairly noisy.

- Transmissivity looks ok, though the latter is low as expected for this sensor.

- The altimetry is noisy at the bottom, but it looks like an estimate can be made for the deep casts.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

A few casts were examined and there is some improvement in mid-depth pressure spikes. Some remain, but those are unlikely to be a problem. If necessary a return will be made to this stage for a more vigorous application of WILDEDIT.

##### CELLTM

Based on the results of 2010-19 CELLTM was run using parameters α = 0.02 & β=7 for both sensors.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The Tau correction was chosen for the DO derivation.

on 1 cast to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. That file was placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Cast #13 was plotted to check for agreement between the pairs of T and C sensors. The differences are included from 2 casts of 2010-19.

| Cast # (CTD#) | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2010-13-0013 | 300 | +0.001 | -0.0002 | -0.0027 | Moderate |
| 2010-19-0051 | 300 | +0.0015 | -0.0003 | -0.0032 | High, steady |
| 2010-19-0064 | 300 | +0.0008 | -0.00025 | -0.003 | High, steady |

The 2010-16 differences are close to those of 2010-19.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number; it was also used to reset the header limits.

Cast #11 was started with the pumps off, then returned to the surface to run again. To ensure that DELETE picks the right data, the first section was removed from the CLN file.

##### Checking Headers

HEADER CHECK was run and there is evidence of many bad data in files #11 and 12. Because the problems are chiefly in the upcasts and those are not needed for the profile, scans 10800 to the end were removed from file #11 and scans 37000 to the end were removed from the file #12. After running CLEAN to fix header limits, the HEADER CHECK showed most problems had been removed. This step will be repeated after DELETE to ensure no problems of significance remain.

The cross-reference check was compared with the log book, and no errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program shows the average surface pressure to be 0.9db for CTD which is low, but the researchers were deliberately trying to sample near-the surface, so this is likely not a sign of poor pressure calibration. However, it is a sign that the +0.8db offset is not too large!

The altimeter readings and water depths were exported from the headers to a spreadsheet. There were readings for only 5 casts because most did not get close to the bottom. One of the readings (cast #7) is clearly wrong, having been caused by one of the many spikes. It was removed from the header.

##### BOTTLE FILE PREPARATION

An unusual scheme was used to assign sample numbers for this cruise. Because many samples were needed at each level, two bottles were fired, but the same sample number was used for both. This was done because the desired output was to see all the data from one level in a single row. There were never samples for the same variable from each of the pair, so using a single sample number can be accommodated, though it will be awkward to identify misfired bottles. Another approach would be to create different sample numbers such as 20b or 920 for the second case of sample #20, but that is very awkward since some programs expect increasing sample numbers. So the first approach will be used, but for future projects this is not recommended. If a researcher wants to combine data in this way, the usual method could be used (1 sample # per bottle) and then special files could be created at the end by bin-averaging on pressure to achieve the single line output desired.

The BOT files were averaged to enable an addsamp file to be created. This file was edited by removing one bottle for each firing; then sample numbers were added.

The ADDSAMP file was used to create SAM files, which were then bin-averaged as SAMAVG files.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Only 4 files had entries and all look reasonable except for cast #12 for which there were many spikes in the signal; the header reading looks too low. The header reading was changed to 6m from 2m.

In the course of studying the altimeter headers, it was noted that bottles #1 and 2 of cast #13 were fired without a wait. There is evidence of a shed wake affecting the data at the time these bottles were closed. This is a particularly bad idea when firing 2 bottles at a level since there is a chance they will be sampling different conditions. In this particular case both seem to be sampling the same shed wake so most CTD values are similar for the two bottles, though the Dissolved Oxygen did not have sufficient time to equilibrate, so it may not reflect the same conditions as the temperature and salinity. The whole stop lasted less than 20s.  Other bottles from that cast were fired after a wait of at least 30s.

SALINITY

The salinity data were delivered in spreadsheet 2010-16.xls which was renamed as 2010-16 sal.xls; there were no sample numbers, just station names and bottle numbers. The spreadsheet was simplified (unneeded columns removed, separate columns were created for station name and sample number and bottle number and standard channel names were inserted) and saved as 2010-16-sal.csv which was then converted to individual SAL files. There was one comment but no flag. A flag “4” was entered since the comment referred to a serious problem. This will be checked later. There were not as many bottles as noted on the rosette sheets. There was one duplicate pair which differed by 0.0149; those were replaced with the average value, flag “6” was attached and a comment with the 2 values was added.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2010-16oxy.xls with quality flags and comments. The spreadsheet was simplified and saved as 2010-16-oxy.csv which was converted into individual ADD files. The quality flags were changed from alphabetic to numeric to be consistent with the nutrient data. There were only 2 duplicates, 1 of which was poor, so no duplicate statistics were not prepared.

NUTRIENTS

The nutrient data were obtained in spreadsheet 2010-16 SOGnuts.xls which included a comparison of duplicates. The spreadsheet was simplified, reordered on sample number and saved as 2010-16-nuts.csv. File 2010-16-nuts.csv was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file 2010-16chlarc.xls with c flags assigned by the analyst where the duplicates differed by >10%. The file was saved as 2010-16-chl.csv, header names were edited, “CHL:” was entered before the comments and event numbers were added. The file was then converted to individual CHL files. (Because all data were the average of duplicates, “6” flags should have been added to all, but this was forgotten – the flags were added later - near the end of processing.)

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files. CLEAN was run again to remove SeaBird headers and comments. (Output:MRGCLN2).

Plots were made of Salinity versus DO from the CTD and Titrated DO. For the full profiles the results look reasonable. A few of the surface-only casts look odd, but the ranges are very small, so this is not a reasonable test.

The data from pumped channels in 2010-16-0004.MRGCLN2 were replaced with pad values because the pumps were turned off. A note was placed in the header to explain this.

##### Compare

Salinity

COMPARE was run. There were only 10 bottle below 10db of which 4 were below 200db. Most of the surface values were outliers with the CTD much higher than the bottles. That is not surprising given the 1m offset in such a high gradient zone. The only significant outlier is the 103db bottle from cast #12, but the CTD data are very noisy there, so the bottle will not be flagged. The bottle from cast #1 that had been flagged earlier does not stand out in this comparison, so the flag was removed. The primary salinity is high by 0.0014 (or 0.0015 below 200db) and the secondary is low by 0.0010 (or 0.0013 below 200db). The standard deviation is ~0.002 for both channels. The difference between the primary and secondary is close to the differences noted in section 7. (See 2010-16-sal-comp1.xls.)

The values found for 2010-19 were that the secondary was low by about 0.002 (~0 below 200db) and the secondary was low by 0.005 (-0.003 below 200db) with standard deviations were 0.004.

Dissolved Oxygen –

COMPARE was run using the SBE DO and the Titrated bottle DO data. Three outliers were investigated:

- Cast #4: Duplicates both differ from the CTD by almost 2mL/L, but the local gradient appears to be extremely high, far higher than at adjacent stations. The CTD pumps were never turned on. No flag was assigned, but it was discovered that the duplicates had not been averaged so that was fixed in the ADD file. The pumped channels will have to be removed.

- Cast #12: The bottle at 103db is off by 0.4mL/L. The CTD data are very noisy. This bottle was also an outlier in salinity. No flag was assigned.

- Cast #1: The surface sample was flagged by the analyst as having low transmissivity (turbid) and it is an outlier in the comparison. The CTD data was not particularly noisy. A flag “4” was assigned.

When those outliers plus a few others (identified by residuals) are excluded the fit is:
                         CTD-BOT = 1.0258 DOX-CTD + 0.076

(See 2010-16-dox-comp1.xls.)

The fit found for 2010-19 was:

CTD-BOT = 1.0378 DOX-CTD + 0.0259

It is not obvious whether to use the 2010-19 fit which was based on a lot more data, or to use this very limited set that is heavily biased towards shallow sampling and has a narrow range of DO. The 2010-16 data obviously suits the particular project, and the values are not far off those found for other cruises using this sensor, so it will be applied.

Fluorescence

COMPARE was run using the CTD CHL and the Extracted Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD Fluorescence. The plots were unusual with CTD Fluorescence ranging from 1.2 to 2.5 while the CHL ranged from 0 to 6.2. Looking at the full files, the fluorescence has low values at depth, peaks at about 40db for casts #12 and 13 and then decreases slightly above that, though the signal is extremely noisy so there are spikes that come close to the maximum. The upcasts are similar but even noisier. The ratio of Fluorescence to Extracted CHL is about 0.3 for the CHL>4, 0.5 for CHL between 3 and 4, and 1 or higher when CHL<1. We usually do see lower ratios for higher CHL though not as extreme as this. However, since the CTD data were particularly noisy for 1 of the 2 deeper casts, we should not expect a great comparison. (See 2010-16-fl-chl-comp1.xls.)

Data from the MRG files were exported to a spreadsheet to check that everything looks ok. The only problem found was a missing sample number in file #12. That was added to the SAMAVG and MRG files.

##### 12. Shift –

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always rough estimate as the upcast data are usually very noisy. There is insufficient data from this cruise to test this, so a shift of +24 records was applied, as is almost always found best. Profile plots show little change because there is so much noise in the data. (Output: SHFFL)

Conductivity

The results of 2010-19 were applied to these files since there was insufficient data for making a judgment about the best settings. SHIFT was run on the primary and secondary conductivity for all casts with a setting of -0.5 records. (Output: *.SHFC1) This shift shows little effect on these data.

Dissolved Oxygen

During 2010-19 a shift of 60 records was found best, so that was chosen for these data as well. Plots from before and after the shift show a very satisfactory result.

pH

We have little experience with how to align pH data, so tests were done on cast #13 using shifts of +60, +80 and +100. The +100 records showed the best results, having the offset between upcast and downcast look most like that of the temperature trace. The setting of +100 was also used for this instrument during 2009-60. All casts were put through a shift of +100 records for the pH channel.

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The following warnings were investigated:

- Casts 1-10 – The warning is that the pressure range is less than 10db which is as it should be.

- Cast #12 has a jump in pressure around scan #4100. This is in the original file. A close examination of the data shows that pressure rose quite smoothly to 70.54db, then abruptly fell to 66.88db. The pressure then started to rise smoothly again. When a plot is made of temperature and conductivity against pressure there is no indication of a reversal of the CTD. Close examination of the CTD file shows that what happened is that a whole block of data (114 records) was duplicated. The file was opened in EXCEL so that temperature and conductivity differences could be calculated and they were exactly 0 for both temperature channels and ~0.00001 for conductivity. There were a series of spikes in the data from record #4101 to 4104 after which the repeated section appears. So in fact, it does not matter which of these blocks DELETE selects!

##### 14. DETAILED EDITING

From 2010-19 the primary salinity is likely closer to the bottles and there are more spikes in the secondary data for both this cruise and 2010-19, so the primary channels were selected for editing. Only casts #11-13 were edited since only the bottle files will be archived for the first 10 casts.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

In the course of editing it was found that cast #13 had an initial soak that had not been removed before running DELETE with the result that the file was missing surface data. The SHFPH file was edited to remove that initial soak and the file was then put through DELETE again.

Casts #11 and 12 needed only surface and/or bottom records removed and some light cleaning of salinity. Cast #13 needed heavier editing due to shed wake corruption.

Plots were made of casts 11-13 and the following were noted:

- Fluorescence has no off-scale values though it is very noisy.

- PAR looks unusable with higher values at depth than at the surface for some casts and too much noise to tell what is happening for others.

- Surface PAR has no signal.

- DO:SBE concentration, pH:SBE and Transmissivity need some editing in cast #12, so a second pass through CTDEDIT was made. A few isolated large spikes were cleaned in all 3 channels and some corrupted sections of DO:SBE concentration data were removed.

##### 15. Initial Recalibration

The primary salinity was found to be high by 0.0014 for this cruise and low by 0.002 for 2010-19. Both comparisons were noisy with a standard deviation of 0.002 even after outliers are excluded. This cruise included only 6 bottles below 100db and 1 of those was an outlier. No recalibration is justified.

File 2010-16-recal.ccf was applied to correct the transmissivity using the same factor as for 2010-19 and to apply the following correction to the Oxygen:Dissolved:SBE channel:

CTD-BOT = 1.0258 DOX-CTD + 0.076

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were. (See 2010-16-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

##### 16. Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are usually compared to bottle data from the same pressure. For this cruise there is not enough data to make that test reliable, but it was run just to see if there is an indication of a problem.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were not large, but there was a lot of scatter; identifying outliers leaves little data upon which to create a fit. It looks as though the downcast values may be slightly low, but by <0.01mL/L. This does not justify a 2nd recalibration. (See 2010-16-dox-comp3.xls.)

##### 17. Special Fluorometer Processing

The fluorescence data is so noisy that special files were not prepared, but just in case the data are wanted later, the COR1 files were clipped to 150db and saved for A. Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### 18. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### 19.Other Comparisons

Previous experience with these sensors –

1. Salinity - CTD#0550: Both conductivity sensors have been used many times since their last recalibrations in 2007. The results are highly variable with none of the comparison thought to be very reliable. In most cases the primary was higher than the secondary with the most reliable in recent times showing the primary to be within 0.001 of the bottles and the secondary to be low by ~0.004.

2. Dissolved Oxygen – There are 4 other cruises since the last factory calibration from which there is enough DO calibration sampling for a reasonable fit. The fits from those were:

CTD-BOT = 1.0414 DOX-CTD + 0.0005 (2009-14 - September)

CTD-BOT = 1.0388 DOX-CTD + 0.0088 (2009-26 – April)

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March)

CTD-BOT = 1.0203 DOX-CTD + 0.0696 (2009-64 - November)

CTD-BOT = 1.0378 DOX-CTD + 0.0259 (2010-19 – late April)

3. Pressure –The sensor for CTD #0550 was recalibrated in August 2007 and an offset of +0.7db has used for 4 recent cruises using offsets of +0.7db for the 3 earlier ones and +0.8db for the latest..

Historic ranges (3 standard deviations) - All data fell within the local climatology.

##### 20. FINAL CTD files steps (REMOVE and HEADEDIT)

Profiles and T-S plots were examined to see what channels should be removed. And Header Check was run to look for odd values. A few problems were noted:

- PAR – the signal is totally unbelievable

- Surface PAR – as noted before there is no signal

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Altimeter, PAR, PAR:Reference, Status:Pump and Flag.

A second set of files were prepared with the PH:SBE channel. These have extensions REMSJ (for Sophie Johannessen.)

A second SBE DO channel (with umol/kg units) was added. REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

a few large Spikes were cleaned in transmissivity and some records were

removed in editing temperature and salinity.

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly

but steadily until July 2009; then a sudden shift occurred, so that maximum values

between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study

was made of transmissivity that led to a decision to apply corrections to all

cruises between March 2008 and June 2010. The data are still considered nominal.

For details see file Transmissometer 1005DR Corrections.doc in

OSD_Data_Library:\Cruise_Data\DOCUMENTS

The pH:SBE channel was removed from the files because sensor performance and

recalibration are being studied and are not yet considered ready for the archive.

PAR data were removed because all data were bad.

Based on cruise 2010-19 the  SBE dissolved oxygen data in the CTD files are

considered, roughly:

•	±0.5  ml/l from  0– 25db

•	±0.2  ml/l from 25– 200db

•	±0.06  ml/l below 200db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD. (CTDSJ for Sophie Johannessen.)

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were about 95% for all 3 files.

##### 21. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The lines for bottles with NO sampling were removed using a text editor.

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Altimeter, pH:SBE, PAR, Status:Pump and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

At this stage it was realized that the CHL data had not been flagged as the average of duplicates, so flag “6” was added to the CHL quality channel in the final files and the CHL files.

A second set of files (with pH:SBE) was prepared for the use of the Chief Scientist with extensions CHESJ.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

Header Check was run and turned up no errors.

The most recent files from the analysts were renamed with the date they were last saved, so that we can keep track of future changes and ensure the most recent version is in the archive.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (including notes from log):

ALL CASTS: Two bottles were fired for each sample #.

1-10. Surface sampling only. No need for profile plots.

11. Spikes on upcast. Turned CTD off and brought it to the surface.

12. Repeat of cast #11. More spikes that got progressively worse.

13. No obvious problems

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 3Mar2010 | IOS |  |  |
| Fluorometer | 2356 |  |  |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |