## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | ### Updated channel names & formats in TOB files. GG |
| ### 24 Nov 2021 | ### Correcting Salinity:Bottle precision lost during HPLC addition. S.H. |
| 14 Dec 2020 | Added HPLC Data. S.H. |
| 21 July 2011 | Analysis comments corrected. G.G. |

## PROCESSING NOTES

Cruise: 2010-20

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: John P. Tully

Date: June 22, 2010 – June 27, 2010

Processed by: Germaine Gatien

Date of Processing: 21 January 2011 – 3 May 2011

Number of original CTD casts: 75			Number of CTD casts processed: 74

Number of bottle casts: 	38 				Number of bottle casts processed: 38

Number of original TSG files: 1 		         	Number of TSG files processed: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), a SBE 43 DO sensor (#1176) on the primary pump), a Seapoint Fluorometer (#2345) with a 10X cable (on the primary pump), a Biospherical QSP-400 PAR sensor (#4694), a Surface PAR sensor, an unpumped pH sensor (#692) and an altimeter (#1024). All casts were run with deck unit #0424. All casts were run with the LARS mid-ship station. Seasave version 7.16 was used. The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The rosette log sheets were generally in good order but the Daily Science Log Book contained no information about the CTD and thermosalinograph equipment. This is dangerous since we are then completely dependent on the configuration file having the right information. It is important to have someone prepare a list independent of what is in the configuration file because that is sometimes wrong; someone should be recording this information based on a visual inspection. If that is impossible and it is known that no sensors were changed after the previous cruise, then a note to that effect is useful, though it is hard to be certain that someone has not changed a sensor.

Cast #9 contained only a little surface data so was not processed. The log does not indicate any problem.

Bottles were fired for casts #45 and 46 but no samples were taken, so no bottle files were prepared for the archive.

The PAR sensor malfunctioned from cast #31 onwards. There may be problems before that with some odd spikes at depth, but the near-surface data look reasonable. Later, spikes became more sever and overwhelmed any signal even when the surface PAR indicates there should be a large signal.

The precision of the SBE dissolved oxygen channel is difficult to estimate because of the complex fits. Roughly, the comparison shows DO to be:

•	±0.3ml/l from 0-200 db (Values tend to be slightly high)

•	±0.2ml/l from 200-600 db (Values tend to be slightly high)

•	±0.04ml/l from 600-2000 db

•	±0.05ml/l from 2000-4000 db (Values tend to be slightly low)

The SBE Fluorescence values are lower compared to extracted chlorophyll than usual. They are also lower than expected compared to the thermosalinograph fluorescence.

The Thermosalinograph performed reasonably well, with none of the download problems seen during its use in two other 2010 cruises. There were only minor variations in the flow rate. There was no loop sampling but there were many CTD casts that were used for calibration information.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Note file 2010-20-0012b was produced just after file 2010-20-0012 just to fire a bottle. There is no CTD file for this event, but the bottle file 2010-20-0012.CHE comes from file 12b.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was no equipment list.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed. Without an equipment list it must be assumed the configuration files are correct.

The history of the pressure sensor, conductivity and DO sensors were obtained.

There were two configurations, differing only in the primary conductivity channel. The calibration constants were checked for all instruments. There were a number of problems:

- The transmissometer calibration date was wrong. There was a post-cruise calibration that is quite close in time and we know this sensor had a sudden change in late 2009, so the July 2010 parameters were chosen.

- The date and parameters for the dissolved oxygen sensor were not from the most recent calibration.

- The slope was entered wrong for the pressure and the offset was out of date.

After those errors were corrected the configuration files were saved as 2010-20-ctd1 (files 1-12) and 2010-20-ctd2.con (files 13-75).

The headers for ROS and CNV files for cast #61 were fixed to correct station name error mentioned in the log book.

##### Conversion of Raw Data

Data were converted using configuration files 2010-20-ctd1.con and 2010-20-ctd1.con. The Tau correction was selected but not the hysteresis correction.

A few casts were examined and all expected channels are present. There is no signal in the secondary temperature for cast #13 – the log explains what went wrong.

For cast #9 there is only a little surface data and the pumps were off. The log indicates sampling went to 82db. The problem is not in conversion as the HEX file is very small. This cast will not be processed further.

The temperature channels seem further apart than usual at times and close at other times; this could be due to a flow problem.

The dissolved oxygen voltage and pH have the usual shape. Transmissivity looks ok – these data were converted using the July 2010 calibration, so no correction will be required. The altimeter appears to have worked well. Fluorescence looks ok. PAR looks fine when readings are high enough – in some low light casts it looks very noisy. The pH traces closely resemble the oxygen voltage channel, but with slightly more hysteresis; it is not known if it was pumped or not.

Rosette files were converted using a start time of -5s and duration of 10s. The Tau correction was chosen, but not the hysteresis correction since there was no sampling below 400db.

The rosette files were then converted to IOS SHELL files. Many of these files contained no data, so were deleted after conversion. Included among those was file 2010-20-0012.ROS. So the file 2010-20-0012b was renamed 2010-20-0012.ROS.

The converted files were saved with output named *.BOT.

Temperature and salinity were plotted for all BOT files and no editing was found necessary.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity, temperature and descent rate channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM. The results were difficult to judge because the data were very noisy and the best setting varied from depth to depth and from cast to cast.

The choice (α = 0.03, β=9) looked best overall for the primary for the first configuration (casts 1-12) and (α = 0.03, β=7) for the second configuration (casts 13-75) and (α = 0.03, β=7) was best for the secondary (all casts).

CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors, though all are shallow so local gradients are significant and data tend to be noisy. The differences are often extremely noisy so these are very rough estimates.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2 | 30 | -0.0007 VN | +0.0007 | +0.007 VN | High, steady |
| 7 | 240 | -0.0003 XN | +0.0007 N | +0.006 XN | Moderate, noisy |
| 46 | 280 | -0.0001 N | +0.00005 N | +0.0005 XN | High, steady |
| 47 | 300 | +0.0001 N | +0.00005 N | +0.0005 N | High, steady |
| 52 | 300 | -0.0001 | +0.0001 | +0.0008 | High, steady |
| 75 | 300 | -0.0001 N | +0.0001 | +0.0011 N | High, steady |

The temperature differences are small but noisier than usual.

The conductivity sensor was changed between casts #12 and 13, after which the conductivity differences are extremely small. The salinity differences are also much smaller after the switch.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers,

From this point onwards file 12b will not be processed, since it contains only surface data .

An initial check shows there were no off-scale fluorescence values.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. A few problems were found, chiefly negative spikes. A text editor was used to replace a few of these, but probably more will turn up.

For cast #57 there are very slightly negative pressures at the end of the cast, with very low conductivity; transmissivity goes to near 0 and then very high, so this does appear to be the surface and not indicative of pressure error.

The cross-reference check was compared with the log book and a few errors in station names were found and corrected in the BOT and CLN files.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run. The average surface pressure was 2.4db.

The altimeter readings from the headers of the CLN and BOT files were exported to a spreadsheet. The readings were checked for the “noisy” casts at the mouth of Juan de Fuca Strait and a few other casts. Where there were entries, they looked right. However, there were casts without an entry, so checks were made that the CTD had not reached within 15m of bottom. For cast #55 it is clear it reached within 0m of the bottom! A header entry was added for that cast. For cast #73 there was a bottle right at the bottom of the cast, so the altimetry entry for the downcast was added to the header of the bottle file. For 2 other casts the CTD did not get close to the bottom. Most of the casts with no entry had only surface sampling so the algorithm worked appropriately.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. For cast #7 one of the 50db bottles was removed from the list since there was only 1 sample number for the 2 bottles fired at that level. There were several cases in which 2 bottles were fired at the surface (usually at about 6 and 1db), but only 1 sample number was assigned. For cast #62 1 bottle was fired in air and not given a sample #. The rosette sheet shows the sample being at 5db, so the first of these pairs was assumed to be the bottle sampled. There are bottle files for casts #43, 45, 46, 55, but no sample numbers, so those will not be processed further.

The addsamp.csv file was converted to CST files (with bottle position # included) to be used as a framework for the bottle files. It was used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572.

The Autosal data were delivered in spreadsheet “2010-20-sal.xls”. For cast #62 the rosette sheet indicates that the salinity sample was #101, but the analysis sheet shows 102. There were comments, but no flag column, so one was added. For the comment “no liner” a flag 3 was added. The other comments indicated that 3 readings were needed; no flag was added at this point. After COMPARE those should be revisited. The file was simplified, event numbers were added and the station name and sample number column was separated into 2 columns. The file was saved as 2010-20-sal.csv and was then converted into individual *.SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2010-20-oxy.xlsx which includes a duplicate study. The spreadsheet was simplified and then saved as 2010-13-oxy.csv. That file was converted into individual *.ADD files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2010-20nuts.xls which included a report on precisions. The file was simplified and saved as 2010-20nuts.csv. Extraneous columns were removed and header names were changed to standard format. Data were sorted on sample number. The file was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2010-20_CHL.xls which included comments and flags. On a separate page the raw data are given and details on calculation of final values. The main file was edited to add event numbers, to remove extraneous lines and columns, and to change header names to standard format; the file was sorted on sample number, and saved as 2010-20-chl.csv which was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in 4 steps. After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files choosing the bottle position # from the MRGCLN1 files.

##### Compare

Salinity

Compare was run. There is one major outlier (cast 13) in the secondary comparison. This is due to a malfunction in the CTD secondary channel for that cast. There are 3 other notable outliers and since they occur in both primary and secondary plots, these are assumed to be due sampling or analysis problems. In two cases the analyst noted that 3 readings were needed, but the differences were not large enough to explain the outliers and the 2nd and 3rd readings were identical. The third outlier had no comments from the analyst. Flag “3” was attached to the following samples:

- Sample #2, cast #3

- Sample #81, cast #19

- Sample #128, cast #38

There is a lot of scatter in the comparison and most shallow samples have much larger differences than the deeper ones.

Sample #179, cast #50 was flagged 4 because the liner was missing, but it looks ok in COMPARE.

Comments were updated by adding results of COMPARE to any analyst’s comments even if flags were not changed.

The results of COMPARE indicate that:

- Primary CTD salinity with the first configuration (casts #1-12) is low by 0.014 with only 4 points of comparison, one of which is an outlier. Including just the 2 bottles below 200db produces roughly the same result.

- Primary CTD salinity with the second configuration is low by an average of 0.0038 or 0.0001 if only bottles below 200db are included.

- Secondary CTD salinity is low by 0.0022 or high by 0.0007 if only bottles below 200db are included.

These results are in reasonable agreement with the results of section 8, given the few bottles available for the first configuration.

There is little time-dependence in these data and there are not enough deep data to assess pressure dependence.

For more detail see 2010-20-sal-comp1.xls.

Dissolved Oxygen –

COMPARE was run for Dissolved Oxygen. There are a few outliers and other points appear to fall into 2 groups as seen during 2010-36 and 2010-14 in July and August 2010.

Cast #1, sample #1 – CTD data were very noisy – no flag added.

Cast #7, sample #36 – slight outlier and leaky Niskin so flagged 4.

Cast #13, samples 54, 55 and 66 – already flagged 4 by analyst

Cast #19, sample #89 – right at surface and only slight outlier – more likely CTD problem. No flag.

Cast #54, sample #196, 198 and 204 already flagged by analyst – first 2 outliers, 3rd only slight outlier.

A few other outliers were then identified that were mostly from near the surface with noisy CTD data.

Next, a plot against file pair number was made to establish whether the two different fits fall into a temporal pattern. It does appear that the group with the higher offset occur for events #13 to #38. We don’t have bottles from every cast, so plots were examined to see if the DO voltage signal varies in the same way as it did for 2010-14 and -36.

Casts 1, 3, 10 and 11 had the “quiet” look with noise levels ~±0.0005 in low-gradient areas. Casts #12, 13, 14, 19, 25, 38, 39, 40 and 41 had the “noisy” look with low-gradient variations ~±0.0010. Casts #42, 53, 54, 60, 69 and 75 had the quiet look. In all cases the upcast data had the same noise level as the downcast data. During 2010-36 that was not the case, but the shifts were probably did seem to be associated with other CTD problems not seen in this case. During 2010-14 all casts near the beginning had one fit and then there was a shift to the other fit.

Generally the noisy casts can be identified by a quick look, so that was done for all casts for which there were no bottles and no case was found with noisy data other than between casts 12 and 41. 
In the course of checking the noise level it was discovered that the DO signal for cast #29 is extremely noisy both up and down, but the range is small and it is an area of active mixing, so the data are probably good.

Dividing the data in that way and excluding outliers described previously produced fits:

Bottle DO = 1.0330 * CTD DO - 0.0465 (Quiet DO casts 1-11 and 42-75)

Bottle DO = 1.0366 * CTD DO - 0.2302 (Noisy DO casts 12-41)

(See 2010-20-dox-comp1.xls.).

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. Note that the only samples from cast #1 were replaced with pad values, so that cast was removed from list. The ratio of SBE fluorescence to extracted chlorophyll is about 20% with slightly higher values for low values of chlorophyll and the lowest ratios for the highest CHL. (See 2010-20-fl-chl-comp1.xls.)

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records, but a few casts were examined to ensure that was appropriate and it was.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on cast #1 and 10 for the first configuration and casts 46 and 75 for the second configuration. The results were examined on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary for both configurations and +0.2s to the secondary.

SHIFT was run using those settings.

Dissolved Oxygen

Tests were run on a few casts for each sensor to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. A shift of +90 records seemed best and has been used for other 2010 cruises on which this sensor was used.

SHIFT was run using +90 records for all casts.

pH

During 2010-01 tests suggested that shifting the pH data was not needed. For this cruise it looks as though a shift of +90 records does improve the data. The same setting was used for 2010-36 and 2010-14 which followed this cruise.

SHIFT was run using +90 records.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was only 1 warning for cast #38 and it pertained to the upcast, so will not affect the downcast file.

For many casts the CTD was first lowered to about 10db for a soak period, then brought to the top and lowered for the full cast. DELETE usually chose the appropriate data (the full cast, not data from the initial soak), but sometimes it did not so that data from the initial soak got patched to the full cast producing sudden jumps in values and unstable profile. For the following casts the initial soak data were removed form the SHFpH files and DELETE was run again: 18, 24, 37, 42, 43, 57, 58, 59, 60, 62 and 75.

##### DETAILED EDITING

The secondary channels look reliable since the results of 2010-20 falls very close to those of 2010-36 and 2010-14 which followed it. The CTD was close to bottles for all three of these cruises. The primary sensors were switched before cast #13. The primary sensor for CTD1 produced salinity that was very low; after the switch the CTD2 primary salinity values are close to bottles, but evidence from later cruises showed significant drift from that sensor, so it seems better to go with the secondary.

The secondary temperature and salinity channels were selected for editing for all casts except for #11 (due to poor performance of the secondary sensors in the top 7db) and 13 (no secondary temperature data available).

All casts required some editing except for cast #68.

Only one cast had more than 1db of data removed from the bottom. Cast #34 had the bottom 5.2db removed, so the altimeter header was changed from 3.9 to 8.1db.

All EDU files were copied to EDT.

##### Initial Recalibration

No salinity recalibration is necessary for either sensor pair.

The SAM files were recalibrated using file 2010-13-recal-bottle.ccf to multiply transmissivity by 2.5, add 0.8db to the pressure, subtract 0.0025 from the secondary salinity for all casts and applied the following correction to CTD DO:

Corrected SBE DO = 1.0330 * CTD DO - 0.0465 (Quiet DO casts 1-11 and 42-75)

Corrected SBE DO = 1.0366 * CTD DO - 0.2302 (Noisy DO casts 12-41)

After this step COMPARE was rerun to ensure the changes were as expected and they were. The average difference when outliers (~7% of data) were excluded was <0.001. (See 2010-20-dox-comp2.xls.)

The MRGCLN2 files were then recalibrated using the same calibration control file.

The EDT files were recalibrated to produce COR1 files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction can be applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When major outliers are excluded the SBE DO values are high by about 0.05mL/L at the low end of the range and by 0.04 at the high end, with an average of 0.045mL/L. The plot against pressure shows a similar story with values high by ~0.06mL/L at the surface and close to zero at the bottom. (See 2010-20-dox-comp3.xls.)

The thinned files were recalibrated by subtracting 0.045mL/L and COMPARE was rerun. The results looked satisfactory with an average difference of 0.0006mL/L when only the major outliers are excluded. (See 2010-20-comp4.xls.)

CALIBRATE was run to apply this 2nd correction to the COR1 files.

##### Special Fluorometer Processing

No values close to off-scale were found.

Special files were prepared for Dr. Peña by clipping the COR1 files to 200db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A special spreadsheet was prepared with data (including sigma-t) to 50db from CTD files for all rosette casts.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The primary sensors used for the first configuration were used for 2 cruises in May and June when they showed considerable temporal drift starting at ~0 and ending low by 0.009.

The primary sensors used for the second configuration were used for cruises in July and August when they were found to be low by 0.002 and 0.003, respectively with time-dependence noted.

The secondary sensors are the same for both configurations and have been used for 4 other cruises. In May and June CTD salinity was found to be high by 0.003 and 0.002, but during 2010-36 and 2010-14 in July and August the secondary salinity was found to be very close to bottles.

2. Dissolved Oxygen

This sensor has been used for 5 other cruises since its last factory recalibration. The first 3 were spring 2010 cruises and the fits showed small and steady drift in slope and small offsets. But the 2 cruises that followed this one had very unusual fits with some casts or parts of casts with noisy dissolved oxygen data having a fit that was offset from the fit for the quieter sections. The cruises in time order and the corrections applied were:

2010-28 - Bottle DO = 1.0296 * CTD DO - 0

2010-12 - Bottle DO = 1.0339 * CTD DO - 0.0043

2010-13 - Bottle DO = 1.0414 * CTD DO - 0.0057

2010-36 “quiet”- Bottle DO = 1.0491 * CTD DO - 0.0237

2010-36 “noisy”- Bottle DO = 1.0347 * CTD DO + 0.2383

2010-14 “quiet”- Bottle DO = 1.0497 * CTD DO - 0.0115

2010-14 “noisy”- used 2010-36 “noisy” result

3. Pressure

Since its latest factory recalibration in May 2006 this sensor has been used mainly on the Ricker. Since May 2010 an offset of +0.8db has been applied.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The temperatures were slightly high for a few near-shore stations in northern Strait of Georgia and salinity was slightly low for a few near-shore casts in Juan de Fuca Strait. These do not look like evidence of instrumental problems.

Repeat Casts – There were no repeat casts.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

REMOVE was run to remove the following channels for all casts except #11 and #13: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run to remove the following channels for casts #11 and #13 only: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts #31 to 75.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

On-screen plots were checked and a few problems were found:

- In the PAR data there are deep spikes and values look very low from about cast #31 onwards, even when the Reference PAR does record high light levels. Even before that there are some odd variations, but the near-surface data look reasonable.

- For cast #14 the dissolved oxygen data looked unbelievable from about 36 to 44db; the fluorescence is also odd. This is an area where problems were noted when CTDEDIT was run, but in the course of aligning the fluorescence and DO channels the noise was likely spread beyond the area of bad T and S; the bad values were replaced with pad values in the REM and REO files between 33 and 38db.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity and spiky

data were removed from cast #14.

PAR data were removed from casts #31 to 75 since the signal was low and spiky.

The precision of the SBE dissolved oxygen channel is difficult to estimate

because the comparison with bottles was noisy, but roughly, the DO should

be considered:

•	±0.6ml/l from	   0- 25 db

•	±0.4ml/l from     25-125 db

•	±0.2ml/l below 125db

******************************************************************************

The SBE Dissolved Oxygen sensor behaved oddly for casts 12-41, with data offset

from those of the other casts when differences between bottles were plotted against

DO values. Similar problems were noted during cruises 2010-36 and 2010-14 in July

and August 2010. Using an appropriate recalibration brought those casts into

reasonable agreement with the other casts. For casts without titrated dissolved

oxygen sampling, a decision on how to recalibrate was based on the noise level

in the raw oxygen voltage. The casts with the offset all had a higher noise level

for all the cruises that had this problem.

******************************************************************************

For details on the processing see processing report: 2010-20-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were mostly between 70% and 90% in Haro Strait and Juan de Fuca Strait. The highest value was, as usual, in Saanich Inlet at ~130%. For the central part of Georgia Strait values were between 90 and 110% while for casts #45 to the end in the northern part of Georgia Strait values were between 110 and 130%. For a few cases where the saturation values were high, 3 SBE DO downcast surface values were checked against bottles. The values were sometimes slightly high, sometimes slightly low, but given the difference in time they are quite close. Comparing the upcast the SBE DO with bottles showed the SBE DO to be a little high for the Saanich Inlet case, but for casts #65 and #73 the SBE DO values were very close to the bottle values.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For the remaining steps there was a second run with an “X” added to the extensions to produce files that include pH channel. These are available from the chief scientist, but are not considered suitable for archiving at this time.

REMOVE was run on all casts except #13 to remove: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on cast #13 to remove: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts #31-75.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers and to add a comment about quality flags and analysis methods. The following note was added to explain how the SBE dissolved oxygen data were recalibrated.

************************************************************************************

The SBE Dissolved Oxygen sensor behaved oddly for casts 12-41, with data offset

from those of the other casts when differences between bottles were plotted against

DO values. Similar problems were noted during cruises 2010-36 and 2010-14 in July

and August 2010. Using an appropriate recalibration brought those casts into

reasonable agreement with the other casts. For casts without titrated dissolved

oxygen sampling, a decision on how to recalibrate was based on the noise level

in the raw oxygen voltage. The casts with the offset all had a higher noise level

for all the cruises that had this problem.

************************************************************************************

A second set of bottle files was prepared for the Chief Scientist with pH and PAR data included. Those files have extension CHEX.

The CHE bottle data were exported to a spreadsheet to check that all data are present. No problems were noted.

Standards check was run on all files and no errors were found.

##### 25. Thermosalinograph Data

Data were provided in 1 hex file. There was no loop sampling.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer settings which had the wrong scale factor. After that correction the CON file was saved as 2010-20-tsg.con.

b.) The file was converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were produced. The flow rate was ~1 with a few minor jumps.

Salinity values were very low on the 25th of June, but the ship was around the Fraser River plume, so is likely right. There is no obvious need to edit.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2010-13-ctd-tsg-comp.xls. All the data came from ~4.5db. For a few casts there were no data available at that level.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for temperature, salinity and fluorescence and the file was then reduced to the times when CTDs were run or loop samples were taken.

Those files were added to the CTD data in file 2010-13-ctd-tsg-comp.xls. There were 70 matches. In some recent uses of this equipment latitude and longitude have become stuck, or times have been wrong due to missing scans. In this file the times from the CTD and TSG are matched, so comparing positions will turn up problems of that sort. The differences in latitude were all <0.0005° and in latitude <0.0013°, and the average differences are 0.00001° and 0.00004°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

T1 vs T2 The intake thermistor was connected throughout the cruise. The average difference was 0.25º when most data were included. That is mid-way between the results for May and July cruises using the same equipment.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. The differences are very noisy as is expected in June in inland waters. Only 3 casts had a mixed-layer depth >10m (judged by the distance below 4m when the salinity changed by >0.005.)

The TSG intake temperature was higher than the CTD by an average of 0.21Cº and a median of.0.073Cº when all data were included and medians of 0.062Cº and 0.061Cº when 23 and 32 outliers were excluded, respectively. When only the 3 well-mixed casts were included the TSG temperature was high by 0.07Cº.

The TSG salinity is lower than the CTD by an average and median of -1.1 and -0.16 all data are included. When only the 48 values with the lowest standard deviation in the TSG salinity are included, the TSG salinity is low by 0.100 and when a further 9 values are excluded the median is -0.071. When only the 3 casts with well-mixed surface waters are included the median difference is -0.064.

The ratio of TSG fluorescence to CTD fluorescence is ranges from 5.8 to 11.4 with an average of 12.4 and median 11.4 when all data are included. Removing 23 casts with the highest standard deviation in the TSG fluorescence shows a median ratio of 11.6 and removing a further 9 values makes only a small difference; the median ratio to 11.9. Restricting the group to well-mixed casts is probably not useful for fluorescence, but the average was 16.82 for those 3 values. This ratio is higher than normally observed, but this may be because of problems in the CTD fluorescence. The latter was found to be low compared to extracted chlorophyll samples with values of SBE Fluorescence about 20% of the extracted CHL. See 2010-20-ctd-tsg-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and were used for 2009-10 (with a different intake thermistor), 2009-11, 2010-01, 2010-28, 2010-12 and 2010-13 before this cruise and 2010-36 afterwards.

The TSG intake temperature was within 0.004Cº for the two 2009 cruises and 0.001Cº during 2010-01. For 2010-13 the TSG intake was high by 0.02Cº and for 2010-36 the differences were 0.007Cº. These are all good results – there are bound to be variations based on how well the surface waters are mixed.

The salinity was low by about 0.02 for the two 2009 cruises. For 2010-01 the results varied with flow rate with salinity low by 0.06 and 0.16. For 2010-28 and 2010-12 no reliable comparison was available because of clock problems. For 2010-13 the salinity was low by 0.06. For 2010-36 the salinity started low by 0.06 but drifted until it was low by 0.4; this was probably due to fouling.

The TSG fluorometer was high by a factor of 2.2 to 6 for 2009-10, by about 8 for 2009-11, ~2 for 2010-01, 2.5 for 2010-12 and 3 for 2010-13. For 2010-36 the average was ~4 but there was a lot of variability and known biological fouling.

Conclusions

1. The TSG clock worked well and there were no problems with the download of position information.

2. The temperature increased by about 0.25º between intake and lab.

3. The CTD intake temperature is within 0.06 or 0.07Cº of the CTD temperature, which is reasonable given the sampling levels are not identical and there is a fairly high surface temperature gradient.

4. Salinity is low by ~0.06 based on the CTD comparison which matches the observations of cruises before and after this one.

5. As usual there is a lot of variability in the ratio of TSG fluorescence to CTD fluorescence. The TSG to CTD fluorescence ratio is ~12 which is higher than the usual 3 or 4. This may be due to CTD fluorescence being unusually low.

f.) Editing

Plots were examined and no editing was deemed necessary.

g.) Recalibration

File 2010-20-recal1.ccf was used to add 0.06 to the salinity.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars – including notes from log, rosette sheets:

.

1. 3X fluorometer cable

3. Up to 100 then back down to 125 to fire bottle.

4. Stopped at 174m due to wire angle and restarted.

7 . Bottle 13 leaking.

9. Log says sampled to 82db but only surface data in file. No files prepared.

10. Bottle 13 leaking and bottle 3 vent not tight enough.

11. Acquisition did not start until 11db. pH sensor cap left on.

12b. Run only to get 5m bottle –combine with 12 for bottle file. Not CTD file to be prepared.

13. Secondary temperature malfunctioned due to a tie-wrap. Use primary sensors.

13. Bottle #13 replaced.

45. 2 bottles fired for water, no sample #s.

46. 2 bottles fired for water, no sample #s.

50. Bottle #1 had open vent.

55. May have touched bottom gently.

56. Reboot – froze during NMEA initializing.

59. High fluorescence

61. Station name was 78 – change to 8.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 06May78 | Factory “ |  |  |
| Conductivity Casts 1-12b | 2280 | 03Jul08 | “ |  |  |
| Conductivity Casts 13 - 75 | 1764 | 10Feb09 |  |  |  |
| Secondary Temp. | 2038 | 06May08 | “ |  |  |
| Secondary Cond. | 3394 | 06Mar09 | “ |  |  |
| Transmissometer | 1005DR | 05Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 10Nov2009 | Factory |  |  |
| PAR | 4565 | 03Mar2010 | IOS |  |  |
| SPAR | 16504 | 3Mar10 |  |  |  |
| Fluorometer | 2345 | ? | IOS |  |  |
| Pressure Sensor | 69698 | 26May2006 | Factory |  |  |
| pH | 180692 |  |  |  |  |
| Altimeter | 1024 | ? | ? |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2010-20

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:secondary | 2416 | 23Dec06 |  |  |  |