## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2010-29

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Discovery Islands

Area: Discovery Islands

Chief Scientist: Tuele D.

Platform: ECHO

Date: 1 April 2010 – 4 April 2010

Processed by: Germaine Gatien

Date of Processing: 16 July 2010 –  19 July 2010

Number of original CTD casts: 32	Number of casts processed: 31

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#1294) was used with pump working. A 1.7L Niskin bottle was attached to the hydro wire about 1 m above the CTD.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard. The names include an event number but they differ from those in the log book by 2. There are notations in the log book beside many of the entries giving the numbers that are found in the file names. When changing the file names to standard names, the original event numbers from the log were used.

The CTD Daily Log was available and includes details on salinity calibration sampling. The times in the log book are said to be UTC, but since all sampling was between 09:00 and 16:30 these are assumed to be in local time, Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

The soak period was generally 1 to 2 minutes, which is probably long enough when surface temperatures are close to air temperature, but a 3-minute soak is recommended for the initial cast.

For two casts, the upcast section was selected for archiving because of poor downcast data near the surface; one was the initial cast for which the soak was very short.

The pressure appears to be good to ±0.2db.

There was good agreement between the salinity samples and CTD, suggesting salinity is good to ±0.002 at 80m. Errors are likely larger in higher gradients due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. For this cruise gradients were generally low so this should not be a major problem.

There have been many poor comparisons between bottles and CTDs over the past year. One explanation is that poor seals (due to using old liners or chipped bottles) may lead to evaporation between collection and analysis. These samples were analyzed less than 2 weeks after the cruise, minimizing evaporation, so this lends weight to that explanation.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The CTD Daily Log was obtained.

The original file names were non-standard but the names and event numbers were in the names, so it was clear which files went with what event.  
The dates in the log book appear to be in Pacific Daylight Time though there is no note indicating that. Work was between 09:00 and 16:30 so UTC makes no sense and it is usual to record local time for this type of cruise.

Configuration file 1294.con contains the most up-to-date pre-cruise calibration coefficients for this instrument; all sensors were calibrated in April 2008. No data has been processed that deployed this instrument since that time. All coefficients were checked and were correct. The file was saved as 2010-29.con

- Conversion of Raw Data

- The raw data were converted using conversion file 2010-29.con.

- File names were changed to standard format.

- The start times in the headers are in fair agreement with the log book times.

- A few casts were checked and no problems were found. The soak time was typically ~190 scans or 95s. The temperature and conductivity seemed to have settled by the end of the soak, but a longer soak is recommended especially for the first cast. After the first few casts, the CTD was usually lowered to ~2m and soaked some more, then returned to the surface. That brings the soak to the recommended length.

- Cast #10 was aborted and the file contains only surface data. It will not be processed further.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

The conductivity was not aligned at this point – that step will be done later using SHIFT.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets. The EDIT HEADERS routine was run with file “hdr-fix.hdr” to add the lines for that information and a text editor was used to add details based on the log entries.

8. Checking Headers

Track plots and header checks turned up a few header errors which were fixed. A header summary was produced and details were checked against the log records. No further errors were found. Track plots were added to the end of this report.

The surface check was run. The average surface pressure is +0.4db. There are some slightly negative values, but the pressure values from the first few scans are probably not reliable. For the upcasts when the pressure reaches zero the conductivity is generally extremely low. The pressure offset applied was 0 and looks appropriate. No further correction is needed.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. No problems were noted. ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

10. SHIFT

Tests were run to see if alignment of conductivity was required. Conductivity shifts from +1.5 to -2.5 records followed by a recalculation of salinity were tried. The best results were with -1.5 though -1.7 also but produced good results.

SHIFT was run using -1.5 records for all casts.

11. DELETE

For most of these casts there was a drop to 1 or 2 metres, then a return to the surface, a short wait, then a full profile. To get the best data, it is necessary to remove the initial drops or DELETE will patch the first drop to part of the later drop.

So a text editor was used to remove the initial drop from most SHF files.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

CTDEDIT was used to clean salinity very slightly, remove bad surface records for some casts and bottom records corrupted by shed wakes for others. 9 casts required no editing.

In the course of editing cast #1 it was seen that the soak period had not been long enough, so the file was put through REVERSE, and then DELETE. The upcast data is ok.

Cast #30 also had poor data in the top 10db of the downcast, so REVERSE and DELETE were run on that as well.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined. There are unstable features, but they occur in well-mixed casts with little variation – the instabilities are very small. No further editing will be done.

14. Intercomparisons

COMPARE – There were 3 samples from about 80m. These were compared with CTD data from either 1db or 1.5db above the maximum pressure. Log notes are contradictory about which distance was used, or perhaps it varied among the 3 samples. The average difference between samples and CTD salinity indicates that the CTD was either high by 0.002 or 0.0001. In either case this is a good result for this instrument.

Previous Use of CTD – No data has been processed from this instrument since its last factory recalibration.

Historic ranges – No local climatology is available.

15. CALIBRATION

There is no indication that the salinity or pressure data need recalibration. None was applied.

16. REMOVE and HEADEDIT

REMOVE was run to remove the following channels from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

HEADEDIT routine was used to add the following comment to the headers:

Salinity calibration samples suggest that salinity is +/- 0.002 in

low temperature gradient areas. Errors are likely to be somewhat larger

in higher gradients.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Particulars:

10 – Cast aborted due to strong current.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise # 2010-29 	Dates: 	Start:	1 Arpil 2010	    End:	4 April 2010

Location: Discovery Islands        Vessel:     ECHO               Party Chief: Tuele D.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/19 SEACAT / 1294       Cruise ID#:	2010-29

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 1 Apr 2008 | Factory |  |  |
| Conductivity | 1294 | 1 Apr. 2008 | Factory |  |  |
| Pressure Sensor | 163223 | 2 Apr. 2008 | Factory |  |  |