## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | ### Updated channel names & formats in TOB files. GG |
| ### 24 Nov 2021 | ### Correcting Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 30 Nov 2020 | ### Added HPLC data. S.H. |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |
| 24 Nov. 2015 | Added missing CHL data to file 2010-36-0215. G.G. |
| 20 Feb 2014 | Corrected errors in file 2010-36-0204. Sample had been assigned to wrong Niskin bottles near surface. CHE G.G. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2010-36

Agency: OSD

Location: North-East Pacific

Project: Ocean Acidification

Party Chief: Ianson D.

Platform: John P. Tully

Date: July 20, 2010 – August 17, 2010

Processed by: Germaine Gatien

Date of Processing: 13 December 2010 –25 February 2011

Number of original HEX files: 149 CTD and 5 TSG (raw files for 3 CTD casts are missing)

Number of CTD files: 149 (many split casts)    	Number of CTD casts processed: 126

Number of bottle casts:	104 (many split casts)    Number of bottle casts processed: 81

Number of original TSG files:  5	         	Number of TSG files processed:  5

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), a SBE 43 DO sensor (#1176) on the primary pump), a Seapoint Fluorometer (#2345) with a 3X cable (on the primary pump), a Biospherical QSP-400 PAR sensor (#4694), an unpumped pH sensor (#692) and an altimeter (#1024). All casts were run with deck unit #0424. All casts were run with the LARS mid-ship station. Seasave version 7.18c was used. The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order but there was no information entered about the thermosalinograph. The sampling notes from Leg 1 were extremely helpful in dealing with problems that occurred at sea. Many rosette casts were run with user input (bottles fired out of order.)

This was a very complex cruise with frequent computer crashes leading to many split casts. This was further complicated by the use of user-driven bottle firing where decisions were made on the fly. When Niskin bottles are fired out of order, good record-keeping is critical. While it was sometimes done well on this cruise, there was often little indication of precisely what was done, and sometimes the information provided was incorrect. The newer (white) rosette sheets are much more conducive to making this clear. The BL files do provide a good record of the order of Niskin firing, but it is useful to have detailed rosette logs to complete the picture. The BL files do not indicate how sample numbers were assigned or what samples were drawn. Usually when table-drive firing is done the sample numbers are assigned in order of firing, but for this cruise they were assigned in order of Niskin bottle #s. There was a good reason for using the method, but it took some trial and error to discover what had been done.

The data from events #43, 151 and 152 are missing.

The salinity analysis spreadsheet has sample number and station name in the same column. It would be much more efficient to record these in separate columns since someone else will need to separate them later and that process is error-prone. Having an event # column would also be helpful, but is not essential.

In cases where data acquisition was interrupted so that there are 2 or 3 files for each deployment, data were combined as required and event numbers were set to match the first of each of the set. The CTD files all come from a single original file, but the CHE files for 17 casts are pieced together.

The secondary CTD salinity was very close to the bottles. For the past few years we have found the CTD to be consistently lower than bottles even when a post-cruise calibration indicated minimal drift. Sampling problems (worn lids, chipped bottles or salt crystals in the bottles) have been suggested as the cause. Lids had been replaced for this cruise and it is possible greater care was given to cleaning. Or, if may just be that the sensor has drifted high enough that it offset any sampling problem. The scatter in the fit is +/-0.002 even after removing many outliers. The primary salinity was within 0.001 of the bottle values at the beginning of the cruise, but drifted so that it was reading low by about 0.004 by the end. Primary temperature and salinity data were chosen for archiving for only 1 cast (#199) because the secondary conductivity data were bad.

Comparisons with 3 Argo floats launched during the cruise confirm that the secondary salinity is the better choice, and suggests that the salinity is reliable within 0.002. Unfortunately the floats were not equipped with dissolved oxygen sensors.

Comparisons with data collected using a CTD from the University of Victoria show reasonable agreement between temperature and salinity along lines of constant sigma-theta in deep water.

Work is ongoing to develop calibration and assessment methods to apply to the data from the SBE pH sensor. These data will not be archived at this time, but they are available from the chief scientist.

The SBE Dissolved Oxygen data for this cruise were unusual and their quality is considered lower than usual. There appear to be 2 stable states for the sensor, one having higher noise level than the other, even in the raw voltage channel. The high-noise segments produce DO data that is approximately 0.24mL/L lower than the low-noise segments (based on comparisons with bottles.) For most casts the whole downcast has one or the other state, but for 6 casts there was a mix of the two. The higher noise level segments were recalibrated to add 0.24mL/L to each. CTDEDIT was used to remove DO data between the 2 sections because there was a spiky and gradual shift between them. A similar process was applied to the bottle files.

The precision of the SBE dissolved oxygen channel is difficult to estimate because the comparison with bottles was complex and noisy, but roughly, the DO should be considered:

•	±0.4ml/l from 0-300 db

•	±0.25ml/l from 300-500 db

•	±0.1ml/l below 500db

The transmissometer used for this cruise had a sudden shift in values in summer 2009 and since then has displayed unexpected decreases in values at high pressures, >500db. There were also more spikes at depth than usual, all towards lower values. As usual, we state that transmissivity values are nominal, but there is less confidence in these values than usual.

The Thermosalinograph data had some sections of bad salinity and fluorescence. The salinity problem may be due to biological fouling or bubbles, while the fluorescence became a problem after a large influx of euphausids in Barkley Sound. Data that are obviously bad have been removed. The salinity data from late in the cruise are considered of lower quality than usual, so only 2 decimal points are reported.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. For the first leg there was a summary of problems and points of interest with reference to processing.

This was a long and very complex cruise, so the first step was to prepare a spreadsheet (2010-36-summary.xls) to keep track of some of the variables. The logs and sampling notes were examined and information was entered in the spreadsheet such as:

- Event #

- What sample #s go with this file?

- Was the PAR and/or pH sensor mounted?

- Were the Niskin bottles fired out of order? (Table-driven)

- Was this a repeat cast? (This refers to full casts, not partial casts due to computer crashes.)

- Were there loop samples during CTD?

- Was there a cast from the Trace Metals CTD at the same site, and if so, what sample numbers go with those data?

- A comment column was used to enter other details likely to affect processing, such as known errors in headers or details on computer crashes leading to multiple files to get a single cast.

- A TSG comment column was used to indicate rough timing of loop samples.

Details were added to this spreadsheet as further problems were discovered or resolved.

In preparing this file it was discovered that there were no data for casts #43, 151 and 152 which are mentioned in the log.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity were obtained in spreadsheet format.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. There were a number of problems:

- The transmissometer calibration parameters were wrong.

- The date and parameters for the dissolved oxygen sensor were not from the most recent calibration. Fine-tuning of factor E was done during 2010-13, so that value, 3.5, was substituted for the factory value 3.6.

- The slope was entered wrong for the pressure.

After those errors were corrected the file configuration file was saved as 2010-36-ctd.con.

##### Conversion of Raw Data

Data were converted using configuration file 2010-36-ctd.con.

A correction was made to the header of cast #76 based on the sampling notes. The other error mentioned in those notes could not be corrected since the log does not have an entry for bottom depth because the sounder was not working.

A few casts were examined and all expected channels are present. The temperature and conductivity pairs are reasonably close, though upcast and downcast differ significantly with more noise in the upcast. The descent rate was often very noisy with some complete reversals of direction and there is evidence of shed wakes in the T and C traces. Transmissivity shows the pressure dependence that has been observed from this particular sensor since late 2009. From about 500m downwards the values go down. For some early casts the transmissivity is noisy below 150db; because the spikes are all towards lower values, the bin-averaged data are apt to be lower than they should be. The dissolved oxygen voltage has some odd traces at depth (ex. #52, 57 and 186). Fluorescence looks ok with dark values <0.03 – 0.05ug/L. PAR looks fine. The pH traces closely resemble the oxygen voltage channel.

##### Rosette File Conversion and Study of Split Casts

Rosette files were converted using a start time of -5s and duration of 10s. The TAU and hysteresis corrections were chosen. Because Niskin bottles were fired out of order on many casts, the Bottle Position in Carousel was converted. There were rosette files for every CTD cast, but in 45 cases the files were empty. Those ROS files were deleted.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers, with output named *.BOT. A preliminary check of headers turned up no problems; fluorescence did not go off-scale. Peak transmissivity is 63%/m which looks reasonable. As noted in the log there were instances of negative dissolved oxygen values that will have to be investigated later.

Temperature and salinity were plotted for all BOT files and a few casts appeared to have outliers. Those casts were examined in CTDEDIT.

- For cast #6 the secondary temperature and salinity data are much noisier than the primary for 3 bottles around 12db, but it is not completely clear whether the noise is real or instrumental. The one salinity sample available would suggest that the problem is with the primary channels and the comparison of upcast with downcast confirms this. It looks like there was a problem with flow to the primary sensors which could suppress variations in those channels. No editing was applied.

- For cast #31 there are large differences in both temperature and salinity above 150db. The pumps were not turned on for all but the last bottle at the surface. The salinity sample shows the data for that bottle (Niskin #21, Firing #9) are useful. For the other bottles it will be necessary to remove the pumped channels.

- For cast #108 the primary salinity looks low for a few points around 20db. Primary temperature and conductivity also look a little off. The affected primary temperature and salinity values were removed in CTDEDIT and the corresponding conductivity values were replaced with pad values in the ED1 file.

- CTDEDIT was used to remove a few points in the primary salinity channel around 300db (Niskin #9).

- For cast #199/200 the salinity channels differ significantly at about 30db of #199; examination of the full files shows that the secondary salinity looks poor, in general, both during the downcast and upcast, as though the flow to the sensors was impeded. It eventually settles to reasonable values during most bottle stops. The problem continued after a restart due to a computer crash, so event #200 is affected.

- For cast #225 bottle #5 has 3 bad secondary salinity points – these were removed using CTDEDIT.

Editing details were added to the header comments of the edited files. The output files were then copied to *.BOT.

The following groups of 2 or 3 casts come from cases where the computer froze and a new file was started to complete the cast. In each case there is a full CTD cast in the first file, so CTD file preparation will be straight-forward. However, the bottle files will require careful merging and the first file # will be the one selected for the final files. In the following description of the problem casts [#-# & #-#] will be used to indicate [Firing # range & Niskin # range] in each file.

- 1/2 File #1 has a complete CTD profile. File #2 has only surface data and there was no sampling, so this file will not be processed further.

- 30/31 – [1-12 & 1-12] / [1-9 & 13-21]. Straight-forward merge possible.

- 33/34/35 - [1-6 & 1-6] / [1-4 & 9-12] / [1-3 & 13-15] – This is the most confusing case.

The BOT files are missing data from Niskin Bottles #7-8 and #13. The rosette log indicates that both Niskin #7 and #8 were intended to be fired at 150m, though #8 has obviously been corrected and probably originally was planned at 125m. The DO samples from bottles #7 and #8 have values that are very close, both looking appropriate for 150db, so the rosette entry change was probably based on that result. While there are no data in the BOT files from 150db, data can be fabricated from the full file #33 for the time just before the computer crash that was likely set off by the bottle firing.

For Niskin #13 at about 20db no data were acquired in the 2nd or the 3rd file, even though the BL file indicates it was fired. There is data from the full file #34 from just before the crash at 20db with DO values that are a reasonable match for the bottle sample. There is also data from the beginning of the next file that are also a reasonable match, so data can be fabricated to go with this bottle.

The biggest mystery is the source of sample #307 said to be from 125m and entered beside Niskin #17; there is an arrow linking the sample to Niskin #8, presumably to indicate that this is to make up for the fact that #8 closed at the wrong level. The DO value looks reasonable for 125m, based on downcast CTD data at that level. But there is no Niskin #17 in any BL file or BOT file, and no CTD data were acquired at 125db during the upcast. Is it possible that there was another file between #34 and 35 that perhaps got overwritten by 35? That would be an easy mistake to make under the circumstances.  Data could be fabricated to go with the analysis results for Niskin #17 by using downcast data, but that is quite a stretch at a level where variability is fairly high. So it may be best to enter only the DO sample value and Niskin # as a header comment.

- 89/90/91- [1-15 +17 & 1-15 +17] / [1-2 & 16-17] / [1-7 & 18-24] While there are data for bottle #17 in the first file, the pressure does not fit the rosette log. Fortunately, salinity bottles make it clear that the data in file #90 is not from the depth from which the samples came. The bottle had fired at ~70db, not the intended 50db, so that should not be used. The data in the file #89 are the ones to pick. Remove Niskin #17 from file #90.

- 94/95 - [1-19 & 1-17 + 19-20] / [1-4 & 17-18+21-22] Niskin #17 is found in both files, but since the first crashed during firing, that data would be the closest to what is in the bottle, if the bottle actually closed then. Remove Niskin #17 from file #95.

- 107/108 - [1-13 & 1-13] / [1-9 & 14-22] Straight-forward merge possible.

- 111/113 - [1-17 & 1-16+20] / [1-8 & 16-19+21-24] Bottle #16 is repeated. Bottle #16 looks ok in both, but first file likely to be the data from bottle closing time. Remove Niskin #16 from file #113.

- 163/164 - [1-10 & 1-10] / [1-13 & 10-22] - Bottle #10 is repeated. Bottle #10 looks ok in both, but first file likely to be the data from bottle closing time. Remove Niskin #10 from file #164.

- 169/170/171 - [1-2+4 & 1-3] / [1-4+6 & 4-6+9-10] / [1-3 & 7-8+11] For file #169 there are no data from the 3rd firing because the BL file has identical entries for the 3rd and 4th firings. There are data for Niskin #3 from the 4th firing at 31db. Similarly, for file #170 there are no data for the 5th firing, but the BL file shows identical entries for the 5th and 6th and the 6th does have data for Niskin #10. All bottles are present so a straight-forward merge should work, though the bottle #s are going to need editing in all 3 files.

- 174/175 – [1-10 & 1-10] / [1-8 & 10-17]. Bottle #10 is in both files. Remove Niskin #10 from file 175.

- 184/185 – [1-5 & 1-4+13] / [1-8 & 4-12]. Bottle #4 is in both files. Remove Niskin #4 from file #185.

- 188/189 – [1-7 & 1-7] / [1-18 & 7-24]. Niskin #7 is in both files. Remove Niskin #7 from file #189.

- 197/198 – [1-19 & 1-16+18-20] / [1-2 & 17+19]. It is unclear why bottle #19 was repeated as it was not the last bottle fired before the crash. The firing of Niskin #19 in the second file is near the surface, while the first was at the intended ~15db level. There is no way to determine what was in the bottle as there was no DO or salinity sampling. Perhaps the HPLC will shed light on this at some point, but it seems most likely that the samples are from ~15db. Niskin #19 will be taken from the first file. Remove Niskin #19 from File #198.

- 199/200 – [1-12 & 1-12] / [1-9 & 12-20].  Niskin #9 is in both files. Remove Niskin #12 from file 200.

- 204/205/206 – [1-9 & 1-9] / [1-3 & 9-10+15] / [1-8 & 10-17]. Use Niskin #9 from 1st file, #15 from 2nd file. Remove Niskin #9 from file #205 and remove Niskins #10 &15 from file #206.

- 209/210/211– [1-8 & 1-8 / [1-7 & 8-11+13-15] / [1-3 & 11+15-16]. Niskin #12 did not fire. Niskin #8 should be removed from file #210 and Niskins #11 and #15 should be removed from file #211.

- 215/216– [1-9 & 1-8+12] / [1-16 & 2-11+13-14]. Niskin #8 is repeated. Remove Niskin #8 from File #216.

- 220/221– [1-2 & 1-2] / [1-16 & 2-17]. Niskin #2 is repeated. Remove Niskin #2 from File #221.

Establishing which files to use for which bottles is complicated because user-driven bottle firing was in use, which required adjustment each time the computer crashed. The Niskin bottle associated with the crash was often fired again in the next file, presumably as there was some fear it had not actually closed during the crash. In a few cases where there was significant change between two files, bottle samples suggest that the bottles did close the first time. In most cases there is no significant difference. The 2nd firing should be removed before the files are merged. When COMPARE is run, if there are outliers at the critical spots, this decision can be revisited.

In summary, there are 17 cases where 2 or 3 files need to be merged to create complete bottle files.

Before merging, 15 files will need to have some data removed and headers adjusted to reflect that. The addition of sample numbers to the bottle files is likely best done before merging. After merging and bin-averaging, the entries for file #33 will need adjusting to add CTD data for 3 bottles and to add a comment for sample #307. For details on how this process was done, see section 11.

There were often good notes about which Niskin Bottle was associated with which consecutive bottle firing. In a few cases the notes were far from clear. While the facts can be determined from other evidence such as the BL files, only the rosette sheet shows what sampling was done; if it is not clear whether a sample came from Niskin #13 or the 13th bottle fired, errors can occur. The newer style white sheets are much clearer than the yellow ones; the Niskin # and Firing # columns were often reversed, but that fact was usually noted on the sheets. User-driven firing requires great care in record keeping.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity, temperature and descent rate channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM. The goal is to make upcasts look closer to downcasts on a T-S surface. Most casts had stops for bottles and few casts without those stops had a steady enough descent rate to make it easy to judge which settings are best. Near-surface data often differed a lot, and those changes may well be real so the results were extremely difficult to judge. It was clear that some adjustment helped, but most choices looked best in some part of the profile.

The choice (α = 0.02, β=7) looked best overall for the primary and (α = 0.0245, β=9.5) was best for the secondary. CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often extremely noisy so these are very rough estimates and if there were a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 20 | 500 1000 1900 | Too noisy -0.0002 -0.0003 | Too noisy 0.0001 0.0001 | Too noisy 0.0015 VN 0.0014 | High, Very Noisy |
| 50 | 500 1000 1900 | -0.0007 -0.0005 -0.0006 | 0.0001 0.0001 0.0001 | 0.0015 0.0017 0.0017 | High |
| 65 | 500 1000 1900 | -0.0008 -0.0007 -0.0007 | 0.00005 0.0001 0.0001 | 0.0015 0.0017 0.0018 | High, noisy |
| 132 | 500 1000 1900 | -0.0004 -0.0002 -0.0003 | 0.00015 0.0002 0.0002 | 0.0025 0.0023 0.0023 | F.High, noisy |
| 162 | 500 1000 1900 | ~0 -0.0007 -0.0007 | 0.0002 0.0002 0.0002 | 0.0025 0.0033 0.0032 | High, very noisy |

The temperature differences are reasonably small; they might be increasing with time, but that is not clear. Conductivity differences are small with a small increase with time. Salinity differences are fairly small but are increasing with time. There is no significant pressure dependence in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers,

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. Negative pressure values were found in one cast with pumps on and data showing one conductivity cell was in water, but the other might not be, so very close to the surface. It may soon be time to increase the offset by a further +0.1 or 0.2db, but the current setting looks ok for this cruise. The transmissivity looks too high for one cast, but that proved to be a single spike in the record. Dissolved oxygen concentration has spikes, but so does temperature. That will have to be checked later when spikes have been removed. The fluorescence did not go off-scale.

The cross-reference check was compared with the log book and a few errors found and corrected:

Station names were inconsistent for the LG line: G 1, G 2… were changed to LG1, LG2 etc.

Station names were also inconsistent for the LC line. Those were fixed.

Cast #203 had been saved as 202 which was actually a NET cast, so that was fixed.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run. The average surface pressure was 3.4db.

The altimeter readings from the headers of the CLN and BOT files were exported to a spreadsheet. A selection of casts was checked and where there were altimetry header entries, they looked fine; in spite of some large spikes the algorithm worked well. There was no altimeter signal for cast #198. There were a number of casts where either the water was much deeper than recorded in the log or the altimeter malfunctioned; for some of them there is a note in the log that the altimeter was not working, but there is a signal, just not a useful one. Bottom depths were missing from some entries. In two cases the information from the log was added to the header of the CLN and MRGCLN2 files, but in the other cases there was either no entry in the log, or the maximum depth in the data file made the log depth look wrong. Frequently the water depth entry is lower than the maximum sampling pressure (even allowing for the pressure being higher than depth).

##### BOTTLE FILE PREPARATION

Before running the usual steps in bottle preparation it was necessary to remove some data from the following BOT files: 90, 95, 113, 164, 175, 185, 189,198, 200, 205, 206, 210, 211, 216 & 221. (See section 4 for details on what was removed.)

First those files were copied to another directory in case an error was made. The adjusted files were saved as BOT1 and put through CLEAN to fix the header limits with output BOT.

The file names for all the split casts (multiple files) were then adjusted so that the partial casts all have the name of the 1st file in the group, but the extensions were named BOTx, BOTy and BOTz.

The BOTy and BOTz files were opened in Ultraedit to adjust the Bottle number to match the rosette sheet number.  Then JOIN was used to produce a single BOT file for the split casts: 30, 33, 89, 94, 107, 111, 163, 169, 174, 184, 188, 197, 199, 204, 209, 215 & 220. The cast file list was adjusted to remove the 2nd and 3rd fragments from further processing.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. An initial attempt to do this was found to be wrong when COMPARE was run. It was discovered that sample numbers had not been assigned in the usual Botttle_Number sequence, but instead they are in Bottle_Position sequence, i.e. in order of Niskin numbers. This would not be a difficult job if both sets of numbers were clear from the rosette sheets, but they are not. For example cast #148 has the columns for Firing # and Niskin # reversed, two bottles that were fired are not recorded and because of that Niskin #10 is associated with the wrong Firing order #. The BL files were used to assist in the process, but that is a slow route. The ADDSAMP file was reordered on Cast # & Bottle_Position, then sample numbers entered, and then it was reordered again on Cast # & Sample_Number and file ADDSAMP3 was produced. Not surprisingly, it was later discovered that there were many errors made in this process!

For cast #33, as noted earlier, there were samples assigned to 3 bottles with no CTD data in BOT files. For now those sample #s will not be included – they will be added later to the SAMAVG files.

Two sample #s were repeated on casts 178 and 182. The samples from #178 look like they do not involve any IOS analyses, the sample numbers were replaced with pad values. If the researcher requires CTD data, there are several other bottles in the file from the same level.

The addsamp3.csv file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

The SAMAVG file for cast #33 was adjusted to add Bottle Numbers 8, 9 and 13. CTD data were obtained for bottles #7 and #8 from last 5s of the full file 2010-36-0033 (just before the bottles are presumed to have closed) and for bottle #13 from 2010-36-0034. Plots of descent rate and temperature versus scan number suggest that the temperature had settled reasonably for each case. The files were opened in EXCEL and 60 scans were averaged. It is assumed that bottles #7 and #8 were fired at about the same time. A warning was put in the headers, and after merging of bottle files “c” flags will be added to the bottle analysis results to indicate some uncertainty in the match to CTD data.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. Comments and were included but no quality flags. There were some odd format errors in the sample label entries for 12 lines in the spreadsheet, but based on rosette log entries it was easy to correct them. The spreadsheet included duplicates and loop samples and samples from the Trace Metal Rosette casts. The TMR samples were not labelled as such, and many were not ticked off on the TMR rosette sheets. There were a few duplicates which had not been averaged. The analysis log sheets were also available.

The spreadsheet was saved as 2010-36-sal.csv. Sample numbers and station names were combined in a single column so those were separated. A column was added for quality flag and where comments suggest the need, flag “3” was added. One comment is hard to understand (sample #1274 -“Odd cond + sample # + Niskin #”) and after COMPARE should be investigated.  Event numbers were added, and a column was added to distinguish the TMR data from other data. The duplicates were copied to spreadsheet 2010-36-sal-duplicates.csv and differences and averages were found. The duplicate values in the main spreadsheet were replaced with the average value and flag “6” was added. Where the difference was >0.005 a flag “3” was added. (After COMPARE one value was selected for sample #209, but both were entered in a comment; the flag was changed from 36 to 3. For the other case of duplicates differing by >0.005 the comparison was not helpful in deciding which was better.)

There were two sample #s repeated in files #22 and 23 – they were at the same site, but one was part of the TMR cast and the other a regular rosette cast. On the rosette sheet the TMR sample #s were changed to 9201 and 9202, but the salinity analysis does not distinguish between the two. However, the order of analysis and the presence of one duplicate makes it fairly clear which sample goes with which event. The two TMR bottles were collected near the surface while the other two were deep samples, so the salinity values confirm that the right bottles have been assigned to the proper casts.

The loop samples were copied to 2010-36-sal-loop.csv and the TMR samples to 2010-36-sal-TMR. The loop and TMR samples were then removed from 2010-36-sal.csv.

The duplicates differed by 0.0077, 0.0006, 0.0058 and 0.0010 for a standard deviation of 0.0035.

File 2010-36-sal.csv was converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2010-36oxy.xls and a duplicate study is in file 2010-36-oxygen-duplicates.xls. The spreadsheet was simplified, flags changed to numeric format and headings changed to standard form. . The spreadsheet was then saved as 2010-36-oxy.csv and that file was converted into individual *.ADD files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2010-36nuts.xls which included a report on precisions. The loop data in a separate page were copied to the TSG folder. There were TMR samples on the same page as the regular ROS cast data.

The file was simplified and saved as 2010-36-nuts.csv. Headings were changed to standard format, event #s were changed to a single # for the split casts and quality flags were changed to numeric values. A few comments were removed such as what rack was used and who the samples were for. Pad values were added to a few blank entries with “9” flags. The TMR data were copied to 2010-36-nuts-tmr.csv and those were then removed from the main spreadsheet. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2010-36chlarc.xls which included comments and flags. An event-number column was added and completed, and in cases of multiple files for a single cast, the number chosen was the first in the group. Header names were changed to standard format, flags were changed to numeric format, the file was sorted on sample number, and saved as 2010-36-chl.csv which was then converted to individual CHL files.

pH The pH bottle data were not yet available.

The SAL, CHL, ADD and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number. So one or the other set needs to be reordered in order to merge them. It was decided to reorder the MRGCLN1 files on Bottle_Number since that is the usual method used. The output files were named MRGx.

Those files were then merged with SAMAVG files choosing the bottle position # from the SAMAVG files.

##### Compare

Salinity  (Some errors were found in salinity bottle sample numbers in the bottle files after the initial run of COMPARE, so the routine was rerun with those corrections. It had only a minor effect on the fits, and none on the choice of outliers.)

Compare was run with pressure as reference channel. Three major outliers turned out to be due to more errors in the ADDSAMP file. All significant outliers were checked to see if similar problems had occurred, but no others were found. Most outliers are either very close to the surface and/or have relatively high standard deviations in the CTD salinity and only those that were severe outliers were investigated further. The following problems were found:

- Event #3 – Niskin #2. The salinity is only a minor outlier against both CTD salinity channels, but given the depth of 235db it does look off. When the plot against file pair number is examined it does seem even more out of line. There is a comment on the rosette log sheet that the bottom spigot was dripping. While the DO from that bottle is slightly higher than from the 3 other bottles at the same depth, the differences are insignificant. An examination of the full file shows that there was a lot of variability while the CTD was stopped at the bottom. That does not show up in the data around bottle firing time but the salinity sample has a value that can be found during the stop. No flag will be attached, but the sample will not be included in the comparison.

- Event #8 – Sample #122 is said to be from Niskin #21 at about 5db in the rosette log, but the value looks like it is from Niskin #2 at 175db. There is already a sample from that bottle. The dissolved oxygen sample shows no evidence of a misfire, so this sample must have been drawn from the wrong bottle accidentally or was intended as a duplicate. The value was replaced with a pad value and flagged “5” with the original value entered in the header comment.

- Event #17 – Sample #156 from Niskin #2 looks out of line in the plot against pressure and even more so in the plot against file pair number. The DO value is also out of line, but that was probably due to stopper being out and a large bubble in the sample. This sample looks enough off to keep it out of the comparison, for now, but it won’t be flagged. This sample and the one from Event #3 are the only two from the early part of the cruise below 200db, so it does affect the fit against time to remove them. It is unlikely that these two outliers are due to real calibration change early in the cruise since the two sensor channels look similar.

- Event #57 – Sample #445 was an outlier, difference >0.2 at 200db and was flagged “3”.

- Event #77 - Sample #652 was an outlier, difference >0.2 at 200db and was flagged “3”.

- Event #79 – The sample from Niskin #2 at 100db is off by 0.3 and does not resemble any other level. DO and nutrients seem ok, so likely not a misfire. Flagged “5” and replaced with pad value.

- Event #80 – sample #693 from 90db was off by 0.18 which does not resemble any other bottle level sampled. The DO and nutrients seem ok. This will be flagged “3”. The value may be ok since it is near the bottom and the standard deviation in the CTD salinity is a little high.

- Event #89 – Sample #788 was an outlier, but looks fine if called #789. If was renamed, but flagged “3” with a comment in the header.

- Event #135 – Sample #1184 was an outlier, difference >0.2 at 200db and was flagged “3”.

- Event #140 – Sample #1274 - the analyst flagged this sample because of concern that it might be mislabelled, but it looks fine. The flag and comments were removed.

- Event #150 – Sample #1332 was an outlier, difference >0.2 at 200db and was flagged “3”.

- Event #186 – Samples #1577 and 1582 were outliers, differences >0.2 at 200db and were flagged “3”.

An initial comparison was made to get a general idea of the magnitude of the differences. This was used to determine what data should be excluded based on differences from bottles – data were excluded that were outside the window (initial average–0.004 to initial average+0.004). Also excluded were all bottles from the top 100db. The primary salinity was found to be low by an average of 0.0023 and the secondary was high by an average of 0.0004; standard deviations were <0.002 for both. The fits against file pair number show that the primary sensors were within 0.001 of the bottles at the beginning of the cruise and about 0.004 lower by the end. The secondary drifted by <0.001 starting slightly high and ending slightly low. The differences at the beginning of the cruise were ~0.0014 and by the end of the cruise they were ~0.004. The differences at ~1000db at about the time of the 5 casts examined in section 8 are:

0.0020, 0.0018, 0.0029, 0.0034 & 0.0037

compared to those reported in section 8:

0.0015, 0.0017, 0.0017, 0.0023 & 0.0033

The differences from COMPARE are a little larger, but the change from beginning to end is very close. The section 8 profiles were mostly very noisy, so a close match is not expected.

A run of COMPARE was done using Niskin # as reference channel to determine if any bottle consistently produced poor results. There are 9 cases of differences >0.2, with only 1 outlier for 5 bottles and 2 each for bottles #2 and #20. Most outliers are shallow samples with fairly high standard deviations in the CTD data. One of the outliers (from Niskin #21) had already been flagged as a mis-sample. There were 2 deeper outliers for Niskin #2 but the standard deviations are a little high in the CTD data in both cases. Given that this bottle was used more than any other, seeing a few outliers is not a surprise.

The 2 duplicate pairs that differed by >0.005 were examined in light of COMPARE. For sample #209 one value is clearly better than the other so the average was replaced with that value, but both values were entered in the header comment; the flag was changed from “36” to “3”. For the other duplicate pair the comparison was not helpful, neither value looking convincing. The average fits best.

Dissolved Oxygen –

COMPARE was run for Dissolved Oxygen. The first run turned up an error in the SAM file for cast #30 – the last three lines had the wrong Bottle_Number. The next run produced very odd results. After eliminating some obvious outliers, the data fall into two groups, each with a reasonable looking trendline. Data were first selected by trial and error to fit the first trendline. Then another page was prepared in the same way to fit the second trendline. This took a few iterations, but when finished, the YES/NO columns from the two fits were copied to adjacent columns on one page and any point which did not fit either was identified as an outlier for both. About 20 of those are extreme outliers, but some lie between the two fits and would be considered only minor outliers in either of them. The biggest difference in the fits is in the offsets; it is possible that a creative choice of outliers would lead to the same slopes.

A test copy of 2010-36-dox-comp1.xls was created as 2010-36-dox-comp1-vs-filepair.xls and used to study the fits against file pair number. When plotted against file pair number it is clear that there are patches where one or the other fit is found, but there is no simple pattern. On one sheet of this file, the outliers from both files were identified and plotted against DO, pressure and file pair number, and they were all found to be from above 150db.

Because some of these files were created by merging other files, there is some risk that sample numbers were assigned incorrectly. The salinity comparison turned up some errors, but there are many cases where there are DO samples, but no salinity sample. There is also a possibility that the computer crashes affected the CTD in some way.

First, the outliers were investigated:

- Cast #3 – The top 4 bottles look odd. The temperature gradient is high. The problems could be due to DO never equilibrating, or a slight mismatch between sample and CTD data. The differences are not large and no flags are justified.

- Cast #6 – The CTD Oxygen sensor did not work properly during the latter part of the upcast (above 50db.). This was noted at sea. These outliers are extreme – the CTD DO values for the last 12 bottles in the SAMAVG file (samples 80-91) were replaced with pad values since they are clearly wrong.

- Cast #30 – Bottles #13 to 20 are extreme outliers. This cast was created by merging two files, so the first step was to ensure a good match. The problem cases are all from the second file. The pumps were off until near the top of the cast. The pumped channels should be replaced with pad values in the SAMAVG file. The downcast data is ok.

- Cast #76 – Sample #641 at ~40db is an outlier. There is sufficient variability in temperature at this level to explain the difference. No flag is justified.

- Cast #77 – The outlier at 16db appears to be due to noisy CTD data, high gradient. No flag justified. The bottle #s are not in pressure order because the CTD was raised to 20db and then returned to 35db before finishing the upcast.

- Cast #79 – The outlier is ~100db and is associated with very noisy CTD DO data. No flag justifed.

- Cast #80 – The 2 bottom samples are only slight outliers for Fit #2. The CTD was probably still equilibrating when the samples were taken. No flags are justified.

- Cast #81 – The outlier is ~35db. This is a case that is between the two fits and would only be a slight outlier in either. No flag is justified.

- Cast #94 – As for #81 this sample from 21db would only be a slight outlier in fit #1. No flag.

- Cast #111 – The 31db sample is out of line, but this was the sample taken at the time the computer crashed, so a small mismatch is not unexpected. There was also a lot of vertical motion during the stop before the crash. No flag is justified.

- Cast #121 – Samples from ~51 and 6db are both outliers and are both in high-gradient zones. No flags are justified.

- Cast #124 – Sample from ~21 is an outlier. There was very high variability during the stop and equilibrium was never achieved in temperature or DO. No flags.

- Cast #137 – The samples from ~30 and 6db are in high variability areas. No flags.

- Cast #138 – Very high gradients ~16db. No flag.

- Cast #145 - Outlier at ~6m in high-variability stop. No flag.

- Cast #167 – There is one sample that looks more like Fit #2 than Fit #1, but would only be a minor outlier in Fit #1. Since all the casts around that time fall into Fit #1, it is reasonable to treat the whole file as Fit #1. No flag is justified.

- Cast #169 – This bottle file involved a complex merge of 3 files with some repetitions in the BL files complicating the issue. There were also some gaps in the data though all the bottle firings were present. The CTD salinity and temperature data are clearly not in the order of bottle_number, yet the pressure and altimetry are. A plot shows that there was a large spike in temperature at the bottle #4 firing time which would explain the CTD values being off, but the samples that go with the bottle all look like they are from shallower water which does suggest a misfire. The only explanation is a combination of spiky CTD data and a bottle that was fired but did not close until it was somewhat higher. In either case it appears that the samples from Niskin #4 (sample #1486) should be flagged “5”. For Niskin #5 there is less evidence but a “3” is suggested by the CHL (Samples #1491). The CTD temperature, salinity, conductivity, pH, fluorescence and DO were replaced with pad values for those two bottles as well since they all seemed affected by spikes.

- Cast #188 – The samples from 151, 101 and 92db are outliers, but not major ones. There is a lot of variability in the CTD data during the stops and the downcast data show a DO reversal near 150db and fairly high variability in the 90-150db range. No flags will be assigned.

- Cast #204 – Outlier at 16db is associated with very noisy CTD DO. No flag.

- Cast #209 – A number of near-surface bottles are outliers, but the CTD DO are very noisy. No flags.

- Cast #220 – For 3 outliers the DO values are bad because the pump was off. Bottles look fine. The CTD pumped channels were replaced with pad values including CTD DO .

The next issue is to try to understand the two calibration fits. Is there some common element? For the casts that all fall into one or the other fit, is this also the case for the downcast data? And what should we do about those casts that have sections of each? The history of this sensor is not terribly helpful, but fit #1 is probably more in line with the expectation of drift. However, the range of DO sampled and offsets found vary from cruise to cruise, so we can’t presume a steady variation in slope of the fits. Moreover this was a long cruise, so there could be drift through the cruise.

The following casts have bottle comparisons that all fall within the Fit #1 group with the exception of some outliers that have been accounted for above: 3-30, 36, 57, 75-76, 85, 121-124, 137-178, 188, 204.

The following casts fall within Fit #2 with the exception of some outliers that have been accounted for above: 33, 41, 49, 66-68, 78-81, 89-119,127-135, 186, 197, 199.

Mixed casts include: 44-47, 60, 72 & 77. Except for #60 these mostly look like fit #2.

To try to understand these fits close examination was made of a pair of casts which were close in time, but had different fits, casts #60 and 66. Downcast values at a few levels were compared with bottles. Because the CTD was in motion (and DO has not been aligned) we expect that above the DO minimum the CTD should read a little high and below the minimum we expect it to be a little low. Keeping that in mind, Cast #60 has values that are lower than the bottles at all levels on the downcast and the differences resemble those expected from Fit #2. The upcast, taken from bottle stops, fell into the Fit 1 group except for the bottom bottle. There was a large shift (0.22mL/L) during the time at the bottom. The downcast data are much noisier than the upcast data. For cast #66, the reverse is seen with quiet data until the 100db level of the upcast. There is no shift at the bottom. The upcast data best suited Fit #2, but the downcast values are lower than the bottles by about the amount expected from Fit #1.

To see if the noise level in the CTD DO is associated with the different fits, a number of casts were examined closely. A noise level in DO of ±0.002 is typical of a quiet section. It is a little tricky to separate “real” spikes from noise given very unsteady descent rates, but there does seem to be a clear separation between two types of traces. Cast #65 looks fairly smooth for the downcast and becomes very noisy during the upcast – there are no bottles for that cast. Cast #107 has a similar noise level throughout, on the noisy side, and the differences are more like Fit #2 than #1. Cast #167 looks quiet both ways and more like Fit #1. Cast #186 is very interesting – the DO data are smooth until about 1220db down, and then they become noisier and stay that way to the end. The change occurs very dramatically with a large spike and shift in DO downwards by 0.23mL/L. That is very close to the change seen at the bottom of cast #60 (0.22mL/L).

The two separate fits found were:

Bottle DO = 1.0492 * CTD DO - 0.0237 (Fit #1)

Bottle DO = 1.0347 * CTD DO + 0.2383 (Fit #2)

The fit achieved with the shift of the Fit #2 bottles by 0.24mL/L was:

Bottle DO = 1.0410 * CTD DO - 0.0156

A shift of 0.22mL/L produced fit

Bottle DO = 1.0408 * CTD DO - 0.0134

However, the slope of Fit #1 alone is higher than expected. When just casts 1-72 are included the slope is 1.0418. Casts #188 to 225 have a slope of 1.0575. Those later casts generally had a smaller DO range and it is likely that there had been continuing drift through the cruise. While it is tempting to do a separate recalibration for the later casts, the noise level and complications in having 2 fits makes this of questionable value.

The combined result is in reasonable agreement with the 2010-13 fit:

Bottle DO = 1.0414 * CTD DO + 0.0057

It is interesting that the combination of fit #1 and shifted fit #2 is closer to previous results than the separate case of Fit #1. This may just be because more bottles reduce the effect of unusual profiles, or it may be of no significance. (See 2010-36-dox-comp1-Jan18.xls.).

It is noteworthy that the shifts and noise are also seen in the Oxygen Voltage channel and not in the temperature channel, so this is not a “derived” problem. No other channel shows the same behaviour. There were some observations of odd DO data from 2010-13, so results from that cruise were reviewed with to see if revisions were required. There was no indication of two fits, but there was less DO sampling for that cruise. The dissolved oxygen channel was removed from 2 casts, one of which has a bottle comparison. The slope looks unusual, but the offset is not as large as seen in 2010-36. The low DO values seen near the oxygen minimum during 2010-13 are not associated with noisy DO except where temperature is also noisy. The CTD DO channel was removed from 2 casts where the data were noisy and suspect; those might have had the same problem as this cruise. There is insufficient sampling to try to apply an offset to those.

So it does appear that an examination of the full profile will provide a means to determine which fit will be more satisfactory, but the noisy casts are probably going to be of lower quality. A systematic examination of casts will need to be done. Recalibration can be done simply where the whole downcast or whole upcast are in one group or the other. For mixed profiles, a more complex scheme will be required.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. When all bottles are included there is tremendous scatter, but when examined in low and high CHL range the picture is a little clearer. For CHL<2ug/L the fluorescence is close in value to the extracted CHL samples. For CHL>2ug/L the fluorescence is about ⅔ that of the CHL. (See 2010-36-chl-fluor-comp.xls.)

Outliers included:
Cast #21, Sample #182: The CTD data show a sharp peak in both dissolved oxygen and fluorescence in both the downcast and upcast. It is likely that the large difference between bottle and fluorometer is due to a slight mismatch in depth. A comment was added.

Cast #107, sample #971. This sample had already been flagged “3” since there was only no replicate. A comment was added.

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

Data were exported to spreadsheet 2010-36-bottles.xls and compared to the rosette sheets to ensure all expected data are present. Problems found included:

- One salinity sample (#58) had been mislabelled; when fixed it was a better match.

- Several salinity samples had been mislabelled (likely errors were made in the separation of sample #s and station names) – cast 12-sample 148, cast #23-sample 58, cast #44- sample 372, Cast #89- Sample #788, Cast #138, Sample #1233,

- Cast #33 was missing data for 150db. That had to be recreated from the end of the full file, just before the crash at 60db. Those data were added to the SAMAVG file and the sample numbers and bottle numbers entered as needed. A note was put in the header and “3” flags were added to all sample to warn users that while there were no problems with the samples, there is some uncertainty about the CTD data.

- Cast #60 had samples out of order. The problem is that the file was made up of 3 fragments. The simplest thing is to fix the SAMAVG file and run Merge again.

- Cast #124 – Sample #1093 should have been associated with Niskin #18, not #14. Niskin #14 was not fired.

- Cast #159 – Sample #1393 is given by the salinity analyst and the result compares well with the 20db CTD salinity in the merged bottle file, but the rosette sheet indicates the sample should be #1394 from 5db. If we assume 1394 is correct then the salinity looks off by about 0.2. It seems likely that the rosette sheet is wrong, so this will not be changed.

- Cast #163 – The nutrients were missing. This was a “fractured” cast. Since no data were acquired in #162, the bottle file was named #163, but the nutrients had been saved as #162, so they were renamed and the merge process rerun.

- Cast #189 – The surface bottles were assigned wrong. This concerns only the near-surface data.

- Casts #215, 222 and 225 were missing Dissolved Oxygen data because the individual ADD files had not been converted (because the number of lines had been set to only 1000 in the conversion routine!). Those data were found and the merge process repeated.

- Cast #220 – The near-surface salinity sample is labelled by the analyst as #1787, but is shown on the rosette sheet to be #1786 at 20db. The value is close to the bottle at 5db, but not at 20db, so the analysis label is assumed to be correct.

The MRG file data were repeatedly exported to a spreadsheet and checked until all the above problems had been resolved. (COMPARE was rerun for DO and Salinity after those changes. The effects were very small.)

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records, but a few casts without stops and quiet descent rate were examined to ensure that was appropriate and it was.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.7s was applied to the primary and no shift for the secondary.

SHIFT was run using those settings.

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. It was harder to judge this than usual, but a shift of +90 records seemed best and was found best for this sensor before its last calibration and for 2010-12 and 2010-13. There is some cause to expect variation among casts given the changes in DO signal described earlier. There are few casts with no bottle stops to use to test this idea, and the problem is not likely due to alignment variation. Moreover, there may be differences between upcast and downcast traces with one having values offset relative to the other because of the problems with noisy DO data. It seems best to go with the history of this sensor, backed up by a few tests that suggest it is ok.

SHIFT was run using +90 records for all casts.

pH

During 2010-01 tests suggested that shifting the pH data was not needed. For this cruise it looks as though a shift of +90 records does improve the data, by reducing the vertical offset between similar features in the up and downcasts so that it is similar to the offset between upcast and downcast temperature features. After the shift there is still a large horizontal offset, so hysteresis is a problem. Increasing the shift reduces the horizontal offset, but has a bad affect on the vertical offset. SHIFT was run on all casts using +90 records. The SHFO files should probably be saved in case future work suggests a better approach.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: Most warnings concerned casts with upcast data only, though 2 of those were not in the list (90and 221) because there was a short reversal of direction so there were data in the file. The only warning that needed checking was for cast #52 where pressure had shifted by more than 2 db between records. This was a pressure spike and DELETE worked well to correct that.

It was discovered in the editing step that cast #56 had an initial lowering before the main cast and that DELETE had not selected the best data, so a text editor was used to remove the initial lowering in the SHFopH file and DELETE was rerun. The results were much better.

At this point the file list was amended to remove upcast-only files.

##### Dissolved Oxygen Calibration Study

At this point casts were examined in detail to see if there is a reasonable method to determine which downcasts have data that suit Fit 1, Fit 2 or a mix of both. Examination of upcast data for which we know the answer from bottle comparisons showed that the Oxygen Voltage channel has a distinctive noise level (~0.002v), with data from Fit #2 upcasts having at least twice the noise level of those found in Fit #1 upcast data.

Every downcast was examined and files were identified as Fit #1, Fit #2 or Mixed. The mixed casts each had a distinct shift at some point in the profile, though the shift sometimes took more than 100db to complete. Some of the shallow casts were a little hard to judge because of the large gradients, but they generally fell into a group, so if it was quite unclear the choice was made that fit the casts before and after it.

To recalibrate the mixed casts we would really like to apply a correction based on DO values (like those found in section 11) that is bi-linear in P. There is no IOS SHELL routine available that will do that. We would have to split the casts, run the calibration and then put them back together. That would be time-consuming and error prone. We are able to apply a bi-linear correction based on pressure. This correction looks like

DO(corrected) = DO +offset +slope*Pressure

with a different offset and slope to be entered for pressures < and > a pressure level entered in the calibration file. We can add an offset to the Fit #2 section only by setting the slope = 0 for both segments, and choosing an offset of +0.24mL/L for the segment with Fit#2 characteristics and an offset of 0 for the Fit #1 section. Then a 2nd calibration will be applied to the full profile treating it like Fit #1. This is certainly a rough approximation, but the quality of these results is already limited and the tests run in COMPARE produced reasonable results by just applying an offset.

There were only 6 mixed downcasts. The following scheme was applied:

- CALIBRATE was applied to the DEL files using equation #12 (a correction factor which is Bi-linear in P) to add 0.24mL/L to Oxygen:Dissolved:SBE in the sections with noisy oxygen voltage.  (Output: CORSP)

- CTDEDIT was then used to delete points in the Oxygen:Dissolved:SBE channel from the  “shift zone” between the two fits. (Output: EDSP)

- A note was placed in the headers of these files to indicate that confidence in the DO values is lower than for the other casts.

That procedure was applied to casts #50, 52, 57, 63, 85 and 186.  When CTDEDIT is used for editing of T-S data, the EDSP files should be used rather than DEL.

For recalibration of the DO data, the Fit #1 scheme can be applied to casts: 1-32, 36, 50, 52, 57, 63, 74, 76-78, 85, 120-124, 137-180, 184-193, 203-225. (Note the list includes the 6 adjusted above.)

Fit #2 will be applied to: 33, 38-49, 51, 56, 60, 61, 64-72, 75, 79-83, 87-89, 94-119, 127-135, 181, 183, 197, 199.

The final step is to analyze the bottle files to determine which fits to use in recalibration. First, the bottle comparison was examined to pick out casts that fell mostly into one fit or the other. If the “odd balls” were in line with neither fit, then they were ignored. But if most bottles were in one fit and a few in another, they were investigated. There were 7 such casts. None of the bottles that stood out from the others had been flagged by the analyst. Some of the odd ones were close to the surface in low DO gradient. At those levels the differences are sometimes smaller than you might expect because the sensor is better equilibrated than in a higher gradient. An examination of the noise level in the DO voltage was made for casts where the picture was not clear.

- Cast #44 bottles mostly fit #2 except near 70db and bottles above from 30db up. The 50db bottle was identified as Fit #2, but is really between the two. There is so much noise at 70db that it is likely just a CTD outlier. But there does look like a shift at about 30db. The upcast surface data have higher values than the downcast which is unusual. So it may be that while #44 should be recalibrated using Fit #2, the bottles above 30db should have 0.24mL/L subtracted from the Oxygen:Dissolved:SBE channel.

- Cast #60 is mostly fit #1 but the top and bottom look more like fit #2. The downcast did look like Fit #2, but there is a shift in DO at the bottom that would appear to mark a change to Fit #1. While DO is a little noisy after the shift, this is a different sort of noise likely related to depth variations. There is a shift around 48db so it is likely that the bottles above that are fit #2. The noise level seems a little higher as well. Use Fit #1 for the cast, but add 0.24mL/L later for bottles above 50db.

- Cast #72 – 90db could be fit #1, rest fit #2. There is no change of noise level around 90db so this is probably just an outlier. Use Fit #1 for all.

- Cast #77 – mostly fit 2, but 20 and 40db look odd. However, this was an unusual cast because the CTD was lowered from 20 to 35db during the upcast to fire a bottle and the standard deviation in the CTD data suggests poor equilibration. The DO noise level looks like fit #2. Use Fit #2 for all.

- Cast #79 – mostly fit #2 but 90db is closer to fit #1. But the noise level looks like Fit #2. Use Fit #2 for all.

- Cast #80 – mostly fit #2 but 50db looks more like Fit #1. The downcast shows this is an unusual profile with some reversals so we can’t expect a simple fit. The noise level is seems a little high, so use Fit #2 for all.

- Cast #167 – mostly fit #1, but bottle at 76db might be fit #2. There is a lot of structure in the temperature and DO, so this is probably just an outlier. This cast occurred during a long period of Fit #1 data, so it seems best to put this in the fit #1 group.

Before proceeding, the MRGCLN2 files were edited for cast #44 to subtract 0.24mL/L to the bottles from 30.3db to the surface and for cast #60 to add 0.24mL/L to the bottles from 25.9db to the surface.

Based on this study the following groups should be used:

The following hydro casts belong to the Fit #1 group: 3-30, 36, 57-60, 72-76, 85, 121-124, 137-178, 188, 199-225

The following casts belong to the Fit #2 group: 33, 41-49, 66-68, 77-81, 89-119, 127-135, 186, 194-197.

##### Argo Comparison

There were 3 Argo floats launched during this cruise. Unfortunately, none of them had oxygen sensors.

The following profiles were found for comparison and secondary sensors were selected for the 911+:
Cast #52 – Argo profile 20100805_4901112 is from 18.3km away and 8 days separation. The T-S curves are similar, and at depth the float is cooler by about ~0.01C° and fresher by ~0.002.

Cast #61 – Argo profile 20100722_4901064 is from the same day and 3.2km away. There are significant differences above 175db, in an area that looks like there is active mixing. At depth the Argo profile is cooler by ~0.01 C° by and fresher by ~0.002.

Cast #188 – Argo profile 20100813_4901118 is from the same day and 2.1km away. It looks quite similar to the 911+ data. At depth the differences are on the order of 0.01C° and 0.001.

While these differences are small, it is interesting that the 911+ is always the warmer and saltier of the CTDs. Our comparison with bottles suggests that the CTD salinity is high by about 0.0006 at the beginning of the cruise and even closer to bottles at the end, but the fit has a standard deviation of 0.002, so it is certainly possible that it is reading slightly high, say by 0.002. The evidence is not considered strong enough to justify recalibration, but does confirm that the secondary sensors are a better choice for archiving than the primary.

##### DETAILED EDITING

The secondary T and S sensors were chosen for archiving, though it is likely that those channels may not be usable for some casts. Both channels showed very little pressure dependence in the bottle comparison, but the primary salinity calibration shows some evidence of drift through the cruise, while the secondary did not. There appears to be a little more noise in the primary salinity than in the secondary.

All casts required some editing.

A few problems were encountered:
	Cast #140 had some sections of bad secondary temperature and salinity. The primary channels look better in those 2 sections, but in other parts of the profile the secondary looks more reliable. The secondary were chosen, but records were removed in some places and just salinity data were removed from others.

Cast #154 has a similar problem with bad secondary salinity between 7 and 21db. The primary salinity is better at those depths, but is worse below with many spiky sections, so the secondary will be used. The secondary salinity does show the water was well mixed to 12db, so secondary salinity was edited to reflect that, but salinity data were removed from 12 to 21db.

For cast #199 the two salinity channels are far apart and the secondary looks unstable; one bottle at 250db confirms the primary is better. The temperature channels do not look bad. Primary channels were chosen for editing.

All EDU files were copied to EDT.

T-S Plots were made of the edited data. Cast #101 was rerun in CTDEDIT and a little more editing was applied to cast #105. Those files were copied to EDT.

##### Initial Recalibration

The SBE Dissolved Oxygen data were recalibrated in two groups depending on which fit worked best in COMPARE.

Fit 1:    Corrected CTD DO = 1.0491 * CTD DO – 0.0237

Fit 2:    Corrected CTD DO = 1.0347 * CTD DO + 0.2383

The SAM files were recalibrated using file 2010-36-recal-fit1.ccf and 2010-36-recal-fit2.ccf to the two groups of casts.

The CTD secondary salinity was very close to the bottle salinity so no recalibration will be applied to that channel. However, the primary channels will be needed for cast #199 and by that time the primary salinity was reading low by about 0.004, so a correction of +0.004 will be applied as a 2nd recalibration.

COMPARE was rerun to see how well the DO recalibration worked. We do not expect great results from casts #44 and 60 because the correction applied to allow for the mixed calibration, were not applied to the SAM files. The first attempt made it clear that 2 casts had been put in the wrong group. The second try showed satisfactory results. (See 2010-36-dox-comp2.xls.)

The MRGCLN2 files were then recalibrated using the same calibration control file.

The same calibrations were applied to the EDT files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction can be applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. There is a lot of scatter in the plot of differences against DO concentration, with a tendency towards values that are high by about 0.07mL/L at the low end of the range and by 0.04mL/L at the high end. The plot against pressure shows values too high above 500db and close to bottles below that. The DO minimum at about 1000db greatly complicates interpretation. When differences versus time are examined the SBE DO is high by about 0.07 at the beginning and by 0.03 at the end, but if one cast is excluded, the range is from 0.06 to 0.04mL/L. There is a lot of noise in the comparison and the way the fits were done does not lead us to expect perfect results. Subtracting 0.06mL/L would lead to better values near the surface, but poorer results at the oxygen minimum and below that. What we have is probably as good as we can expect from these data, so no further recalibration will be applied. (See 2010-36-dox-comp3.xls.)

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The primary sensors have not been used since they were last recalibrated at the factory. The secondary sensors produced salinity higher than the CTD by 0.0033 and 0.002 during 2010-12 and 2010-13.

2. Dissolved Oxygen

This sensor has been used for 3 other cruises since its last factory recalibration, but only 2 of those has been processed. The corrections found were:

Bottle DO = 1.0339 * CTD DO - 0.0043 (2010-12)

Bottle DO = 1.0414 * CTD DO + 0.0057 (2010-13)

3. Pressure

Since its latest factory recalibration in May 2006 this sensor has been used mainly on the Ricker. The factory offset was used for the Ricker cruises (~-0.8db). It is likely since it is an older sensor that there is some drift, but it is harder to judge from Ricker data since the CTD usually is started fairly deep. During 2010-12 pressure looked low by at least 0.7db and for that cruise and 2010-13 a net offset of 0db was used (an addition of 0.8db to the factory offset.)

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The temperatures were all within the historic ranges except for a little near-surface data in shallow waters. Salinity had more excursions from the climatology but these were also mostly close to shore in shallow waters. However, there was a persistent feature with high salinity around 40db at stations LG3 to LG5, LG7 and LJ3 to LJ4. Sometimes it is possible to get a salinity overshoot at the base of a strong gradient, probably an artefact from the CELLTM correction, but the location of these excursions does not fit that pattern. It looks like these are reflections of real conditions.

Repeat Casts – There were some repeat casts but there were no cases of 2 casts below 300db, so these are not useful to test repeatability.

TMR CTD

Data were made available by researchers from the University of Victoria who used a different CTD as part of a trace metal study. That CTD did not have a DO sensor. A few quick comparisons were done by selecting data from below 1000db from one cast and then finding the level at which the same sigma-theta value was found in the other; then temperature and salinity were compared. The casts were separated by an hour in one case and almost 3 hours in another, and the comparisons were far from systematic, so the significance of the results is limited. However, the T-S data show reasonable agreement with temperature differences ranging from 0.0004C° to 0.0017 C° and salinity differences ranging from 0.0004 to 0.0007.

Profiles

Profiles of DO turned up 3 problems.

- Cast #15 had spikes around 850db. This is in the original data, not created in processing. It is not a simple spike, being spread over 2.5db. So replacing the bad values with pad values in the file just before bin-averaging is a reasonable approach. The FIL file was edited and bin-averaged.

- Cast #51 has surface spikes the FIL file was edited to replace the bad data with pad values and the file was once again bin-averaged.

- Cast #52 looks bad – the “mixed” recalibration was obviously done wrong. The steps described in section 15 were repeated and it looked fine then.

There was no fluorescence signal for cast #33. This channel should be removed.

The transmissivity looks odd near the bottom of cast #3 but the upcast agrees with the downcast.

PAR looked fine.

T-S Plots

After bin-averaging a few T-S plots show small instabilities, but these are in shallow casts near the surface close to Brooks Peninsula where such results are often seen.

##### Final CTD File Steps (REMOVE and HEADEDIT)

This step is more complex than usual because there are casts with either PAR or pH or both or neither. To ensure that the right channels are removed from each cast, 4 cast lists were prepared. Plots were then made of PAR and pH using the 4 lists. A few errors were found in the lists and corrected. In a few cases there appeared to be PAR or pH signal where they were not expected, but on closer examination these are not realistic values; and log notes (and notes from Marie Robert for the early part of the cruise) are clear on what sensors were mounted. There were a few cases where it looks like there was no PAR when it was supposedly mounted, but as it was night time this is possible, though we usually see some slight variation. The channel will not be removed in those cases.

REMOVE was run 4 times using the 4 lists to removing the following channels in every case:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

And channels pH:SBE and/or PAR depending on which list was being used. For cast #33 the Fluorescence:SBE channel was also removed.

REMOVE was run a fifth time on cast #199 only to remove channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

Those files were named REMDI. They will be provided to the chief scientist.

REMOVE was then run again to remove pH:SBE from all casts. Those files were named REM and will be prepared for the archive.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Transmissivity, fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

The pH:SBE channel was removed from the files because sensor performance and

recalibration are being studied and are not yet considered ready for the archive.

The precision of the SBE dissolved oxygen channel is difficult to estimate

because the comparison with bottles was complex and noisy, but roughly,

the DO should be considered:

•	±0.4ml/l from 0-300 db

•	±0.25ml/l from 300-500 db

•	±0.1ml/l below 500db

For details on the processing see processing report: 2010-36-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

A separate set of files that include the pH:SBE channel were prepared for the use of the Chief Scientists and saved as *.CTDDI.

Profile plots were made and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 60% to 150%. The lowest values were in the area of Brooks Peninsula and the mouth of Juan de Fuca Strait. The highest were in Saanich Inlet and near the coast in the LG line and at LC5. High values were also seen in Hecate Strait and LC and LG lines. Most casts were between 90 and 120%.

A few extreme values were checked to see if bottles confirmed the surface CTD values. Cast #121 has a surface bottle value of 9.19mL/L and CTD value of 9.05mL/L. For cast #101 those values were 4.70mL/L for the bottle and 4.84mL/L for the CTD. This confirms that the extremes are reasonable.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run in the same 4 groups as the CTD files as described in section 22.

Casts #33 and 199 were run through separately to remove Fluorescence:SBE from the first and secondary T and S from the latter.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods.

Separate files were prepared for the use of the Chief Scientist with pH:SBE channel included. The relevant files are named as *.CHEDI,

The CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets to check that all data are present and in the correct place. A few discrepancies were noted:

- The chlorophyll from cast #7 had been mislabelled as being cast #8. That was fixed and CHE files for casts 7 and 8 were recompiled.

- Cast #33 was missing the CTD data from 22db. The data were fabricated from the end of file 2010-36-0034.IOS and added to the SAMAVG file. The CHE file was recompiled and “3” flags were added to all samples from the levels with fabricated CTD data since the pressure is not certain.

- Cast #41 – The deep salinity sample was identified as #336 on the rosette and sample, but the label also said Niskin #3 which corresponds to is #335 and the CTD data from 335 makes more sense. No note had been entered in the header about this, so one was added.

- Cast #154 – The surface CHL sample is named as #1364 in rosette log and 1360 in the CHL file, but they are both from bottles fired at 6.2db so the discrepancy is of no significance.

- Cast #159 – The two salinity samples had different sample #s than indicated on the rosette sheet, but the values fit the sample labels much better.

- Cast #167 –Salinity sample labelled #1445 but rosette sheet #1446 – the sample label give much better results in comparison with CTD salinity.

- Cast #168 – Like #167 deep salinity sample labelled #1472 said to be #1471 on rosette sheet; label appears correct as salinity matches CTD salinity much better.

- Cast #220 – Deep salinity sample labelled #1787 said to be #1786 on rosette sheet; label appears correct as salinity matches CTD salinity much better.

- Cast #225 – Similar to the last case, but the two possibilities are from the same level so it is not completely obvious which is right, but #1806 (as on the label) will be used instead of #1807 from the rosette log.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no outliers were identified.

Standards check was run on all files and no errors were found.

##### 25. Thermosalinograph Data

Data were provided in 5 hex files. There were errors in the first 3 file names, so those were corrected (ex. 2010-36.0001 changed to 2010-36-0001.) There was loop sampling for Salinity, Nutrients and Extracted CHL at 43 locations. Some of the loop samples were not recorded in the log and some that were entered were missing times and/or positions. The dissolved oxygen analysis spreadsheet did include station names, times and positions for all but one sample. For samples taken during CTD casts, the times from the DO file were sometimes taken from the beginning of the cast but the analyst says the samples were taken at the end for comparison to the rosette samples. So the times were changed to agree with the end time of the CTD casts.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong date and scale factor. After that correction the CON file was saved as 2010-36-tsg.con.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were produced. A few problems were noted:

- Flow interruption – The flow rate is steady at ~1.1 but there are two gaps in file #4 on August 10 from 0147 to 0441 and 0738 to 1043. There is a note in the log about the first gap indicating that the flow stopped after the flow meter was cleaned. There is no note about the 2nd gap. The data from those sections should obviously be removed.

- The salinity looks a little odd before the first gap, from about 2330 on Aug. 9 until about 0500 on Aug. 12, well after the flow was back on. The early “odd” salinity data could be real, but the later part looks unbelievable. A few quick checks against CTD salinity data suggests that the sensor was performing well up to July 28th and then starts to have much larger differences than expected and the data looks very noisy on August 12th, but somewhat noisy even before that. This will need a closer examination at the editing and/or recalibration stage.

- Flow stops at 1900 on Aug. 15 so the data will need to be removed from the end of file #5.

- Bad fluorescence – The fluorescence data looks bad for the end of file #4 and all of file #5. Checks against loop and CTD data may help establish which data need to be removed.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2010-36-ctd-tsg-comp.xls. All the data came from ~4.5db. For some CTD casts there was no data at that level.

The 5 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence. TSG values were recorded for times of underway loop sampling for the TSG-LOOP comparison. Then the file was then reduced to the times when CTDs were run to provide data for the CTD-TSG comparison.

Those data were added to the CTD data in file 2010-36-ctd-tsg-comp.xls. There were 111 matches. In some recent uses of this equipment latitude and longitude have become stuck, or times have been wrong due to missing scans. In this file the times from the CTD and TSG are matched, so comparing positions will turn up problems of that sort. The differences in latitude were all <0.0003° and in latitude <0.0009°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

Files of nutrients, chlorophyll, dissolved oxygen and salinity loop samples were found and combined in spreadsheet 2010-36-loop.xls. Those data were simplified and columns created for TSG data which will be added later based on matching times. That file was saved as 2010-36-loop-TSG-comp.xls.

File 2010-36-ROS-LOOP-comp.xls was prepared combining with all loop samples and corresponding data from rosette bottles at 6db.

- T1 vs T2 The intake thermistor was connected throughout the cruise. For the first file the temperature differences were extremely noisy early in the file and the flow rate was also noisy. Towards the end of the file the average differences was 0.27Cº (standard deviation 0.08) and the temperatures ranged from 9.8 to 10.9Cº (median 10.4ºC). For file #4 (the largest file), the average difference was 0.19Cº excluding the sections with zero flow rate; the standard deviation was 0.23. The averages were 0.18 for 2 quiet sections with standard deviations of 0.03 and 0.07. The intake temperature ranged from 8.8 to 15.6ºC (median 11.8 º C). The temperature differences for file #4 are a little lower than for 2010-13 in June, but the intake temperatures were higher than in June and the flow rate slightly higher, so we expect less heating in the loop. The data from late in file #1 are more like the June results. Files 2, 3 and 5 are similar to file #4 with differences ~0.17Cº in quieter sections.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. When all data were included the TSG intake temperature was higher than the CTD by 0.04Cº while the median was 0.0075Cº. When outliers were excluded based on a standard deviation in the TSG being >0.008 plus 2 outliers with differences >0.2Cº, the average difference was 0.007Cº and the median 0.006Cº. This is an excellent match. Nonetheless, the outliers need explaining. They are higher than usually observed, although there were a few large outliers in 2010-13 in June. Some are associated with high standard deviation in the TSG data and most of the others are from near-shore casts with local temperature gradients that are fairly high. For one outlier the CTD data shows a small inversion which could either be the result of bad CTD data, or a real inversion that the TSG did not sample. There does not seem to be any systematic problem with temperature.

The TSG salinity is lower than the CTD by an average of 1.05 and a median of 0.157 using all the data, but there are some very large differences. Almost all the large outliers involved TSG salinity being much lower than the CTD salinity; the outliers are mostly in areas of relatively low standard deviations in salinity calculated over 5 records (2 minutes). A few CTD casts were examined to see if there was a strong salinity gradient at the intake level when the large outliers occurred and there was not. The outliers have salinity values that are much too low, sometimes <20 and these were not casts in inlets. It is noted that in some cases the fluorescence also looks odd when this happens, but the temperature differences are not notable outliers. These cases did not occur during the two short breaks in flow. In the past low salinity was thought to be caused by the introduction of fresh water, but if that is happening, it seems odd that fluorescence is also high. The only major outlier before cast #150 was for cast #3 which had high near-surface gradients. After cast #181 the differences start to get smaller again, but never as small as before #150. The differences start at about 0.06 which is similar to the results of 2010-13 in June. By cast #78 the differences are ~0.08, then start to increase a little faster and then become more variable and larger. Up until cast #50 it is reasonable to think this was caused by calibration drift, but it gradually becomes less believable. The sharp drop from #150 to #178 is followed by a sudden rise and from cast #192 to the end it is possible to believe the data, though the differences are ~0.4 to ~0.5 which is much higher than seen early in the cruise.

Plots were made of TSG data collected during 4 periods when the ship was stopped to see how much variation occurred.

- The stop at W12 was long and included CTD casts #56 and 57. The TSG and CTD were in good agreement for both (allowing for the usual salinity offset). There was some variability during the stop but not the sort of jumps seen later.

- The stop at Stn BP9 included CTD cast #119. There was a step up in T and down in S shortly before the CTD cast – could this be related to ship positioning? The ship was not moving much during that time but perhaps thrusters were on. The CTD temperature agrees well with the TSG value and the two salinity values are within 0.01.

- The next stop studied was cast #150 at stn BC1. The variability in T and S are similar in scale. The gradients were quite high at this site and there was some ship drift, so we expect some variation. The TSG temperature is higher than the CTD by ~0.2Cº and the salinity is low by ~0.4 units. This could just be water from a shallower level than measured by the CTD. There is no evidence of a big problem.

- Finally the stop at A13 was examined. Cast #188 was early in the stop. The temperature is close to that of the CTD and the salinity is within 0.1 units. After the CTD started there are a series of jumps up and down in the TSG salinity. They look suspicious. There is one spike in temperature too. The surface water was well mixed and the upcast is very close to the downcast, so the source of this variability is not obvious.

The ratio of TSG fluorescence to CTD fluorescence ranges from 1.6 to 311, with the ratios >20 all occurring after cast #156. If only casts #4 to #50 are included, the ratios vary from 2.4 to 10.1 with an average of 4; this is similar to results in the past with this equipment. The largest ratios are associated with the lowest standard deviations in TSG fluorescence. This could suggest a poor flow through the fluorometer. It is notable that the high fluorescence occurs at roughly the same time that the TSG salinity appears to be bad. When a plot of TSG and CTD fluorescence against cast # is arranged so that the two traces are close for the first 50 casts, it shows that the fluorescence values are reasonable until cast #148. The extreme TSG fluorescence values are associated with CTD fluorescence <5ug/L but that is probably more a factor of time than related to the fluorescence itself. (See 2010-36-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons  A comparison was done in spreadsheet 2010-36-loop-TSG-comp.xls of loop salinity and chlorophyll samples with TSG salinity and fluorescence from the same time. There are 39 points of comparison for salinity and 42 for chlorophyll. All loop samples came from before the time the fluorescence and salinity became very bad. The ratio of the TSG fluorescence to the loop extracted CHL ranges from 0.9 to 5.2 with an average of 2.45. The difference between the loop salinity and TSG salinity is highly variable with a range of +0.1 to -1.9. The average is -0.21 with a median value of -0.13. The differences are growing with time. (See file 2010-36-loop-TSG-comp.xls.).

- Loop Bottle - Rosette Comparisons

File 2010-36-ROS-loop-comp.xls contains only 2 rosette casts with associated loop sampling for salinity and chlorophyll and 3 for nutrients, but there are many more with dissolved oxygen sampling. The loop salinity was higher than the CTD salinity from the rosette files by 0.009 and 0.008 for the 2 casts, but it was very close to the rosette bottle samples. During cast #6 the salinity was very noisy and at cast #80 it gradually rose in value through the stop without ever settling down. So it would be unwise to place much confidence in this comparison. The nutrients and chlorophyll agree reasonably well.

For dissolved oxygen the loop samples were lower than the rosette bottle samples by an average of 0.021mL/L and lower than the Rosette CTD DO by 0.022mL/L. There was 1 outlier in the comparison with the SBE DO value, but it is close to the bottle DO value. One sample was taken at station A10 because the TSG fluorescence looked bad and it was thought it might be due to growth in the loop. For that case the loop was lower than the rosette bottle by 0.03mL/L. So the loop sample had slightly lower DO than the average, but does not seem particularly out of line. The loop sample was lower than the SBE DO sensor data from the rosette file by 0.06mL/L but the sensor calibration is not considered reliable enough to judge a difference that small. (See 2010-36-ROS-Loop-COMP.xls)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and were used for 2009-10 (with a different intake thermistor), 2009-11, 2010-01, 2010-12 and 2010-13. The salinity was low by about 0.02, 0.02, 0.06 / 0.16, 0.07 and 0.06 for those cruises with the change for 2010-01 being associated with a change of flow rate; there are problems with the 2010-12 data due to clock problems. The TSG intake temperature was within 0.004 of the CTD for the two 2009 cruises and 0.001 during 2010-01, but differed by 0.04Cº for 2010-12 and 0.02Cº for 2010-13. The TSG fluorometer was high by a factor of 2.2 to 6 for 2009-10, by about 8 for 2009-11, ~2 for 2010-01, 2.5 for 2010-12 and ~3 for 2010-13 with a lot of variability.

Conclusions

1. The TSG clock appears to have worked well; matching times in the TSG file with those of CTD casts, the positions were mostly very close and none were farther apart than 70m.

2. The flow rate was steady except for two brief interruptions. Comparing features in primary and secondary temperature shows that water spends about 4 minutes in the loop.

3. The loop and rosette samples compared very well. The dissolved oxygen was lower by an average of 0.21mL/L or 0.4%.

4. The temperature in the loop warms by from 0.17Cº to 0.27Cº, with the highest values associated with the cooler waters early in the cruise.

5. Salinity is low by ~0.06 at the beginning of the cruise, rises to about 0.08 by cast #50 and 0.14 by cast #76. The rate of drift increases at about cast #80 and starting with cast #138 the results become much noisier and differences become larger and often look extremely bad with salinity sometimes low by 15 units. The loop comparisons show a similar trend, but there are no loop salinity samples after the data began to look very bad.

6. There is always a lot of variability in the ratio of TSG fluorescence to either CTD fluorescence or loop CHL, but this cruise was exceptional in having fluorescence start rising rapidly until it went off-scale. The source seems likely to be related to the “euphausiid” invasion. Before that event the TSG fluorometer had values higher than the CTD by a factor of about 4 and higher than the loop CHL by a factor of about 2.5. Given that the CTD fluorometer was found to be about 2/3 the extracted chlorophyll samples from the rosette for CHL>2ug/L, and most of the loop CHL samples were in that range, then these two comparisons are in reasonable agreement.

7. The intake temperature is within 0.01 Cº of the CTD temperature from 4.5m.This is an excellent agreement.

8. It has been noted on other cruises that underway samples show much more variability in comparison of TSG to loops, than those taken on station. There were few on-station loop samples for this cruise, but there is certainly a lot of noise in the underway loops. Unfortunately, the odd behaviour of the salinity and fluorescence during this cruise means we cannot distinguish the noise due to equipment problems from the noise due to ship effects. Whether the underway samples are affected by more bubbles or whether the surface waters are pulled into the loop when the ship is moving are issues we would like to understand better. The latter would introduce greater variability so the matching of the loop sample and the reading from the TSG would be more sensitive.

9. A quick look at Line P results shows that salinity was too low during the early part of that trip and then it crashed to extremely low values just as during this cruise. Scott Rose worked on the system and the de-bubbler since he felt bubbles were the problem. However, SeaBird advise that an extremely dirty conductivity cell will give lower salinity than usual. They also say that debris can cause extremely noisy salinity. So while a bubble problem may account for what we are seeing, a dirty TSG is also possible and the things Scott did to clear up bubbles may have cleaned the system as well. There was no sign of fluorescence problems in the Line P TSG data.

10. A TSG log book would be helpful, or entries in the CTD log book, mentioning any problems noted and steps taken to resolve them.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. Because we can only see 3 channels at a time, the files were opened several times to check for flow rate, to see if the two temperature channels tracked each other reasonably and finally to look at details in the salinity and fluorescence data. The secondary temperature shows more variation than the primary which is assumed to be due to mixing in the loop or effects of the de-bubbler which can affect temperature.

File #1 was very short. The flow was fine and just a few points were removed from the salinity.

File #2 had no problems in flow or temperature; salinity was cleaned lightly where spikes appeared to have been produced by slight misalignment of conductivity and temperature.

File #3 has no problems in flow or temperature. Salinity errors began to grow during the period of this file, becoming worse towards the end, but no evidence is seen of this through the editor. There are no obvious sharp drops. Salinity was cleaned lightly.

File #4 has zero flow periods, so a first edit was done to remove those.

Then a run was done to remove salinity where conductivity and temperature diverged in a suspicious fashion. Mostly this was at the end of the file. Then a run was done to edit the bad fluorescence data removing many records from the end of the file. There was a short section of what seemed ok in the middle of what was removed.

File #5 had zero flow at the end and the fluorescence is bad throughout. The first edit was used to remove all data from the zero flow area and fluorescence from the whole file. A second edit was run to remove salinity points where salinity and conductivity looked bad.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

The recalibration scheme was developed based on the TSG-CTD comparison. A value was chosen that looked reasonable for the start and end of the file and fitted to the length of the file in days. This is a crude way to arrive at a correction but seemed to work better than any fits based on the data itself due to the arbitrariness of picking outliers amid all the variability.

File 2010-36-tsg-recal1.ccf was prepared to adjust salinity as follows:

For file #1 - 0.06 was added

For file #2 the following equation was applied:

Salinity (corrected) = Salinity (uncorrected) +0.001992* Time:Julian – 0.34627

For file 3, there is a split with the first part looking like file #2 and the second like file #3, so a bi-linear fit was applied with Julian Day 209 as the division point..

For files # 4 and 5 the following equation was applied:

Salinity (corrected) = Salinity (uncorrected) +0.018804* Time:Julian – 3.86162

A few values were checked before and after this step to ensure it was applied correctly and after an error was fixed, the results looked right.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

For files #4 and 5 the salinity format was set to F7.2 instead of the usual F9.4 as a warning that quality is lower than usual. Comments were added to files #3, 4 and 5 to mention problems in salinity calibration.

File #3: The salinity is considered +/- 0.02 until July 29th, +/- 0.06 from July 30 to the end of the file.

File #4: Many obviously bad salinity data have been removed, but since some may remain, the salinity is considered +/- 0.5.

File #5: Some obviously bad salinity data have been removed, but since some may remain, the salinity is considered +/- 0.5.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars –

2. There was a rosette file, but no sampling. The firing was done at the surface just to test the bottles. The file was not processed.

33. Acquisition started late – CTD was at 100db of downcast. There are files from 3 partial casts to complete this station, but there is a gap, so one file may have been overwritten.

43. Mentioned in log book as a ROS cast at W4 but there are no raw files for that event.

52. Argo Float launched at same site.

61. Argo Float launched at same site.

66. Niskin #1 did not fire; no samples for sample #529.

77. CTD lowered from 20 to 35db to get another sample below CHL max.

119. Altimeter problem.

145. pH cap left on.

151. Mentioned in log book as a ROS cast at C7 but there are no raw files for that event.

152. Mentioned in log book as a ROS cast at C8 but there are no raw files for that event.

159. Altimeter problem.

167. Altimeter problem.

193. Altimeter problem.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 06May08 | Factory “ |  |  |
| Conductivity | 1764 | 10Feb09 | “ |  |  |
| Secondary Temp. | 2038 | 06May08 | “ |  |  |
| Secondary Cond. | 3394 | 06Mar09 | “ |  |  |
| Transmissometer | 1005DR | 20Jul10 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 10Nov2009 | Factory |  |  |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Fluorometer | 2345 |  | IOS |  |  |
| Pressure Sensor | 69698 | 26May2006 | Factory |  |  |
| pH | 180692 |  |  |  |  |
| Altimeter | 1024 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2010-36

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 2416 | 23Dec06 |  |  |  |