## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |

# PROCESSING NOTES

Cruise: 2010-41

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: West Coast Vancouver Island

Project: WCVI Sardine Survey

Chief Scientist: McFarlane S.

Platform: W.E. Ricker

Date: July 26, 2010 – August 6, 2010

Processed by: Germaine Gatien			Date of Processing: May 13, 2011 – May 18, 2011

Number of original CTD casts: 	26		Number of CTD casts processed: 27

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with an SBE43 Dissolved Oxygen sensor (#997). One transmissometer (#983) was mounted on the CTD until cast #419 when 2 transmissometers (#983 and #953) were used. The salinometer used was a model 8400B Autosal, serial # 69086.

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard.

There was no log book, but there was a list of casts with positions, times, maximum depth samples and bottom depths, as well as the number of salinity bottles.

Salinity was recalibrated by adding 0.001 based on a post-cruise calibration drift report for the conductivity sensor plus the comparison of salinity samples with the CTD salinity.

Pressure was recalibrated based on near-surface CTD data plus a post-cruise sensor calibration drift report.

There was no dissolved oxygen sampling, so data were recalibrated based on the results of cruise 2009-09. There was a post-cruise calibration that showed less drift than the 2009-09 results would suggest, but the factory report indicates that the membrane was torn, so these data must be considered less reliable than usual.

Cast #22 did not have the pumps turned on until the CTD was at

# PROCESSING SUMMARY

##### 1. Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names were non-standard. They were changed using program CKRename.

##### 2. Preliminary Steps

There was no log book, but there was a spreadsheet with event numbers, dates, positions, times, bottom depth, max sampling depth, # of salinity bottles and approximate surface temperature.

There were salinity samples from every cast, but no note of the depth, so it will be assumed they are from about 5m above the maximum pressure given the values are quite high.

Bottle data were provided in spreadsheet 2010-41-xls which was renamed 2010-41-sal.xls. This sheet was simplified by removing unnecessary columns and the event numbers and station names were entered based on notes in the Sample Label column. There were no salinity duplicates. There were Autosal analysis log sheets. No flags had been entered – 1 sample required 3 readings, but looked stable.

The cruise summary sheet was completed.

The histories of the conductivity, dissolved oxygen and pressure sensors were obtained.

Only one configuration file was used and it was saved as 2010-41-ctd.con.

Calibration constants were checked. The pressure sensor calibrations were wrong, so those were corrected. The DO sensor parameters entered were for the Owens-Millard algorithm, so they were replaced by the Sea-Bird algorithm parameters. Nominal values were used for E, H1 and H3.

##### 3. Conversion of Raw Data

All data were converted using files 2010-41-ctd.con.

Plots were made for a few casts.

The temperature traces are reasonably close on the downcast, but the upcast secondary temperature is very noisy. There is fine-scale noise in the secondary conductivity. The two channels are close during downcasts, further apart for upcasts. The dissolved oxygen looks as usual. The descent rates were generally high; some were very noisy, others quiet.

##### 4. WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### 5. CELLTM

Tests were run on 3 casts to decide on the best parameters for CELLTM.

CELLTM was run using (0.03, 9) for the primary and the secondary conductivity for all casts.

##### 6. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 7. Test Plots and Channel Check

Two 500db casts were studied. The same sensors were used for 2009-46 and 2010-03 so data from two casts from each of those cruises are included in the comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| Early 2009-46 | 240 | +0.001 | -0.0005 | -0.006 | High, steady |
| Late 2009-46 | 240 | +0.001 | -0.0006 | -0.008 | High, steady |
| 2010-03-0191 | 400 | +0.0008 | -0.0005 | -0.0065 | Mod, noisy |
| 2010-03-0365 | 240 | +0.0010 | -0.0005 | -0.0065 | High, noisy |
| 2010-41-0003 | 465 | +0.0007 | -0.0004 | -0.006 | High, moderate |
| 2010-41-0015 | 495 | +0.0008 | -0.0004 | -0.0065 | High, moderate |

The differences are slightly smaller than for the earlier cruises with no evidence of pressure-dependence.

##### 8. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### 9. Checking Headers

The header check was run and no problems were found.

A few minor inconsistencies in the station name format were corrected.

A cross-reference report was produced and checked against the spreadsheet file “CTD log Sardine Survey 2010-41.xls”. The only major discrepancy is the position of event #5, but the positions in the spreadsheet are almost exactly the same for events #4 and #5, so this is presumed to be an error and the CTD headers are assumed to be correct.

Track plots (using event #s and station names) were produced and added to the end of this report.

The surface report shows an average surface pressure of 1.6db and many readings around 1db are associated with very low salinity suggesting that the CTD was very close to the surface. However, examination of the files shows the pumps were not turned on for those low readings, so the surface pressure may not be reading quite as low as those results suggest. There a few cases where the upcast data look like the CTD reached the surface with pressures of 0.4, 0.5 and 0.9db, but again the pumps were off so the surface might be a little deeper than those results suggest. It looks as though the pressure is low by something between 0.4db and 1db.

##### 10. SHIFT

Conductivity

Tests were run on both conductivity channels to see if salinity noise can be reduced. The best results for both channels was with a shift of +0.5 records, although the effect on the secondary was a little hard to understand with either positive or negative shifts both improving the salinity.

All casts were put through SHIFT twice using +0.5s for the primary and the secondary conductivity channels.

Dissolved Oxygen

In 2009-46 this sensor was shifted by +60 records, while +50 looked better for 2010-03. Tests were run using shifts from +30 to +70 records and the best results were with +60 records.

SHIFT was run on all casts using an offset of +60.

##### 11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### 12. COMPARISON WITH BOTTLES

There was a salinity sample for every cast, but no record of the depth from which it was collected. The values suggest they are bottom samples, but this still leaves doubt as to the separation between the Niskin bottle and the CTD.

The DEL files were thinned to just bottom values and those data were added to the 2010-41-sal.csv file to create 2010-41-salinity-comp.xls. The primary salinity ranged from being lower than the bottles by about 0.01 to being high by 0.19.The secondary looked similar but the CTD values were slightly lower relative to the bottles. A plot of differences against pressure shows that the deep samples have small differences for both channels, so this supports the presumption that these are deep samples. The cases of large differences were examined to see if the gradient over the 5db above the bottom of the casts can explain the differences. For cast #20 has primary CTD salinity is higher than the bottle by 0.19. The Niskin bottles are typically mounted 5m above the CTD. The difference in salinity over the bottom 5m of cast #20 is about 0.16, thus explaining a lot of the difference. For the 6 deep casts the salinity differences ranged from -0.004 to 0.004. Only one case had bottom salinity lower than 5m higher, and that was cast with the smallest difference between CTD and bottle.

When only the 7 casts that sampled below 250m are included in the average, the primary salinity is found to be low by an average of 0.0022 and the secondary is higher than the bottle by 0.0033. When the effect of the gradient is subtracted from the differences, the primary salinity is found to be high by <0.001 and the secondary low by ~0.005.

These results are remarkably close to the results of 2009-46 and 2010-03.

##### 14. DETAILED EDITING

All DEL files were copied to *.EDT.

The primary channels were selected for archiving based on the salinity comparison, the history of the sensors and the noise in the secondary conductivity.

CTDEDIT was used to clean the primary temperature and salinity data for all downcasts.

In the course of editing it was discovered that the pumps had not been turned on for the first lowering and raising of the CTD during cast #16. However, once the CTD was back at the surface, the pumps were turned on and the cast was rerun. Both casts are in a single file. A text editor was used to remove the first from file SHFO and then DELETE was rerun.

The pumps did not come until the CTD was below 4db, generally close to 5 or 6db. Surface records were removed where pumps were off and for the first few records after they were turned on if the data looked unstable. Records were also removed where there was evidence of corruption by shed wakes or near-surface noise that looks likely to be due to ship effects. Salinity was cleaned – in most cases lightly. Notes were added to the headers briefly describing the editing done to each cast.

For cast #22 the pumps did not come on until the CTD was at about 21.5db, so data were removed to just below that level. A file was prepared with upcast data and named 2010-41-9022, but the quality of the upcast data are lower than from the upcast.

On-screen plots of descent rate and pump status were used to guide editing.

T-S plots were examined and a few casts were found to need a few more records removed from the surface.

All casts required some editing.

The edited files were copied to *.EDT files.

##### 14. Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors have been used on many other cruises since their last factory service, but there were few salinity samples available and those were mostly from the surface. For 2010-03 and 2009-46 the primary salinity was higher than bottles by 0.0015 and by <0.001, respectively. The secondary salinity from those two cruises were lower than bottles by 0.006 and 0.005.

- The pressure sensor has been used many times but with an incorrect configuration. For 2008-04 and 2010-03 a correct configuration was used for conversion and the offset was adjusted by adding 0.9db for a net offset of -0.2db.

- The dissolved oxygen sensor has been used for 7 other cruises since its last factory service. There was no calibration sampling for 1 cruise, for another the sampling was from a narrow DO range and looked odd. The last reliable calibration was from 2009-09 in June 2009.

Historic ranges –The only excursions from the 3-standard deviation plots of temperature and salinity ranges were from casts 9, 11, 13 and 18 which had salinity above the maxima between 20 and 50db. All of those casts come from the southern half of the cruise and sites with bottom depths between 110 and 160db, suggesting that this reflects real conditions. Only 1 cast from that area and depth range did not have data outside the ranges and it was at the low end of the bottom depth range. The excursions do not look like evidence of equipment or calibration problems.

Repeat casts – There were no repeat casts. There are a few near-by casts but they are in regions of high variability, so there is little to learn about calibrations by comparing them.

Post-cruise calibrations –

The pressure offset based on post-cruise calibration was about 0.65db, with only slight variations with pressures between 0 and 500db.

The dissolved oxygen sensor was found to have drifted little, ~0.1mL/L at 7mL/L. However, the membrane was found to be torn.

The primary conductivity sensor was found to produce salinity high by about 0.0018 if we assume drift was linear with time, and the temperature drift would have had the effect of lowering salinity slightly, so that the net effect would be ~0.0014.

The secondary conductivity sensor was found to produce salinity low by about 0.0018 if we assume drift was linear with time, and the temperature drift would have had the effect of raising salinity slightly, so that the net effect would be low by ~0.0009.

##### 15. Initial Recalibration

The post-cruise calibration suggests that pressure is low by 0.65db and that fits the observations of near-surface data. So pressure was recalibrated by adding +0.65db, for a net offset of ~-0.4db.

The estimate based on post-cruise factory checks is that the primary salinity is high by 0.0014 while COMPARE suggested it was high by 0.0007. Neither estimate is perfect as the first assumes linear drift with time and the latter made assumptions about the distance between Niskin and CTD and calculated the gradient in a rough way. Averaging these two estimates looks reasonable so primary salinity will be adjusted by subtracting 0.001. (It is noted that the estimates from COMPARE for the secondary were that is was low by 0.005, but the post-cruise results suggest it was low by only 0.001. However, the temperature sensor needed repair at that time, so there may have been a problem that did not show up in the calibration test, but might affect data collected during a cast.)

There are no dissolved oxygen samples, so recalibration of the DO channel can only be based on other cruises and/or the post-cruise calibration. The latter shows little drift in the sensor since March 2008. When one cast was reprocessed using the new parameters, the fit of old DO values against new values suggests that the data should be recalibrated using the following fit:

Dissolved Oxygen (Corrected) = 1.005 * Dissolved Oxygen - 0.01

But that last good bottle calibration obtained for this sensor during 2009-09 indicated that:

Dissolved Oxygen (Corrected) = 1.0528 * Dissolved Oxygen - 0.0305

Why would the correction be lower now? Most likely this is due to the torn membrane, and using the 2009-03 recalibration is the best option available. But it should be noted that the DO data is not as reliable as usual.

File 2010-41-recal1.ccf  was used to add 0.65db to the pressure, to subtract 0.001 from the primary salinity and to apply the following correction to the DO channel:

Dissolved Oxygen (Corrected) = 1.0528 * Dissolved Oxygen - 0.0305

##### 16. Fluorescence Processing

There was no fluorometer mounted on the CTD.

##### 17.  BIN AVERAGE of CTD files

The following Bin Average values were applied to the COR1 files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots and profiles were examined on screen. No problems were seen.

##### 18. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  Status:Pump, Descent_Rate and Flag.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. Values ranged from 85% to 170% with most values 95% and 120%. The values are very high for some casts and we have no titrated samples to check and little faith in the recalibration scheme. Luckily, there was another cruise in that area at the same time and it used a different DO sensor and had many titrated samples for comparison. The table below shows 3 pairs of casts that were reasonably close in time and space with the maximum DO measured. The maximum values measured are similar showing that the final DO values for this cruise are not unrealistic.

| 2010-36 Event # | 2010-41 Event # | Days apart | 2010-36 DO max | 2010-41 DO max |
| --- | --- | --- | --- | --- |
| 146 | 23 | 4 | 7.6 | 7.7 |
| 150 | 25 | 4 | 7.5 | 6.9 |
| 120 | 12 | 7 | 8.8 | 9.0 |

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Dissolved Oxygen data were recalibrated based on the results of cruise 2009-09

because there was no oxygen calibration sampling during this cruise. The DO sensor

was found to have a damaged membrane after this cruise. However, comparisons with

data collected during cruise 2010-36 in the same region and time suggest the DO value are reasonable.

##### 19. Producing final files

The Standards Check routine was run and no problems were found. The final files were named CTD.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 24Dec08 | Factory “ |  |  |
| Conductivity | 3321 | 16Jan09 | “ |  |  |
| Secondary Temp. | 4700 | 24Dec08 | “ |  |  |
| Secondary Cond. | 1766 | 16Jan09 | “ |  |  |
| SBE 43 DO sensor | 0997 | 01Mar2008 | Factory |  |  |
| Pressure Sensor | 77511 | 13/Mar/2000 | Factory |  |  |