## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |
| 4-Nov-2013 | Files 2010-57-0011.CHE and 2010-57-0013.CHE: CTD temperature and salinity data were corrupted for the upper 75 m so they were replaced with the downcast values. See note in the Comment section of the Headers. |

## PROCESSING NOTES

Cruise: 2010-57

Agency: OSD

Location: Strait of Georgia

Project: SoG Moorings

Party Chief: Johannessen

Platform: Vector

Date: 6 August 2010 – 7 August 2010

Processed by: Germaine Gatien

Date of Processing: 17 February 2011 – 13 May 2011

Number of original CTD casts:  13 (includes 1 test cast)		Number of CTD casts processed: 3

Number of bottle casts:	13 (includes 1 test cast)			Number of bottle casts processed: 12

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs transmissometer (#498DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2356) with a 10X cable and an altimeter (#1024).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book contained many notes about problems encountered during the cruise. Pressure spikes were a problem as they were for many cruises in spring/summer 2010.

The nutrient analyst noted that the majority of sample caps were found to be loose; tube caps need to be retightened after freezing.

There was insufficient dissolved oxygen sampling to determine how well the SBE sensor performed. Many of the samples had been flagged due to bubbles; half of the unflagged samples came from ~2db and the rest came from below 750b. Calibration of DO was based on cruise 2010-19.

Because many samples were needed 2 bottles were fired at each level (and occasionally 4). In order to have all data appear in a single line, the lower of the two sample numbers was selected in each cast, but the samples did not all come from the same Niskin bottle. CTD data were taken from the last bottle fired, but since all bottles were fired quickly there is no significant difference among them.

The CTD data from event #6 was accidentally overwritten at sea. Because there were surface samples from that cast, a file was created using nominal depth instead of pressure.

The upcasts for events #11 and 13 were heavily corrupted by spikes. Conversion of these files to IOS HEADER format was only possible by removing the data from 75db upwards for the first and 15db upwards for the other. Lines were created for the missing levels with pad values for CTD channels in order to enter the bottle sample results. Nominal depths were entered based on log notes.

Fluorescence and Transmissivity data are nominal and unedited, except that some records were removed in editing temperature and salinity.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained; no rosette log sheets were available but there were extensive notes in the log book about bottle firing.

The surface PAR was not listed in the log but was listed in the configuration file. A check was done on a few casts to ensure there was no signal in that channel and there was not, so it will not be converted.

Dissolved oxygen, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The CON file used during the cruise had a few errors in it:

- The most recent pressure offset had not been entered, so that was changed from 0.1 to 0.8db.

- The dissolved oxygen sensor entries were for the Owens-Millard algorithm. We should be using the SeaBird algorithm, so those were entered.

After those corrections were made the file was saved as 2010-57-ctd.con.

##### Conversion of Raw Data

Rosette files were converted using a start time of -5s and duration of 10s. The Tau feature was used but the hysteresis correction was turned off.

The files were then converted to IOS HEADERS, but the last 2 files could not be converted on the first attempt. This was clearly because they are full of spikes. This has happened for many recent cruises and the only solution was to either remove bottle data that are corrupted by pressure spikes at the end of files, or use Ultraedit to fix bad values. The latter approach does not look possible for these data – there are just too many spikes. So, the final sections were removed. Lines will be fabricated later for the bottle samples, using a Depth:Nominal channel to identify the level.

CLEAN was run to add event numbers to the headers and those files were named *.BOT.

All BOT files were plotted to check for outliers. The secondary salinity data were all bad in cast #11 and the secondary temperature had one bad value. For all other casts there were no outliers.

CNV files were converted using the Tau correction. The bad records from the end of files #11 and 13 were removed using a text editor. Plots were made of the two deep casts:

- The temperature channels are reasonably close on the downcasts, but the differences are larger on the upcasts with a lot of noise in both channels - sometimes one sensor looking better, sometimes the other. The differences between upcast and downcast traces seem larger than usual. Perhpas the package was swinging a lot or the plumbing was not working well on the upcasts.

- Conductivity traces looked much like temperature.

- The fluorescence values look reasonable and not close to going off-scale.

- Dissolved oxygen voltage and transmissivity channels look reasonable.

- The descent rate looks steady.

- The altimetry looks like it has reasonable values, but there are spikes, so header entries will need to be checked against plots.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

There is insufficient data to do reasonable tests on the CELLTM setting, so the one used for 2010-19 was applied.

CELLTM was applied using settings (α = 0.02, β=7) for both primary and secondary channels.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The Tau correction was chosen for the DO derivation.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number; it was also used to reset the header limits.

##### Checking Headers

Track plots were produced. An inconsistency in station names was found; the entries for casts #1 and 2 were corrected.

Event #7 had accidentally been saved in File #6, 2010-57-0006 was renamed as 2010-57-0007 and the event number was changed in the file.

The cruise track plots were rerun and added to the end of this report.

HEADER CHECK was run and the only problems noted are unrealistic values in the header limits. These may disappear at the DELETE stage; otherwise an editor will be used to remove bad values.

The cross-reference check was compared with the log book, and no errors were found.

The surface values program shows the average surface pressure to be 1.2db which looks reasonable for the Vector in quiet waters.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Plots were made of all casts and for the deep casts the header entries were found to be accurate. However, 5 of the surface only casts had an erroneous header entry because spikes were interpreted as signals. The headers were corrected by removing the altimeter entry if the CTD did not get within 15m of the bottom. For casts #3 and 4, the water depths were adjusted to match the log book entries.

##### BOTTLE FILE PREPARATION

Altimeter header entries and water depths were exported to file 2010-57-altimeter-bot.csv. The water depths were corrected for files #3 and 4 and the altimetry headers were removed from files #1 and 10 because there were only surface samples and the CTD was not within 15m of the bottom for those samples.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Some adjustments had to be made to accommodate the fact that there are cases where bottle samples were taken, but no CTD data exits:

- For event #6, the cast which was overwritten, a line was entered for the surface samples

- For event #11 lines were entered for the bottles at 50, 25, 15, 10, 5 and 0m.

- For event #13 lines were entered for the bottles at 15, 10, 5 and 0.

There are double bottle entries (and sometimes 4 entries) for each level but the bottles were fired very close together, so that the bottle file data overlap; this results in truncated records for the 1st bottle. The last bottle fired at a level was the one chosen as the source for the CTD data. The sample number assigned was always the first # used for that level.

The ADDSAMP file was used to create SAM files. Those were bin-averaged on bottle number to create SAMAVG files. For casts #11 and #13 lines were added for the missing bottles with pad values entered for CTD data and the header limits were adjusted.

The ADSAMP file was used to create CST files to be used in the merging process.

SALINITY

The salinity data were delivered in spreadsheet 2010-57.xls which was renamed as 2010-57-sal.xls. There are a pair of samples from 2 bottles fired at the surface of cast #6. There are also duplicates from casts #11 and 13, but unfortunately both are from the surface, so there is little expectation of good repeatability. Nonetheless, the values were copied to file 2010-57-sal-duplicates.csv for later study. The duplicates in the main spreadsheet were replaced with the average value. There were a number of cases where the analyst noted unstable readings, so flag “4” was assigned to those. There were no sample numbers but bottle numbers were given so sample numbers were deduced from those, choosing the lower of the sample numbers for each level. For cast #6 the salinity values from the 2 bottles were averaged. The main spreadsheet was simplified (unneeded columns removed, separate columns were created to replace the one with sample #s and station names combined, event #s were added and headers changed to standard format) and saved as 2010-57-sal.csv which was then converted to individual SAL files.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2010-57oxy.xlsx with quality flags and comments. There were no sample numbers. Those numbers were added based on log book entries, but there were 2 bottles fired at each level. The lower sample number was chosen in each case. This will allow all samples to appear on a single line of the bottle file. There was an analysis of duplicates. The spreadsheet was simplified and saved as 2010-57-oxy.csv which was converted into individual ADD files. Sample #47 was flagged “2” by the analyst with the comment Bad Curve. This was changed to a “4” but will be checked after COMPARE to see if it looks ok, but being a surface sample it is unlikely that there will be strong enough evidence to decide no flag is needed.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2010-57nuts.xls which included a report on precision. There were sample numbers entered. The spreadsheet was simplified, reordered on sample number and saved as 2010-57-nuts.csv. File 2010-57-nuts.csv was then converted to individual NUT files. It was noted by the analyst that caps were not tight, so had probably not been retightened after freezing.

The SAL, ADD and NUT files were merged with CST files in three steps. (Output: MRG1, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files. MERGE was run choosing the MRGCLN1 files as primary and the SAMAVG files as secondary. This is opposite to the usual scheme, but was done to get rid of the extra bottles, choosing only those with entries in the MRGCLN1 files.

Because there are no pressures available to go with some of the bottle data, a channel was added for nominal depths. (Output MRGDEP).

Plots were made of Salinity versus DO from the CTD and Titrated DO. No outliers were found.

##### Compare

Salinity

COMPARE was run and there are a lot of severe outliers at the surface, but an examination of the downcast files indicates that the surface gradients were very high, so this is not indicative of instrumental problems. There are only 6 bottles from below the surface and 3 of those look like outliers for the primary channels; 2 having a high standard deviation in the CTD data and 1 out of line with the others. For the secondary just the 2 with a high standard deviation in CTD salinity look like outliers. The primary is low by an average of ~0.001 and the secondary low by ~0.009. The secondary differences have a higher standard deviation. There is no indication that outliers need flags since the problems are much more likely to be due to CTD problems. There is insufficient data to study pressure or time dependence. (See 2010-57-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run using the SBE DO and the Titrated bottle DO data. There are too few data to conclude much, but the general shape of the plot of differences against titrated values looks reasonable. Many of the bottle data had been flagged, but whether or not flagged values are included makes little difference to the fit which is approximately

CTD-BOT = 1.12 DOX-CTD

Most of the samples are from the surface where we expect a lot of noise in these sort of comparisons, and there are no samples between 3db and 70db, hence the range of DO values represented is very limited. No changes to quality flags are justified. (See 2010-57-dox-comp1.xls.)

##### 12. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence is divided by the average descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always a rough estimate as the upcast data are usually very noisy. For this cruise there is only 1 cast for which there are near-surface data for both upcast and downcast. The usual shift of +24 records (1s) was found to improve the alignment of that file, so it was applied to all files. (Output: SHFFL)

Conductivity

Tests were run on a few casts and while the differences were not large, the best choice was -0.2 for the primary conductivity channel while the secondary was best with no adjustment. SHIFT was run on the primary conductivity for all casts with a setting of -0.2 records. (Output: *.SHFC0)

Dissolved Oxygen

Only cast #1 is suitable for determining the best SHIFT value to apply to the Dissolved Oxygen channel. This is judged by how the vertical offset between downcast and upcast traces compare with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. The setting that has been used on other recent cruises using this sensor is +60 records and that setting produces good results for cast #1, so it was applied to all casts.

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The first time this routine was run there were a few warnings about jumps in pressure values. The SHFO files were examined and for casts #1 and #6 there were some bad pressure values before the pumps were turned on. The bad values were replaced with pad values and DELETE was rerun. The only warnings on the 2nd run were notifications that some files contain less than 10db of data, which is correct.

Profile plots were examined to see if there remain problems from pressure spikes.

The transmissivity values are very low near the surface of some casts, but this is likely real.

The pumps did not come on until the CTD was at 2 or 3db, so the surface-only casts contain little useful data and will not be processed further.

Cast #12 never had the pumps on, so will not be processed further.

##### 14. DETAILED EDITING

COMPARE indicates that the primary salinity is closer to the bottles as does the history of these sensors. Moreover, secondary data were often bad. The primary channels were selected for editing.

Graphical editing was done to casts #1, 11 and 13 using program CTDEDIT. On-screen plots of descent rate and pump status were also used. Surface records were removed where the pumps had not been turned on, bottom records were removed near the bottom of #13 due to shed wake corruption. Salinity was cleaned lightly in all 3 files.

##### 15. Initial Recalibration

No recalibration will be applied to the salinity channel since the results of COMPARE and the history of the instrument suggest it is within ±0.002.  It is harder to decide what to do with the SBE Dissolved Oxygen data. The comparison from this cruise suggests a larger correction than other recent cruises have indicated. Many of the titrated samples had bubbles and were flagged by the analyst, and many that remain are surface samples or from below 70db so there is a limited range. The results of 2010-19 will be applied.

File 2010-57-recal.ccf was prepared to correct the transmissivity by multiplying by 2.382 and to recalibrate the Dissolved Oxygen by applying equation:

CTD-BOT = 1.0378 DOX-CTD + 0.0259

This was applied to the SAM and MRGDEP files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were, though the CTD values still look low. Fits forced through the origin have a slope of 0.11 before recalibration and 0.07 afterwards. (See 2010-57-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

A second recalibration was applied to the COR1 files, using the 2010-19 correction:

DOX(corrected) =  DOX +0.0001*Pressure -0.0329

This correction would reduce surface values slightly, and have no affect at 300db, so will reduce the slope very slightly.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### 18. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were a few unstable features, but these are small and near the surface where it is not clear what, if any, data to remove. No further editing was applied.

##### 19. Other Comparisons

Previous experience with these sensors –

1. Salinity - Both conductivity sensors have been used many times since their last recalibrations in 2007. The results are highly variable with none of the comparisons thought to be very reliable. In the 2 most recent uses the primary was higher than the bottles by 0.0001 and lower by 0.002, though in the latter case, the differences were very small below 200db. The secondary salinity has been found to be lower than bottles by from 0.001 to 0.005 with the last two uses being -0.004 and -0.005.

2. Dissolved Oxygen – The 4 most recent cruises that used this sensor since the last factory calibration had the following fits:

CTD-BOT = 1.0203 DOX-CTD + 0.0696 (2009-64 - November)

CTD-BOT = 1.0258 DOX-CTD + 0.0760 (2010-16 – April)

CTD-BOT = 1.0259 DOX-CTD + 0.0602 (2010-24 – April)

CTD-BOT = 1.0378 DOX-CTD + 0.0259 (2010-19 – April)

3. Pressure –The sensor was recalibrated in August 2007 and has been used with an offset of +0.8db for 3 recent cruises.

Historic ranges (3 standard deviations) - All data fell within the local climatology.

##### 20. FINAL CTD files steps (REMOVE and HEADEDIT)

Header Check was run to look for odd values and none were found.

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Altimeter, Status:Pump and Flag.

A second SBE DO channel (with umol/kg units) was added. REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

There is insufficient information to make an estimate of the precision of the SBE dissolved oxygen data in the CTD files.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 110% to 120%.

##### 21. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Specific comments were added to explain how the files were prepared for casts #6, 11 and 13.

Standards check was run on all files and a problem was found. Because the files were merged in an unusual way, some header information was lost. It exists in the SAMAVG files, so was copied from SAMAVG to the MRGREO files. CLEAN was rerun to ensure the header limits are correct. Then HEADER EDIT was rerun and the header checks found no further problems in the CHE files.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files and no problems were noted.

The sensor history was updated.

Particulars:

6. File overwritten by event #7, renamed as 7.

11. Downcast fine, but major electronic glitch during upcast. Pressure unavailable at 75 to 50m of upcast. Many flashes and beeps for rest of cast.

12. Test cast – same site as #11 – no bottles fired and pumps off.

13. Electronic glitch again at 15m level of upcast.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 498DR | 1Nov10 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| Reference PAR | 16504 | 3Mar2010 | IOS |  |  |
| Fluorometer | ? |  |  |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1024 |  |  |  |  |