## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2010-72

Agency: IOS, IOS, Ocean Sciences Division, Sidney, B.C.

Party Chief: Tuele D.				Location: Discovery Islands

Project: Discovery Islands Circulation		Platform: Nucleus

Date: 29 August 2010 –3 September 2010

Processed by: Germaine Gatien

Date of Processing: 18 May 2011 – 21 May 2011

Number of original CTD casts: 38		Number of CTD casts processed: 38

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0334) was used.

# SUMMARY OF QUALITY AND CONCERNS

The pressure sensor is considered ±0.2db.

The times in the log are in PDT; the times in the headers of the archived files are in UTC.

Comparison of CTD salinity with samples plus post-cruise calibration of temperature and conductivity sensors show that salinity accuracy is within ±0.001.

The CTD salinity precision is considered ±0.001psu when stopped. In motion the salinity is considered ±0.005psu. (In areas of rapid temperature change the SBE25 does not perform this well, but for this data set the gradients were not high enough to cause significant problems.)

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data is in a single HEX file.

2. Preliminary Steps

The Daily Log was obtained. There was note of a few small problems and record of where salinity samples came from.

File names were non-standard, but contain the date and the event number, so it was easy to rename them in standard format. However, the file names were out of line with the log event numbers, with the first file ending in 0 instead of 1 and the final ending in 37 instead of 38. A test conversion showed that the “0” file was not just a test, but really is event #1, so the hex file names were adjusted by making the last 2 digits match the log event numbers.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

The configuration file provided was saved as 2010-72-ctd.con. Conversion did not work using that file until the serial number of the conductivity sensor was changed to -9 and the sampling rate changed to 8 samples per second.

A few casts were checked to see if it is necessary to add an offset to the pressure. For cruises in 2008 and 2009 an offset of +0.7db was found appropriate. For these casts the operator typically let the CTD soak very close to the surface because the conditions were calm, and usually he let the CTD sit for a while just below the surface at the end of the cast, then ran the CTD as it came through the surface. So there is a lot of evidence on which to base a decision. Typically, the initial pressure looks close to 0. On the way up the pressure as the CTD breaks through the surface ranges from 0 to -0.9db, with most being ~-0.6db. This is likely due partly to hysteresis since the casts with maximum pressure <100db show values closer to 0 (though that is not always the case), and partly because the system may take a short time to clear of seawater. This sensor has resolution of ±0.2db. Given that most casts are quite shallow and upcast data are not likely to be used, no offset will be applied at this stage. This will be reviewed after conversion to IOS header format.

Data from a few casts were plotted. The temperature and conductivity traces are similar for downcast and upcast, except for a few cases for which surface variations likely reflect real change.

The descent rate appears to have been kept reasonably high most of the time and about as steady as possible for this type of deployment. It was mostly >0.5m/s and for the deeper casts, the rate was above 1m/s for most of the profile.

The soak period was generally 1.5 to 2 minutes, and seems to have been long enough to reach equilibrium. However, for casts #35-38 there was a shorter soak and for cast #37 the CTD started down before the pumps could have come on. The weather may have been a little rougher at that point and staying at the surface may have been judged to be too dangerous.

4. WILDEDIT

This step is usually skipped and since there is no evidence of spikes it was not run.

5. WFILTER

WFILTER was run using cosine filter, size 5 on the pressure, temperature and conductivity channels. Plots were made before and after this step. Most reversals disappear from the pressure after this step but the trace is not overly smooth. When salinity was derived on filtered files it was much smoother than the unfiltered version, but temperature and salinity still have some reversals as expected.

6. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. A variety of settings were tested with results varying from feature to feature, but the recommended setting was best overall.

CELLTM was run on all casts using (α, 1/β) = (0.04, 8)

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

9. Checking Headers

A spreadsheet was prepared containing file name, latitude, longitude, station name and water depth with the information taken from the log book. The file was saved as 2010-72-headers.csv.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the CSV file to the IOS files with output MRH. 

Track plots were produced to turn up obvious errors in the spreadsheet, and those were corrected and the merge and plot routines rerun until all problems were resolved.

Then a cross-reference list was produced and checked against the log. The times in the headers are close to the log entries for the end of each cast; the differences are likely due to the computer clock and the clock used on deck being slightly offset, and the time being recorded at slightly different stages of the operation.

A problem was found in checking the entries for log event #s 19 to 24 with the cast numbers and times not corresponding to log entries in the same way as they do for events 1-18 and 25-38. There is no file with data from the time of the log entry #19 and there is a note by that event saying “Not ON”, so perhaps the CTD had not been turned on, and no data were acquired. The cast was rerun the next morning and is noted in the log as Event #20. If we assume that event #19 is the one given in the log as #20, and 20-23 are the ones written in the log as 21-24, then the 2nd cast in “Pryce Ch”, which had no event # recorded in the log would become #24 and all makes sense.

One check on this assumption is to look at plots, checking maximum pressures and whether there are 2 casts that look close enough to both be at “Pryce Ch”. The cast currently labeled as #19 is deep enough for “Calm Ch” with the right time for the 2nd “Calm Ch” cast. Casts #18 and 20 do not look like #19 and have different maximum pressures from the log entries for “Calm Ch”, so there definitely were not 2 “Calm Ch” casts logged. Cast #20 has the right pressure range for “Ra Pass”. The next bunch of casts all have roughly the same maximum depth, but the only ones that look alike are #23 and 24. They are not identical, which is expected since while both in Pryce Channel, they are some distance apart and one was in much deeper water than the other. The party chief agrees that the event #s are entered wrong in the log for events #19 to 24, so notes were placed in the log book about what the event numbers should be.

The time zone is PDT, so ADD TIME CHANNEL was run to add 7 hours to the header times, so they are in UTC.

HEADER CHECK was run, and no problems were found.

The surface check showed all casts started between -0.5 and +0.8. No recalibration will be applied to the pressure, as it looks as close as the resolution allows.

##### 10. Test plots

Plots were made. The near-surface data looks unreliable, so before editing (but after DELETE) it would be efficient to clip the data as is often done with SBE25 data. While the top 1 to 2db are usually poor, for this cruise the CTD was soaked very close to the surface. Clipping to +0.5db would leave some bad surface data, but clipping any deeper might lead to the loss of some good data.

11. Salinity bottles

Salinity

There were 4 salinity samples, one each from casts 6, 18, 21 and 28. The CTD salinity values were recorded from the files at the maximum pressure and from 1db above the maximum pressure in the downcast data. The salinity data were not smoothed, but the variability was low. The differences show good agreement, especially for the readings 1db above the maximum depth, and the party chief confirms the Niskin was about 1m above the CTD sensors. The average difference was only 0.0001 with a range from the CTD being low by 0.002 to high by 0.003.  See file 2010-72-sal-compl.xls for details.

12. SHIFT

Conductivity

Tests were run using values from -1 to +0.2. During 2007-13 and -14 when this sensor was used a setting of -0.6s looked best, for 2008-20 a setting of -0.25s looked best overall and -0.8 records was used for 2009-69. For this cruise the salinity is pretty smooth and settings of -0.4 to -1 made things worse. There was little difference among -0.2, 0 and +0.2 records, but overall, no change seemed best. So this routine was not run.

13. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: The only warning pertained to cast #25, which had less than 10db of data.

- DETAILED EDITING

CLIP was used to remove data above 0.5db. All CLIP files were copied to *.EDT.

All files needed a little editing except cast #15. Usually some surface records were removed and sometimes bottom records. CTDEDIT was used to clean salinity where unstable features looked likely to be caused by misalignment of T and C. Not all unstable features were removed, since they were not obviously due to instrumental problems and in this region could well be real.

Notes of editing details were made in the headers.

All edited files were copied to *.EDT.

- Inter-comparisons

Sensor History – The temperature and conductivity sensors were used on 2 cruises in late 2009 when there was no bottle comparison available. The pressure sensor had been used for many other cruises that have been processed since it was last calibrated at the factory. An offset of +0.7db was applied to all but one of them. For one cruise in 2008 no offset was applied.

Comparison of repeat casts –The only repeat casts are very shallow, so comparisons are not useful.

Historic Ranges – There is no local climatology available for this area.

Post-cruise calibrations –  The conductivity and temperature sensors were recalibrated in December 2010, at which time there was no drift in conductivity at low salinity values, though possibly a little above salinity values of 32. The temperature drift was not significant.

- Bin Average and REMOVE

No recalibration is required.

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Descent Rate and Flag channels.

- HEADER EDIT and final checks

Header Edit was used to fix formats and channel names and to add the following note to the headers:

The pressure sensor is considered ±0.2db

The CTD salinity is within 0.001 of bottles and post-cruise calibration

showed no signficiant drift in temperature and conductivity.

The CTD salinity precision is considered ±0.001psu when stopped. In motion

the salinity is considered ±0.005psu.

The final files were named CTD.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors found.

The sensor history was updated.

Particulars:

7. CTD may have hit bottom.

15. Down to 5m then up to reposition.

17. Very rough

19. Not on – don’t know what this refers to

Salinity samples at casts 18, 20, 28.

## Institute of Ocean Sciences

## CRUISE SUMMARY

Cruise ID#: 2010-72

Dates: Start: 29 August 2010 End: 3 September 2010

Location: Discovery Islands

Vessel: Nucleus

Party Chief: Tuele D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2010-72

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5013 | 30 Dec 08 | Factory | 30Dec2011 | Factory |
| Conductivity | 3531 | 07 Jan 09 | Factory | 29Dec2011 | Factory |
| Pressure | 0464 | 23 Mar 06 | Factory | 16march2011 | Factory |