## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2010-79

Agency: IOS, IOS, Ocean Sciences Division, Sidney, B.C.

Party Chief: Tuele D.				Location: Discovery Islands

Project: Discovery Islands Circulation		Platform: Big Winds

Date: 5 December 2010 – 10 December 2010

Processed by: Germaine Gatien

Date of Processing: 24 May 2011 – 27 May 2011

Number of original CTD casts: 20		Number of CTD casts processed: 20

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0334) was used.

# SUMMARY OF QUALITY AND CONCERNS

The pressure sensor is considered ±0.2db.

The times in the log are in PST while the computer times are in PDT. The computer clock may be running fast, but it is impossible to check this accurately and the differences from the log may be due to the stage in the operation at which each was recorded. The times in the headers of the archived files are in UTC.

The configuration file provided with the data was clearly not the one with which the data were acquired. After some alterations it was possible to convert the files, but not when the PAR channel was included. There was a file converted in the field that did contain PAR data, but there was no signal, so it is assumed that it malfunctioned. Saving configuration files with different names, rather than adjusting them under the same name, would help figure out problems like this.

There was no salinity sampling for this cruise, but there was a tank test. A post-cruise factory report showed that salinity calibration was within 0.001.

CTD salinity precision is considered ±0.001psu when stopped. In motion the salinity is considered ±0.005psu.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data is in a single HEX file.

2. Preliminary Steps

The Daily Log was obtained and was in good order. No equipment problems were noted.

File names were non-standard but contain something that looks like an event number. File names were changed to standard format using those numbers, but they do not correspond to the numbers in the log book. The file names will be corrected after conversion.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

The configuration file provided was saved as 2010-79-ctd.con. Conversion did not work using that file until the serial number of the conductivity sensor was changed to -9 and the sampling rate changed to 8 samples per second. After those changes a CNV file was produced, but it contained no data. If the PAR sensor was dropped from the configuration file, with only 2 external voltage channels, then data were produced and all looked reasonable. There is a file with the RAW data that was converted with both external sensors, so conversion must be possible somehow, but repeated experiments proved fruitless. However, in that 1 converted file the PAR channel does not have a signal, just a constant high value at all depths. So, it must be assumed that use of the PAR sensor was abandoned and the configuration changed for these casts. This does not explain how they were converted in the field, but it must be assumed the configuration file used does not have the same settings as it does now. Other evidence for that is the fact that no conversions could be done with the hex file provided until the serial number of the conductivity sensor was changed to -9 and the sampling rate changed from 1/second to 8/second.

A few casts were checked to see if it is necessary to add an offset to the pressure. This cruise and the one in August provide more evidence for this judgment than usual.  The operator typically let the CTD soak very close to the surface, and usually he let the CTD sit for a while just below the surface at the end of the cast, then ran the CTD as it came through the surface. Judging by the pressure at which the conductivity shifts rapidly up or down, the surface was typically at -0.2db on the downcast and about -0.5db for the upcasts. The larger upcast offset would be partly because the system takes a short time to clear of seawater. This sensor has a resolution of ±0.2db. An offset of +0.2db looks advisable, so conversion was rerun with that setting.

After conversion the files were examined to see which files refer to which log events. This is not as simple as it might appear since the computer is probably on PDT while the log is in PST. It also looks as though the computer clock is running faster than that used for the log, though it is possible the two times may record a different stage in the acquisition process, such as the log recording when the boat arrived on station and the file when the CTD acquisition was turned on. Many of these casts were very close together, which also makes it hard to be sure which event goes with which file. However, the pattern seems clear that log event #1 goes with file #5 and so on, with log event #20 going with file #24. This makes sense, so the file names were changed to match the log by subtracting 4 from the final 2 digits.

Data from a few casts were plotted. The temperature and conductivity traces are similar for downcast and upcast, except for a few cases for which surface variations likely reflect real change. The transmissivity traces look reasonable. The descent rate was low but very steady. DELETE will remove most data from some casts unless the drop rate cut-off criteria are set lower than usual.

The soak period was generally 1.5 to 2 minutes, and long enough to reach equilibrium.

4. WILDEDIT

This step is usually skipped and since there is no evidence of spikes it was not run.

5. WFILTER

WFILTER was run using cosine filter, size 5 on the pressure, temperature and conductivity channels. Plots were made before and after this step. Most reversals disappear from the pressure after this step but the trace is not overly smooth. When salinity was derived on filtered files it was much smoother than the unfiltered version, but temperature and salinity still have some reversals as expected.

6. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. A variety of settings were tested with results varying from feature to feature, but the recommended setting was best overall.

CELLTM was run on all casts using (α, 1/β) = (0.04, 8)

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

9. Checking Headers

A spreadsheet was prepared containing file name, latitude, longitude, station name and water depth with the information taken from the log book. The file was saved as 2010-79-headers.csv.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the CSV file to the IOS files with output MRH.

A cross-reference list was produced and checked against the log. The times in the headers are approximately 1 hour ahead of the log entries and a few minutes later; the differences in minutes are likely due to the computer clock and the clock used on deck being slightly offset, and/or the time being recorded at slightly different stages of the operation. One error in position was found and corrected in the CSV file and the merge was rerun for that cast.

Track plots produced no obvious errors.

The time zone in the file headers is PDT, so ADD TIME CHANNEL was run to add 7 hours to the header times, so they are in UTC.

HEADER CHECK was run, and no problems were found.

The surface check had an average value of -0.29 with a range of -0.6 to -0.8. The associated salinity values were extremely low. Examining a few files shows transmissivity going to ~0 between -0.2 and +0.2db for both upcast and downcast, so no further correction will be applied.

##### 10. Test plots

Plots were made. The near-surface data looks unreliable, so before editing (but after DELETE) it would be efficient to clip the data as is often done with SBE25 data. While the top 1 to 2db are usually poor, for this cruise the CTD was soaked very close to the surface. Clipping to +0.7db would leave some bad surface data, but clipping any deeper might lead to the loss of some good data. This will be done after SHIFT.

11. Salinity bottles

Salinity

There was no salinity sampling, but there was a tank test. A bottle sample from the tank was analyzed on an Autosal and had salinity 23.7321. The data collected in the tank were averaged over 100 scans and salinity varied from 23.7366 to 23.7396, and temperature varied from 15.6803 to 15.7013°C.This would suggest that the CTD was reading low by from 0.004 to 0.007. Looking at the full file, there were some data with salinity lower than the bottle values, but they look like spikes. There was a lot of noise in the CTD readings.

12. SHIFT

Conductivity

Tests were run using values from -1 to +0.2. During 2007-13 and -14 when this sensor was used a setting of -0.6s looked best, for 2008-20 a setting of -0.25s looked best overall and -0.8 records was used for 2009-69. For 2010-72 all settings made things worse. For this cruise the results were different, perhaps because there is more temperature variation. A setting of +0.4s looked best for most features.

SHIFT was run using an advancement of +0.4s on the conductivity.

13. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.2m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

NOTE: An initial run using the usual 0.3m/s cut off was found to remove too much data. Some of the casts had a very steady but low descent rate and using 0.2m/s preserved more data that look good.

- DETAILED EDITING

CLIP was used to remove data above +0.7db. All CLIP files were copied to *.EDT.

All files needed a little editing except cast #18. CTDEDIT was sometimes used to remove surface records and bottom records and records corrupted by shed wakes. It was also used to clean salinity where unstable features looked likely to be caused by misalignment of T and C. Not all unstable features were removed, since they were not obviously due to instrumental problems and could well be real.

Notes of editing details were made in the headers.

All edited files were copied to *.EDT.

- Inter-comparisons

Sensor History – The temperature and conductivity sensors were used on 2 cruises in late 2009 when there was no bottle comparison available. During 2010-72 there was salinity sampling and the sensor was found to be within 0.001 of deep bottles. The pressure sensor had been used for many other cruises that have been processed since it was last calibrated at the factory. An offset of +0.7db was applied to most but no offset was applied to 2 summer cruises, possibly because the sensor calibration is temperature sensitive.

Comparison of repeat casts –There were no repeat casts and nearby casts are too shallow and variable for a useful comparison.

Historic Ranges – There is no local climatology available for this area.

Post-cruise calibrations – The conductivity and temperature sensors were recalibrated in December 2010, at which time there was no drift in conductivity at low salinity values, though there might be a slight drift for salinity values >32. The temperature drift was not significant.

- CALIBRATE

The post-cruise drift report plus the comparison from 2010-72 show that there is little or no drift in the salinity calibration. The tank test shows the CTD salinity to be a little low, but with only one bottle and noisy CTD data this is not strong evidence. Salinity will not be recalibrated. 
Pressure was corrected by redoing the conversion with an offset, so CALIBRATE will not be run.

- Bin Average and REMOVE

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Descent Rate and Flag channels.

- HEADER EDIT and final checks

Header Edit was used to fix formats and channel names and to add the following note to the headers:

The pressure sensor is considered ±0.2db

Based on a bottle comparison from another cruise and a post-cruise

factory report on sensor calibration drift, the CTD salinity

calibration is considered to be accurate to within 0.001.

The CTD salinity precision is considered ±0.001psu when stopped.

In motion the salinity is considered ±0.005psu.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors found.

The sensor history was updated.

## Institute of Ocean Sciences

## CRUISE SUMMARY

Cruise ID#: 2010-79

Dates: Start: 5 December 2010 End: 10 December 2010

Location: Discovery Islands

Vessel: Big Winds

Party Chief: Tuele D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2010-79

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5013 | 30 Dec 08 | Factory | 30Dec2011 | Factory |
| Conductivity | 3531 | 07 Jan 09 | Factory | 29Dec2011 | Factory |
| Pressure | 0464 | 23 Mar 06 | Factory | 16march2011 | Factory |
| Transmissometer | 498DR | 1Nov0211 | IOS |  |  |
| PAR | 4615 | 15dec2000 | ? |  |  |