## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | Updated channel names & formats in TOB & loop files. GG & SH |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 6 April 2021 | ### Corrected dmsp comments event 4, changed event 38 to 39 on QF dmsp. S.H. |
| ### 30 August 2020 | ### Added HPLC data. S.H. |
| 27 July 2020 | Corrections to DMSP-T comments & precision G.G. |
| 10 Nov 2015 | Added DMSP data to CHE files. G.G. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2011-01

Agency: OSD

Location: North-East Pacific		Project: Line P

Party Chief: Robert M.			Platform: John P. Tully

Date: February 8, 2011 – February 22, 2011

Processed by: Germaine Gatien

Date of Processing: 19 August 2011 – 28 October 2011

Number of original HEX files: 39 CTD and 5 TSG

Number of CTD files:   34	Number of bottle casts:	33

Number of TSG files:  5

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2345) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4694) and an altimeter (#1024). All casts were run with deck unit #0424.

The deck unit was #0508.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had no equipment list, but a printed configuration file noted that the serial numbers of sensors had been checked by a visual inspection. However, there was no record of the TSG equipment being checked. There were many equipment problems and cases of split casts and irregular bottle firings; these were noted in the log and summarized in notes from the Chief Scientist. The rosette logs were in good order and there was a sheet for every cast, even when only surface samples are taken, which is very helpful.

This was a complex cruise with computer crashes, multiple files for some casts and corrupted data. The multiple files were joined as needed and event numbers were set to match the first of each of the set. Where bottles were collected from a separate upcast file, the event number and file names for the CHE files were changed to match the corresponding CTD file.

There was a cable problem that led to loss of signal from the primary conductivity sensor and the dissolved oxygen sensors for sections of Events #10, 13, 17, 18, 19 and 21. For cast #21 the conductivity sensors were reversed and the primary conductivity still had many zero values, thus proving that the problem was with the cable. It was replaced and the problem disappeared.

The fluorometer malfunctioned during the downcast of event #1, so was removed from the CTD file.

The secondary salinity was found to be higher than bottles by about 0.001 with no pressure dependence. The primary salinity was slightly closer to the bottles, but had more pressure dependence and there was no signal from the primary conductivity sensor for large parts of several casts, so the secondary salinity was chosen for archiving. This cruise was the 2nd to take place with both new bottle liners and bottles washed using the new bottle washer. Salinity analysis took place within a month of the cruise end. It is encouraging that there were fewer major outliers in the fits and that the comparisons fit well with the post-cruise calibration results. The secondary sensors were selected for archiving, but note that for cast #21 the secondary conductivity sensor is a different one from all the other casts.

A shift in the DO standardization parameters occurred between events #32 and #44; this was not investigated at the time, so it cannot be determined if the resultant change in Thiosulfate normality was real, or the result of a shift in lab temperature. Checks of the fit of SBE DO against titrated DO suggest that the titrated values from the first part of the cruise are lower than those of the latter part by about 0.02mL/L. Since the SBE DO was recalibrated using data from the whole cruise, the resultant error in SBE DO is likely on the order of 0.01mL/L.

A new approach was taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel was compared with bottles to find the slope and offset to enter in the configuration files. This method is a little slower and less convenient, but it is the standard approach and is recommended by SeaBird.

The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder:

OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

The fluorometer used on the Thermosalinograph is long overdue for recalibration, it gives significantly higher values than both loop samples and the fluorometer mounted on the CTD.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as a report from the Chief Scientist summarizing problems and points of interest with reference to processing. There were many problems with signal drop out, spikes, fuzzy traces and interrupted casts. A number of steps were taken to try to resolve problems in the primary salinity and dissolved oxygen channels, including switching conductivity sensors for cast #21. The results of that test led to a cable change before cast #22 and no further drop-outs were reported. However, problems with fuzzy traces and/or spikes continued in the transmissivity and fluorescence channels.

There were three casts that were interrupted by spikes or computer crashes so that 2 or 3 files were created to cover the cast.

- For the cast at station SI03, the downcast data were all in event #1 and the upcast all in event #2. The bottle file will be renamed as 2011-01-0001.

- For the 2nd cast at P4 there are 3 files and #10 and 11 will have to be combined to create a full downcast and files #11 and 12 will need to be combined to create the bottle file; both will be named 2011-01-0010.

- File #46 contains part of the downcast while #47 contains the rest of the downcast plus the upcast. So the downcasts of 46 and 47 will need to be combined and named 46, while #the bottle file is all in #47 so no join is needed but it will be renamed as 46, to match the downcast and the rosette log record.

A special configuration file will need to be created for cast #21.

File names for event #17 were corrected.

Extracted chlorophyll, nutrients, dissolved oxygen, salinity and DMS data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. There were a number of errors in the con files:

- The pressure slope was wrong, so was fixed.

- The data entered for the dissolved oxygen sensor were from an earlier calibration; a more recent one in February 2008 should have been used and the Sea-Bird parameters should have been selected instead of the Owens-Millard, so new data were entered.

- The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

Because post-cruise calibrations are available for many of the sensors, a second configuration file was prepared. A few test conversions were done and the pressure seemed a little low with the post-cruise calibration, but the salinity channels were closer. The pressure can be adjusted later, if necessary. The temperature and conductivity calibrations from March 2011 were selected since they are much closer in time to this cruise. Those calibrations were the “as received” values, not those found after the new connectors were added. The membrane of the dissolved oxygen sensor was found to be damaged when it was returned to the factory so no “as received” calibration was done. The sensor was repaired and then calibrated, but those values will not be applicable to the data acquired in February. There is only one calibration available for the pressure, but only the slope and offset are changed, so that is probably fine. The configuration file using the new T, C and P values was saved as 2011-01-ctd-post-cal.con.

Configuration file 2011-01-ctd-post-cal-21.con was created with the conductivity channels reversed; this will be used for cast #21 only. It was necessary to leave the serial numbers unchanged in this special con file because they were not changed for the acquisition. So the header calibrations are right, but not the serial number identification. The calibration parameters listed in the headers have the wrong conductivity sensor numbers for cast #21. This will be changed in the final files to reflect what was actually mounted.

##### Initial Rosette File Conversion and DO Calibration Study

Rosette files were converted from HEX files.

File 2011-01-001.ros contains only surface data with pumps off, so it was deleted.

File 2011-01-0001.ros looks ok, but the 2 bottles were fired by accident during the downcast and were not sampled, so that file was deleted.

File 2011-01-0037 has 2 bottles fired, but there was no sampling and the cast was rerun as #38, so the ROS file was deleted.

The cast at SI03 was interrupted due to bottle problems, but there is no need to merge files as all the downcast data are in file #1 and the upcast are in file #2. The ROS file from SI03was renamed as 2011-01-0001.ROS to match the downcast file at the same site.

The upcast files 11 and 12 will be merged and be renamed as 2011-01-0010. The rosette file merge looks to be straight-forward.

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files for files 11 and 12 were joined; first the names were changed to 11.botx and 11.boty and then the bottle numbers in the BOTY file were changed from 1-6 to 14-19 to match the rosette sheet). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

There were a few problems encountered in preparing the ADDSAMP file:

- Two sample numbers were used for both events #1 and 3. This will not interfere with creation of the bottle files but care will be needed when adding sample data to ensure they go to the right file.

- For cast #22 Niskin bottles 1-5 were closed but no samples were taken so they were removed from the ADDSAMP file. The sample numbers that would have been used for those bottles were instead assigned to deep bottles of cast #23, so sample numbers are in irregular order. This will add a complication when analysis results are merged with cast #23.

- Cast #32 has errors in the rosette log sheet Firing #s. The notes from the Chief Scientist make it clear what happened. A bottle was added at 1000db. This should be Firing #9, Niskin #24. The Niskin #s are correct on the log sheet, but not the Firing #s. I entered the correct numbers.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-01-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. The data for most of casts #11 and #21 were also removed because there were bad oxygen solubility values. A calculations was made of  Ф  using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.5928 and B = -0.3058 with an R2 value of 0.9989. While R2 is quite tight, 5 points do stand out, so those were investigated. Among those outliers were 3 that came from the Saanich Inlet cast which is always a challenge for the DO sensor, and 1 was near the surface. The other outlier is associated with a higher standard deviation in the SBE DO voltage channel. When those 5 outliers were removed from the fit, M = 0.5907 and B = -0.3034 with an R2 value of 0.9998.

During a post-cruise review of the DO data it was noted that while precision was excellent, there was a potential problem with the standardization (see “note re DO shift.doc”.) This might lead to a shift in values starting at cast #44. So the data were divided into 2 groups to see if that happened. The M and B values for the two groups were (0.5910 & -0.3028) and (0.5907 & -0.3034). The net difference this would make to DO concentration for 2.2 volts (the maximum voltage included in the fit) is 0.02mL/L.

Difference in concentration = ((Soc1-Soc2)*Voltage + (Voffset1/Soc1 – Voffset2/Soc2) *phi

or, 	Difference in concentration = ((Soc1-Soc2)*Voltage + (B1 – B2) *phi

If after derivation of concentration it appears that one of these two sets looks notably better than the other then it is possible that the analysis may be reconsidered, but at this point there is no obvious way to decide which set is better.

The value found for the cruise as a whole was entered in the configuration files.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. Those files were converted to IOS header format. They were put through CLEAN to create BOT files.

The following corrections were applied to the BOT files:

- File 2011-01-001.ros was deleted.

- File 2011-01-0001.ros was deleted.

- File 2011-01-0037 was deleted.

- File 2011-01-0002.ROS was renamed as 2011-01-0001.ROS to match the downcast file at the same site.

Temperature and salinity were plotted for all BOT files and a few outliers were found in the secondary salinity in casts 42 and 52. Secondary salinity was cleaned lightly using CTDEDIT and the output files were then copied to BOT. There were many bad primary salinity values as noted in the log.

The BOT files 11 and 12 were renamed as 2011-01-0010.BOTX and 2011-01-0010.BOTY. Ultraedit was used to fix the bottle numbers in the BOTY file to match the rosette log sheet; the header was adjusted to match that. They were then joined, with the output file names 2011-01-0010.BOT.

It was noted that the bad primary salinity values do not have normal pad values, so another run of CLEAN was used to replace values between 1000 and 9000 with -99.000 in that channel only.

The addsamp.csv file prepared in the DO calibration step was ordered on event # and then sample number. It was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT2 files. The SAM files were then bin-averaged. Some lines were removed from the SAMAVG files for casts #1 and 23 for bottles fired with no sampling done.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-01-bot-hdr.txt; it may need further editing to reflect problems found during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2011-01chl.xls. The name of the file was changed by adding the date of creation. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed, data were sorted on sample number and the file was saved as 2011-01-chl.csv. Loop samples were placed in a new file, 2011-01-loop-chl.csv, and removed from the main spreadsheet which was then converted to individual CHL files.

DMS

DMS data were obtained in file DMS 2011-01 summary.xls. The file was saved as 2011-01-dms.csv and edited. There were no flags or comments, but since there were explanations of flags at the bottom of the page, it is assumed that there were no problems to report. All entries “<” were replaced with “0”; a note in the header will explain that the minimum detectable level is 0.1. There was a set of duplicates for each cast sampled and they were both reported. The pairs were replaced by a single line with the average value and flag “6” was added. Headers were changed to standard format and unnecessary columns were removed; an event number column was added and filled in with information from the log book. The file was then converted to individual DMS files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-01nuts.xls which included a report on precisions. Loop data were copied to a separate spreadsheet (2011-01-loop-nuts.csv) and removed from the rosette spreadsheet with the final data which was saved as 2011-01-nuts.csv. Headers were changed to standard format. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. In a few cases the comment appeared to affect all sampling, so “ALL” was placed in front instead of “Nuts”. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-01oxy.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified by removing a few unnecessary columns or rows, “DO:” was entered in front of comments and the file was then saved as 2011-01-oxy.csv. That file was converted into individual *.ADD files.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. Comments and were included but no quality flags. The sheet from the analyst was saved as 2011-01-sal.csv. Loop samples were copied to a separate spreadsheet (2010-14-loop-sal.csv) and removed from the main spreadsheet. Some columns were removed and a column was added for the event number – the latter information was found in the log book. The station name and sample numbers were combined in a single column; those had to be separated. The data were ordered on event number and sample number.

The spreadsheet included duplicates which had not been averaged. Both values were removed to a separate sheet. The duplicates were averaged and calculations done to apply Chauvenet’s criterion for rejection of outliers. Then Sp (pooled standard deviation of pairs of samples) was calculated for the whole set and the set without outliers, using:

Sp = SQRT{sum (d*d)/2k}

Sp was found to be 0.0008 when the one pair was rejected that failed the Chauvenet test. There had been 2 Niskins fired at the depth from which those particular duplicates came, so the values from sample #31 were compared with sample #32; the differences between the duplicates of sample #31 and sample #32 were 0.03 and 0.003. So the value that differed by 0.03 was rejected and the other was used to replace the duplicates in the main spreadsheet for sample #31. In all other cases the average was used to replace the duplicates in the main spreadsheet together with flag “6”.  The spreadsheet was then converted to individual SAL files.

The SAL, CHL, ADD, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

Because this process was complex, all bottle data from the MRGCLN2 files were exported to a spreadsheet and checked against the rosette sheets to ensure no data were lost. A few discrepancies were investigated.

- Sample #328 is indicated on the rosette sheet but was not in the spreadsheet from the analyst. It is not listed on the Autosal Analysis log sheet either, so it is assumed to have not been taken or lost. It was added to the SAL file so it will be known that it was not missed in the processing.

- Sample #292 nutrient and chlorophyll samples were missing. This was a cast that had been restarted with another event number and some were labelled as event #46 and some as #47. All were changed to #47.

- Event #10 was missing entirely, again due to mislabelling of some files. This was a split cast and all files should be #10 though some data were logged as #11 and #12.

- DMS data were missing for cast #39 because they had been identified as being from cast #38. The file name was adjusted and the merge process rerun for that cast.

After those corrections the merge steps were repeated and the data exported to a spreadsheet to check all changes were done properly.

All the file names that were 2011-01-0047 were changed to 2011-01-0046 to match the downcast file and the rosette log entry.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When bottles below 200db are used and a few outliers (differences >0.004) are excluded the primary and secondary salinity were both high by about 0.001. The secondary fit was very flat with pressure. The primary had a little more pressure dependence and there are fewer data because the conductivity sensor data was lost from many casts in the early part of the cruise. This was the second cruise for which the new bottle washer was used and there were new liners for the bottles. The analysis was done within a month of the end of the cruise. Since the post-cruise calibration was used to derive the salinity we would expect the CTD to be very close to the bottles, and it is. The results are encouraging. If good sampling protocols are used and analysis is done quickly, we appear to be able to get good results.

The outliers were investigated and all were either associated with high standard deviation in the CTD data or were near the surface where the distance between sensor and rosette might well explain the differences. Samples previously flagged were checked and at most they were slight outliers. Notes were added to the header comments where appropriate.

For details see 2011-01-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

This was the first time that DO voltage had been recalibrated prior to COMPARE. As expected the fit of differences against DO concentration is quite flat and most differences are within ±0.07mL/L. No further recalibration is warranted.

The fits against pressure suggest some hysteresis below 2000db. Normally the fine-tuning of factors E, H1 and H3 in the DO calibration can reduce hysteresis. This sensor was never fine-tuned due to lack of sufficient deep-water data. Some tests will be run to see if this can be improved by adjusting the value of E.

The flagged values were checked in COMPARE and 2 were outliers and 2 looked ok. Notes were added to the headers about these results.

Other outliers were investigated. There were several from the Saanich Inlet cast; this is normal as this cast usually is very challenging for the sensor with a range of 0 to 6mL/L in 100m. Two had already been flagged by the analyst, 1 had a high standard deviation in the SBE DO channel, 1 was a very deep bottle so that is probably due to sensor hysteresis and 1 was close to the bottom where the gradient between bottle and sensor might be large. Sample #289 (7.474mL/L) from cast #44 does appear to be an outlier so was flagged “3”.

When the titrated DO data were reviewed after the cruise, the following was noted:

Precision on this cruise was excellent indicating sampling and analysis was carried out well. However, there was a problem with the standardizations. Only two sets of standards were run on the cruise and the two did not agree with each other, giving 2 different Thiosulfate normalities. A third set should have been run to confirm whether the change in Thio N was real or perceived. Given a two week cruise I would assume standards would be run at least 3 times even with the relatively few samples for this cruise, particularly when a large change is noted. Since this was not done there is no way of assessing whether the Thio did indeed change or whether the temperature shifted dramatically in the lab leading to this perceived change or whether there was just a problem with the standardization. As a result there is a potential shift in the DO values starting at P16.

So groups of casts were compared: Events 1-32 and 44-58. Both produce fairly flat fits against pressure with the first group having CTD look low by 0.012mL/L and the second group, high by 0.011mL/L for a net offset of ~0.02mL/L. This is very likely to be due to the shift noted in the comments above. The effect of this on recalibration of the SBE DO should be minimized by the use of all the titrated values in the fit that was used.

(See 2011-01-dox-comp1*.xls.).

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. The ratio of CTD Fluorescence to extracted CHL was high, but this is typical of low CHL conditions. CHL decreases as the distance from shore increases, while the ratio of CTD Fluorescence to Extracted CHL increases. (See 2011-01-chl-fluor-comp.xls.)

The merge process was repeated to capture the new flags and comments.

Once again data were exported to spreadsheet 2011-01-bottles.xls and compared to the rosette sheets and all expected data are present.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### DO sensor hysteresis study

At this stage tests were run on the DO sensor hysteresis parameters.

A deep cast was selected and converted using E=3.5 and E=3.7 to see if the deep values are more in line than those examined in COMPARE when E=3.6 was used. An initial test looking at the deepest bottle shows that increasing the value of E increases the DO concentration by roughly the right amount, so E=3.65 and E=3.7 were studied for 4 deep bottles.

The way to check for hysteresis is to compare bottles from below 2000db with some above 1000db with similar DO concentration. So from cast #58 we have the bottom 4 bottles with values around 2.6 to 3.3mL/L. The differences from bottles are ~-0.04

| Pressure | DO E=3.6 | SBE-Bot | E=3.65 | DO-Bottle | DO E=3.7 | SBE-bot |
| --- | --- | --- | --- | --- | --- | --- |
| 4308 | 3.187 | -0.060 | 3.212 | -0.035 | 3.237 | -0.010 |
| 4009 | 3.160 | -0.007 | 3.184 | 0.017 | 3.207 | 0.040 |
| 3500 | 2.974 | 0.002 | 2.994 | 0.022 | 3.012 | 0.040 |
| 3000 | 2.597 | 0.003 | 2.611 | 0.017 | 2.626 | 0.032 |

For comparison a few shallow bottles were found where concentrations were close to 3mL/L and they all had sensor DO slightly lower than bottles.

| Pressure | DO E=3.6 | DO-bot |
| --- | --- | --- |
| 200 | 3.179 | -0.011 |
| 150 | 3.088 | -0.030 |
| 250 | 3.112 | -0.008 |

E=3.7 looks like it works well at 4300db, but overdoes the correction above that. E=3.65 is not enough for the deepest bottles, but still looks a little overdone from 3000-4000. It may be that since the membrane is damaged, we cannot fix this level of hysteresis well enough to make a change. The errors are not large, but a warning can be put in the headers that values below 2500db may be low by up to 0.06mL/L.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-01-ctd-post-cal.con, except for cast #21 for which file 2011-01-ctd-post-cal-21.con was used.

A few casts were examined and all expected channels are present, though as mentioned in the sampling notes and log the primary conductivity and oxygen channels frequently had no signals. The two temperature channels are fairly close during the downcasts but the upcast traces differ much more, and at times they differ from the downcasts significantly. There are more spikes in the primary than secondary but both are very noisy during the upcasts.

The fluorescence and DO voltage look as expected with a dark value of about 0.02 for the former.

PAR looks fine. The altimetry looks ok where the CTD got close to the bottom. The transmissivity data look ok.

The downcast files for casts #10 and 11, and 47 and 48 will be joined after DELETE is run.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 3 casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge in these data because of noisy upcast temperatures and frequent stops for bottles, but 5s looks reasonable and is what was found appropriate for 2010-22 when this sensor was used.

ALIGNCTD was sued to advance the DO Voltage by 5s relative to the pressure.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using 4 casts. The goal is to make upcasts look closer to downcasts on a T-S surface. Most casts had stops for bottles and the upcasts were generally very noisy so it is difficult to judge which settings are best. The choice (α = 0.03, β=9) looked best overall for the both sensors. CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-01-0019 | 1000 1800 | +0.0003 +0.0004 | +0.00010 +0.00012 | +0.0007 ~0 | Noisy, High |
| 2011-01-0022 | 1000 1800 | +0.0003 +0.0003 | +0.00014 +0.00010 | +0.0014 +0.0012 | V. noisy, High |
| 2011-01-0026 | 1000 1800 | +0.0002 +0.0003 | +0.00013 +0.00010 | +0.0014 +0.0010 | X Noisy, High |

The differences are all small suggesting there was not much drift between the cruise and the post-cruise calibrations in March 2011.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. The bad primary conductivity and oxygen voltage do not have the normal pad values, producing large oxygen concentration values and negative primary salinity values.

CLEAN was run to add event numbers and to replace salinity values between 40 and 10000 with the normal pad value, -99.

A second run of CLEAN was used to put pad values in the SBE Dissolved Oxygen channel where values were between -1000 and 0. Some bad DO values remained that had values >10mL/L, so a 3rd run was done to replace with pad values any with DO>10mL/L and to reset the header limits.

There probably remain a few bad values in DO and there are some zero values in both the primary and secondary salinity. The latter should be corrected in the CTDEDIT phase, but plots will need to be examined carefully later to find the DO bad values.

##### Checking Headers

The header check was run.  There were no off-scale fluorescence values.

There are values of dissolved oxygen that are probably too high. Later in the processing it will be necessary to examine carefully the casts with some bad SBE DO data, to see if what remains is reliable.

Surface check was run and shows an average surface pressure for the cruise was 1.4db. This is a little low for the Tully, but the corresponding salinity values are very low, so this looks ok.

The cross-reference check was compared with the log book and the only problem was a format error in the station name for 2 casts; those were fixed in both bottle files and full profile files.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. Most casts did not have a header entry because they did not get within 15m of the bottom. There were 6 cases with a header entry, and the values appear to be correct.

The Water Depth header was also examined. There are deviations from the log book entries, but given there might have been some drift between the readings these were ignored unless the differences were large. The differences were large for events #18, 23, and 58. For the first of these the header entry is much lower than the usual depth of P6. For the second the depth is a little lower than the log and the usual depth for P10. For P26 the header depth was 4097, the log entry was 4252. The combination of maximum pressure plus altimeter reading is 4325db, which converts to a depth close to the log entry. So for all 3 cases the log entry in the CLN3 files was used to replace the header entry.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. As noted above few casts got within 15m of the bottom. Of the 6 casts with header entries all appeared to be correct. The same header water depth entry corrections were made to the MRGCLN2 files as were applied to the CLN3 files.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise there were few casts without stops and with a steady descent rate, and the upcast data were extremely noisy, so it is hard to make a judgment about the best setting. A few casts were examined for each fluorometer and the setting seems appropriate.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.2s was applied to the primary and a shift of -0.4s to the secondary conductivity.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concern casts #2 (a test) and #12 (upcast only), so of no concern.

At this point the two split casts were combined:

Files 10.DEL and 11.DEL were renamed 10.DELx and 10.DELy. There was a slight overlap in pressures so a few records were removed from the end of 10.DELx. The two files were then put through JOIN to produce 10.DEL. The output file was examined and looks fine.

The same process was applied to 46.DEL and 47.DEL with a few records removed from the end of the first file.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The post cruise-calibrations were used for conversion. The drifts noted at the factory were small, and would be negligible for the short time between February and March unless there had been some damage during this cruise. There is no evidence of that being the case.

2. Dissolved Oxygen

The DO sensor was repaired and recalibrated in April 2011. The membrane was found to be wrinkled due to impact at that time. No post-cruise calibration parameters are available.

3. Pressure

The post-cruise calibration parameters were used, and no significant drift is likely over 3 months, so no correction was applied.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursions were a few cases of salinity looking slightly high for a small section in the halocline; this does not look like evidence of instrumental problems, but a matter of a slightly shallower halocline than usual.

Repeat Casts –

There were two deep casts at each of stations P4 and P12. The first pair went to 1300db and the second to 2000db. The differences near the bottom for the first pair were <0.01C° and <0.001 salinity units for the first pair which occurred 2 hours apart. The differences for the second pair were <0.02C° and <0.002 salinity units and those casts were 6 hours apart. This shows good repeatability.

Post-Cruise Calibration

Post-cruise calibrations were available for the pressure, temperature and conductivity sensors. Those were used for the conversion of these data.

##### DETAILED EDITING

The secondary T and S sensors were chosen for archiving since the primary conductivity was heavily corrupted for many casts.

Note that for cast #21 the secondary are the better choice, but the secondary conductivity sensor is a different one than for the other casts.

The calibration parameters in the header of file 2011-01-0021.DEL were corrected so that the sensor numbers are those that were actually mounted.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as they appear corrupted by either ship effects or too short a wait at the top.

The descent rate of the CTD was extremely noisy for many of the casts resulting in heavy corruption due to shed wakes.

All EDU files were copied to EDT.

Another run of CTDEDIT was made to remove bad DO values from casts: 7, 10, 15, 17, 18, 19, 21 with output EDOX; those files were copied to EDT.

##### Initial Recalibration

Based on the results of COMPARE runs and other observations, there is no need to recalibrate the pressure, salinity and pressure. Dissolved oxygen requires no further recalibration to correct for sensor calibration drift.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure there is a lot of scatter above 500db. Below 2000db the CTD values are a little low. As noted in section 5 there is hysteresis for which no suitable correction was found. When the differences are plotted against DO concentration and outliers (including all data below 2000db) excluded, a flat fit is found with SBE DO values high by about 0.023mL/L.

File 2011-01-recal.ccf was prepared to subtract 0.023mL/L from the Oxygen:Dissolved:SBE channel.

This was first applied to the thinned files and COMPARE was rerun. The differences between 0 and 1000db look better, but below 1000db the SBE DO values look low by up to 0.1mL/L. A pressure-dependent recalibration could be attempted, but given the fact there was damage to this sensor and that it likely led to hysteresis problems, it seems unwise to put weight on these deep values. The data will be left in the file since the likely errors are <0.1mL/L and the profile shape may be useful. However, a warning will be entered in the headers.

CALIBRATE was run on the EDT files.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db, except for file #1 for which the fluorometer malfunctioned. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A problem was found in file #1. The steps from editing to the end were repeated for this file.

Lists were produced of files with and without PAR sensor channel.

Plots were made of those casts to ensure they were accurate and they were.

Plots of the SBE DO channel were examined carefully to find any bad values not removed earlier and one cast was put through CTDEDIT again, as well as CALIBRATE, FILTER and BIN AVERAGE.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

For cast #1, the channel Fluorescence:Seapoint was also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the CTD serial number, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

A shift in the DO standardization parameters occurred between events #32 and

#44; this was not investigated to determine if the resultant change in

Thiosulfate normality was real, or the result of a shift in lab temperature.

Checks of the fit of SBE DO against titrated DO suggest that the titrated

values from the first part of the cruise are lower than those of the latter

part by about 0.02mL/L. The SBE DO was recalibrated using titrated DO data

from the whole cruise, so any resultant error in the SBE DO is likely

negligible compared to other sources of error (<0.01mL/L).

The Oxygen:Dissolved:SBE channel was corrupted for all or parts of

Events #7, 10, 15, 17, 18, 19 and 21.

WARNING: The dissolved oxygen sensor was found to have a damaged membrane

when it was examined at the factory shortly after this cruise. There is evidence of

hysteresis which could not be corrected by adjusting the calibration parameters.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4mL/L from 0 to 125db

±0.3mL/L from 125db to 400db

±0.1mL/L below 400db

Below 1000db the SBE DO reads consistently low by <0.1mL/L.

For details on the processing see processing report: 2011-01-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were ~85% in Saanich Inlet. A check of the bottle files shows that the SBE values were a little higher than bottles, not lower, so this is not due to an underestimate by the CTD sensor. For all other casts the saturation was between 95% and 102%, with most between 100% and 102%.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was used on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was used on casts without a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, PAR, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

A second REMOVE was run on casts #7, 15, 17, 18 and 19 to remove the 2 Oxygen:Dissolved:SBE channels because all the SBE DO data in those files was corrupted.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data..

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no discrepancies were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no outliers were identified.

Standards check was run on all files and a format error was found; that was fixed and no further errors were found.

##### 25. Thermosalinograph Data

Data were provided in 5 hex files. There was loop sampling for Salinity, Chlorophyll and Nutrient samples; those were combined in file 2011-01-tsg-loop-comp.xls. Times were added to the spreadsheet based on log records. The log indicates there was no sampling for Loop 14, but there are samples for all except CHL, so it is assumed the comment was just referring to CHL.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong date and scale factor. After that correction the CON file was saved as 2011-01-tsg.con.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. Overall the records look good, but a few problems were noted:

- Flow interruption – The flow rate varies occasionally. is mostly steady at ~1 but it was higher during file #1 which was very short and dropped to zero for about 10 hours during file #4.

- The salinity data look oddly flat but that is likely due to some zero values in the 4th and 5th files leading to a very large range of values in the plots. Setting a manual range for those files produces results that seem ok.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2011-01-ctd-tsg-comp.xls. All the data came from ~4.5db.

There is no overlap between TSG file #5 and CTD casts.

The 4 TSG files that overlap CTD casts were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 33 matches since there was no overlap for cast #34. Those data were added to 2011-01-ctd-tsg-comp.xls. TSG values were found in the same way for times of underway loop sampling and added to file 2011-01-tsg-loop-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.0004° and in longitude <0.0007°. The median differences were 0.00002° and 0.00004°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. The differences between the two temperatures were mostly between 0.25 and 0.3Cº, but the traces are very noisy.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

When all data were included the TSG intake temperature was higher than the CTD by 0.008Cº while the median difference was 0.018Cº. When the differences and standard deviations in the TSG temperature were plotted against longitude, it is clear that the outliers in both are close to shore. When only samples from west of 127W are included, the average and median differences are both 0.017Cº. The waters were very well mixed offshore, so this is likely a good comparison, although the two most westerly casts show differences of only 0.005, so there may be other factors such as ship noise during stops and noisy descent rates affecting the comparison.  The CTD temperature came from a post-cruise calibration, so is more likely accurate than the TSG thermistor, but since we cannot be sure they were measuring the same water this evidence is probably not sufficiently clear to justify a recalibration.

The TSG salinity is lower than the CTD by an average of 0.094 and a median of 0.079. The range of differences was from -0.059 to 0.408. When only casts west of 127º are included the average is -0.080 and the median is -0.078. A plot made of differences versus standard deviation in TSG salinity shows values approaching a value close to the median.

The ratio of TSG fluorescence to CTD fluorescence ranges from 1.2 to 3.4 and a median of 1.5. When plotted together the patterns are similar. The ratio is close to 1.5 when the CTD Fluorescence >1, but for <1 the ratio increases as CTD fluorescence decreases. (See 2011-01-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons  A comparison was done in spreadsheet 2011-01-tsg-loop-comp.xls of loop salinity and chlorophyll samples with TSG salinity and fluorescence from the same time. TSG salinity was lower than bottle salinity by between an average of 0.083 and a median of 0.079. When 3 outliers are excluded the average is -0.079. One of the outliers had the salinity bottle flagged.  The TSG fluorescence ranged from 2.5 to 8.5 times the extracted chlorophyll. When the ratio of TSG Fluorescence to Loop CHL is plotted against Loop CHL there is a clear inverse relationship. The performance of the TSG fluorometer looks a little better than seen in August and early Sept. 2010 (2010-14 and 2010-15) when some very high ratios were seen, and there was some concern about contamination of the loop. During the mid-August cruise, 2010-69, the problem appeared to have been resolved and this cruise confirms that. (See file 2011-01-loop-TSG-comp.xls.).

- Loop Bottle - Rosette Comparisons

There were no cases of loop bottles being taken during a CTD cast.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and were used for 2009-10 (with a different intake thermistor), 2009-11, 2010-01 2010-13, 2010-22, 2010-14 and 2010-15 and 2010-69 (and 2 other cruises with clock problems so that data were not useful for recalibration). For 2010-14 the TSG temperature was lower than the CTD by 0.018C The salinity was low by about 0.02, 0.02, 0.06/0.16, 0.06, 0.06 up to June 2010 with the change for 2010-01 being associated with a change of flow rate. During 2010-36 the differences started at 0.06 but were at ~0.4 by the end of the cruise, likely due to a fouling problem. For the last 3 cruises the difference was found to be +0.11 but the flow rate was higher than usual.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was steady except for a high section at the beginning of file #1 and 2 and interruptions in file #4.

3. The temperature in the loop increases by an average of 0.27Cº between intake and lab, with little variation, which is as expected given the narrow range of intake temperatures.

4. The TSG intake temperature appears to be higher than the CTD by no more than 0.02C, based on median values, but there is a lot of variability. The CTD temperature conversion was done using post-cruise calibrations making recalibration of TSG temperature tempting. But there is sufficient noise in the comparison that this looks unwise.

5. The TSG Salinity is lower than the CTD salinity by a median value of ~0.08 with little variation between from P14 to P26. Salinity is lower than loop salinity bottles by 0.079. These two comparisons are remarkably close. Using a correction of +0.08 for salinity looks appropriate.

6. The TSG fluorometer had values higher than the CTD fluorometer by a median factor of 1.5 (range of ~1.2 to ~3.4) and was higher than the loop CHL samples by a factor of 2.5 to 8.5. This is consistent with the results of the CTD fluorometer comparison to extracted CHL. These ratios confirm that the loop problems have been resolved that led to some extremely high fluorescence values relative to CHL in the cruises of August and September 2010. Nonetheless, they are still too high and recalibration of the TSG sensor is long overdue.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. Because only 3 channels can be displayed, the files were opened several times to remove data where the flow rate was zero – position and time channels need to be kept, so simple removal of records won’t work. Data were removed form the beginning of file #1 and from 2 sections of file #4. One salinity point was also removed from file #1.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

File 2011-01-tsg-recal1.ccf was prepared to adjust salinity by adding 0.08. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; the track is odd, but is supported by the log.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

The final loop file 2011-01-loop.csv was prepared by the chief scientist including data from the final CTD files and samples from the loop or from 5m bottles. That spreadsheet was simplified, date calculated in DD/MM/YYYY format. A 6-line header was added and the file was saved as 2011-01-loop-6linehdr.csv. It was converted to IOS format, put through CLEAN and HEADEDIT to get start and stop times and positions, and to add general comments and specific comments for flagged values. The final file was named 2011-01-surface.loop. (One comment had to be truncated to enable conversion, but it was restored in the final file.) A track plot looks reasonable and a plot of salinity versus date looks right.

Particulars (mostly from logs and sample notes)

1&2 – 2 is upcast for 1. Niskins 1 and 2 closed on way down by accident. Niskins #12 and 13 were fired at bottom-10m and 200m, resp. and then 3 to 11 at 175 and up. Bottles 14-24 were fired just to test.

1. Fluorometer very low on downcast.

3. down/up very different

7. Primary salinity and oxygen lost at 544db downwards and for upcast. Winch lost power briefly at 370db.

8. Winch lost power at 175, 100, 40, 15db of upcast

10-11-12. Problems with primary salinity and oxygen. Stop at 505db down, long wait then start file 11 at 506db until freeze at 100db upcast, start file 12 at 101upcast.

13. Problems with primary sal and DO – sometimes present sometimes not.

17. BL and HDR files called 16 – should be 17.

18. Continuing problems with sal/DO. Problems with fuzzy transmissivity – continued through rest of cruise, sometimes also spiky.

19. Continuing problems with sal/DO.

21. Switched conductivity sensors but not in con file but did not help with salinity/DO problem.

22. Back to original conductivity sensor arrangement. Bottles 1-5 for Keith were not sampled – should be dropped from file. Winch stopped briefly at 1874db on upcast. Problems with fluorometer spikes – that continued through rest of cruise at least until file #44.

23. Sample numbers out of order because reran 5 bottles for Keith.

24. 29kt winds.

31. Sample #s out of order, but no IOS samples.

32. Niskins 3 and 4 had lanyards crossed – potential contamination although oxygen value looks ok. Niskin 24 at 1000db closed out of order and sample number out of sequence.

34. Sample taken for Mike, no value but has a sample number – to be included in file. TSG turned off for valve repair.

37. Intended as rosette cast but only 2 bottles closed and no samples. Cast 38 is a rerun.

39. Retake of 36

47/48. Stop at 307db down, next file starts at 308 down.

57. Cast run 20milese west of P20.

58. 30kt winds, “bumpy”

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory | 18Mar11 | Factory |
| Conductivity | 1763 | 11Oct07 | Factory | 16Mar11 | Factory |
| Secondary Temp. | 2038 | 06May08 | Factory | 16Mar11 | Factory |
| Secondary Cond. | 3394 | 06Mar09 | Factory | 16Mar11 | Factory |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory | 4Aug11 | Factory |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Fluorometer | 2345 |  | IOS |  |  |
| Pressure Sensor | 69698 | 26May2006 | Factory | 15Apr11 | Factory |
| Altimeter | 1024 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2011-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 2416 | 23Dec06 |  |  |  |