## Notes for Germaine
Two sets of standards were run during the cruise, one at the beginning and one about half way through.
The calculated THIO normalities based on these 2 standardizations did not agree to within specs (0.0005 N)
A third set of standards should have been run to confirm a new normality, this was not done by the analyst
As a result there is no way of assessing whether the shift in THIO normality was real or perceived (the result of a poor standardization or change in temp)
Thus, although precision for this cruise was very good, accuracy is in question since proper analysis protocols were not followed.
There is potentially an offset in values from Event #44 station P16 onwards
If we can figure out which set looks better, they can potentially be re-adjusted, otherwise this will have to go into the header for the whole cruise

## 2011-01 QF_FINAL
OXY HEADER: | Oxygen samples were analyzed at sea using an automated Winkler titration system (Brinkmann Dosimat model 765 and a PC 950 Colorimeter controlled by AutoOxy software V3.4)
with modifications based on Carpenter (1965) and adhering to WOCE protocols (Culberson 1991). For details including a duplicate analysis, see document 2011-01oxy.xls.
**See Notes for potential data problesm that may need to be in Header, or cruise comments***
Analysis: | Hugh Maclean at sea using Kit #2
Cruise_Number | EVENT_NUMBER | Sample_number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | comments
2011-01 | 1 | 1 | 0
2011-01 | 1 | 2 | 0
2011-01 | 1 | 3 | 0
2011-01 | 1 | 4 | 0.0029732
2011-01 | 1 | 5 | 0.050013
2011-01 | 1 | 6 | 0.42992
2011-01 | 1 | 7 | 4.0464 | 6
2011-01 | 1 | 8 | 4.7827
2011-01 | 1 | 9 | 4.7434
2011-01 | 1 | 10 | 5.386
2011-01 | 1 | 11 | 5.8794
2011-01 | 6 | 29 | 3.0288
2011-01 | 6 | 30 | 3.2042
2011-01 | 6 | 31 | 5.2485 | 6
2011-01 | 6 | 32 | 5.1718
2011-01 | 6 | 33 | 6.5902
2011-01 | 6 | 34 | 6.5485
2011-01 | 6 | 35 | 6.9372
2011-01 | 6 | 36 | 6.9543
2011-01 | 10 | 56 | 0.47269 | 3 | operator error, colorimeter not properly zeroed; outlier in comparison with SBE sensor
2011-01 | 10 | 57 | 0.42284
2011-01 | 10 | 58 | 0.42869 | 6
2011-01 | 10 | 59 | 0.26464
2011-01 | 10 | 60 | 0.2204
2011-01 | 10 | 61 | 0.40685
2011-01 | 10 | 62 | 0.91634
2011-01 | 10 | 63 | 1.3769
2011-01 | 10 | 64 | 1.6053
2011-01 | 10 | 65 | 2.0862
2011-01 | 10 | 66 | 2.2843
2011-01 | 10 | 67 | 2.4588
2011-01 | 10 | 68 | 2.6623
2011-01 | 10 | 69 | 3.9137
2011-01 | 10 | 70 | 5.2618
2011-01 | 10 | 71 | 6.5963
2011-01 | 10 | 72 | 6.6851 | 6
2011-01 | 10 | 73 | 6.808
2011-01 | 10 | 74 | 6.8428
2011-01 | 21 | 93 | 1.5937 | 3 | Niskin valve not tight; outlier in comparison with SBE sensor
2011-01 | 21 | 93 | 1.5333
2011-01 | 21 | 94 | 0.75612
2011-01 | 21 | 95 | 0.49795
2011-01 | 21 | 96 | 0.28896
2011-01 | 21 | 97 | 0.23189 | 6
2011-01 | 21 | 98 | 0.31768
2011-01 | 21 | 100 | 0.93746
2011-01 | 21 | 101 | 1.7193
2011-01 | 21 | 102 | 2.2938
2011-01 | 21 | 103 | 2.8802
2011-01 | 21 | 104 | 3.0593
2011-01 | 21 | 105 | 3.3153
2011-01 | 21 | 106 | 3.6812
2011-01 | 21 | 107 | 4.3068
2011-01 | 21 | 108 | 6.765
2011-01 | 21 | 109 | 6.806
2011-01 | 21 | 110 | 6.8193 | 6
2011-01 | 21 | 111 | 6.8124
2011-01 | 21 | 112 | 6.8174
2011-01 | 32 | 168 | 2.2892
2011-01 | 32 | 169 | 2.2735
2011-01 | 32 | 170 | 2.278
2011-01 | 32 | 171 | 1.8832
2011-01 | 32 | 172 | 1.4389
2011-01 | 32 | 173 | 0.67633
2011-01 | 32 | 174 | 0.37668
2011-01 | 32 | 175 | 0.2617
2011-01 | 32 | 176 | 0.33363 | 6
2011-01 | 32 | 177 | 0.46717
2011-01 | 32 | 178 | 1.2926
2011-01 | 32 | 179 | 1.9004
2011-01 | 32 | 180 | 2.1416
2011-01 | 32 | 181 | 2.6955
2011-01 | 32 | 182 | 2.9866
2011-01 | 32 | 183 | 3.147
2011-01 | 32 | 184 | 3.5353 | 6
2011-01 | 32 | 185 | 4.1369
2011-01 | 32 | 186 | 5.905
2011-01 | 32 | 187 | 6.8303
2011-01 | 32 | 188 | 6.8598
2011-01 | 32 | 189 | 6.858
2011-01 | 32 | 190 | 6.8611
2011-01 | 44 | 267 | 2.7484
2011-01 | 44 | 268 | 2.6949
2011-01 | 44 | 269 | 2.7117
2011-01 | 44 | 270 | 2.3475 | 6
2011-01 | 44 | 271 | 1.8641
2011-01 | 44 | 272 | 1.3191
2011-01 | 44 | 273 | 0.56771
2011-01 | 44 | 274 | 0.31323
2011-01 | 44 | 275 | 0.28523
2011-01 | 44 | 276 | 0.40206 | 6
2011-01 | 44 | 277 | 0.60803
2011-01 | 44 | 278 | 1.4769
2011-01 | 44 | 279 | 2.4333
2011-01 | 44 | 280 | 3.1202
2011-01 | 44 | 281 | 3.8819
2011-01 | 44 | 282 | 4.1828
2011-01 | 44 | 283 | 4.4529
2011-01 | 44 | 284 | 5.1165
2011-01 | 44 | 285 | 6.3188
2011-01 | 44 | 286 | 6.7551 | 6
2011-01 | 44 | 287 | 6.7722
2011-01 | 44 | 288 | 6.7783 | 3 | suspect open air vent but value looks ok in comparison to SBE DO sensor.
2011-01 | 44 | 289 | 7.4737 | 3 | outlier in comparison to SBE DO sensor with no known reason
2011-01 | 44 | 290 | 6.7743
2011-01 | 55 | 319 | 2.9841
2011-01 | 55 | 320 | 2.831
2011-01 | 55 | 321 | 2.845
2011-01 | 55 | 322 | 2.5436 | 6
2011-01 | 55 | 323 | 1.9527
2011-01 | 55 | 324 | 1.2493
2011-01 | 55 | 325 | 0.53814
2011-01 | 55 | 326 | 0.36202
2011-01 | 55 | 327 | 0.29262 | 6
2011-01 | 55 | 328 | 0.37972
2011-01 | 55 | 329 | 0.49359
2011-01 | 55 | 330 | -99 | 5 | flask broke
2011-01 | 55 | 331 | 2.0033
2011-01 | 55 | 332 | 2.559
2011-01 | 55 | 333 | 3.1897
2011-01 | 55 | 334 | 3.8052
2011-01 | 55 | 335 | 4.2588
2011-01 | 55 | 336 | 4.8505
2011-01 | 55 | 337 | 6.7911 | 6
2011-01 | 55 | 338 | 6.8139
2011-01 | 55 | 339 | 6.8085
2011-01 | 55 | 340 | 6.8126
2011-01 | 55 | 341 | 6.8066
2011-01 | 55 | 342 | 6.812
2011-01 | 58 | 387 | 3.2472
2011-01 | 58 | 388 | 3.1666 | 6
2011-01 | 58 | 389 | 2.9715
2011-01 | 58 | 390 | 2.594
2011-01 | 58 | 391 | 2.0394
2011-01 | 58 | 392 | 1.3601
2011-01 | 58 | 393 | 0.62248
2011-01 | 58 | 394 | 0.39527
2011-01 | 58 | 395 | 0.36322 | 6
2011-01 | 58 | 396 | 0.40371 |  | spigot drip; value looks ok in comparison to SBE DO sensor.
2011-01 | 58 | 397 | 0.44028
2011-01 | 58 | 398 | 0.77993
2011-01 | 58 | 399 | 1.1496
2011-01 | 58 | 400 | 1.3344
2011-01 | 58 | 401 | 1.9066
2011-01 | 58 | 402 | 2.3173
2011-01 | 58 | 403 | 3.1185
2011-01 | 58 | 404 | 4.158
2011-01 | 58 | 405 | 6.5179
2011-01 | 58 | 406 | 6.9748 | 6
2011-01 | 58 | 407 | 6.9827
2011-01 | 58 | 408 | 6.9805
2011-01 | 58 | 409 | 6.989
2011-01 | 58 | 410 | 6.988

## raw data
File Name | LOC:EVENT_NUMBER | LOC:LATITUDE | LOC:LONGITUDE | STN:STATION | Sample_number | Depth | Oxygen:Dissolved | Thiosulfate_titer | Flask_Number | Flask_Volume | Blank | Standard | KIO3 Std volume ml | KIO3 Normality | DO final (ml/l) | flag | comments
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 1 | 215 | 0.014 | 0.006 | 2056 | 142.1 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 2 | 200 | 0.045 | 0.011 | 1386 | 134.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 3 | 175 | 0.02 | 0.007 | 3061 | 143.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 4 | 150 | 0.003 | 0.004 | 3178 | 143.12 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.0029732
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 5 | 125 | 0.05 | 0.012 | 1201 | 135.96 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.050013
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 6 | 100 | 0.43 | 0.075 | 1414 | 132.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.42992
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 7 | 75 | 4.041 | 0.685 | 1327 | 134.41 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.0407
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 7 | 75 | 4.052 | 0.724 | 3107 | 141.59 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.0522
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 8 | 50 | 4.783 | 0.821 | 1326 | 136.19 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.7827
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 9 | 25 | 4.743 | 0.808 | 1335 | 135.15 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.7434
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 10 | 10 | 5.386 | 0.955 | 570 | 140.69 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.386
2011-01-0001.oxy | 1 | 48.5928 | -123.4998 | SI03 | 11 | 5 | 5.88 | 1.043 | 3118 | 140.8 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.8794
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 29 | 106 | 3.029 | 0.544 | 2056 | 142.1 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.0288
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 30 | 100 | 3.204 | 0.544 | 1386 | 134.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.2042
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 31 | 75 | 5.254 | 0.948 | 3061 | 143.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.2539 | 3 | messy endpoint
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 31 | 75 | 5.243 | 0.946 | 3178 | 143.12 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.2432
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 32 | 75 | 5.172 | 0.886 | 1201 | 135.96 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.1718
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 33 | 50 | 6.59 | 1.1 | 1414 | 132.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.5902
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 34 | 25 | 6.548 | 1.108 | 1327 | 134.41 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.5485
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 35 | 10 | 6.937 | 1.237 | 3107 | 141.59 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.9372
2011-01-0006.oxy | 6 | 48.6002 | -126.001 | P2 | 36 | 5 | 6.954 | 1.183 | 1335 | 135.15 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.9543
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 56 | 1300 | 0.473 | 0.084 | 559 | 135.75 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.47269 | 3 | operator error, colorimeter not properly zeroed
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 57 | 1250 | 0.423 | 0.08 | 2088 | 144.09 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.42284
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 58 | 1250 | 0.428 | 0.079 | 1491 | 140.64 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.4277
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 58 | 1250 | 0.43 | 0.079 | 3094 | 140 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.42969
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 59 | 1000 | 0.265 | 0.047 | 1487 | 131.14 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.26464
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 60 | 800 | 0.22 | 0.04 | 1465 | 132.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.2204
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 61 | 600 | 0.407 | 0.073 | 632 | 136.17 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.40685
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 62 | 400 | 0.916 | 0.155 | 552 | 131.82 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.91634
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 63 | 300 | 1.377 | 0.251 | 1175 | 143.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.3769
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 64 | 250 | 1.605 | 0.285 | 1496 | 139.68 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.6053
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 65 | 200 | 2.086 | 0.345 | 1158 | 130.52 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.0862
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 66 | 175 | 2.284 | 0.383 | 1483 | 132.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.2843
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 67 | 150 | 2.459 | 0.417 | 1306 | 134.03 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.4588
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 68 | 125 | 2.662 | 0.456 | 1368 | 135.44 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.6623
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 69 | 100 | 3.914 | 0.66 | 620 | 133.69 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.9137
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 70 | 75 | 5.262 | 0.885 | 1333 | 133.52 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.2618
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 71 | 50 | 6.596 | 1.059 | 1105 | 127.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.5963
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 72 | 25 | 6.685 | 1.112 | 1085 | 132.17 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.6853
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 72 | 25 | 6.685 | 1.108 | 611 | 131.71 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.6848
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 73 | 10 | 6.808 | 1.096 | 1126 | 127.98 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.808
2011-01-0010.oxy | 10 | 48.6512 | -126.6638 | P4 | 74 | 5 | 6.843 | 1.132 | 1206 | 131.47 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8428
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 93 | 2000 | 1.594 | 0.275 | 559 | 135.75 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.5937 | 3 | air valve not tight
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 93 | 2000 | 1.533 | 0.281 | 2088 | 144.09 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.5333
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 94 | 1500 | 0.756 | 0.137 | 1491 | 140.64 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.75612
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 95 | 1250 | 0.498 | 0.091 | 3094 | 140 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.49795
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 96 | 1000 | 0.289 | 0.051 | 1487 | 131.14 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.28896
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 97 | 800 | 0.232 | 0.042 | 1465 | 132.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.23246
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 97 | 800 | 0.231 | 0.043 | 632 | 136.17 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.23132
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 98 | 600 | 0.318 | 0.056 | 552 | 131.82 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.31768
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 100 | 400 | 0.937 | 0.172 | 1175 | 143.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.93746
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 101 | 300 | 1.719 | 0.305 | 1496 | 139.68 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.7193
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 102 | 250 | 2.294 | 0.379 | 1158 | 130.52 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.2938
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 103 | 200 | 2.88 | 0.482 | 1483 | 132.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.8802
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 104 | 175 | 3.059 | 0.518 | 1306 | 134.03 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.0593
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 105 | 150 | 3.315 | 0.567 | 1368 | 135.44 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.3153
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 106 | 125 | 3.681 | 0.621 | 620 | 133.69 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.6812
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 107 | 100 | 4.307 | 0.725 | 1333 | 133.52 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.3068
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 108 | 75 | 6.765 | 1.086 | 1105 | 127.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.765
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 109 | 50 | 6.806 | 1.132 | 1085 | 132.17 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.806
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 110 | 25 | 6.812 | 1.129 | 611 | 131.71 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8119
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 110 | 25 | 6.827 | 1.099 | 1126 | 127.98 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8267
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 111 | 10 | 6.813 | 1.127 | 1206 | 131.47 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8124
2011-01-0021.oxy | 21 | 48.8173 | -128.6752 | P8 | 112 | 5 | 6.817 | 1.149 | 510 | 133.91 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8174
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 168 | 3000 | 2.289 | 0.412 | 2056 | 142.1 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.2892
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 169 | 3000 | 2.273 | 0.387 | 1386 | 134.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.2735
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 170 | 3000 | 2.278 | 0.413 | 3061 | 143.13 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.278
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 171 | 2500 | 1.883 | 0.342 | 3178 | 143.12 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.8832
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 172 | 2000 | 1.439 | 0.249 | 1201 | 135.96 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.4389
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 173 | 1500 | 0.676 | 0.116 | 1414 | 132.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.67633
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 174 | 1250 | 0.377 | 0.067 | 1327 | 134.41 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.37668
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 175 | 1000 | 0.262 | 0.05 | 3107 | 141.59 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.2617
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 176 | 800 | 0.333 | 0.06 | 1335 | 135.15 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.33332
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 176 | 800 | 0.334 | 0.061 | 1134 | 137.25 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.33395
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 177 | 600 | 0.467 | 0.086 | 570 | 140.69 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 0.46717
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 178 | 400 | 1.293 | 0.232 | 3118 | 140.8 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.2926
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 179 | 300 | 1.9 | 0.347 | 2060 | 143.91 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 1.9004
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 180 | 250 | 2.142 | 0.381 | 3135 | 140.39 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.1416
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 181 | 200 | 2.696 | 0.453 | 1459 | 132.92 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.6955
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 182 | 175 | 2.987 | 0.513 | 612 | 135.93 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 2.9866
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 183 | 150 | 3.147 | 0.561 | 3116 | 141.08 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.147
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 184 | 125 | 3.538 | 0.634 | 3164 | 141.92 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.5376
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 184 | 125 | 3.533 | 0.659 | 1460 | 147.66 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 3.533
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 185 | 100 | 4.137 | 0.686 | 2000 | 131.52 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 4.1369
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 186 | 75 | 5.905 | 0.986 | 1302 | 132.62 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 5.905
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 187 | 50 | 6.831 | 1.234 | 1260 | 143.43 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8303
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 188 | 25 | 6.86 | 1.164 | 2075 | 134.81 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8598
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 189 | 10 | 6.858 | 1.163 | 3194 | 134.73 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.858
2011-01-0032.oxy | 32 | 48.969 | -130.674 | P12 | 190 | 5 | 6.861 | 1.226 | 2027 | 141.88 | 0.0013 | 0.7208 | 10.0158 | 0.010074 | 6.8611
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 267 | 3692 | 2.748 | 0.497 | 2056 | 142.1 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.7484
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 268 | 3500 | 2.695 | 0.461 | 1386 | 134.43 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.6949
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 269 | 3500 | 2.712 | 0.494 | 3061 | 143.13 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.7117
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 270 | 3000 | 2.346 | 0.428 | 3178 | 143.12 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.3461
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 270 | 3000 | 2.349 | 0.407 | 1201 | 135.96 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.3489
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 271 | 2500 | 1.864 | 0.316 | 1414 | 132.62 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.8641
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 272 | 2000 | 1.319 | 0.228 | 1327 | 134.41 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.3191
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 273 | 1500 | 0.568 | 0.106 | 3107 | 141.59 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.56771
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 274 | 1250 | 0.313 | 0.058 | 1335 | 135.15 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.31323
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 275 | 1000 | 0.285 | 0.054 | 1134 | 137.25 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.28523
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 276 | 800 | 0.402 | 0.076 | 570 | 140.69 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.40222
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 276 | 800 | 0.402 | 0.076 | 3118 | 140.8 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.4019
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 277 | 600 | 0.608 | 0.115 | 2060 | 143.91 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.60803
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 278 | 400 | 1.477 | 0.266 | 3135 | 140.39 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.4769
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 279 | 300 | 2.433 | 0.412 | 1459 | 132.92 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.4333
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 280 | 250 | 3.12 | 0.539 | 612 | 135.93 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.1202
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 281 | 200 | 3.882 | 0.695 | 3116 | 141.08 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.8819
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 282 | 175 | 4.183 | 0.753 | 3164 | 141.92 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 4.1828
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 283 | 150 | 4.453 | 0.834 | 1460 | 147.66 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 4.4529
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 284 | 125 | 5.117 | 0.852 | 2000 | 131.52 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 5.1165
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 285 | 100 | 6.319 | 1.06 | 1302 | 132.62 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.3188
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 286 | 75 | 6.754 | 1.226 | 1260 | 143.43 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7538
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 286 | 75 | 6.756 | 1.152 | 2075 | 134.81 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7564
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 287 | 50 | 6.772 | 1.154 | 3194 | 134.73 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7722
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 288 | 25 | 6.778 | 1.217 | 2027 | 141.88 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7783 | 3 | suspect open air vent
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 289 | 10 | 7.474 | 1.287 | 1326 | 136.19 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 7.4737
2011-01-0044.oxy | 44 | 49.2595 | -134.6665 | P16 | 290 | 5 | 6.774 | 1.22 | 1142 | 142.31 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7743
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 319 | 4024 | 2.984 | 0.515 | 559 | 135.75 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.9841
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 320 | 3500 | 2.831 | 0.519 | 2088 | 144.09 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.831
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 321 | 3500 | 2.845 | 0.509 | 1491 | 140.64 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.845
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 322 | 3000 | 2.53 | 0.451 | 3094 | 140 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.5295
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 322 | 3000 | 2.558 | 0.427 | 1487 | 131.14 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.5577
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 323 | 2500 | 1.953 | 0.332 | 1485 | 133.1 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.9527
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 324 | 2000 | 1.249 | 0.219 | 632 | 136.17 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.2493
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 325 | 1500 | 0.538 | 0.094 | 552 | 131.82 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.53814
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 326 | 1250 | 0.362 | 0.07 | 1175 | 143.13 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.36202
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 327 | 1000 | 0.297 | 0.057 | 1496 | 139.68 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.29724
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 327 | 1000 | 0.288 | 0.052 | 1158 | 130.52 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.288
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 328 | 800 | 0.38 | 0.068 | 1483 | 132.43 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.37972
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 329 | 600 | 0.494 | 0.088 | 1306 | 134.03 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.49359
|  |  |  | P20 | 330 | [9 empty] | -99 | 5 | flask broke
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 331 | 300 | 2.003 | 0.342 | 620 | 133.69 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.0033
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 332 | 250 | 2.559 | 0.435 | 1333 | 133.52 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.559
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 333 | 200 | 3.19 | 0.517 | 1105 | 127.62 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.1897
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 334 | 175 | 3.805 | 0.638 | 1085 | 132.17 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.8052
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 335 | 150 | 4.259 | 0.711 | 611 | 131.71 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 4.2588
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 336 | 125 | 4.85 | 0.786 | 1126 | 127.98 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 4.8505
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 337 | 100 | 6.786 | 1.128 | 1206 | 131.47 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7857
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 337 | 100 | 6.796 | 1.151 | 510 | 133.91 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.7966
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 338 | 75 | 6.814 | 1.242 | 1218 | 144.02 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.8139
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 339 | 50 | 6.809 | 1.143 | 396 | 132.76 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.8085
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 340 | 25 | 6.813 | 1.126 | 505 | 130.73 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.8126
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 341 | 10 | 6.807 | 1.101 | 1301 | 127.97 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.8066
2011-01-0055.oxy | 55 | 49.5675 | -138.668 | P20 | 342 | 5 | 6.812 | 1.159 | 1374 | 134.53 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.812
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 387 | 4309 | 3.247 | 0.56 | 559 | 135.75 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.2472
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 388 | 4000 | 3.167 | 0.58 | 2088 | 144.09 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.1667
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 388 | 4000 | 3.167 | 0.566 | 1491 | 140.64 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.1665
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 389 | 3500 | 2.972 | 0.529 | 3094 | 140 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.9715
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 390 | 3000 | 2.594 | 0.433 | 1487 | 131.14 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.594
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 391 | 2500 | 2.039 | 0.344 | 1465 | 132.13 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.0394
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 392 | 2000 | 1.36 | 0.238 | 632 | 136.17 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.3601
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 393 | 1500 | 0.622 | 0.108 | 552 | 131.82 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.62248
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 394 | 1250 | 0.395 | 0.076 | 1175 | 143.13 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.39527
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 395 | 1000 | 0.365 | 0.069 | 1496 | 139.68 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.36541
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 395 | 1000 | 0.361 | 0.064 | 1158 | 130.52 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.36102
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 396 | 800 | 0.404 | 0.072 | 1483 | 132.43 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.40371 | 3 | spigot drip
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 397 | 600 | 0.44 | 0.079 | 1306 | 134.03 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.44028
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 398 | 400 | 0.78 | 0.136 | 620 | 133.69 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 0.77993
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 399 | 300 | 1.15 | 0.198 | 1333 | 133.52 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.1496
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 400 | 250 | 1.334 | 0.219 | 1105 | 127.62 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.3344
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 401 | 200 | 1.907 | 0.322 | 1085 | 132.17 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 1.9066
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 402 | 175 | 2.317 | 0.389 | 611 | 131.71 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 2.3173
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 403 | 150 | 3.118 | 0.507 | 1126 | 127.98 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 3.1185
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 404 | 125 | 4.158 | 0.693 | 1206 | 131.47 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 4.158
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 405 | 100 | 6.518 | 1.104 | 510 | 133.91 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.5179
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 406 | 75 | 6.974 | 1.271 | 1218 | 144.02 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.9736
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 406 | 75 | 6.976 | 1.171 | 396 | 132.76 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.976
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 407 | 50 | 6.983 | 1.154 | 505 | 130.73 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.9827
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 408 | 25 | 6.98 | 1.129 | 1301 | 127.97 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.9805
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 409 | 10 | 6.989 | 1.189 | 1374 | 134.53 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.989
2011-01-0058.oxy | 58 | 49.9953 | -144.9997 | P26 | 410 | 5 | 6.988 | 1.174 | 3157 | 132.87 | 0.0025 | 0.7247 | 10.0158 | 0.010074 | 6.988

## duplicates
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded
Duplicate samples from a single bottle
Event # | Sample# | DO1 | DO2 | Flask1 | Flask2 | Average | Difference | Diff^2 | Chauvenet | reject? | Critical values for Chauvenet statistic
1 | 7 | 4.0407 | 4.0522 |  |  | 4.0464 | 0.011492 | 0.00013207 |  |  | n | Chauvenet
6 | 31 | 5.2539 | 5.2432 |  |  | 5.2485 | -0.010753 | 0.00011564 |  |  | 3 | 1.37
10 | 58 | 0.4277 | 0.42969 |  |  | 0.42869 | 0.0019835 | 3.9345e-06 |  |  | 4 | 1.53
10 | 72 | 6.6853 | 6.6848 |  |  | 6.6851 | -0.00050001 | 2.5001e-07 |  |  | 5 | 1.64
21 | 93 | 1.5937 | 1.5333 |  |  | 1.5635 | -0.060395 | 0.0036476 | -3.4145 | yes | 6 | 1.73
21 | 97 | 0.23246 | 0.23132 |  |  | 0.23189 | -0.0011488 | 1.3197e-06 |  |  | 7 | 1.8
21 | 110 | 6.8119 | 6.8267 |  |  | 6.8193 | 0.014746 | 0.00021744 |  |  | 8 | 1.87
32 | 176 | 0.33332 | 0.33395 |  |  | 0.33363 | 0.00062893 | 3.9555e-07 |  |  | 9 | 1.91
32 | 184 | 3.5376 | 3.533 |  |  | 3.5353 | -0.004671 | 2.1818e-05 |  |  | 10 | 1.97
44 | 270 | 2.3461 | 2.3489 |  |  | 2.3475 | 0.0027924 | 7.7975e-06 |  |  | 20 | 2.24
44 | 276 | 0.40222 | 0.4019 |  |  | 0.40206 | -0.00031876 | 1.0161e-07 |  |  | 30 | 2.39
44 | 286 | 6.7538 | 6.7564 |  |  | 6.7551 | 0.0025821 | 6.6671e-06 |  |  | 50 | 2.57
55 | 322 | 2.5295 | 2.5577 |  |  | 2.5436 | 0.028194 | 0.00079491 | 2.6163 | yes | 80 | 2.74
55 | 327 | 0.29724 | 0.288 |  |  | 0.29262 | -0.0092417 | 8.5409e-05 |  |  | 100 | 2.81
55 | 337 | 6.7857 | 6.7966 |  |  | 6.7911 | 0.010848 | 0.00011769 |  |  | 150 | 2.93
58 | 388 | 3.1667 | 3.1665 |  |  | 3.1666 | -0.00017406 | 3.0295e-08 |  |  | 200 | 3.02
58 | 395 | 0.36541 | 0.36102 |  |  | 0.36322 | -0.0043833 | 1.9213e-05
58 | 406 | 6.9736 | 6.976 |  |  | 6.9748 | 0.0023936 | 5.7292e-06
[6 empty] | mean | -0.00088476
[6 empty] | stdev | 0.017429
[8 empty] | 18 | 16 | with 2 outliers removed
[8 empty] | 0.005178 | 0.00073549
[7 empty] | Sp | 0.0019988 | 0.0008475
Multiple bottles from the same level
Event # | Sample# | DO1 | DO2 | Flask1 | Flask2 | Average | Difference | Diff^2 | Chauvenet | reject?
10 | 57/58 | 0.42284 | 0.4277 |  |  | 0.42527 | 0.0048599 | 2.3619e-05
32 | 168/169 | 2.2892 | 2.2735 |  |  | 2.2814 | -0.015615 | 0.00024382
44 | 268/269 | 2.6949 | 2.7117 |  |  | 2.7033 | 0.016745 | 0.00028038
55 | 320/321 | 2.831 | 2.845 |  |  | 2.838 | 0.014035 | 0.000197
[6 empty] | mean | 0.0050063
[6 empty] | stdev | 0.014658
[7 empty] | count | 4
[7 empty] | SUM | 0.00074482
[7 empty] | Sp | 0.009649
Pooled standard deviation of pairs of samples (Sp) was calculated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs