## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 30 Nov 2020 | Added HPLC data. S.H. |

## PROCESSING NOTES

Cruise: 2011-08

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdeF Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: April 13, 2011 – April 17, 2011

Processed by: Germaine Gatien

Date of Processing: 27 September 2011 – 5 October 2011

Number of original HEX files: 78   	Number of original ROS files: 22

Number of CTD files:   74	Number of bottle casts:  19

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1483), a Seapoint Fluorometer (#2745) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4694), a surface PAR sensor (#16504) and an altimeter (#1252).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had no equipment list. There was a Rinko III Temperature sensor listed in the configuration file. The chief scientist knew nothing about this and tests showed no signal present, so this was not processed. An equipment list would have saved some time.

There were two errors in the configuration file – the pressure offset (-0.9db) had not been entered and the Surface PAR conversion factor had a small error and was entered in the wrong place. Because most of the equipment came from a different group who were mostly in the field at the time, it was impossible to get confirmation of the parameters early in the processing job. This is awkward.

There were many problems with computer crashes and misnamed files and one repeated sample number. There were good notes in the log about most of these problems, though the rosette log was very confusing for cast #45. There were also many format errors in the file names that were not noted, but generally it was obvious which files corresponded to which events.

Some of the computer crashes occurred at the beginning of a cast, so the full cast did get recorded in a single file. For some others the crash occurred during the upcast of casts where no rosette sampling was intended so there was no need to create a 2nd file for the cast. In 4 cases there were 2 files that had to be merged. This was made a little awkward in 2 cases because the 2nd file was named by adding a “b”. It is easier to process if the next available event # is assigned (as was done in 2 cases).

Where 2 files were joined the event numbers were set to match the first of each of the set. The CHE file names always match the corresponding CTD file.

This was the first cruise processed with a draw temperature taken from Niskin bottles, allowing for the calculation of dissolved oxygen concentration in mass units for the bottle samples.

Traditionally, salinity calibration sampling for this project involves taking one sample about 10m off the bottom for each rosette cast. Many of these samples are not useful since they come from areas where local variability is very high even near the bottom, so that the bottles can easily contain water quite different from what the CTD is measuring. It would be better to take more samples (say Niskins #2, 3, 4 and 5) from 4 to 6 casts where the salinity gradient is lower, such as in the middle of the Strait of Georgia and near the mouth of the Juan de Fuca Strait. Samples from the eastern part of Juan de Fuca and from Haro Strait are not likely to be useful.

The salinity comparison was very noisy but suggested that both salinity sensors were reading low by approximately 0.0045. This seems unlikely in sensors that were recently serviced and which agreed with each other very well. The noise in the comparison and the small number of samples did not support recalibration. There is a possibility that non-linearity in the Autosal might have occurred since the bottles and CTD appear in much better agreement for salinity >32 than below that. Again, the evidence is weak. No correction was made to the CTD salinity, but if information from summer cruises using these sensors shows significant drift, a correction can be applied later.

Starting with 2011 cruises a new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel was compared with bottles to find the slope and offset to enter in the configuration files. While a little less convenient than calibrating oxygen concentration against bottles, it is the standard approach and is recommended by SeaBird, and should produce better results.

Tests were run to see what the effect would be on the accuracy of SBE dissolved oxygen data if oxygen sampling were reduced. For these data using samples from all bottles on alternate rosette casts would lead to errors on the order of ±0.015mL/L for DO=2.8 to ±0.01mL/L at DO=8.5mL/L. Choosing alternate bottles or every 4th bottle had slightly poorer results. Given the noise level in the calibration data, the error caused by reducing sampling is probably acceptable. However, further tests are recommended because this sensor had just been recalibrated so that the corrections were small, and there was a limited range of DO values with none less than 1.9mL/L during this cruise.

Starting with 2011, the transmissivity conversion has been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder:

OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were many problems with computer crashes; sometimes the cast was restarted and for others there were split casts with 2 files. The numbering of the split casts was non-standard in 2 cases with a “b” added to the file name. This is a little awkward to deal with. It is better to use the next available event number as was done in 2 cases.

Of the many casts interrupted by a computer crash only 4 require merging since the others either occurred upon initialization or during the upcast when no bottle sampling was done. The files that require merges are:

- 16/16b – Downcast ok; upcast requires merge. Save as 16.

- 30/31 – Downcast ok; upcast requires merge. Save as 30.

- 36/36b - Downcast ok; upcast requires merge. Save as 36.

- 58/59 – This case is unclear from the log note, but a test conversion shows there was an interruption during the downcast. There was no rosette sampling, so only the downcast needs merging. Save as 58.

The rosette log sheet is very unclear for event #45. It suggests that Niskin #2 did not fire, but shows oxygen titration values for that bottle. (Later examination shows that Niskin #2 did fire, but not at the intended depth of 200m. It fired at 175m and then Niskins #3 and 4 were both fired at 150m.)

From event #10 to the end the file names had the wrong format (an extra 0); this was fixed using program CKRename. Two other files had been incorrectly named, but had been fixed at sea and put in a special folder; those were copied to the same folder as the other files. Files were missing for a few casts when a different computer was used, but they were obtained from the chief scientist.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity were obtained in spreadsheet format from the analysts. The file creation date was added to the name of the salinity file to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

Only one configuration was used for the cruise. The calibration constants were checked.

- When processing began no calibration information was available for the pressure, temperature and conductivity sensors. It was later obtained and the only error found was in the offset for the pressure sensor. This will be corrected at the calibration stage.

- The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

- There is an entry for a Rinko III Temperature sensor, but the chief scientist knows nothing about this and a test showed no signal in that channel, so it must be an artefact from another cruise.

The file was saved as 2011-08-ctd.cor.xmlcon.

(NOTE: File 2011-08-ctd-cor2.xmlcon was later prepared with corrections to the pressure offset and the Surface PAR parameters; this was not used in processing 2011-08, but was created for use in other cruises using this equipment.)

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files for files 16 & 16b, 39 & 40 and 36 & 36b were joined; first the names were changed to 16.botx and 16.boty etc and then the bottle numbers in the BOTY file were changed to match the rosette sheet and data from bottles that were fired a second time just to make the bottle numbers correct were removed.

The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

There was a problem encountered in preparing the various analyses data to be merged. Sample #11 was used for cast #8 and again for cast #10. The rosette log shows the samples for cast #10 as 12-27, but the samples were labelled as 11-26. The dissolved oxygen analysis has sample numbers that agree with the rosette log sheets, but the others are identified with what was on the labels. One analyst chose to rename #11 as #911. While that is a good approach for a repeated sample number, it does not seem best in this case where the log and labels are not in agreement.  It was decided to change the analyst’s spreadsheets so that the sample numbers agree with the rosette log sheets; that required the following changes:

- CHL: Samples labelled as 23, 24 and 26 were changed to 24, 25 and 27.

- DO: No change

- Salinity: Sample labelled as 12 was changed to 13.

- Nutrients: Sample labelled as 911 changed to 11 and 11-26 changed to 12-27.

Those changes were made in a Sept. 29 version of the spreadsheets. Only the nutrient analyst was available to approve the changes, so this will be considered provisional until the salinity and chlorophyll analysts give their approval.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-08-DO-cal.csv. The titrated DO values and flags were added to that file; there were entries for every line so none needed to be removed. A calculation was made of Ф using equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P, T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

When all data were included M= 0.4313 and B= -0.2052. R2 =0.9823

When 5 outliers were excluded M = 0.4346 and B= - 0.2135. R2 =.9993

The outliers were samples 96 from cast 23, 146 from cast 36, and samples 250, 259 and 260 from cast 76. These will be examined later when the full bottle files have been created and the noise in the SBE DO can be assessed; then a decision will be made whether flags are justified or not.

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are defined as Soc = M and Voffset = B/M.

| Summary of Soc & Voffset including the original values in the factory calibration |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | M | B | Soc | Voffset |
| original | 0.4329 | -0.2218 | 0.4329 | -0.5123 |
| fit-all | 0.4313 | -0.2052 | 0.4313 | -0.4758 |
| fit-5 outliers | 0.4346 | -0.2135 | 0.4346 | -0.4913 |

These are small corrections as is expected since this sensor had been recalibrated in late 2010. New values of Soc and Voffset were entered and the corrected file was saved as 2011-08-ctd-cor.con.

Tests were run using a subset of the available bottles to try to determine what would happen if sampling was reduced during future cruises in the SoG/JdeF program. Runs were done using every other cast, every other bottle, every 4th bottle and a selection of casts thought to include a good variety of DO profiles for the survey area. The bottles excluded for the optimal fit were also excluded from the test runs. The results were as follows:

| Test runs using subsets of bottles |  | Effect of runs on min and max for cast #21 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | DO Concentration Errors |  |  |  |  |
|  |  | DO Conc. Min | Do Conc. Max | min-opt | max-opt | % error min |  | % error max |
| Fit 1 | All data | 2.7125 | 8.3726 | -0.0730 | -0.0957 | -2.69 |  | -1.14 |
| Optimal | Excluding 5 outliers | 2.7855 | 8.4683 | 0.0000 | 0.0000 | 0.00 |  | 0.00 |
| Run1 | every other cast-group a | 2.7985 | 8.4774 | 0.0130 | 0.0091 | 0.46 |  | 0.11 |
| Run2 | every other cast- group b | 2.7697 | 8.4575 | -0.0158 | -0.0108 | -0.57 |  | -0.13 |
| Run3 | every other bottle-group a | 2.7992 | 8.4821 | 0.0137 | 0.0137 | 0.49 |  | 0.16 |
| Run4 | every other bottle-group b | 2.7705 | 8.4557 | -0.0150 | -0.0126 | -0.54 |  | -0.15 |
| Run5 | every 4th bottle group a | 2.8096 | 8.4990 | 0.0241 | 0.0307 | 0.86 |  | 0.36 |
| Run6 | every 4th bottle group b | 2.7877 | 8.4639 | 0.0022 | -0.0044 | 0.08 |  | -0.05 |
| Run7 | selection a | 2.7705 | 8.4557 | -0.0150 | -0.0126 | -0.54 |  | -0.15 |
| Run8 | selection b | 2.7992 | 8.4821 | 0.0137 | 0.0137 | 0.49 |  | 0.16 |

The reduced sets never cause an error of >0.6% or >0.31mL/L in concentration. The largest errors came from one of the runs using every 4th bottle, but the run using a different selection of every 4th bottle had the best results. The runs for which casts were carefully selected to represent the area did not perform as well as using every other cast and were about the same as choosing every other bottle. The best choice appears to be using all bottles from alternate casts. Using set depths was not tested since the effectiveness of the scheme is likely to vary greatly from one cruise to another.

This is not the best data set for analyzing the effect of reducing sampling because there are no DO values <1.9mL/L and the corrections are small to start with. Further tests in different conditions would be wise before deciding on a new sampling scheme. (See file 2011-08-do-test-xls for details.)

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. Those files were converted to IOS header format. They were put through CLEAN to create BOT files. Again, the BOT files for files 16 & 16b, 39 & 40 and 36 & 36b were joined; first the names were changed to 16.botx and 16.boty etc and then the bottle numbers in the BOTY file were changed to match the rosette sheet and data from bottles that were fired a second time just to make the bottle numbers correct were removed (cast 16 only).

Temperature and salinity were plotted for all BOT files and a few outliers were found in the secondary salinity in casts 18, 56 and 69. Secondary salinity was cleaned lightly using CTDEDIT and the output files were then copied to BOT.

A header check was run and shows that there are no off-scale fluorescence values in the BOT files.

The addsamp.csv file prepared in the DO calibration step was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files.

The SAM files were then bin-averaged. The files for 16, 30 and 36 were checked to make sure the JOIN was done right and all looks well.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-08-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file 2011-08_chl.xls. The name of the file was changed by adding the date of creation. The file included comments and flags and duplicate analysis, but no event-number column. The data were already in sample number order. A simplified version of the spreadsheet was prepared in which some columns were removed or reordered and headers were changed to standard format names, event numbers were added and the file was saved as 2011-08-chl.csv.  The file was then converted to individual CHL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-08nuts.xls which included a report on precisions. The file was saved as 2011-08-nuts.csv. Headers were changed to standard format. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-08oxy.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified by removing a few unnecessary columns or rows, adding event numbers, “DO:” was entered in front of comments and the file was then saved as 2011-08-oxy.csv. That file was converted into individual *.ADD files.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. This file was saved as 2011-08-sal.csv. There were no comments or quality flags, but the only problem noted on the log sheet was 1 case of 3 readings which did not seem to settle down. A comment was entered about that but no flag. If it is an outlier in COMPARE then it should be flagged. The sheet from the analyst was saved as 2011-08-sal.csv. Some columns were removed and a column was added for the event number – the latter information was entered using VLOOKUP and the data from the ADDSAMP file. The station name and sample numbers were combined in a single column; those had to be separated. There were no duplicates. The file was converted to individual SAL files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When 3 outliers are excluded (differences >0.02) the primary salinity is low by an average of 0.0041 with a standard deviation of 0.0032, and the secondary is low by 0.0034 with a standard deviation of 0.0044. The major outliers are at casts #56, 64 and 69; the first two have standard deviations in the CTD data >0.01; the 3rd has a low standard deviation during the 10s around firing, but examination of the whole bottle stop shows high variability so that the bottle contents could well be quite different from what the CTD measured during that time. There is more noise in these comparisons than usual. When plotted against file pair number there appears to be time dependence but it is weak and more likely due to regional variability than calibration drift. When more outliers are removed in order to produce a flat fit against pressure, both sensors appear to be low by about 0.0045, but this is based on only 7 casts for the primary and 6 for the secondary.

None of the outliers suggest a need for flagging the samples.

For details see 2011-08-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. When the same 5 outliers are excluded as in section 3, the average difference between bottles and SBE DO was very low (0.0002mL/L), as expected, and the standard deviation of 0.043mL/L. The noise level is higher than found for 2011-01 when the standard deviation was <0.03mL/L. No further recalibration is required.

The flagged values were checked in COMPARE. None stood out as outliers. A note was added to the comments about that. A note about the depth being unknown for sample #164 was removed, since it is clear that it was from 175m.

The 5 outliers were examined, and for 4 of them the SBE DO data either had a high standard deviation or looked out of line, or was in an area of high gradient. For sample #250 the bottle looks out of line in profile and there is no evidence that the SBE DO data is noisy or the gradient high. Flag “3” was attached to that bottle. The spreadsheet was reconverted to individual files.

(See 2011-08-dox-comp1*.xls.)

Fluorescence

COMPARE was run using the CTD Fluorescence and the Extracted Chlorophyll from bottles. The ratio of CTD Fluorescence to extracted CHL was between 0.6 and 2.5 for low CHL values (<1.5ug/L) then fell steeply with values between 0.2 and 0.8 for higher CHL values (>1.5ug/L). (See 2011-08-chl-fluor-comp.xls.)

The merge process was repeated to capture the new flags and comments.

Data were exported to spreadsheet 2011-08-bottles.xls and compared to the rosette sheets and all expected data are present.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### DO sensor hysteresis study

This sensor has not been fine-tuned for hysteresis parameters, but there is no deep sampling from this cruise so the tests cannot be done. This does not matter for this cruise since sampling was fairly shallow.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-08-ctd-cor.xmlcon.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are in good agreement, although the secondary channels are noisier. Upcasts are noisier than downcasts but in reasonable agreement.

The Surface PAR values are extremely low even when the PAR is high. There does seem to be a signal but it is small, ~2 when PAR is ~2000.

The altimetry looks useful near the bottom.

The downcast files for casts #58 and 59 will be joined after DELETE is run.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 3 casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge in these data because of noisy upcast temperatures, but 4.4s looks best overall.

ALIGNCTD was used to advance the DO Voltage by 4.4s relative to the pressure.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using 3 casts. The goal is to make upcasts look closer to downcasts on a T-S surface. There is little temperature variation and for most of the profiles temperature is increasing with depth, so the usual settings often make things worse. No useful setting was found.

CELLTM was not run.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy so these are very rough estimates.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-08-0004 | 280 | +0.0002 | <+0.0001 | +0.0005 | Noisy, High |
| 2011-08-0025 | 380 | +0.0003 | <+0.0001 | +0.0003 | V. noisy, High |
| 2011-08-0028 | 200 | +0.0005 | <+0.0001 | +0.0006 | X Noisy, High |

The differences are all small.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace any pad values in pressure using linear interpolation based on scan number.

##### Checking Headers

The header check was run.  There were no off-scale fluorescence values.

Surface check was run and shows an average surface pressure for the cruise was 3.9 db. This is a little high for the Vector. When pressure calibrations were obtained it was found that the offset had been entered incorrectly in the configuration file so that all values are too high by 0.9039.An average of 3db does look more reasonable for the Vector. The pressure will be corrected later.

The cross-reference check was compared with the log book. There were 3 station name errors, one of which was mentioned in the log, and STA had been entered before some station names. The errors were corrected, and “STA” was removed in the CLN and MRGCLN2 and SAMAVG files

Cast #31 CLN file was deleted since it has only upcast data in it.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. Plots were made of 14 casts and the readings were fit the plots reasonably well.

Water Depth readings were also examined. One was missing so the log entry was checked and looks reasonable so that was entered in the file header. Most deviations from the log book entries were small and likely due to drift between the two readings. The differences were larger for 4 events; based on the maximum sampling level and altimetry, the log entries looked better than the header entry, so the latter were replaced in the headers by the log entries. None of the changes affected rosette files.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. A few casts were examined and the readings look fine.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise the upcast temperature data were quite noisy, making a judgment difficult, but +24 looked best for more casts than any other setting.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of -0.6s to the secondary conductivity.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

At this point the split casts (58/59) were combined:

Files 58.DEL and 59.DEL were renamed 58.DELx and 59.DELy. There was a 4.4m overlap in pressures so a few records were removed from the end of 58.DELx and a few from the beginning of 59.DELY, choosing 79db as the dividing value. This captured the best possible data from each file since the descent rate was a little low at the end of one and the beginning of the other. The two files were then put through JOIN to produce 10.DEL. The output file was examined and looks fine.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

These sensors were recalibrated in Jan. 2011 and used for one cruise just before this one and some Arctic cruises after, but none of those have been processed yet.

2. Dissolved Oxygen

This was the first known use of the oxygen sensor since it was recalibrated in Dec. 2010, though it was likely used for 2011-24.

3. Pressure

The sensor was recalibrated in Dec. 2009; it was used for one cruise before this one and some Arctic cruises after, but none of those have been processed yet.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were some excursions with temperature low around 60 to 80db at casts on the mainland side of the northern end of the Strait of Georgia and salinity low near the top and bottom at the mouth of Juan de Fuca and at a few other sites on the south side of that Strait. Since these are seen in both channels it does not look like evidence of instrumental problems.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

No post-cruise calibrations were available.

##### DETAILED EDITING

The choice of primary or secondary channels was not obvious. There appears to be slightly more noise in the primary channels, but neither channel had many large spikes. The secondary is slightly closer to the bottles. The secondary T and S sensors were chosen for archiving.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as they appear corrupted by either ship effects or too short a wait at the top.

The descent rate of the CTD was extremely noisy for some of the casts resulting in heavy corruption due to shed wakes.

All EDU files were copied to EDT.

##### Initial Recalibration

Based on the results of COMPARE runs and other observations, there is no need to recalibrate the dissolved oxygen data.

Pressure will be increased by adding 0.9039db.

It is difficult to decide how to treat the salinity channel. The sensors had been recalibrated only a few months before this cruise, so it is surprising that the CTD values appear to be low by ~0.0045. There is a lot of noise in the comparison, but the deepest bottles suggest that value. There was 1 previous cruise, 2011-24, that used the same sensors but there was very little deep sampling from that cruise. The only samples from below 150db were examined and showed the CTD salinity values to be high by 0.0004 and low by 0.0014.

One possible problem could be that the salinity from this cruise was low overall, even at depth. The deepest bottles have salinity <31. The highest salinity values are from the casts near the mouth of Juan de Fuca Strait. For casts #45, 48 and 51 the differences are -0.0001, +0.0007 and -0.0093. The cast #51 value came from a stop during which there was a steady decrease in salinity, so it is not surprising that the bottle might contain higher values than were seen by the CTD. A plot of differences against salinity values suggests that there might be an Autosal linearity problem with the 2011-08 salinity analysis; this would explain why the differences are smaller for high salinity values. The deepest casts do not have as high salinity. Linearity problems have occurred in the past, but the evidence here is not strong.

Given the doubts salinity will not be recalibrated now. The sensors were used for some 2011 Arctic cruises; when processed, that data may lead us to reconsider recalibration of salinity.

One other issue to be investigated is the reference PAR channel which has values that are too low. It was discovered that the conversion factor and the ratio multiplier entries were reversed in the configuration file. A test shows that given a conversion factor of 1, it didn’t matter what was entered in the ratio multiplier. So correcting for the conversion factor error is all that is required. Tests were done to see if a simple recalibration by multiplying by that factor was sufficient; there was a slight difference, <0.01%, which is not large enough to justify the more error-prone method of rederivation and merging of the reference PAR channel with the final files.

CALIBRATE was run on all EDT files using file 2011-08-recal.ccf to subtract 0.9039db from the pressure channel and to multiply the PAR:Reference channel by the conversion factor, 1527.282.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but it is sometimes found appropriate to do a further correction for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure there is a lot of scatter above 300db. The SBE DO values appear to be low by an average of 0.01mL/L but the standard deviation is 0.08mL/L, with values somewhat low at the surface and at depth, but a little high at mid-depths. No further recalibration will be applied. Error estimates were based on this plot.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 200db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

Spreadsheet 2011-08-SOG.csv was prepared by clipping COR1 files to 50db for only casts with bottle data, deriving sigma-t and exporting start time, event #, latitude, longitude, station name, pressure, temperature, fluorescence, salinity and sigma-t to the spreadsheet.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

For cast #3, channel PAR was also removed because the cap was left on the sensor.

A second SBE DO channel (with umol/kg units) was added.

A second titrated DO channel (with umol/kg units) was added using the draw temperature.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following header comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.25mL/L from 0 to 100db

±0.15mL/L from 100db to 200db

±0.1mL/L below 200db

For details on the processing see processing report: 2011-08-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

T-S plots were examined and the only one that looks odd is cast #69 at station 56. The temperature shows very little variation so that small salinity variations create unstable features. The downcast and upcast are quite different. Primary and secondary salinity are close. There is no indication of instrumental problems. This cast looks quite different from cast #70 with higher salinity and lower surface dissolved oxygen. The instabilities are likely real.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 85% to 140%, with the lowest values in the eastern part of Juan de Fuca Strait where there was active mixing, and highest values in the central part of the Strait of Georgia. A check of titrated DO values at a few casts with high saturation showed good agreement between bottle values and SBE DO.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was used to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

For cast #3, channel PAR was also removed because the cap was left on the sensor.

A second SBE DO channel (with umol/kg units) was added.

A second titrated DO channel (with umol/kg units) was added using the draw temperature.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data..

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no discrepancies were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no outliers were identified. There are two odd features, but in one case both titrated value and SBE DO value agree, and in the other the SBE DO looks out of line and was excluded in the comparison. The bottle value looks ok.

Standards check was run on all files and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

3. Computer crash – cast restarted. Fluorescence high. PAR cap left on.

9/10. Sample #11 was used twice.

16. Computer crash just before tripping bottle #10; second file started with event # 16b. Station name is wrong – should be 9.

24. Computer crash – cast restarted.

28. Station number wrong in raw file.

30/31. Computer crash during upcast.

33. Computer crash during upcast. No restart since no rosette sampling.

36. Computer crash just before 10m bottle; new file saved as 36b.

41. Computer changed before this cast.

44. Computer crash in initialization.

45. Bottle 2 not fired, Niskin 4 fired twice – rosette sheet not clear about what happened – there is a DO value listed for Niskin 2 and 3/4 is listed by Niskin 4.

46. Computer crashed twice on initialization and froze at the end.

47. Computer changed before this cast.

58/59. Reboot at 25m off bottom. New file saved as 59. Not clear if this was during upcast or downcast.

64. Computer crash on initialization.

72. Computer crash at 80m; cast started again from surface.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 06Jan11 | Factory |  |  |
| Conductivity | 3579 | 08Jan11 | Factory |  |  |
| Secondary Temp. | 5073 | 06Jan11 | Factory |  |  |
| Secondary Cond. | 3581 | 08Jan11 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 24Dec2010 | Factory |  |  |
| PAR | 4694 | 03Mar2010 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| Fluorometer | 2745 |  | IOS |  |  |
| Pressure Sensor | 0941 | 7Dec09 | Factory |  |  |
| Rinko III Temperature | 005 | 2009 |  |  |  |
| Altimeter | 1252 |  |  |  |  |