## 2011-08-do-cal
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle_Number | Bottle:Position | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved
8 | 156.754 | 8.9407 | 8.9408 | 30.6062 | 30.6065 | 6.6491 | 1.7012 | 1 | 1 | 1 | 6.7628 | 0.52996 | 3.584 |  | Summary of Soc Voffset including the original values in the factory calibration
8 | 125.406 | 8.838 | 8.8382 | 30.5252 | 30.5255 | 6.6679 | 1.7573 | 2 | 2 | 2 | 6.7541 | 0.55537 | 3.751 |  |  | m | b | Soc | Voffset
8 | 100.198 | 8.6691 | 8.6671 | 30.2896 | 30.2861 | 6.7034 | 1.8515 | 3 | 3 | 3 | 6.767 | 0.59657 | 4.037 |  | original | 0.4329 | -0.22177 | 0.4329 | -0.5123
8 | 73.515 | 8.5193 | 8.5201 | 29.9863 | 29.9872 | 6.7393 | 1.9436 | 4 | 4 | 4 | 6.779 | 0.62679 | 4.249 |  | fit-all | 0.4313 | -0.2052 | 0.4313 | -0.47577
8 | 47.305 | 8.2914 | 8.2915 | 29.5151 | 29.5145 | 6.7949 | 2.2763 | 5 | 5 | 5 | 6.8105 | 0.78511 | 5.347 | 6 | fit-5 outliers | 0.4346 | -0.2135 | 0.4346 | -0.49126
8 | 40.092 | 8.245 | 8.2427 | 29.4096 | 29.4043 | 6.8067 | 2.3971 | 6 | 6 | 6 | 6.8158 | 0.83586 | 5.697
8 | 31.586 | 8.1708 | 8.1702 | 29.2508 | 29.2504 | 6.8252 | 2.537 | 7 | 7 | 7 | 6.8264 | 0.89476 | 6.108
8 | 23.798 | 8.1418 | 8.1405 | 29.2113 | 29.2097 | 6.8315 | 2.5633 | 8 | 8 | 8 | 6.8257 | 0.90907 | 6.205
8 | 9.899 | 8.0123 | 8.0098 | 29.0686 | 29.0679 | 6.858 | 2.7169 | 9 | 9 | 9 | 6.8392 | 0.97103 | 6.641
8 | 4.703 | 7.9971 | 7.9972 | 29.0631 | 29.0636 | 6.8606 | 2.7201 | 10 | 10 | 10 | 6.8371 | 0.97511 | 6.667
8 | 3.701 | 7.9941 | 7.9943 | 29.0628 | 29.0633 | 6.8611 | 2.725 | 11 | 11 | 11 | 6.8367 | 0.97327 | 6.654
10 | 308.264 | 9.1851 | 9.1856 | 30.7238 | 30.7244 | 6.608 | 1.554 | 1 | 1 | 12 | 6.8539 | 0.47155 | 3.232
10 | 299.905 | 9.1617 | 9.1597 | 30.7148 | 30.7144 | 6.6118 | 1.5692 | 2 | 2 | 13 | 6.8504 | 0.47107 | 3.227
10 | 249.69 | 8.9525 | 8.9532 | 30.6214 | 30.6222 | 6.6467 | 1.6791 | 3 | 3 | 14 | 6.8411 | 0.52009 | 3.558
10 | 200.042 | 8.8967 | 8.8965 | 30.5819 | 30.5835 | 6.6567 | 1.716 | 4 | 4 | 15 | 6.8077 | 0.53425 | 3.637
10 | 175.195 | 8.836 | 8.8363 | 30.4975 | 30.4981 | 6.6694 | 1.7694 | 5 | 5 | 16 | 6.7986 | 0.55599 | 3.78
10 | 149.869 | 8.627 | 8.6258 | 30.3994 | 30.3996 | 6.705 | 1.8597 | 6 | 6 | 17 | 6.8114 | 0.59063 | 4.023 | 6
10 | 124.915 | 8.5182 | 8.5322 | 30.2877 | 30.2906 | 6.7263 | 1.8983 | 7 | 7 | 18 | 6.8105 | 0.6139 | 4.181
10 | 100.394 | 8.545 | 8.5454 | 30.0692 | 30.0697 | 6.7318 | 1.9001 | 8 | 8 | 19 | 6.7949 | 0.61781 | 4.198
10 | 74.767 | 8.3272 | 8.3273 | 29.5963 | 29.5987 | 6.7858 | 2.1899 | 9 | 9 | 20 | 6.8256 | 0.74075 | 5.056
10 | 50.119 | 8.0039 | 8.0039 | 29.0078 | 29.0064 | 6.862 | 2.6581 | 10 | 10 | 21 | 6.8785 | 0.94032 | 6.468
10 | 39.987 | 7.9943 | 7.9945 | 28.818 | 28.8205 | 6.872 | 2.9746 | 11 | 11 | 22 | 6.8795 | 1.0874 | 7.481
10 | 30.333 | 7.9177 | 7.9181 | 28.7758 | 28.7764 | 6.886 | 3.0057 | 12 | 12 | 23 | 6.8845 | 1.0958 | 7.544 | 6
10 | 20.233 | 7.9278 | 7.9274 | 28.7756 | 28.7763 | 6.8844 | 3.0019 | 13 | 13 | 24 | 6.8741 | 1.0948 | 7.526
10 | 9.936 | 7.9223 | 7.9226 | 28.7707 | 28.7715 | 6.8855 | 3.0319 | 14 | 14 | 25 | 6.866 | 1.1078 | 7.606
10 | 5.146 | 7.9247 | 7.9248 | 28.7703 | 28.7709 | 6.8851 | 3.0418 | 15 | 15 | 26 | 6.8615 | 1.1122 | 7.631
10 | 3.559 | 7.9327 | 7.9297 | 28.7707 | 28.7714 | 6.8838 | 3.0416 | 16 | 16 | 27 | 6.8589 | 1.1124 | 7.63
12 | 351.702 | 9.2585 | 9.259 | 30.7591 | 30.7598 | 6.5957 | 1.5001 | 1 | 1 | 28 | 6.8797 | 0.44638 | 3.071
12 | 299.468 | 9.0173 | 9.0174 | 30.6815 | 30.682 | 6.6345 | 1.6379 | 2 | 2 | 29 | 6.8725 | 0.50098 | 3.443
12 | 250.434 | 8.7778 | 8.7784 | 30.6082 | 30.6081 | 6.6733 | 1.7372 | 3 | 3 | 30 | 6.8679 | 0.54456 | 3.74
12 | 199.958 | 8.3318 | 8.3319 | 30.4652 | 30.4657 | 6.7469 | 1.9091 | 4 | 4 | 31 | 6.896 | 0.61485 | 4.24 | 36
12 | 174.444 | 8.2269 | 8.2255 | 30.3798 | 30.3801 | 6.7666 | 1.9713 | 5 | 5 | 32 | 6.893 | 0.64326 | 4.434 | 3
12 | 150.233 | 8.2064 | 8.2067 | 30.3289 | 30.3294 | 6.772 | 1.992 | 6 | 6 | 33 | 6.877 | 0.65464 | 4.502
12 | 124.709 | 8.3141 | 8.3143 | 30.2695 | 30.2703 | 6.7581 | 1.9678 | 7 | 7 | 34 | 6.8413 | 0.6471 | 4.427
12 | 99.783 | 8.4487 | 8.4486 | 30.0911 | 30.0929 | 6.7454 | 1.9316 | 8 | 8 | 35 | 6.8075 | 0.63782 | 4.342
12 | 74.665 | 8.2918 | 8.2926 | 29.7906 | 29.7918 | 6.7826 | 2.0219 | 9 | 9 | 36 | 6.8221 | 0.65435 | 4.464
12 | 49.859 | 8.0041 | 8.0038 | 28.9359 | 28.9366 | 6.8652 | 2.8876 | 10 | 10 | 37 | 6.8814 | 1.0345 | 7.119
12 | 40.161 | 8.0086 | 8.0084 | 28.7976 | 28.7962 | 6.8707 | 3.0075 | 11 | 11 | 38 | 6.8784 | 1.0962 | 7.54
12 | 29.8 | 7.9974 | 7.9977 | 28.7302 | 28.7299 | 6.8755 | 3.0625 | 12 | 12 | 39 | 6.874 | 1.117 | 7.678
12 | 20.122 | 7.995 | 7.9951 | 28.6854 | 28.6861 | 6.8779 | 3.0971 | 13 | 13 | 40 | 6.8679 | 1.137 | 7.809
12 | 10.003 | 7.9978 | 7.9979 | 28.68 | 28.6808 | 6.8777 | 3.1198 | 14 | 14 | 41 | 6.8588 | 1.1455 | 7.857 | 6
12 | 4.686 | 8.0006 | 8.0007 | 28.6793 | 28.6765 | 6.8773 | 3.1246 | 15 | 15 | 42 | 6.8537 | 1.149 | 7.875
12 | 4.07 | 8.0059 | 8.0069 | 28.6782 | 28.6792 | 6.8765 | 3.1349 | 16 | 16 | 43 | 6.8524 | 1.1501 | 7.881
16 | 164.164 | 8.6463 | 8.6467 | 30.4896 | 30.4906 | 6.6982 | 1.7919 | 1 | 1 | 44 | 6.817 | 0.5699 | 3.885
16 | 149.956 | 8.6378 | 8.638 | 30.4686 | 30.4697 | 6.7004 | 1.8029 | 2 | 2 | 45 | 6.8068 | 0.57413 | 3.908
16 | 124.622 | 8.5675 | 8.5678 | 30.2824 | 30.2803 | 6.7191 | 1.8735 | 3 | 3 | 46 | 6.8033 | 0.60618 | 4.124
16 | 99.975 | 8.545 | 8.5452 | 30.0405 | 30.0411 | 6.733 | 1.9205 | 4 | 4 | 47 | 6.7958 | 0.62465 | 4.245
16 | 74.119 | 8.282 | 8.2809 | 29.7466 | 29.7459 | 6.7861 | 2.0256 | 5 | 5 | 48 | 6.825 | 0.67135 | 4.582 | 6
16 | 49.4 | 8.003 | 8.0031 | 29.4214 | 29.4226 | 6.8437 | 2.1978 | 6 | 6 | 49 | 6.8595 | 0.75195 | 5.158
16 | 40.324 | 7.9512 | 7.9502 | 29.2659 | 29.2658 | 6.8587 | 2.3586 | 7 | 7 | 50 | 6.8662 | 0.82418 | 5.659
16 | 30.302 | 7.9542 | 7.9513 | 29.1734 | 29.1772 | 6.8624 | 2.489 | 8 | 8 | 51 | 6.8611 | 0.88572 | 6.077
16 | 19.776 | 7.986 | 7.9866 | 29.0786 | 29.0762 | 6.8617 | 2.6832 | 9 | 9 | 52 | 6.8513 | 0.96755 | 6.629
16 | 9.917 | 8.0027 | 8.0033 | 28.9526 | 28.9534 | 6.8647 | 2.9215 | 10 | 10 | 53 | 6.8458 | 1.0586 | 7.247
16 | 5.293 | 8.0098 | 8.0091 | 28.9602 | 28.9611 | 6.8632 | 2.9165 | 11 | 11 | 54 | 6.8403 | 1.0641 | 7.279
16 | 4.334 | 8.0063 | 8.0067 | 28.9588 | 28.9594 | 6.8638 | 2.9285 | 12 | 12 | 55 | 6.8401 | 1.068 | 7.305
18 | 185.731 | 8.3544 | 8.3545 | 30.4636 | 30.464 | 6.7435 | 1.8953 | 1 | 1 | 56 | 6.8801 | 0.61583 | 4.237
18 | 174.681 | 8.216 | 8.2151 | 30.4022 | 30.4024 | 6.7673 | 1.9555 | 2 | 2 | 57 | 6.8938 | 0.6384 | 4.401
18 | 149.866 | 8.0665 | 8.0668 | 30.3218 | 30.322 | 6.7938 | 2.041 | 3 | 3 | 58 | 6.898 | 0.67541 | 4.659
18 | 125.034 | 7.9264 | 7.929 | 30.1503 | 30.1519 | 6.8232 | 2.122 | 4 | 4 | 59 | 6.905 | 0.71253 | 4.92
18 | 99.925 | 8.4631 | 8.4633 | 29.9233 | 29.9234 | 6.7506 | 1.9863 | 5 | 5 | 60 | 6.813 | 0.65184 | 4.441
18 | 74.731 | 8.1979 | 8.1989 | 29.67 | 29.6711 | 6.8024 | 2.0449 | 6 | 6 | 61 | 6.8414 | 0.67413 | 4.612 | 6
18 | 49.93 | 7.9399 | 7.9401 | 29.152 | 29.1526 | 6.8656 | 2.5174 | 7 | 7 | 62 | 6.8815 | 0.87655 | 6.032
18 | 39.705 | 7.9112 | 7.9109 | 28.9003 | 28.9008 | 6.8814 | 2.7961 | 8 | 8 | 63 | 6.8881 | 1.0029 | 6.908
18 | 29.874 | 8.0557 | 8.0559 | 28.5642 | 28.5646 | 6.8738 | 3.25 | 9 | 9 | 64 | 6.8727 | 1.2048 | 8.28
18 | 20.082 | 8.0732 | 8.0748 | 28.1456 | 28.1378 | 6.8898 | 3.3756 | 10 | 10 | 65 | 6.8802 | 1.2585 | 8.659
18 | 9.928 | 8.089 | 8.0891 | 28.0382 | 28.0393 | 6.8921 | 3.4013 | 11 | 11 | 66 | 6.8737 | 1.2689 | 8.722
18 | 4.431 | 8.0911 | 8.0905 | 28.0268 | 28.0275 | 6.8923 | 3.4079 | 12 | 12 | 67 | 6.8691 | 1.2722 | 8.739
18 | 3.671 | 8.0816 | 8.0817 | 28.0215 | 28.0217 | 6.894 | 3.3905 | 13 | 13 | 68 | 6.8701 | 1.2716 | 8.736
21 | 317.599 | 9.2832 | 9.2834 | 30.9406 | 30.9408 | 6.5844 | 1.433 | 1 | 1 | 69 | 6.8383 | 0.41107 | 2.811
21 | 299.699 | 9.2533 | 9.2535 | 30.9198 | 30.9201 | 6.5897 | 1.4575 | 2 | 2 | 70 | 6.8279 | 0.42033 | 2.87
21 | 249.704 | 9.1458 | 9.1513 | 30.8543 | 30.8568 | 6.6082 | 1.5324 | 3 | 3 | 71 | 6.8029 | 0.45628 | 3.104
21 | 199.778 | 8.4644 | 8.4671 | 30.5634 | 30.5649 | 6.7224 | 1.8354 | 4 | 4 | 72 | 6.8717 | 0.58734 | 4.036
21 | 174.94 | 8.139 | 8.1392 | 30.3707 | 30.3712 | 6.7805 | 1.9962 | 5 | 5 | 73 | 6.9071 | 0.65657 | 4.535
21 | 150.175 | 7.9356 | 7.936 | 30.2236 | 30.2239 | 6.8185 | 2.099 | 6 | 6 | 74 | 6.9225 | 0.69945 | 4.842
21 | 124.839 | 7.871 | 7.8716 | 30.1075 | 30.108 | 6.8337 | 2.1488 | 7 | 7 | 75 | 6.9151 | 0.72363 | 5.004
21 | 100.057 | 7.939 | 7.9399 | 30.0085 | 30.0088 | 6.8275 | 2.1398 | 8 | 8 | 76 | 6.8874 | 0.72001 | 4.959
21 | 74.81 | 8.1937 | 8.1937 | 29.6722 | 29.6723 | 6.803 | 2.0653 | 9 | 9 | 77 | 6.842 | 0.68225 | 4.668
21 | 50.009 | 7.8461 | 7.8461 | 29.1522 | 29.1526 | 6.8803 | 2.3777 | 10 | 10 | 78 | 6.8958 | 0.81528 | 5.622
21 | 39.958 | 7.8476 | 7.849 | 28.9888 | 28.9645 | 6.8874 | 2.5109 | 11 | 11 | 79 | 6.894 | 0.87148 | 6.008 | 6
21 | 29.689 | 7.8856 | 7.892 | 28.773 | 28.7669 | 6.8911 | 2.6917 | 12 | 12 | 80 | 6.8889 | 0.94326 | 6.498
21 | 19.746 | 8.0468 | 8.0472 | 27.9257 | 27.9247 | 6.9038 | 3.2879 | 13 | 13 | 81 | 6.8938 | 1.2112 | 8.35
21 | 10.12 | 8.0579 | 8.0591 | 27.8448 | 27.8453 | 6.9057 | 3.3047 | 14 | 14 | 82 | 6.8872 | 1.2253 | 8.439
21 | 4.598 | 8.0527 | 8.053 | 27.8445 | 27.845 | 6.9065 | 3.3042 | 15 | 15 | 83 | 6.8832 | 1.2224 | 8.414
21 | 3.652 | 8.0501 | 8.0426 | 27.845 | 27.8458 | 6.9069 | 3.3047 | 16 | 16 | 84 | 6.8827 | 1.2213 | 8.406
23 | 274.498 | 9.2207 | 9.222 | 30.9198 | 30.9206 | 6.5944 | 1.4533 | 1 | 1 | 85 | 6.8107 | 0.42183 | 2.873
23 | 249.389 | 8.9194 | 8.9089 | 30.7672 | 30.7625 | 6.6454 | 1.6209 | 2 | 2 | 86 | 6.8392 | 0.48909 | 3.345
23 | 199.718 | 8.3597 | 8.3595 | 30.5162 | 30.5163 | 6.7404 | 1.8881 | 3 | 3 | 87 | 6.8893 | 0.60935 | 4.198
23 | 174.647 | 8.1922 | 8.1856 | 30.4054 | 30.4028 | 6.7708 | 1.9737 | 4 | 4 | 88 | 6.8973 | 0.64721 | 4.464
23 | 149.946 | 8.0287 | 8.0332 | 30.2503 | 30.2524 | 6.8029 | 2.0629 | 5 | 5 | 89 | 6.907 | 0.6796 | 4.694
23 | 124.972 | 7.9012 | 7.9015 | 30.0794 | 30.0797 | 6.8302 | 2.1386 | 6 | 6 | 90 | 6.9119 | 0.7176 | 4.96 | 6
23 | 100.043 | 7.9064 | 7.9057 | 29.9303 | 29.9307 | 6.8361 | 2.1603 | 7 | 7 | 91 | 6.8958 | 0.72537 | 5.002
23 | 75.155 | 7.926 | 7.9277 | 29.7814 | 29.7788 | 6.8396 | 2.1757 | 8 | 8 | 92 | 6.8776 | 0.73572 | 5.06
23 | 49.546 | 7.9743 | 7.9744 | 29.5901 | 29.5913 | 6.8407 | 2.1972 | 9 | 9 | 93 | 6.8564 | 0.74208 | 5.088
23 | 40.019 | 7.8954 | 7.8957 | 29.4311 | 29.4321 | 6.8601 | 2.2387 | 10 | 10 | 94 | 6.867 | 0.76191 | 5.232
23 | 30.156 | 7.7423 | 7.7462 | 28.9315 | 28.9339 | 6.9066 | 2.4447 | 11 | 11 | 95 | 6.904 | 0.8217 | 5.673
23 | 9.794 | 7.9435 | 7.9435 | 28.1753 | 28.1761 | 6.9089 | 3.2127 | 13 | 13 | 97 | 6.8894 | 1.1869 | 8.177
23 | 5.057 | 7.877 | 7.8765 | 27.7685 | 27.7741 | 6.9378 | 3.2706 | 14 | 14 | 98 | 6.9136 | 1.212 | 8.379
23 | 3.843 | 7.8777 | 7.8783 | 27.7387 | 27.734 | 6.939 | 3.2826 | 15 | 15 | 99 | 6.9138 | 1.2184 | 8.424
27 | 343.639 | 9.3506 | 9.3499 | 30.9918 | 30.9919 | 6.5724 | 1.3804 | 1 | 1 | 100 | 6.849 | 0.40473 | 2.772
27 | 300.448 | 9.2263 | 9.2269 | 30.9211 | 30.922 | 6.5936 | 1.4593 | 2 | 2 | 101 | 6.8324 | 0.42386 | 2.896
27 | 249.911 | 9.0395 | 9.0367 | 30.8224 | 30.8217 | 6.6252 | 1.5606 | 3 | 3 | 102 | 6.8198 | 0.46776 | 3.19
27 | 200.103 | 8.2511 | 8.2523 | 30.4997 | 30.5006 | 6.7576 | 1.9104 | 4 | 4 | 103 | 6.9066 | 0.61723 | 4.263
27 | 174.774 | 8.1842 | 8.1843 | 30.4366 | 30.4374 | 6.7707 | 1.9585 | 5 | 5 | 104 | 6.8972 | 0.64244 | 4.431 | 6
27 | 149.75 | 8.2746 | 8.2752 | 30.4233 | 30.4241 | 6.7574 | 1.9413 | 6 | 6 | 105 | 6.8622 | 0.63143 | 4.333
27 | 125.083 | 7.955 | 7.9557 | 30.2776 | 30.2782 | 6.8131 | 2.0857 | 7 | 7 | 106 | 6.895 | 0.69137 | 4.767
27 | 100.086 | 7.7707 | 7.7713 | 30.1236 | 30.1241 | 6.8486 | 2.1891 | 8 | 8 | 107 | 6.9078 | 0.73917 | 5.106
27 | 75.395 | 7.9622 | 7.9621 | 30.0016 | 30.0026 | 6.8242 | 2.1483 | 9 | 9 | 108 | 6.8625 | 0.72204 | 4.955
27 | 49.778 | 8.072 | 8.0718 | 29.7091 | 29.7094 | 6.8202 | 2.1393 | 10 | 10 | 109 | 6.8366 | 0.70985 | 4.853
27 | 40.009 | 7.9438 | 7.9438 | 29.5187 | 29.5209 | 6.8486 | 2.2012 | 11 | 11 | 110 | 6.8558 | 0.73311 | 5.026
27 | 29.748 | 7.7839 | 7.7829 | 29.2549 | 29.2551 | 6.8855 | 2.3136 | 12 | 12 | 111 | 6.8827 | 0.77353 | 5.324 | 6
27 | 19.567 | 7.6798 | 7.6761 | 28.8891 | 28.899 | 6.9185 | 2.505 | 13 | 13 | 112 | 6.906 | 0.8633 | 5.962
27 | 9.606 | 7.8443 | 7.8431 | 28.2779 | 28.2789 | 6.92 | 3.3719 | 14 | 14 | 113 | 6.8997 | 1.2453 | 8.592
27 | 4.878 | 7.8289 | 7.8299 | 28.2751 | 28.277 | 6.9225 | 3.3716 | 15 | 15 | 114 | 6.898 | 1.2525 | 8.64
27 | 3.85 | 7.8312 | 7.8304 | 28.2724 | 28.2732 | 6.9223 | 3.3826 | 16 | 16 | 115 | 6.8968 | 1.2561 | 8.663
30 | 373.495 | 9.342 | 9.3423 | 30.9834 | 30.9841 | 6.574 | 1.3595 | 1 | 1 | 116 | 6.8768 | 0.38143 | 2.623
30 | 349.772 | 9.2754 | 9.276 | 30.9441 | 30.9448 | 6.5854 | 1.419 | 2 | 2 | 117 | 6.8674 | 0.40394 | 2.774
30 | 300.88 | 9.0843 | 9.0834 | 30.8542 | 30.8545 | 6.6173 | 1.5259 | 3 | 3 | 118 | 6.8563 | 0.45009 | 3.086
30 | 249.804 | 8.6972 | 8.7023 | 30.6931 | 30.6955 | 6.6817 | 1.6986 | 4 | 4 | 119 | 6.8754 | 0.51735 | 3.557
30 | 198.372 | 8.1164 | 8.1161 | 30.4203 | 30.4205 | 6.7818 | 1.9621 | 5 | 5 | 120 | 6.929 | 0.63169 | 4.377 | 6
30 | 173.328 | 7.8886 | 7.8892 | 30.3005 | 30.3014 | 6.8223 | 2.0814 | 6 | 6 | 121 | 6.9468 | 0.68852 | 4.783
30 | 151.049 | 7.8254 | 7.8258 | 30.1852 | 30.1861 | 6.8373 | 2.1462 | 7 | 7 | 122 | 6.9418 | 0.72056 | 5.002
30 | 124.999 | 8.0076 | 8.0076 | 30.0591 | 30.0598 | 6.8146 | 2.116 | 8 | 8 | 123 | 6.8968 | 0.70903 | 4.89
30 | 99.086 | 7.9687 | 7.9654 | 29.8402 | 29.838 | 6.8304 | 2.1659 | 9 | 9 | 124 | 6.8896 | 0.72863 | 5.02
30 | 74.6 | 7.6994 | 7.6973 | 29.6632 | 29.6626 | 6.8805 | 2.4095 | 10 | 10 | 125 | 6.9168 | 0.83376 | 5.767
30 | 50.203 | 8.0305 | 8.0286 | 29.6029 | 29.6024 | 6.8313 | 2.1636 | 11 | 11 | 126 | 6.8479 | 0.72664 | 4.976
30 | 41.964 | 7.9926 | 7.989 | 29.5138 | 29.5113 | 6.8412 | 2.1962 | 12 | 12 | 127 | 6.8504 | 0.74419 | 5.098 | 6
30 | 31.072 | 7.8963 | 7.8955 | 29.3748 | 29.3756 | 6.8625 | 2.2644 | 13 | 13 | 128 | 6.8615 | 0.77476 | 5.316
30 | 19.629 | 7.8148 | 7.8146 | 29.1002 | 29.0998 | 6.8876 | 2.477 | 14 | 14 | 129 | 6.8761 | 0.85223 | 5.86
30 | 10.239 | 7.7912 | 7.7938 | 28.2303 | 28.226 | 6.9306 | 3.2208 | 15 | 15 | 130 | 6.9105 | 1.1804 | 8.157
30 | 4.253 | 7.7473 | 7.7247 | 25.3802 | 24.8828 | 7.0679 | 3.414 | 16 | 16 | 131 | 7.0418 | 1.284 | 9.042
30 | 3.611 | 7.6364 | 7.6344 | 23.1661 | 23.0641 | 7.1892 | 3.3926 | 17 | 17 | 132 | 7.1614 | 1.2648 | 9.058
36 | 317.099 | 9.1891 | 9.188 | 30.9055 | 30.9055 | 6.5997 | 1.4389 | 1 | 1 | 133 | 6.853 | 0.42536 | 2.915
36 | 299.941 | 9.0522 | 9.052 | 30.8387 | 30.8392 | 6.6227 | 1.5192 | 2 | 2 | 134 | 6.8609 | 0.4444 | 3.049
36 | 249.395 | 8.8839 | 8.8835 | 30.7621 | 30.7627 | 6.6509 | 1.6143 | 3 | 3 | 135 | 6.8446 | 0.48871 | 3.345
36 | 199.863 | 8.4537 | 8.4539 | 30.5716 | 30.5718 | 6.7237 | 1.8146 | 4 | 4 | 136 | 6.873 | 0.55696 | 3.828
36 | 174.802 | 8.2087 | 8.2082 | 30.4495 | 30.4505 | 6.7664 | 1.9424 | 5 | 5 | 137 | 6.8929 | 0.62528 | 4.31
36 | 149.577 | 8.0083 | 8.01 | 30.2745 | 30.2761 | 6.8049 | 2.0686 | 6 | 6 | 138 | 6.9087 | 0.68204 | 4.712
36 | 124.772 | 7.9461 | 7.9497 | 30.084 | 30.0855 | 6.823 | 2.1444 | 7 | 7 | 139 | 6.9047 | 0.72066 | 4.976
36 | 99.762 | 7.7611 | 7.7597 | 29.9179 | 29.9154 | 6.8593 | 2.302 | 8 | 8 | 140 | 6.9182 | 0.78416 | 5.425
36 | 74.06 | 7.8382 | 7.8381 | 29.8429 | 29.8438 | 6.8506 | 2.2441 | 9 | 9 | 141 | 6.8871 | 0.77376 | 5.329
36 | 50.01 | 7.8459 | 7.8462 | 29.6151 | 29.6165 | 6.8596 | 2.2633 | 10 | 10 | 142 | 6.875 | 0.76596 | 5.266
36 | 39.768 | 7.7925 | 7.7921 | 29.4969 | 29.4972 | 6.8733 | 2.3149 | 11 | 11 | 143 | 6.8794 | 0.79281 | 5.454 | 6
36 | 29.943 | 7.7523 | 7.751 | 29.3385 | 29.3297 | 6.8867 | 2.3891 | 12 | 12 | 144 | 6.8839 | 0.82584 | 5.685
36 | 20.046 | 7.7441 | 7.7449 | 29.0944 | 29.0901 | 6.899 | 2.5124 | 13 | 13 | 145 | 6.8874 | 0.87203 | 6.006
36 | 5.073 | 8.0217 | 8.0376 | 28.6039 | 28.6028 | 6.8773 | 3.6235 | 15 | 15 | 147 | 6.8543 | 1.3673 | 9.372
36 | 3.705 | 8.051 | 8.0688 | 28.5994 | 28.5983 | 6.8729 | 3.7273 | 16 | 16 | 148 | 6.8489 | 1.4017 | 9.6
41 | 230.725 | 7.8883 | 7.8884 | 31.3141 | 31.3146 | 6.7774 | 2.1941 | 1 | 1 | 149 | 6.952 | 0.73922 | 5.139
41 | 200.271 | 7.8854 | 7.8856 | 31.2974 | 31.3052 | 6.7786 | 2.2015 | 2 | 2 | 150 | 6.9261 | 0.74299 | 5.146 | 6
41 | 175.091 | 7.8829 | 7.883 | 31.1745 | 31.1746 | 6.7844 | 2.2428 | 3 | 3 | 151 | 6.9097 | 0.75835 | 5.24
41 | 149.754 | 7.8805 | 7.8806 | 31.1071 | 31.1092 | 6.7878 | 2.2594 | 4 | 4 | 152 | 6.8907 | 0.7693 | 5.301
41 | 125.014 | 7.8782 | 7.8783 | 31.0788 | 31.0781 | 6.7894 | 2.2771 | 5 | 5 | 153 | 6.8705 | 0.7784 | 5.348
41 | 100.382 | 7.8816 | 7.8816 | 30.3781 | 30.3811 | 6.82 | 2.472 | 6 | 6 | 154 | 6.8797 | 0.85977 | 5.915
41 | 74.843 | 7.8856 | 7.8856 | 30.2231 | 30.2249 | 6.8262 | 2.5445 | 7 | 7 | 155 | 6.8636 | 0.89297 | 6.129
41 | 49.662 | 7.8878 | 7.8878 | 30.117 | 30.12 | 6.8306 | 2.596 | 8 | 8 | 156 | 6.8459 | 0.91281 | 6.249
41 | 39.796 | 7.8878 | 7.8878 | 30.0959 | 30.1005 | 6.8316 | 2.6096 | 9 | 9 | 157 | 6.8382 | 0.92246 | 6.308
41 | 29.9 | 7.8884 | 7.8884 | 30.0672 | 30.0676 | 6.8328 | 2.6362 | 10 | 10 | 158 | 6.8307 | 0.93489 | 6.386
41 | 19.731 | 7.8924 | 7.8924 | 29.9179 | 29.9207 | 6.8388 | 2.6903 | 11 | 11 | 159 | 6.8279 | 0.9571 | 6.535
41 | 9.814 | 7.8915 | 7.8913 | 29.8397 | 29.8381 | 6.8424 | 2.7233 | 12 | 12 | 160 | 6.8228 | 0.97071 | 6.623
41 | 4.918 | 7.883 | 7.881 | 29.8133 | 29.812 | 6.8449 | 2.7327 | 13 | 13 | 161 | 6.821 | 0.97478 | 6.649 | 6
41 | 3.785 | 7.8813 | 7.883 | 29.8161 | 29.8183 | 6.8451 | 2.7307 | 14 | 14 | 162 | 6.8201 | 0.97403 | 6.643
45 | 252.158 | 6.7945 | 6.7947 | 33.9061 | 33.907 | 6.832 | 1.1208 | 1 | 1 | 163 | 7.0226 | 0.27241 | 1.913
45 | 175.419 | 7.1716 | 7.1725 | 33.7977 | 33.7982 | 6.778 | 1.2652 | 2 | 2 | 164 | 6.9 | 0.3313 | 2.286 | 3
45 | 147.864 | 7.301 | 7.3012 | 33.7115 | 33.7124 | 6.7618 | 1.3391 | 3 | 3 | 165 | 6.8598 | 0.37669 | 2.584
45 | 147.402 | 7.3005 | 7.3009 | 33.711 | 33.712 | 6.7619 | 1.3396 | 4 | 4 | 166 | 6.8595 | 0.36883 | 2.53 | 6
45 | 122.618 | 7.4867 | 7.4866 | 33.5665 | 33.5676 | 6.7396 | 1.455 | 5 | 5 | 167 | 6.816 | 0.42341 | 2.886
45 | 104.34 | 7.9751 | 7.9862 | 32.9116 | 32.9051 | 6.694 | 1.9021 | 6 | 6 | 168 | 6.7566 | 0.63804 | 4.311
45 | 74.433 | 7.7592 | 7.7593 | 31.5786 | 31.579 | 6.7856 | 2.1802 | 7 | 7 | 169 | 6.8217 | 0.73853 | 5.038
45 | 52.696 | 7.8686 | 7.8707 | 30.9207 | 30.9194 | 6.7979 | 2.4157 | 8 | 8 | 170 | 6.8156 | 0.84424 | 5.754
45 | 41.695 | 7.9261 | 7.9259 | 30.7508 | 30.7508 | 6.7965 | 2.4983 | 9 | 9 | 171 | 6.805 | 0.87568 | 5.959 | 6
45 | 30.386 | 7.9753 | 7.9758 | 30.6682 | 30.6684 | 6.7926 | 2.5684 | 10 | 10 | 172 | 6.7915 | 0.91143 | 6.19
45 | 19.368 | 8.1278 | 8.1274 | 30.3967 | 30.398 | 6.7811 | 2.6377 | 11 | 11 | 173 | 6.7714 | 0.93511 | 6.332
45 | 10.084 | 8.3332 | 8.3347 | 29.9744 | 29.9746 | 6.7682 | 2.7263 | 12 | 12 | 174 | 6.7518 | 0.98669 | 6.662
45 | 4.552 | 8.3983 | 8.3984 | 29.1366 | 29.1356 | 6.7951 | 2.8281 | 13 | 13 | 175 | 6.7743 | 1.0193 | 6.905
45 | 3.894 | 8.3998 | 8.3999 | 29.1513 | 29.1511 | 6.7942 | 2.8333 | 14 | 14 | 176 | 6.7729 | 1.0183 | 6.897
48 | 223.094 | 6.891 | 6.8913 | 33.8792 | 33.8802 | 6.8181 | 1.1623 | 1 | 1 | 177 | 6.9824 | 0.28471 | 1.988
48 | 204.651 | 6.8959 | 6.8963 | 33.8763 | 33.8775 | 6.8174 | 1.1663 | 2 | 2 | 178 | 6.9653 | 0.28714 | 2
48 | 172.085 | 6.9218 | 6.9224 | 33.8672 | 33.8679 | 6.8138 | 1.1781 | 3 | 3 | 179 | 6.9325 | 0.29412 | 2.039
48 | 154.982 | 7.194 | 7.1969 | 33.7294 | 33.7291 | 6.7775 | 1.3159 | 4 | 4 | 180 | 6.8816 | 0.35602 | 2.45 | 6
48 | 125.485 | 7.5368 | 7.5437 | 33.4964 | 33.4915 | 6.735 | 1.5226 | 5 | 5 | 181 | 6.8141 | 0.44657 | 3.043
48 | 98.525 | 8.0385 | 8.0395 | 32.8157 | 32.8148 | 6.6886 | 1.9475 | 6 | 6 | 182 | 6.7465 | 0.63841 | 4.307
48 | 77.084 | 7.8169 | 7.8164 | 31.4954 | 31.4984 | 6.7804 | 2.2258 | 7 | 7 | 183 | 6.8191 | 0.75919 | 5.177 | 3
48 | 50.818 | 8.0434 | 8.0436 | 30.8997 | 30.8966 | 6.7719 | 2.4802 | 8 | 8 | 184 | 6.7889 | 0.86449 | 5.869
48 | 40.687 | 8.0663 | 8.067 | 30.7894 | 30.7901 | 6.7732 | 2.5216 | 9 | 9 | 185 | 6.7816 | 0.87766 | 5.952
48 | 28.307 | 8.126 | 8.1298 | 30.4482 | 30.4447 | 6.7791 | 2.635 | 10 | 10 | 186 | 6.7772 | 0.92871 | 6.294 | 6
48 | 19.638 | 8.4821 | 8.486 | 29.5643 | 29.5377 | 6.7635 | 2.8414 | 11 | 11 | 187 | 6.7564 | 1.0238 | 6.917
48 | 10.106 | 8.5245 | 8.5246 | 29.2216 | 29.2244 | 6.772 | 2.8834 | 12 | 12 | 188 | 6.757 | 1.0358 | 6.999
48 | 5.334 | 8.5416 | 8.5407 | 29.1642 | 29.1651 | 6.7719 | 2.8991 | 13 | 13 | 189 | 6.7529 | 1.0397 | 7.021
48 | 4.054 | 8.5417 | 8.5417 | 29.1643 | 29.1649 | 6.7719 | 2.8911 | 14 | 14 | 190 | 6.7518 | 1.0394 | 7.018
51 | 185.802 | 7.0694 | 7.0693 | 33.8098 | 33.8108 | 6.7933 | 1.242 | 1 | 1 | 191 | 6.9245 | 0.32031 | 2.218
51 | 174.175 | 7.0771 | 7.0764 | 33.8034 | 33.805 | 6.7924 | 1.2465 | 2 | 2 | 192 | 6.9132 | 0.32286 | 2.232
51 | 149.728 | 7.1532 | 7.1532 | 33.7491 | 33.7503 | 6.783 | 1.2908 | 3 | 3 | 193 | 6.8823 | 0.34204 | 2.354
51 | 124.167 | 7.7406 | 7.7448 | 33.1175 | 33.1119 | 6.7206 | 1.6973 | 4 | 4 | 194 | 6.7994 | 0.51902 | 3.529
51 | 100.782 | 8.013 | 8.0132 | 32.7008 | 32.7017 | 6.6974 | 1.9596 | 5 | 5 | 195 | 6.7572 | 0.63458 | 4.288 | 6
51 | 74.292 | 7.9959 | 7.9961 | 31.9047 | 31.9052 | 6.7349 | 2.1675 | 6 | 6 | 196 | 6.7719 | 0.72195 | 4.889
51 | 49.972 | 7.9338 | 7.9342 | 31.2022 | 31.2037 | 6.7754 | 2.3448 | 7 | 7 | 197 | 6.791 | 0.79679 | 5.411
51 | 39.458 | 7.9171 | 7.9174 | 30.9735 | 30.9747 | 6.7881 | 2.4011 | 8 | 8 | 198 | 6.7945 | 0.82567 | 5.61
51 | 29.536 | 8.0076 | 8.01 | 30.7719 | 30.7703 | 6.783 | 2.5506 | 9 | 9 | 199 | 6.7814 | 0.89746 | 6.086 | 6
51 | 20.41 | 8.0985 | 8.0964 | 30.6864 | 30.6889 | 6.7728 | 2.62 | 10 | 10 | 200 | 6.7639 | 0.92329 | 6.245
51 | 8.993 | 8.1194 | 8.1188 | 30.6772 | 30.6782 | 6.77 | 2.6426 | 11 | 11 | 201 | 6.7513 | 0.93197 | 6.292
51 | 4.844 | 8.1174 | 8.1176 | 30.6794 | 30.6803 | 6.7702 | 2.6465 | 12 | 12 | 202 | 6.7479 | 0.93125 | 6.284
51 | 3.765 | 8.1164 | 8.1166 | 30.6798 | 30.6808 | 6.7704 | 2.6459 | 13 | 13 | 203 | 6.7471 | 0.93151 | 6.285
56 | 115.47 | 7.8056 | 7.8053 | 32.1124 | 32.1221 | 6.7549 | 1.9926 | 1 | 1 | 204 | 6.8268 | 0.65243 | 4.454
56 | 100.467 | 7.8624 | 7.8631 | 31.5387 | 31.5337 | 6.7715 | 2.1762 | 2 | 2 | 205 | 6.8308 | 0.73022 | 4.988
56 | 74.974 | 7.9033 | 7.9036 | 31.0344 | 31.0348 | 6.7875 | 2.3093 | 3 | 3 | 206 | 6.8248 | 0.78756 | 5.375
56 | 50.395 | 7.9624 | 7.9628 | 30.6498 | 30.6497 | 6.7954 | 2.4133 | 4 | 4 | 207 | 6.8117 | 0.83107 | 5.661
56 | 39.928 | 7.9892 | 7.9905 | 30.3766 | 30.3699 | 6.8034 | 2.475 | 5 | 5 | 208 | 6.8107 | 0.86203 | 5.871
56 | 29.166 | 7.9585 | 7.9592 | 30.1532 | 30.1543 | 6.818 | 2.5589 | 6 | 6 | 209 | 6.8158 | 0.89381 | 6.092
56 | 20.495 | 7.9578 | 7.958 | 30.1383 | 30.1387 | 6.8188 | 2.5638 | 7 | 7 | 210 | 6.809 | 0.89558 | 6.098
56 | 10.238 | 7.9587 | 7.9589 | 30.1378 | 30.1381 | 6.8187 | 2.5684 | 8 | 8 | 211 | 6.7999 | 0.89662 | 6.097 | 6
56 | 4.881 | 7.9556 | 7.9557 | 30.1356 | 30.1361 | 6.8193 | 2.5725 | 9 | 9 | 212 | 6.7958 | 0.8979 | 6.102
56 | 3.392 | 7.9613 | 7.9635 | 30.1349 | 30.1269 | 6.8184 | 2.5721 | 10 | 10 | 213 | 6.7937 | 0.8998 | 6.113
61 | 126.452 | 7.8121 | 7.8124 | 32.1246 | 32.1239 | 6.7533 | 2.0019 | 1 | 1 | 214 | 6.8349 | 0.65692 | 4.49
61 | 100.172 | 7.8627 | 7.863 | 31.4616 | 31.4615 | 6.7748 | 2.1932 | 2 | 2 | 215 | 6.8339 | 0.73647 | 5.033
61 | 74.787 | 7.9088 | 7.9092 | 31.0005 | 31.0011 | 6.7881 | 2.3191 | 3 | 3 | 216 | 6.8254 | 0.79029 | 5.394
61 | 50.321 | 7.9014 | 7.9017 | 30.755 | 30.7553 | 6.8002 | 2.379 | 4 | 4 | 217 | 6.816 | 0.81646 | 5.565
61 | 40.94 | 7.9292 | 7.9294 | 30.5787 | 30.5796 | 6.8037 | 2.4193 | 5 | 5 | 218 | 6.8115 | 0.83579 | 5.693
61 | 30.021 | 7.9747 | 7.975 | 30.3461 | 30.3463 | 6.807 | 2.4764 | 6 | 6 | 219 | 6.8056 | 0.85944 | 5.849 | 6
61 | 20.129 | 7.9754 | 7.9764 | 30.1845 | 30.1861 | 6.814 | 2.5397 | 7 | 7 | 220 | 6.804 | 0.88477 | 6.02
61 | 9.973 | 7.9511 | 7.9516 | 30.1692 | 30.1698 | 6.8185 | 2.5515 | 8 | 8 | 221 | 6.7994 | 0.89007 | 6.052
61 | 4.627 | 7.9473 | 7.9478 | 30.1703 | 30.1708 | 6.819 | 2.554 | 9 | 9 | 222 | 6.7953 | 0.89106 | 6.055
61 | 3.664 | 7.9473 | 7.9487 | 30.1695 | 30.1699 | 6.819 | 2.5544 | 10 | 10 | 223 | 6.7945 | 0.89337 | 6.07
64 | 149.14 | 7.8619 | 7.8621 | 31.8153 | 31.8159 | 6.7593 | 2.0817 | 1 | 1 | 224 | 6.8612 | 0.68647 | 4.71
64 | 125.465 | 7.8602 | 7.8604 | 31.7954 | 31.7961 | 6.7605 | 2.0907 | 2 | 2 | 225 | 6.8416 | 0.69224 | 4.736
64 | 99.802 | 7.8747 | 7.875 | 31.4314 | 31.4299 | 6.7743 | 2.1974 | 3 | 3 | 226 | 6.8332 | 0.73612 | 5.03
64 | 74.784 | 7.8983 | 7.8976 | 30.9982 | 31.0128 | 6.7899 | 2.3048 | 4 | 4 | 227 | 6.827 | 0.77882 | 5.317
64 | 50.223 | 7.9248 | 7.9252 | 30.5238 | 30.5185 | 6.8068 | 2.4621 | 5 | 5 | 228 | 6.8227 | 0.85244 | 5.816 | 6
64 | 40.426 | 7.927 | 7.9273 | 30.4649 | 30.4653 | 6.8091 | 2.5103 | 6 | 6 | 229 | 6.8165 | 0.87303 | 5.951
64 | 30.039 | 7.9232 | 7.9244 | 30.4509 | 30.4475 | 6.8103 | 2.5122 | 7 | 7 | 230 | 6.8086 | 0.8808 | 5.997 | 3
64 | 19.566 | 7.9279 | 7.9279 | 30.4515 | 30.4522 | 6.8095 | 2.5004 | 8 | 8 | 231 | 6.7987 | 0.87031 | 5.917
64 | 9.406 | 7.9186 | 7.9176 | 30.4273 | 30.4247 | 6.8121 | 2.5372 | 9 | 9 | 232 | 6.7923 | 0.88026 | 5.979
64 | 4.923 | 7.9225 | 7.9229 | 30.4257 | 30.4274 | 6.8115 | 2.5335 | 10 | 10 | 233 | 6.7879 | 0.87862 | 5.964
64 | 3.668 | 7.9224 | 7.9224 | 30.4292 | 30.4304 | 6.8114 | 2.5248 | 11 | 11 | 234 | 6.7867 | 0.8801 | 5.973
69 | 211.685 | 7.9037 | 7.904 | 30.4325 | 30.4315 | 6.8141 | 2.4097 | 1 | 1 | 235 | 6.9727 | 0.82967 | 5.785
69 | 201.106 | 7.9045 | 7.9047 | 30.3163 | 30.3213 | 6.8192 | 2.4311 | 2 | 2 | 236 | 6.9684 | 0.83434 | 5.814
69 | 175.596 | 7.9033 | 7.9034 | 30.2692 | 30.2708 | 6.8214 | 2.4518 | 3 | 3 | 237 | 6.948 | 0.84514 | 5.872
69 | 150.297 | 7.9002 | 7.9003 | 30.2775 | 30.2786 | 6.8216 | 2.4543 | 4 | 4 | 238 | 6.9256 | 0.84975 | 5.885 | 6
69 | 125.64 | 7.8977 | 7.8978 | 30.2089 | 30.2113 | 6.825 | 2.4744 | 5 | 5 | 239 | 6.9072 | 0.85896 | 5.933
69 | 100.008 | 7.8959 | 7.8961 | 30.1811 | 30.1843 | 6.8265 | 2.4886 | 6 | 6 | 240 | 6.8861 | 0.86377 | 5.948
69 | 74.534 | 7.8975 | 7.8975 | 30.1014 | 30.0996 | 6.8298 | 2.523 | 7 | 7 | 241 | 6.867 | 0.88204 | 6.057
69 | 49.629 | 7.9026 | 7.9026 | 30.033 | 30.0342 | 6.8321 | 2.5501 | 8 | 8 | 242 | 6.8474 | 0.8907 | 6.099
69 | 39.956 | 7.912 | 7.9117 | 29.9808 | 29.9835 | 6.833 | 2.5782 | 9 | 9 | 243 | 6.8399 | 0.90177 | 6.168
69 | 29.631 | 7.9152 | 7.9152 | 29.9708 | 29.9706 | 6.8329 | 2.5831 | 10 | 10 | 244 | 6.8308 | 0.90575 | 6.187
69 | 19.778 | 7.9165 | 7.9165 | 29.9566 | 29.9574 | 6.8333 | 2.5951 | 11 | 11 | 245 | 6.8226 | 0.90845 | 6.198 | 6
69 | 9.794 | 7.9223 | 7.9214 | 29.9534 | 29.9542 | 6.8326 | 2.598 | 12 | 12 | 246 | 6.8132 | 0.91323 | 6.222
69 | 5.407 | 7.9208 | 7.9213 | 29.9598 | 29.9602 | 6.8325 | 2.6027 | 13 | 13 | 247 | 6.8093 | 0.91493 | 6.23
69 | 3.766 | 7.9533 | 7.9543 | 29.9469 | 29.9454 | 6.828 | 2.6113 | 14 | 14 | 248 | 6.8036 | 0.9157 | 6.23
76 | 179.384 | 7.8958 | 7.8961 | 30.1843 | 30.1848 | 6.8264 | 2.4139 | 1 | 1 | 249 | 6.9563 | 0.8309 | 5.78
76 | 124.698 | 7.8915 | 7.8918 | 30.0705 | 30.071 | 6.8321 | 2.4605 | 3 | 3 | 251 | 6.9136 | 0.8518 | 5.889
76 | 100.26 | 7.8822 | 7.8826 | 29.9781 | 29.9775 | 6.8377 | 2.4733 | 4 | 4 | 252 | 6.8975 | 0.85929 | 5.927
76 | 75.218 | 7.8832 | 7.8818 | 29.9398 | 29.9328 | 6.8393 | 2.4982 | 5 | 5 | 253 | 6.877 | 0.86724 | 5.964
76 | 49.938 | 7.8045 | 7.8046 | 29.6957 | 29.6962 | 6.8625 | 2.4927 | 6 | 6 | 254 | 6.8776 | 0.86324 | 5.937 | 3
76 | 40.283 | 7.7957 | 7.7955 | 29.6488 | 29.6491 | 6.866 | 2.5008 | 7 | 7 | 255 | 6.8725 | 0.86766 | 5.963 | 6
76 | 30.267 | 7.804 | 7.8034 | 29.6224 | 29.6227 | 6.8658 | 2.517 | 8 | 8 | 256 | 6.8636 | 0.87243 | 5.988
76 | 19.985 | 7.7923 | 7.7917 | 29.5937 | 29.5946 | 6.869 | 2.5195 | 9 | 9 | 257 | 6.8576 | 0.87698 | 6.014
76 | 10.076 | 7.865 | 7.8767 | 29.4905 | 29.4815 | 6.8622 | 2.6425 | 10 | 10 | 258 | 6.8426 | 0.9131 | 6.248
23 | 19.873 | 7.9456 | 7.9453 | 28.6344 | 28.635 | 6.8879 | 3.0033 | 12 | 12 | 96 | 6.8774 | 1.0521 | 7.236
36 | 10.006 | 7.9066 | 7.9068 | 28.6249 | 28.6253 | 6.8945 | 2.3454 | 14 | 14 | 146 | 6.875 | 1.2669 | 8.71
76 | 150.054 | 7.893 | 7.8932 | 30.1634 | 30.1633 | 6.8278 | 2.4268 | 2 | 2 | 250 | 6.9316 | 0.75985 | 5.267
76 | 4.916 | 8.1325 | 8.1294 | 27.833 | 27.8348 | 6.8945 | 3.4745 | 11 | 11 | 259 | 6.8719 | 1.15 | 7.903
76 | 3.751 | 8.377 | 8.358 | 27.3006 | 27.3229 | 6.88 | 3.3954 | 12 | 12 | 260 | 6.8581 | 1.2119 | 8.311

## Sheet1
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle_Number | Bottle:Position | Sample_Number | Phi | DObot/Phi | sort | Oxygen:Dissolved | Flag:Oxygen:Dissolved
10 | 308.264 | 9.1851 | 9.1856 | 30.7238 | 30.7244 | 6.608 | 1.554 | 1 | 1 | 12 | 6.8539 | 0.66852 | 0 | 4.582 |  |  | Summary of Soc Voffset including the original values in the factory calibration
10 | 299.905 | 9.1617 | 9.1597 | 30.7148 | 30.7144 | 6.6118 | 1.5692 | 2 | 2 | 13 | 6.8504 | 0.96768 | 0 | 6.629 | 6 |  |  | m | b | Soc | Voffset
10 | 249.69 | 8.9525 | 8.9532 | 30.6214 | 30.6222 | 6.6467 | 1.6791 | 3 | 3 | 14 | 6.8411 | 0.61935 | 0 | 4.237 |  |  | original | 0.4329 | -0.22177 | 0.4329 | -0.5123
10 | 200.042 | 8.8967 | 8.8965 | 30.5819 | 30.5835 | 6.6567 | 1.716 | 4 | 4 | 15 | 6.8077 | 0.65235 | 0 | 4.441 |  |  | fit-all | 0.4313 | -0.2052 | 0.4313 | -0.47577
10 | 175.195 | 8.836 | 8.8363 | 30.4975 | 30.4981 | 6.6694 | 1.7694 | 5 | 5 | 16 | 6.7986 | 1.2179 | 0 | 8.28 |  |  | fit-5 outliers | 0.4346 | -0.2135 | 0.4346 | -0.49126
10 | 149.869 | 8.627 | 8.6258 | 30.3994 | 30.3996 | 6.705 | 1.8597 | 6 | 6 | 17 | 6.8114 | 1.2826 | 0 | 8.736 |  | Test runs using subsets of bottles
10 | 124.915 | 8.5182 | 8.5322 | 30.2877 | 30.2906 | 6.7263 | 1.8983 | 7 | 7 | 18 | 6.8105 | 0.59261 | 0 | 4.036 | [9 empty] | DO Concentration cast #21 |  |  |  | min | max
10 | 100.394 | 8.545 | 8.5454 | 30.0692 | 30.0697 | 6.7318 | 1.9001 | 8 | 8 | 19 | 6.7949 | 0.72981 | 0 | 4.959 | [7 empty] | soc-opt | voffset-opt | Min | Max | min-opt | max-opt | % error | % error
10 | 74.767 | 8.3272 | 8.3273 | 29.5963 | 29.5987 | 6.7858 | 2.1899 | 9 | 9 | 20 | 6.8256 | 0.95201 | 0 | 6.498 |  | All data | original | 0.4329 | -0.2218 | 0.4329 | -0.5123 | -0.0017 | -0.02104 | 2.7125 | 8.3726 | -0.07303 | -0.09575 | -2.6924 | -1.1436
10 | 50.119 | 8.0039 | 8.0039 | 29.0078 | 29.0064 | 6.862 | 2.6581 | 10 | 10 | 21 | 6.8785 | 1.2221 | 0 | 8.406 |  | Optimal | all-5 | 0.4346 | -0.2135 | 0.4346 | -0.49126 | 0 | 0 | 2.7855 | 8.4683 | 0 | 0 | 0 | 0
10 | 39.987 | 7.9943 | 7.9945 | 28.818 | 28.8205 | 6.872 | 2.9746 | 11 | 11 | 22 | 6.8795 | 0.64888 | 0 | 4.464 |  | Run1 | every other cast-a | 0.4343 | -0.2112 | 0.4343 | -0.4863 | -0.0003 | 0.0049602 | 2.7985 | 8.4774 | 0.01296 | 0.0091 | 0.46311 | 0.10734
10 | 30.333 | 7.9177 | 7.9181 | 28.7758 | 28.7764 | 6.886 | 3.0057 | 12 | 12 | 23 | 6.8845 | 0.73499 | 0 | 5.06 |  | Run2 | every other cast-b | 0.435 | -0.2164 | 0.435 | -0.49747 | 0.0004 | -0.0062113 | 2.7696 | 8.4575 | -0.01585 | -0.01079 | -0.57227 | -0.12758
10 | 20.233 | 7.9278 | 7.9274 | 28.7756 | 28.7763 | 6.8844 | 3.0019 | 13 | 13 | 24 | 6.8741 | 0.40326 | 0 | 2.772 |  | Run3 | every other bottle-a | 0.4346 | -0.2115 | 0.4346 | -0.48665 | 0 | 0.0046056 | 2.7992 | 8.482 | 0.0137 | 0.01375 | 0.48943 | 0.16211
10 | 9.936 | 7.9223 | 7.9226 | 28.7707 | 28.7715 | 6.8855 | 3.0319 | 14 | 14 | 25 | 6.866 | 0.64535 | 0 | 4.431 | 6 | Run4 | every other bottle-b | 0.4348 | -0.216 | 0.4348 | -0.49678 | 0.0002 | -0.0055201 | 2.7705 | 8.4557 | -0.01504 | -0.01259 | -0.54287 | -0.14889
10 | 5.146 | 7.9247 | 7.9248 | 28.7703 | 28.7709 | 6.8851 | 3.0418 | 15 | 15 | 26 | 6.8615 | 0.72215 | 0 | 4.955 |  | Run5 | every 4th bottle a | 0.4351 | -0.2107 | 0.4351 | -0.48426 | 0.0005 | 0.0070035 | 2.8096 | 8.499 | 0.02407 | 0.03071 | 0.85671 | 0.36134
10 | 3.559 | 7.9327 | 7.9297 | 28.7707 | 28.7714 | 6.8838 | 3.0416 | 16 | 16 | 27 | 6.8589 | 0.86924 | 0 | 5.962 |  | Run6 | every 4th bottle b | 0.4341 | -0.2125 | 0.4341 | -0.48952 | -0.0005 | 0.0017415 | 2.7877 | 8.4639 | 0.00217 | -0.00437 | 0.077843 | -0.051631
16 | 164.164 | 8.6463 | 8.6467 | 30.4896 | 30.4906 | 6.6982 | 1.7919 | 1 | 1 | 44 | 6.817 | 0.35939 | 0 | 2.45 |  | Run7 | selection a | 0.4348 | -0.216 | 0.4348 | -0.49678 | 0.0002 | -0.0055201 | 2.7705 | 8.4557 | -0.015 | -0.0126 | -0.54142 | -0.14901
16 | 149.956 | 8.6378 | 8.638 | 30.4686 | 30.4697 | 6.7004 | 1.8029 | 2 | 2 | 45 | 6.8068 | 0.86222 | 0 | 5.869 |  | Run8 | selection b | 0.4346 | -0.2115 | 0.4346 | -0.48665 | 0 | 0.0046056 | 2.7992 | 8.482 | 0.0137 | 0.01375 | 0.48943 | 0.16211
16 | 124.622 | 8.5675 | 8.5678 | 30.2824 | 30.2803 | 6.7191 | 1.8735 | 3 | 3 | 46 | 6.8033 | 1.0288 | 0 | 6.999 |  |  |  | max-min |  | 0.0008 | 0.013215
16 | 99.975 | 8.545 | 8.5452 | 30.0405 | 30.0411 | 6.733 | 1.9205 | 4 | 4 | 47 | 6.7958 | 0.32844 | 0 | 2.232 | 6
16 | 74.119 | 8.282 | 8.2809 | 29.7466 | 29.7459 | 6.7861 | 2.0256 | 5 | 5 | 48 | 6.825 | 0.71634 | 0 | 4.889
16 | 49.4 | 8.003 | 8.0031 | 29.4214 | 29.4226 | 6.8437 | 2.1978 | 6 | 6 | 49 | 6.8595 | 0.91042 | 0 | 6.245 | 6
16 | 40.324 | 7.9512 | 7.9502 | 29.2659 | 29.2658 | 6.8587 | 2.3586 | 7 | 7 | 50 | 6.8662 | 0.64868 | 0 | 4.454
16 | 30.302 | 7.9542 | 7.9513 | 29.1734 | 29.1772 | 6.8624 | 2.489 | 8 | 8 | 51 | 6.8611 | 0.85569 | 0 | 5.871
16 | 19.776 | 7.986 | 7.9866 | 29.0786 | 29.0762 | 6.8617 | 2.6832 | 9 | 9 | 52 | 6.8513 | 0.89063 | 0 | 6.102
16 | 9.917 | 8.0027 | 8.0033 | 28.9526 | 28.9534 | 6.8647 | 2.9215 | 10 | 10 | 53 | 6.8458 | 0.78793 | 0 | 5.394
16 | 5.293 | 8.0098 | 8.0091 | 28.9602 | 28.9611 | 6.8632 | 2.9165 | 11 | 11 | 54 | 6.8403 | 0.88007 | 0 | 6.02
16 | 4.334 | 8.0063 | 8.0067 | 28.9588 | 28.9594 | 6.8638 | 2.9285 | 12 | 12 | 55 | 6.8401 | 0.68859 | 0 | 4.71
21 | 317.599 | 9.2832 | 9.2834 | 30.9406 | 30.9408 | 6.5844 | 1.433 | 1 | 1 | 69 | 6.8383 | 0.94585 | 0 | 6.468
21 | 299.699 | 9.2533 | 9.2535 | 30.9198 | 30.9201 | 6.5897 | 1.4575 | 2 | 2 | 70 | 6.8279 | 1.114 | 0 | 7.606
21 | 249.704 | 9.1458 | 9.1513 | 30.8543 | 30.8568 | 6.6082 | 1.5324 | 3 | 3 | 71 | 6.8029 | 0.50611 | 0 | 3.443 | 6
21 | 199.778 | 8.4644 | 8.4671 | 30.5634 | 30.5649 | 6.7224 | 1.8354 | 4 | 4 | 72 | 6.8717 | 0.65515 | 0 | 4.502 | 6
21 | 174.94 | 8.139 | 8.1392 | 30.3707 | 30.3712 | 6.7805 | 1.9962 | 5 | 5 | 73 | 6.9071 | 1.0307 | 0 | 7.119
21 | 150.175 | 7.9356 | 7.936 | 30.2236 | 30.2239 | 6.8185 | 2.099 | 6 | 6 | 74 | 6.9225 | 1.135 | 0 | 7.857
21 | 124.839 | 7.871 | 7.8716 | 30.1075 | 30.108 | 6.8337 | 2.1488 | 7 | 7 | 75 | 6.9151 | 0.56514 | 0 | 3.908
21 | 100.057 | 7.939 | 7.9399 | 30.0085 | 30.0088 | 6.8275 | 2.1398 | 8 | 8 | 76 | 6.8874 | 0.74891 | 0 | 5.158
21 | 74.81 | 8.1937 | 8.1937 | 29.6722 | 29.6723 | 6.803 | 2.0653 | 9 | 9 | 77 | 6.842 | 1.0592 | 0 | 7.247
21 | 50.009 | 7.8461 | 7.8461 | 29.1522 | 29.1526 | 6.8803 | 2.3777 | 10 | 10 | 78 | 6.8958 | 0.63822 | 0 | 4.401
21 | 39.958 | 7.8476 | 7.849 | 28.9888 | 28.9645 | 6.8874 | 2.5109 | 11 | 11 | 79 | 6.894 | 0.66899 | 0 | 4.612
21 | 29.689 | 7.8856 | 7.892 | 28.773 | 28.7669 | 6.8911 | 2.6917 | 12 | 12 | 80 | 6.8889 | 1.257 | 0 | 8.659
21 | 19.746 | 8.0468 | 8.0472 | 27.9257 | 27.9247 | 6.9038 | 3.2879 | 13 | 13 | 81 | 6.8938 | 0.40776 | 0 | 2.811
21 | 10.12 | 8.0579 | 8.0591 | 27.8448 | 27.8453 | 6.9057 | 3.3047 | 14 | 14 | 82 | 6.8872 | 0.65847 | 0 | 4.535
21 | 4.598 | 8.0527 | 8.053 | 27.8445 | 27.845 | 6.9065 | 3.3042 | 15 | 15 | 83 | 6.8832 | 0.67818 | 0 | 4.668
21 | 3.652 | 8.0501 | 8.0426 | 27.845 | 27.8458 | 6.9069 | 3.3047 | 16 | 16 | 84 | 6.8827 | 1.2132 | 0 | 8.35
27 | 343.639 | 9.3506 | 9.3499 | 30.9918 | 30.9919 | 6.5724 | 1.3804 | 1 | 1 | 100 | 6.849 | 0.77807 | 0 | 5.329
27 | 300.448 | 9.2263 | 9.2269 | 30.9211 | 30.922 | 6.5936 | 1.4593 | 2 | 2 | 101 | 6.8324 | 0.87904 | 0 | 6.006
27 | 249.911 | 9.0395 | 9.0367 | 30.8224 | 30.8217 | 6.6252 | 1.5606 | 3 | 3 | 102 | 6.8198 | 0.75354 | 0 | 5.139
27 | 200.103 | 8.2511 | 8.2523 | 30.4997 | 30.5006 | 6.7576 | 1.9104 | 4 | 4 | 103 | 6.9066 | 0.77433 | 0 | 5.348
27 | 174.774 | 8.1842 | 8.1843 | 30.4366 | 30.4374 | 6.7707 | 1.9585 | 5 | 5 | 104 | 6.8972 | 0.91458 | 0 | 6.308
27 | 149.75 | 8.2746 | 8.2752 | 30.4233 | 30.4241 | 6.7574 | 1.9413 | 6 | 6 | 105 | 6.8622 | 0.96893 | 0 | 6.649
27 | 125.083 | 7.955 | 7.9557 | 30.2776 | 30.2782 | 6.8131 | 2.0857 | 7 | 7 | 106 | 6.895 | 0.37477 | 0 | 2.584
27 | 100.086 | 7.7707 | 7.7713 | 30.1236 | 30.1241 | 6.8486 | 2.1891 | 8 | 8 | 107 | 6.9078 | 0.72933 | 0 | 5.038
27 | 75.395 | 7.9622 | 7.9621 | 30.0016 | 30.0026 | 6.8242 | 2.1483 | 9 | 9 | 108 | 6.8625 | 0.9227 | 0 | 6.332
27 | 49.778 | 8.072 | 8.0718 | 29.7091 | 29.7094 | 6.8202 | 2.1393 | 10 | 10 | 109 | 6.8366 | 0.29079 | 0 | 1.988
27 | 40.009 | 7.9438 | 7.9438 | 29.5187 | 29.5209 | 6.8486 | 2.2012 | 11 | 11 | 110 | 6.8558 | 0.44386 | 0 | 3.043
27 | 29.748 | 7.7839 | 7.7829 | 29.2549 | 29.2551 | 6.8855 | 2.3136 | 12 | 12 | 111 | 6.8827 | 0.86478 | 0 | 5.952
27 | 19.567 | 7.6798 | 7.6761 | 28.8891 | 28.899 | 6.9185 | 2.505 | 13 | 13 | 112 | 6.906 | 1.0166 | 0 | 7.021 | 6
27 | 9.606 | 7.8443 | 7.8431 | 28.2779 | 28.2789 | 6.92 | 3.3719 | 14 | 14 | 113 | 6.8997 | 0.34117 | 0 | 2.354
27 | 4.878 | 7.8289 | 7.8299 | 28.2751 | 28.277 | 6.9225 | 3.3716 | 15 | 15 | 114 | 6.898 | 0.78443 | 0 | 5.411
27 | 3.85 | 7.8312 | 7.8304 | 28.2724 | 28.2732 | 6.9223 | 3.3826 | 16 | 16 | 115 | 6.8968 | 0.9123 | 0 | 6.292
36 | 317.099 | 9.1891 | 9.188 | 30.9055 | 30.9055 | 6.5997 | 1.4389 | 1 | 1 | 133 | 6.853 | 0.51919 | 0 | 3.558 | 6
36 | 299.941 | 9.0522 | 9.052 | 30.8387 | 30.8392 | 6.6227 | 1.5192 | 2 | 2 | 134 | 6.8609 | 0.6094 | 0 | 4.181
36 | 249.395 | 8.8839 | 8.8835 | 30.7621 | 30.7627 | 6.6509 | 1.6143 | 3 | 3 | 135 | 6.8446 | 1.093 | 0 | 7.481
36 | 199.863 | 8.4537 | 8.4539 | 30.5716 | 30.5718 | 6.7237 | 1.8146 | 4 | 4 | 136 | 6.873 | 1.1103 | 0 | 7.631
36 | 174.802 | 8.2087 | 8.2082 | 30.4495 | 30.4505 | 6.7664 | 1.9424 | 5 | 5 | 137 | 6.8929 | 0.54258 | 0 | 3.74
36 | 149.577 | 8.0083 | 8.01 | 30.2745 | 30.2761 | 6.8049 | 2.0686 | 6 | 6 | 138 | 6.9087 | 0.64079 | 0 | 4.427
36 | 124.772 | 7.9461 | 7.9497 | 30.084 | 30.0855 | 6.823 | 2.1444 | 7 | 7 | 139 | 6.9047 | 1.092 | 0 | 7.54
36 | 99.762 | 7.7611 | 7.7597 | 29.9179 | 29.9154 | 6.8593 | 2.302 | 8 | 8 | 140 | 6.9182 | 1.1383 | 0 | 7.875
36 | 74.06 | 7.8382 | 7.8381 | 29.8429 | 29.8438 | 6.8506 | 2.2441 | 9 | 9 | 141 | 6.8871 | 0.5988 | 0 | 4.124
36 | 50.01 | 7.8459 | 7.8462 | 29.6151 | 29.6165 | 6.8596 | 2.2633 | 10 | 10 | 142 | 6.875 | 0.82313 | 0 | 5.659
36 | 39.768 | 7.7925 | 7.7921 | 29.4969 | 29.4972 | 6.8733 | 2.3149 | 11 | 11 | 143 | 6.8794 | 1.0581 | 0 | 7.279
36 | 29.943 | 7.7523 | 7.751 | 29.3385 | 29.3297 | 6.8867 | 2.3891 | 12 | 12 | 144 | 6.8839 | 0.6768 | 0 | 4.659 | 6
36 | 20.046 | 7.7441 | 7.7449 | 29.0944 | 29.0901 | 6.899 | 2.5124 | 13 | 13 | 145 | 6.8874 | 0.8758 | 0 | 6.032
36 | 5.073 | 8.0217 | 8.0376 | 28.6039 | 28.6028 | 6.8773 | 3.6235 | 15 | 15 | 147 | 6.8543 | 1.2725 | 0 | 8.722
36 | 3.705 | 8.051 | 8.0688 | 28.5994 | 28.5983 | 6.8729 | 3.7273 | 16 | 16 | 148 | 6.8489 | 0.41905 | 0 | 2.87
48 | 223.094 | 6.891 | 6.8913 | 33.8792 | 33.8802 | 6.8181 | 1.1623 | 1 | 1 | 177 | 6.9824 | 1.0051 | 0 | 7.018
48 | 204.651 | 6.8959 | 6.8963 | 33.8763 | 33.8775 | 6.8174 | 1.1663 | 2 | 2 | 178 | 6.9653 | 0.50666 | 0 | 3.529
48 | 172.085 | 6.9218 | 6.9224 | 33.8672 | 33.8679 | 6.8138 | 1.1781 | 3 | 3 | 179 | 6.9325 | 0.80923 | 0 | 5.61
48 | 154.982 | 7.194 | 7.1969 | 33.7294 | 33.7291 | 6.7775 | 1.3159 | 4 | 4 | 180 | 6.8816 | 0.91316 | 0 | 6.284
48 | 125.485 | 7.5368 | 7.5437 | 33.4964 | 33.4915 | 6.735 | 1.5226 | 5 | 5 | 181 | 6.8141 | 0.7888 | 0 | 5.375
48 | 98.525 | 8.0385 | 8.0395 | 32.8157 | 32.8148 | 6.6886 | 1.9475 | 6 | 6 | 182 | 6.7465 | 0.90388 | 0 | 6.098
48 | 77.084 | 7.8169 | 7.8164 | 31.4954 | 31.4984 | 6.7804 | 2.2258 | 7 | 7 | 183 | 6.8191 | 0.65845 | 0 | 4.49 | 6
48 | 50.818 | 8.0434 | 8.0436 | 30.8997 | 30.8966 | 6.7719 | 2.4802 | 8 | 8 | 184 | 6.7889 | 0.83857 | 0 | 5.693
48 | 40.687 | 8.0663 | 8.067 | 30.7894 | 30.7901 | 6.7732 | 2.5216 | 9 | 9 | 185 | 6.7816 | 0.89285 | 0 | 6.055
48 | 28.307 | 8.126 | 8.1298 | 30.4482 | 30.4447 | 6.7791 | 2.635 | 10 | 10 | 186 | 6.7772 | 0.7422 | 0 | 5.03 | 6
48 | 19.638 | 8.4821 | 8.486 | 29.5643 | 29.5377 | 6.7635 | 2.8414 | 11 | 11 | 187 | 6.7564 | 0.88761 | 0 | 5.997
48 | 10.106 | 8.5245 | 8.5246 | 29.2216 | 29.2244 | 6.772 | 2.8834 | 12 | 12 | 188 | 6.757 | 0.88397 | 0 | 5.973
48 | 5.334 | 8.5416 | 8.5407 | 29.1642 | 29.1651 | 6.7719 | 2.8991 | 13 | 13 | 189 | 6.7529 | 0.87148 | 0 | 5.885
48 | 4.054 | 8.5417 | 8.5417 | 29.1643 | 29.1649 | 6.7719 | 2.8911 | 14 | 14 | 190 | 6.7518 | 0.90332 | 0 | 6.099
51 | 185.802 | 7.0694 | 7.0693 | 33.8098 | 33.8108 | 6.7933 | 1.242 | 1 | 1 | 191 | 6.9245 | 0.89855 | 0 | 6.222
51 | 174.175 | 7.0771 | 7.0764 | 33.8034 | 33.805 | 6.7924 | 1.2465 | 2 | 2 | 192 | 6.9132 | 0.85879 | 0 | 5.937
51 | 149.728 | 7.1532 | 7.1532 | 33.7491 | 33.7503 | 6.783 | 1.2908 | 3 | 3 | 193 | 6.8823 | 0.90784 | 0 | 6.248
51 | 124.167 | 7.7406 | 7.7448 | 33.1175 | 33.1119 | 6.7206 | 1.6973 | 4 | 4 | 194 | 6.7994 | 0.59373 | 0 | 4.037 | 6
51 | 100.782 | 8.013 | 8.0132 | 32.7008 | 32.7017 | 6.6974 | 1.9596 | 5 | 5 | 195 | 6.7572 | 0.90392 | 0 | 6.108
51 | 74.292 | 7.9959 | 7.9961 | 31.9047 | 31.9052 | 6.7349 | 2.1675 | 6 | 6 | 196 | 6.7719 | 0.98259 | 0 | 6.654
51 | 49.972 | 7.9338 | 7.9342 | 31.2022 | 31.2037 | 6.7754 | 2.3448 | 7 | 7 | 197 | 6.791 | 0.53556 | 0 | 3.637
51 | 39.458 | 7.9171 | 7.9174 | 30.9735 | 30.9747 | 6.7881 | 2.4011 | 8 | 8 | 198 | 6.7945 | 0.61785 | 0 | 4.198
51 | 29.536 | 8.0076 | 8.01 | 30.7719 | 30.7703 | 6.783 | 2.5506 | 9 | 9 | 199 | 6.7814 | 1.1125 | 0 | 7.544
51 | 20.41 | 8.0985 | 8.0964 | 30.6864 | 30.6889 | 6.7728 | 2.62 | 10 | 10 | 200 | 6.7639 | 1.1281 | 0 | 7.63
51 | 8.993 | 8.1194 | 8.1188 | 30.6772 | 30.6782 | 6.77 | 2.6426 | 11 | 11 | 201 | 6.7513 | 0.62803 | 0 | 4.24
51 | 4.844 | 8.1174 | 8.1176 | 30.6794 | 30.6803 | 6.7702 | 2.6465 | 12 | 12 | 202 | 6.7479 | 0.64346 | 0 | 4.342
51 | 3.765 | 8.1164 | 8.1166 | 30.6798 | 30.6808 | 6.7704 | 2.6459 | 13 | 13 | 203 | 6.7471 | 1.138 | 0 | 7.678
61 | 126.452 | 7.8121 | 7.8124 | 32.1246 | 32.1239 | 6.7533 | 2.0019 | 1 | 1 | 214 | 6.8349 | 1.231 | 0 | 8.414
61 | 100.172 | 7.8627 | 7.863 | 31.4616 | 31.4615 | 6.7748 | 2.1932 | 2 | 2 | 215 | 6.8339 | 0.61429 | 0 | 4.198 | 6
61 | 74.787 | 7.9088 | 7.9092 | 31.0005 | 31.0011 | 6.7881 | 2.3191 | 3 | 3 | 216 | 6.8254 | 0.73286 | 0 | 5.002 | 3
61 | 50.321 | 7.9014 | 7.9017 | 30.755 | 30.7553 | 6.8002 | 2.379 | 4 | 4 | 217 | 6.816 | 0.8323 | 0 | 5.673
61 | 40.94 | 7.9292 | 7.9294 | 30.5787 | 30.5796 | 6.8037 | 2.4193 | 5 | 5 | 218 | 6.8115 | 1.2367 | 0 | 8.424 | 6
61 | 30.021 | 7.9747 | 7.975 | 30.3461 | 30.3463 | 6.807 | 2.4764 | 6 | 6 | 219 | 6.8056 | 0.6264 | 0 | 4.263
61 | 20.129 | 7.9754 | 7.9764 | 30.1845 | 30.1861 | 6.814 | 2.5397 | 7 | 7 | 220 | 6.804 | 0.75044 | 0 | 5.106
61 | 9.973 | 7.9511 | 7.9516 | 30.1692 | 30.1698 | 6.8185 | 2.5515 | 8 | 8 | 221 | 6.7994 | 0.78301 | 0 | 5.324
61 | 4.627 | 7.9473 | 7.9478 | 30.1703 | 30.1708 | 6.819 | 2.554 | 9 | 9 | 222 | 6.7953 | 1.2748 | 0 | 8.663
61 | 3.664 | 7.9473 | 7.9487 | 30.1695 | 30.1699 | 6.819 | 2.5544 | 10 | 10 | 223 | 6.7945 | 0.52351 | 0 | 3.557
69 | 211.685 | 7.9037 | 7.904 | 30.4325 | 30.4315 | 6.8141 | 2.4097 | 1 | 1 | 235 | 6.9727 | 0.4139 | 0 | 2.886
69 | 201.106 | 7.9045 | 7.9047 | 30.3163 | 30.3213 | 6.8192 | 2.4311 | 2 | 2 | 236 | 6.9684 | 0.85515 | 0 | 5.959
69 | 175.596 | 7.9033 | 7.9034 | 30.2692 | 30.2708 | 6.8214 | 2.4518 | 3 | 3 | 237 | 6.948 | 0.99382 | 0 | 6.905
69 | 150.297 | 7.9002 | 7.9003 | 30.2775 | 30.2786 | 6.8216 | 2.4543 | 4 | 4 | 238 | 6.9256 | 0.29442 | 0 | 2.039 | 6
69 | 125.64 | 7.8977 | 7.8978 | 30.2089 | 30.2113 | 6.825 | 2.4744 | 5 | 5 | 239 | 6.9072 | 0.74951 | 0 | 5.177
69 | 100.008 | 7.8959 | 7.8961 | 30.1811 | 30.1843 | 6.8265 | 2.4886 | 6 | 6 | 240 | 6.8861 | 1.0045 | 0 | 6.917
69 | 74.534 | 7.8975 | 7.8975 | 30.1014 | 30.0996 | 6.8298 | 2.523 | 7 | 7 | 241 | 6.867 | 0.32299 | 0 | 2.218
69 | 49.629 | 7.9026 | 7.9026 | 30.033 | 30.0342 | 6.8321 | 2.5501 | 8 | 8 | 242 | 6.8474 | 0.62622 | 0 | 4.288
69 | 39.956 | 7.912 | 7.9117 | 29.9808 | 29.9835 | 6.833 | 2.5782 | 9 | 9 | 243 | 6.8399 | 0.88978 | 0 | 6.086
69 | 29.631 | 7.9152 | 7.9152 | 29.9708 | 29.9706 | 6.8329 | 2.5831 | 10 | 10 | 244 | 6.8308 | 0.9201 | 0 | 6.285 | 3
69 | 19.778 | 7.9165 | 7.9165 | 29.9566 | 29.9574 | 6.8333 | 2.5951 | 11 | 11 | 245 | 6.8226 | 0.82974 | 0 | 5.661
69 | 9.794 | 7.9223 | 7.9214 | 29.9534 | 29.9542 | 6.8326 | 2.598 | 12 | 12 | 246 | 6.8132 | 0.89488 | 0 | 6.097
69 | 5.407 | 7.9208 | 7.9213 | 29.9598 | 29.9602 | 6.8325 | 2.6027 | 13 | 13 | 247 | 6.8093 | 0.73914 | 0 | 5.033
69 | 3.766 | 7.9533 | 7.9543 | 29.9469 | 29.9454 | 6.828 | 2.6113 | 14 | 14 | 248 | 6.8036 | 0.8597 | 0 | 5.849
8 | 156.754 | 8.9407 | 8.9408 | 30.6062 | 30.6065 | 6.6491 | 1.7012 | 1 | 1 | 1 | 6.7628 | 0.62829 | 1 | 4.249
8 | 125.406 | 8.838 | 8.8382 | 30.5252 | 30.5255 | 6.6679 | 1.7573 | 2 | 2 | 2 | 6.7541 | 0.91871 | 1 | 6.205 | 36
8 | 100.198 | 8.6691 | 8.6671 | 30.2896 | 30.2861 | 6.7034 | 1.8515 | 3 | 3 | 3 | 6.767 | 0.47761 | 1 | 3.232
8 | 73.515 | 8.5193 | 8.5201 | 29.9863 | 29.9872 | 6.7393 | 1.9436 | 4 | 4 | 4 | 6.779 | 0.5576 | 1 | 3.78
8 | 47.305 | 8.2914 | 8.2915 | 29.5151 | 29.5145 | 6.7949 | 2.2763 | 5 | 5 | 5 | 6.8105 | 0.74238 | 1 | 5.056
8 | 40.092 | 8.245 | 8.2427 | 29.4096 | 29.4043 | 6.8067 | 2.3971 | 6 | 6 | 6 | 6.8158 | 1.1042 | 1 | 7.526
8 | 31.586 | 8.1708 | 8.1702 | 29.2508 | 29.2504 | 6.8252 | 2.537 | 7 | 7 | 7 | 6.8264 | 0.44987 | 1 | 3.071
8 | 23.798 | 8.1418 | 8.1405 | 29.2113 | 29.2097 | 6.8315 | 2.5633 | 8 | 8 | 8 | 6.8257 | 0.64961 | 1 | 4.434
8 | 9.899 | 8.0123 | 8.0098 | 29.0686 | 29.0679 | 6.858 | 2.7169 | 9 | 9 | 9 | 6.8392 | 0.65271 | 1 | 4.464
8 | 4.703 | 7.9971 | 7.9972 | 29.0631 | 29.0636 | 6.8606 | 2.7201 | 10 | 10 | 10 | 6.8371 | 1.1421 | 1 | 7.809
8 | 3.701 | 7.9941 | 7.9943 | 29.0628 | 29.0633 | 6.8611 | 2.725 | 11 | 11 | 11 | 6.8367 | 0.56825 | 1 | 3.885
12 | 351.702 | 9.2585 | 9.259 | 30.7591 | 30.7598 | 6.5957 | 1.5001 | 1 | 1 | 28 | 6.8797 | 0.38127 | 1 | 2.623
12 | 299.468 | 9.0173 | 9.0174 | 30.6815 | 30.682 | 6.6345 | 1.6379 | 2 | 2 | 29 | 6.8725 | 0.63689 | 1 | 4.377
12 | 250.434 | 8.7778 | 8.7784 | 30.6082 | 30.6081 | 6.6733 | 1.7372 | 3 | 3 | 30 | 6.8679 | 0.73094 | 1 | 5.02 | 3
12 | 199.958 | 8.3318 | 8.3319 | 30.4652 | 30.4657 | 6.7469 | 1.9091 | 4 | 4 | 31 | 6.896 | 0.77088 | 1 | 5.316
12 | 174.444 | 8.2269 | 8.2255 | 30.3798 | 30.3801 | 6.7666 | 1.9713 | 5 | 5 | 32 | 6.893 | 1.3141 | 1 | 9.058
12 | 150.233 | 8.2064 | 8.2067 | 30.3289 | 30.3294 | 6.772 | 1.992 | 6 | 6 | 33 | 6.877 | 0.55664 | 1 | 3.828
12 | 124.709 | 8.3141 | 8.3143 | 30.2695 | 30.2703 | 6.7581 | 1.9678 | 7 | 7 | 34 | 6.8413 | 0.79298 | 1 | 5.425
12 | 99.783 | 8.4487 | 8.4486 | 30.0911 | 30.0929 | 6.7454 | 1.9316 | 8 | 8 | 35 | 6.8075 | 0.8351 | 1 | 5.685
12 | 74.665 | 8.2918 | 8.2926 | 29.7906 | 29.7918 | 6.7826 | 2.0219 | 9 | 9 | 36 | 6.8221 | 1.4072 | 1 | 9.6
12 | 49.859 | 8.0041 | 8.0038 | 28.9359 | 28.9366 | 6.8652 | 2.8876 | 10 | 10 | 37 | 6.8814 | 0.77033 | 1 | 5.301
12 | 40.161 | 8.0086 | 8.0084 | 28.7976 | 28.7962 | 6.8707 | 3.0075 | 11 | 11 | 38 | 6.8784 | 0.9085 | 1 | 6.249
12 | 29.8 | 7.9974 | 7.9977 | 28.7302 | 28.7299 | 6.8755 | 3.0625 | 12 | 12 | 39 | 6.874 | 0.96348 | 1 | 6.623
12 | 20.122 | 7.995 | 7.9951 | 28.6854 | 28.6861 | 6.8779 | 3.0971 | 13 | 13 | 40 | 6.8679 | 0.33285 | 1 | 2.286
12 | 10.003 | 7.9978 | 7.9979 | 28.68 | 28.6808 | 6.8777 | 3.1198 | 14 | 14 | 41 | 6.8588 | 0.62854 | 1 | 4.311
12 | 4.686 | 8.0006 | 8.0007 | 28.6793 | 28.6765 | 6.8773 | 3.1246 | 15 | 15 | 42 | 6.8537 | 0.90316 | 1 | 6.19
12 | 4.07 | 8.0059 | 8.0069 | 28.6782 | 28.6792 | 6.8765 | 3.1349 | 16 | 16 | 43 | 6.8524 | 1.0065 | 1 | 6.897
18 | 185.731 | 8.3544 | 8.3545 | 30.4636 | 30.464 | 6.7435 | 1.8953 | 1 | 1 | 56 | 6.8801 | 0.84533 | 1 | 5.816
18 | 174.681 | 8.216 | 8.2151 | 30.4022 | 30.4024 | 6.7673 | 1.9555 | 2 | 2 | 57 | 6.8938 | 0.86729 | 1 | 5.979
18 | 149.866 | 8.0665 | 8.0668 | 30.3218 | 30.322 | 6.7938 | 2.041 | 3 | 3 | 58 | 6.898 | 0.84285 | 1 | 5.814
18 | 125.034 | 7.9264 | 7.929 | 30.1503 | 30.1519 | 6.8232 | 2.122 | 4 | 4 | 59 | 6.905 | 0.86141 | 1 | 5.948
18 | 99.925 | 8.4631 | 8.4633 | 29.9233 | 29.9234 | 6.7506 | 1.9863 | 5 | 5 | 60 | 6.813 | 0.90812 | 1 | 6.187
18 | 74.731 | 8.1979 | 8.1989 | 29.67 | 29.6711 | 6.8024 | 2.0449 | 6 | 6 | 61 | 6.8414 | 0.91063 | 1 | 6.23
18 | 49.93 | 7.9399 | 7.9401 | 29.152 | 29.1526 | 6.8656 | 2.5174 | 7 | 7 | 62 | 6.8815 | 0.8613 | 1 | 5.927 | 3
18 | 39.705 | 7.9112 | 7.9109 | 28.9003 | 28.9008 | 6.8814 | 2.7961 | 8 | 8 | 63 | 6.8881 | 0.86932 | 1 | 5.988
18 | 29.874 | 8.0557 | 8.0559 | 28.5642 | 28.5646 | 6.8738 | 3.25 | 9 | 9 | 64 | 6.8727 | 0.52148 | 1 | 3.584
18 | 20.082 | 8.0732 | 8.0748 | 28.1456 | 28.1378 | 6.8898 | 3.3756 | 10 | 10 | 65 | 6.8802 | 0.77716 | 1 | 5.347
18 | 9.928 | 8.089 | 8.0891 | 28.0382 | 28.0393 | 6.8921 | 3.4013 | 11 | 11 | 66 | 6.8737 | 0.96615 | 1 | 6.641
18 | 4.431 | 8.0911 | 8.0905 | 28.0268 | 28.0275 | 6.8923 | 3.4079 | 12 | 12 | 67 | 6.8691 | 0.46979 | 1 | 3.227
18 | 3.671 | 8.0816 | 8.0817 | 28.0215 | 28.0217 | 6.894 | 3.3905 | 13 | 13 | 68 | 6.8701 | 0.58558 | 1 | 4.023 | 6
23 | 274.498 | 9.2207 | 9.222 | 30.9198 | 30.9206 | 6.5944 | 1.4533 | 1 | 1 | 85 | 6.8107 | 0.42183 | 1 | 2.873
23 | 249.389 | 8.9194 | 8.9089 | 30.7672 | 30.7625 | 6.6454 | 1.6209 | 2 | 2 | 86 | 6.8392 | 0.68634 | 1 | 4.694
23 | 199.718 | 8.3597 | 8.3595 | 30.5162 | 30.5163 | 6.7404 | 1.8881 | 3 | 3 | 87 | 6.8893 | 0.73854 | 1 | 5.088
23 | 174.647 | 8.1922 | 8.1856 | 30.4054 | 30.4028 | 6.7708 | 1.9737 | 4 | 4 | 88 | 6.8973 | 1.1855 | 1 | 8.177
23 | 149.946 | 8.0287 | 8.0332 | 30.2503 | 30.2524 | 6.8029 | 2.0629 | 5 | 5 | 89 | 6.907 | 0.41928 | 1 | 2.896
23 | 124.972 | 7.9012 | 7.9015 | 30.0794 | 30.0797 | 6.8302 | 2.1386 | 6 | 6 | 90 | 6.9119 | 0.62689 | 1 | 4.333
23 | 100.043 | 7.9064 | 7.9057 | 29.9303 | 29.9307 | 6.8361 | 2.1603 | 7 | 7 | 91 | 6.8958 | 0.70376 | 1 | 4.853 | 6
23 | 75.155 | 7.926 | 7.9277 | 29.7814 | 29.7788 | 6.8396 | 2.1757 | 8 | 8 | 92 | 6.8776 | 1.2493 | 1 | 8.592
23 | 49.546 | 7.9743 | 7.9744 | 29.5901 | 29.5913 | 6.8407 | 2.1972 | 9 | 9 | 93 | 6.8564 | 0.40459 | 1 | 2.774
23 | 40.019 | 7.8954 | 7.8957 | 29.4311 | 29.4321 | 6.8601 | 2.2387 | 10 | 10 | 94 | 6.867 | 0.69652 | 1 | 4.783
23 | 30.156 | 7.7423 | 7.7462 | 28.9315 | 28.9339 | 6.9066 | 2.4447 | 11 | 11 | 95 | 6.904 | 0.83532 | 1 | 5.767
23 | 9.794 | 7.9435 | 7.9435 | 28.1753 | 28.1761 | 6.9089 | 3.2127 | 13 | 13 | 97 | 6.8894 | 0.85058 | 1 | 5.86
23 | 5.057 | 7.877 | 7.8765 | 27.7685 | 27.7741 | 6.9378 | 3.2706 | 14 | 14 | 98 | 6.9136 | 0.42163 | 1 | 2.915
23 | 3.843 | 7.8777 | 7.8783 | 27.7387 | 27.734 | 6.939 | 3.2826 | 15 | 15 | 99 | 6.9138 | 0.62339 | 1 | 4.31
30 | 373.495 | 9.342 | 9.3423 | 30.9834 | 30.9841 | 6.574 | 1.3595 | 1 | 1 | 116 | 6.8768 | 0.72534 | 1 | 4.988 | 6
30 | 349.772 | 9.2754 | 9.276 | 30.9441 | 30.9448 | 6.5854 | 1.419 | 2 | 2 | 117 | 6.8674 | 0.88709 | 1 | 6.092
30 | 300.88 | 9.0843 | 9.0834 | 30.8542 | 30.8545 | 6.6173 | 1.5259 | 3 | 3 | 118 | 6.8563 | 0.89158 | 1 | 6.113 | 6
30 | 249.804 | 8.6972 | 8.7023 | 30.6931 | 30.6955 | 6.6817 | 1.6986 | 4 | 4 | 119 | 6.8754 | 0.80941 | 1 | 5.565
30 | 198.372 | 8.1164 | 8.1161 | 30.4203 | 30.4205 | 6.7818 | 1.9621 | 5 | 5 | 120 | 6.929 | 0.87343 | 1 | 6.052
30 | 173.328 | 7.8886 | 7.8892 | 30.3005 | 30.3014 | 6.8223 | 2.0814 | 6 | 6 | 121 | 6.9468 | 0.68176 | 1 | 4.736
30 | 151.049 | 7.8254 | 7.8258 | 30.1852 | 30.1861 | 6.8373 | 2.1462 | 7 | 7 | 122 | 6.9418 | 0.85727 | 1 | 5.951
30 | 124.999 | 8.0076 | 8.0076 | 30.0591 | 30.0598 | 6.8146 | 2.116 | 8 | 8 | 123 | 6.8968 | 0.86475 | 1 | 5.964
30 | 99.086 | 7.9687 | 7.9654 | 29.8402 | 29.838 | 6.8304 | 2.1659 | 9 | 9 | 124 | 6.8896 | 0.8523 | 1 | 5.872
30 | 74.6 | 7.6994 | 7.6973 | 29.6632 | 29.6626 | 6.8805 | 2.4095 | 10 | 10 | 125 | 6.9168 | 0.87569 | 1 | 6.057
30 | 50.203 | 8.0305 | 8.0286 | 29.6029 | 29.6024 | 6.8313 | 2.1636 | 11 | 11 | 126 | 6.8479 | 0.90509 | 1 | 6.198
30 | 41.964 | 7.9926 | 7.989 | 29.5138 | 29.5113 | 6.8412 | 2.1962 | 12 | 12 | 127 | 6.8504 | 0.84375 | 1 | 5.78
30 | 31.072 | 7.8963 | 7.8955 | 29.3748 | 29.3756 | 6.8625 | 2.2644 | 13 | 13 | 128 | 6.8615 | 0.8692 | 1 | 5.964
30 | 19.629 | 7.8148 | 7.8146 | 29.1002 | 29.0998 | 6.8876 | 2.477 | 14 | 14 | 129 | 6.8761 | 0.87463 | 1 | 6.014 | 6
30 | 10.239 | 7.7912 | 7.7938 | 28.2303 | 28.226 | 6.9306 | 3.2208 | 15 | 15 | 130 | 6.9105 | 0.5428 | 1 | 3.751
30 | 4.253 | 7.7473 | 7.7247 | 25.3802 | 24.8828 | 7.0679 | 3.414 | 16 | 16 | 131 | 7.0418 | 0.80903 | 1 | 5.697
30 | 3.611 | 7.6364 | 7.6344 | 23.1661 | 23.0641 | 7.1892 | 3.3926 | 17 | 17 | 132 | 7.1614 | 0.93096 | 1 | 6.667
41 | 230.725 | 7.8883 | 7.8884 | 31.3141 | 31.3146 | 6.7774 | 2.1941 | 1 | 1 | 149 | 6.952 | 0.69649 | 1 | 4.842
41 | 200.271 | 7.8854 | 7.8856 | 31.2974 | 31.3052 | 6.7786 | 2.2015 | 2 | 2 | 150 | 6.9261 | 0.81171 | 1 | 5.622
41 | 175.091 | 7.8829 | 7.883 | 31.1745 | 31.1746 | 6.7844 | 2.2428 | 3 | 3 | 151 | 6.9097 | 1.2213 | 1 | 8.439 | 6
41 | 149.754 | 7.8805 | 7.8806 | 31.1071 | 31.1092 | 6.7878 | 2.2594 | 4 | 4 | 152 | 6.8907 | 0.48544 | 1 | 3.345
41 | 125.014 | 7.8782 | 7.8783 | 31.0788 | 31.0781 | 6.7894 | 2.2771 | 5 | 5 | 153 | 6.8705 | 0.72193 | 1 | 4.96
41 | 100.382 | 7.8816 | 7.8816 | 30.3781 | 30.3811 | 6.82 | 2.472 | 6 | 6 | 154 | 6.8797 | 0.76049 | 1 | 5.232
41 | 74.843 | 7.8856 | 7.8856 | 30.2231 | 30.2249 | 6.8262 | 2.5445 | 7 | 7 | 155 | 6.8636 | 1.2208 | 1 | 8.379
41 | 49.662 | 7.8878 | 7.8878 | 30.117 | 30.12 | 6.8306 | 2.596 | 8 | 8 | 156 | 6.8459 | 0.46597 | 1 | 3.19 | 6
41 | 39.796 | 7.8878 | 7.8878 | 30.0959 | 30.1005 | 6.8316 | 2.6096 | 9 | 9 | 157 | 6.8382 | 0.69711 | 1 | 4.767
41 | 29.9 | 7.8884 | 7.8884 | 30.0672 | 30.0676 | 6.8328 | 2.6362 | 10 | 10 | 158 | 6.8307 | 0.73579 | 1 | 5.026
41 | 19.731 | 7.8924 | 7.8924 | 29.9179 | 29.9207 | 6.8388 | 2.6903 | 11 | 11 | 159 | 6.8279 | 1.2654 | 1 | 8.64
41 | 9.814 | 7.8915 | 7.8913 | 29.8397 | 29.8381 | 6.8424 | 2.7233 | 12 | 12 | 160 | 6.8228 | 0.4523 | 1 | 3.086
41 | 4.918 | 7.883 | 7.881 | 29.8133 | 29.812 | 6.8449 | 2.7327 | 13 | 13 | 161 | 6.821 | 0.73332 | 1 | 5.002
41 | 3.785 | 7.8813 | 7.883 | 29.8161 | 29.8183 | 6.8451 | 2.7307 | 14 | 14 | 162 | 6.8201 | 0.7296 | 1 | 4.976
45 | 252.158 | 6.7945 | 6.7947 | 33.9061 | 33.907 | 6.832 | 1.1208 | 1 | 1 | 163 | 7.0226 | 1.1615 | 1 | 8.157
45 | 175.419 | 7.1716 | 7.1725 | 33.7977 | 33.7982 | 6.778 | 1.2652 | 2 | 2 | 164 | 6.9 | 0.44188 | 1 | 3.049 | 6
45 | 147.864 | 7.301 | 7.3012 | 33.7115 | 33.7124 | 6.7618 | 1.3391 | 3 | 3 | 165 | 6.8598 | 0.6869 | 1 | 4.712
45 | 147.402 | 7.3005 | 7.3009 | 33.711 | 33.712 | 6.7619 | 1.3396 | 4 | 4 | 166 | 6.8595 | 0.7677 | 1 | 5.266
45 | 122.618 | 7.4867 | 7.4866 | 33.5665 | 33.5676 | 6.7396 | 1.455 | 5 | 5 | 167 | 6.816 | 0.75498 | 1 | 5.146 | 3
45 | 104.34 | 7.9751 | 7.9862 | 32.9116 | 32.9051 | 6.694 | 1.9021 | 6 | 6 | 168 | 6.7566 | 0.87544 | 1 | 5.915
45 | 74.433 | 7.7592 | 7.7593 | 31.5786 | 31.579 | 6.7856 | 2.1802 | 7 | 7 | 169 | 6.8217 | 0.93613 | 1 | 6.386
45 | 52.696 | 7.8686 | 7.8707 | 30.9207 | 30.9194 | 6.7979 | 2.4157 | 8 | 8 | 170 | 6.8156 | 0.97467 | 1 | 6.643
45 | 41.695 | 7.9261 | 7.9259 | 30.7508 | 30.7508 | 6.7965 | 2.4983 | 9 | 9 | 171 | 6.805 | 0.37179 | 1 | 2.53
45 | 30.386 | 7.9753 | 7.9758 | 30.6682 | 30.6684 | 6.7926 | 2.5684 | 10 | 10 | 172 | 6.7915 | 0.84723 | 1 | 5.754
45 | 19.368 | 8.1278 | 8.1274 | 30.3967 | 30.398 | 6.7811 | 2.6377 | 11 | 11 | 173 | 6.7714 | 0.98384 | 1 | 6.662
45 | 10.084 | 8.3332 | 8.3347 | 29.9744 | 29.9746 | 6.7682 | 2.7263 | 12 | 12 | 174 | 6.7518 | 0.29622 | 1 | 2
45 | 4.552 | 8.3983 | 8.3984 | 29.1366 | 29.1356 | 6.7951 | 2.8281 | 13 | 13 | 175 | 6.7743 | 0.63578 | 1 | 4.307
45 | 3.894 | 8.3998 | 8.3999 | 29.1513 | 29.1511 | 6.7942 | 2.8333 | 14 | 14 | 176 | 6.7729 | 0.92929 | 1 | 6.294
56 | 115.47 | 7.8056 | 7.8053 | 32.1124 | 32.1221 | 6.7549 | 1.9926 | 1 | 1 | 204 | 6.8268 | 1.1544 | 1 | 7.881
56 | 100.467 | 7.8624 | 7.8631 | 31.5387 | 31.5337 | 6.7715 | 2.1762 | 2 | 2 | 205 | 6.8308 | 0.62145 | 1 | 4.245 | 6
56 | 74.974 | 7.9033 | 7.9036 | 31.0344 | 31.0348 | 6.7875 | 2.3093 | 3 | 3 | 206 | 6.8248 | 0.89042 | 1 | 6.077
56 | 50.395 | 7.9624 | 7.9628 | 30.6498 | 30.6497 | 6.7954 | 2.4133 | 4 | 4 | 207 | 6.8117 | 1.0724 | 1 | 7.305
56 | 39.928 | 7.9892 | 7.9905 | 30.3766 | 30.3699 | 6.8034 | 2.475 | 5 | 5 | 208 | 6.8107 | 0.72239 | 1 | 4.92
56 | 29.166 | 7.9585 | 7.9592 | 30.1532 | 30.1543 | 6.818 | 2.5589 | 6 | 6 | 209 | 6.8158 | 1.0135 | 1 | 6.908
56 | 20.495 | 7.9578 | 7.958 | 30.1383 | 30.1387 | 6.8188 | 2.5638 | 7 | 7 | 210 | 6.809 | 1.2835 | 1 | 8.739
56 | 10.238 | 7.9587 | 7.9589 | 30.1378 | 30.1381 | 6.8187 | 2.5684 | 8 | 8 | 211 | 6.7999 | 0.45647 | 1 | 3.104
56 | 4.881 | 7.9556 | 7.9557 | 30.1356 | 30.1361 | 6.8193 | 2.5725 | 9 | 9 | 212 | 6.7958 | 0.73633 | 1 | 5.004
56 | 3.392 | 7.9613 | 7.9635 | 30.1349 | 30.1269 | 6.8184 | 2.5721 | 10 | 10 | 213 | 6.7937 | 0.88435 | 1 | 6.008
64 | 149.14 | 7.8619 | 7.8621 | 31.8153 | 31.8159 | 6.7593 | 2.0817 | 1 | 1 | 224 | 6.8612 | 0.7127 | 1 | 4.89
64 | 125.465 | 7.8602 | 7.8604 | 31.7954 | 31.7961 | 6.7605 | 2.0907 | 2 | 2 | 225 | 6.8416 | 0.74515 | 1 | 5.098
64 | 99.802 | 7.8747 | 7.875 | 31.4314 | 31.4299 | 6.7743 | 2.1974 | 3 | 3 | 226 | 6.8332 | 1.3233 | 1 | 9.042
64 | 74.784 | 7.8983 | 7.8976 | 30.9982 | 31.0128 | 6.7899 | 2.3048 | 4 | 4 | 227 | 6.827 | 0.48996 | 1 | 3.345
64 | 50.223 | 7.9248 | 7.9252 | 30.5238 | 30.5185 | 6.8068 | 2.4621 | 5 | 5 | 228 | 6.8227 | 0.72933 | 1 | 4.976
64 | 40.426 | 7.927 | 7.9273 | 30.4649 | 30.4653 | 6.8091 | 2.5103 | 6 | 6 | 229 | 6.8165 | 0.80012 | 1 | 5.454
64 | 30.039 | 7.9232 | 7.9244 | 30.4509 | 30.4475 | 6.8103 | 2.5122 | 7 | 7 | 230 | 6.8086 | 1.3765 | 1 | 9.372 | 6
64 | 19.566 | 7.9279 | 7.9279 | 30.4515 | 30.4522 | 6.8095 | 2.5004 | 8 | 8 | 231 | 6.7987 | 0.77073 | 1 | 5.24
64 | 9.406 | 7.9186 | 7.9176 | 30.4273 | 30.4247 | 6.8121 | 2.5372 | 9 | 9 | 232 | 6.7923 | 0.90234 | 1 | 6.129
64 | 4.923 | 7.9225 | 7.9229 | 30.4257 | 30.4274 | 6.8115 | 2.5335 | 10 | 10 | 233 | 6.7879 | 0.96274 | 1 | 6.535
64 | 3.668 | 7.9224 | 7.9224 | 30.4292 | 30.4304 | 6.8114 | 2.5248 | 11 | 11 | 234 | 6.7867 | 0.28187 | 1 | 1.913
76 | 179.384 | 7.8958 | 7.8961 | 30.1843 | 30.1848 | 6.8264 | 2.4139 | 1 | 1 | 249 | 6.9563 | 0.87259 | 1 | 6.07
76 | 124.698 | 7.8915 | 7.8918 | 30.0705 | 30.071 | 6.8321 | 2.4605 | 3 | 3 | 251 | 6.9136 | 0.76907 | 1 | 5.317
76 | 100.26 | 7.8822 | 7.8826 | 29.9781 | 29.9775 | 6.8377 | 2.4733 | 4 | 4 | 252 | 6.8975 | 0.85784 | 1 | 5.917
76 | 75.218 | 7.8832 | 7.8818 | 29.9398 | 29.9328 | 6.8393 | 2.4982 | 5 | 5 | 253 | 6.877 | 0.84121 | 1 | 5.785
76 | 49.938 | 7.8045 | 7.8046 | 29.6957 | 29.6962 | 6.8625 | 2.4927 | 6 | 6 | 254 | 6.8776 | 0.86266 | 1 | 5.933
76 | 40.283 | 7.7957 | 7.7955 | 29.6488 | 29.6491 | 6.866 | 2.5008 | 7 | 7 | 255 | 6.8725 | 0.89749 | 1 | 6.168 | 6
76 | 30.267 | 7.804 | 7.8034 | 29.6224 | 29.6227 | 6.8658 | 2.517 | 8 | 8 | 256 | 6.8636 | 0.90768 | 1 | 6.23
76 | 19.985 | 7.7923 | 7.7917 | 29.5937 | 29.5946 | 6.869 | 2.5195 | 9 | 9 | 257 | 6.8576 | 0.85875 | 1 | 5.889 | 6
76 | 10.076 | 7.865 | 7.8767 | 29.4905 | 29.4815 | 6.8622 | 2.6425 | 10 | 10 | 258 | 6.8426 | 0.87145 | 1 | 5.963
23 | 19.873 | 7.9456 | 7.9453 | 28.6344 | 28.635 | 6.8879 | 3.0033 | 12 | 12 | 96 | 6.8774 | 1.0521 |  | 7.236
76 | 150.054 | 7.893 | 7.8932 | 30.1634 | 30.1633 | 6.8278 | 2.4268 | 2 | 2 | 250 | 6.9316 | 0.75985 |  | 5.267
76 | 3.751 | 8.377 | 8.358 | 27.3006 | 27.3229 | 6.88 | 3.3954 | 12 | 12 | 260 | 6.8581 | 1.2119 |  | 8.311
36 | 10.006 | 7.9066 | 7.9068 | 28.6249 | 28.6253 | 6.8945 | 2.3454 | 14 | 14 | 146 | 6.875 | 1.2669 |  | 8.71
76 | 4.916 | 8.1325 | 8.1294 | 27.833 | 27.8348 | 6.8945 | 3.4745 | 11 | 11 | 259 | 6.8719 | 1.15 |  | 7.903

## Sheet2
Summary of Soc & Voffset including the original values in the factory calibration
|  | m | b | Soc | Voffset
| original | 0.4329 | -0.22177 | 0.4329 | -0.5123
| fit-all | 0.4313 | -0.2052 | 0.4313 | -0.47577
| fit-5 outliers | 0.4346 | -0.2135 | 0.4346 | -0.49126
Test runs using subsets of bottles | [7 empty] | Effect of runs on min and max for cast #21
[8 empty] | DO Conc. |  | DO Concentration Errors
[6 empty] | soc-opt | voffset-opt | Min | Max | min-opt | max-opt | % error | % error
All data | original | 0.4329 | -0.2218 | 0.4329 | -0.5123 | -0.0017 | -0.02104 | 2.7125 | 8.3726 | -0.07303 | -0.09575 | -2.6924 | -1.1436
Optimal | all-5 | 0.4346 | -0.2135 | 0.4346 | -0.49126 | 0 | 0 | 2.7855 | 8.4683 | 0 | 0 | 0 | 0
Run1 | every other cast-a | 0.4343 | -0.2112 | 0.4343 | -0.4863 | -0.0003 | 0.0049602 | 2.7985 | 8.4774 | 0.01296 | 0.0091 | 0.46311 | 0.10734
Run2 | every other cast-b | 0.435 | -0.2164 | 0.435 | -0.49747 | 0.0004 | -0.0062113 | 2.7696 | 8.4575 | -0.01585 | -0.01079 | -0.57227 | -0.12758
Run3 | every other bottle-a | 0.4346 | -0.2115 | 0.4346 | -0.48665 | 0 | 0.0046056 | 2.7992 | 8.482 | 0.0137 | 0.01375 | 0.48943 | 0.16211
Run4 | every other bottle-b | 0.4348 | -0.216 | 0.4348 | -0.49678 | 0.0002 | -0.0055201 | 2.7705 | 8.4557 | -0.01504 | -0.01259 | -0.54287 | -0.14889
Run5 | every 4th bottle a | 0.4351 | -0.2107 | 0.4351 | -0.48426 | 0.0005 | 0.0070035 | 2.8096 | 8.499 | 0.02407 | 0.03071 | 0.85671 | 0.36134
Run6 | every 4th bottle b | 0.4341 | -0.2125 | 0.4341 | -0.48952 | -0.0005 | 0.0017415 | 2.7877 | 8.4639 | 0.00217 | -0.00437 | 0.077843 | -0.051631
Run7 | selection a | 0.4348 | -0.216 | 0.4348 | -0.49678 | 0.0002 | -0.0055201 | 2.7705 | 8.4557 | -0.015 | -0.0126 | -0.54142 | -0.14901
Run8 | selection b | 0.4346 | -0.2115 | 0.4346 | -0.48665 | 0 | 0.0046056 | 2.7992 | 8.482 | 0.0137 | 0.01375 | 0.48943 | 0.16211
|  | max-min |  | 0.0008 | 0.013215