## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 Nov 2020 | ### Added HPLC data. S.H. |
| 31 May 2018 | Added DIC and Alkalinity data to 4 CHE casts. For details see document Carbon_Data_Addition.docx. |

## PROCESSING NOTES

Cruise: 2011-09

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdF Water Properties Survey

Party Chief: Chandler P.

Platform: John P. Tully

Date: June 21, 2011 – June 25, 2011

Processed by: Germaine Gatien

Date of Processing: 6 December 2011 – 9 January 2012

Number of original HEX files: 64

Number of CTD files:   64

Number of bottle casts:	20

Number of original TSG files:   1

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1176), a Wet Labs Eco-AFL/FL Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4601), a Surface PAR (no serial # listed but believed to be 16504) and an altimeter (no serial # recorded).

The deck unit is assumed to have been a Seabird model 11, serial #0471 since that was used on 2011-26.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 3363) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had no equipment list. The cruise immediately followed cruise 2011-26 with the same equipment except that a surface PAR was in use. The serial number for surface PAR was not recorded in the log or configuration file. An assumption was made that it was sensor #16504 because it is the only such sensor in use at IOS. The altimeter serial number was also missing from the configuration file. There was no log note that the TSG was in use. The rosette logs were in good order with notes about problems encountered; Niskin #3 failed to close or seal properly on several occasions.

Errors in format in the entries of station name and water depth prevented that information from appearing in the IOS headers, so the hex files had to be edited. The colon is critical in entries like STATION: B1 and Depth (m): 125. Entries that apply to the whole cruise are easy to fix, but those that vary from cast to cast are not.

Acquisition was usually turned on during the initial lowering to 10m. The data from that lowering has to be removed with an editor or DELETE (the routine that removes data from times when the CTD is going up) will join the initial lowering from 0 to 10m to the later lowering from 10m to the bottom of the cast. To ensure the highest quality data is chosen for the top 10m, a text editor is used to remove the initial lowering. This time-consuming job can be avoided if acquisition is not started until after the soak period when the CTD is back at the surface.

The Surface PAR data have been removed. The serial number was not known during the cruise, so dummy values were entered at sea. When the correct calibration parameters were entered, the resulting values were extremely low, with odd spikes. During a cruise in November problems found in the Surface PAR disappeared when the cable was changed. That is likely the cause of the problem with these data.

There was a lot of scatter in the comparison of bottle and CTD salinity, almost all showing the CTD salinity to be lower than that of the samples. The CTD channels were in good agreement and 4 of the samples were close to the CTD values, so the scatter is likely due to sampling or analysis problems. There were many bottles rejected: 6 had no liners and another 2 had no liners or caps. It is possible that other samples were not sealed well, which might lead to higher salinity values due to evaporation. The samples were analyzed about 2 months after collection.

Secondary temperature and salinity channels were selected for archiving except for cast #1, during which the primary channels performed better near the surface. For cast #2 upcast data were used because of problems near the surface.

The WetLabs ECO fluorometer was used for this cruise. It does not need to be pumped but has a poorer time response (though Sea-Bird believe it should perform better). The traces are thus much smoother than we obtain from SeaPoint fluorometers and the alignment needs a larger correction.

The differences between CTD Dissolved Oxygen and titrated samples show some variation through the cruise, but they are similar at the beginning and end, and variations are likely related to the changes in DO range rather than calibration drift. Most significant outliers were from cases with high titrated DO values near the surface where the CTD response is challenged.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1.0 mL/L from 0 to 10db

±0.4 mL/L from 10db to 200db

±0.1 mL/L below 200db

The new TSG system has the remote temperature sensor attached serially which means those configuration parameters are not checkable, except by connecting the instrument to a computer and using SEATERM. They do not appear in the configuration file, nor in the file headers. This means the values cannot be confirmed to be correct in processing. So great care will be needed to check that the parameters are entered correctly at the beginning of the cruise, or to ensure that no-one changes the TSG configuration at any time without making a note about exactly what was done. If there is a way to get a printout of the calibration parameters through SEATERM that would be helpful.

The TSG positions stopped updating early in the cruise. Positions were obtained from a record of times and positions that was provided by the chief scientist. Unfortunately, that record had uneven time spacing and many gaps, so it was a very time-consuming process to match to the TSG record. The positions after 13:44UTC on June 22, 2011 must be considered approximate. Most matches of time are within 30s, but a few errors may have been made that would lead to larger differences.

Recalibration of the TSG salinity was based on the sensor history since there were no loop samples and the near-surface CTD salinity was noisy in the area where the best comparisons with TSG were possible.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names were non-standard. They were fixed by adding a zero to the raw file names.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was no equipment list, but it is believed to have had the same equipment as for 2011-26 which immediately preceded it. There were many notes in the rosette log sheets about problems with samples. There were several instances when Niskin #3 did not close or had a poor seal.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. There were a number of errors in the con files:

- The PAR sensor parameters are not correct for the serial number given, #4601. They would fit sensor #4565, but the chief scientist made it clear that the serial number was correct. The parameters were changed to match #4601.

- The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

- The calibration for the Wetlabs Eco fluorometer could not be confirmed. The parameters entered are close to those on the characterization sheet from the factory, but the slight difference in dark value suggests that a field test was done, but the test date is not entered in the configuration file. This value will be assumed correct, since no one is available to confirm this.

- A Surface PAR was in use throughout this cruise, but the serial number is unknown and the multiplier was set to 1 because it was unknown. It will be assumed that it is sensor #16504 which had been recalibrated on March 16, 2011 and was used on other recent cruises.

After these corrections the configuration file was saved as 2011-09-ctd.con.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. (There is a ROS file for cast #29 to collect a phyto sample, but no sample # was assigned, so the bottle file will not be processed. There is also a file for cast #35 with no sample # and no note of any intention to fire a bottle) Sample numbers were added to the ADDSAMP file based on rosette log records.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-09-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. In one case duplicates were both entered since they were outliers. The average was entered, but a check will be made to see if one is clearly better than the other.

A calculations was made of  Ф  using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.4418 and B = -0.1921 with an R2 value of 0.9946. There were no severe outliers. In order to gradually remove outliers, the initial estimates of M and B were used to calculate a difference between each point and the fit:
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. When removing a little more data has little effect on the fit, it may be judged that a reasonable value has been found unless, in so doing, a whole class of points has been removed such as all high values, or all values from late in the cruise.

|  | Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | m | b | Soc | Voffset | R2 |
|  | # of points | factory setting | 0.4453 | -0.2321 | 0.4453 | -0.5212 |  |
| run 1 | 251 | all | 0.4418 | -0.1921 | 0.4418 | -0.4348 | 0.9946 |
| run 2 | 241 | <.05 | 0.4570 | -0.2189 | 0.4570 | -0.4790 | 0.9978 |
| run 3 | 225 | <.02 | 0.4622 | -0.2281 | 0.4622 | -0.4935 | 0.9989 |
| run 3 | 215 | diff<0.015 | 0.4614 | -0.2264 | 0.4614 | -0.4907 | 0.9989 |
| run 3 | 189 | diff<0.01 | 0.4599 | -0.2237 | 0.4599 | -0.4864 | 0.9993 |
| run 3 | 172 | diff<0.008 | 0.4597 | -0.2232 | 0.4597 | -0.4855 | 0.9995 |
|  | The results of 3 previous uses in time order |  |  |  |  |  |  |
|  | 2011-44 | diff<0.015 | 0.4604 | -0.2352 | 0.4604 | -0.5109 | 0.9997 |
|  | 2011-16 | diff<0.02 | 0.4661 | -0.2471 | 0.4661 | -0.5301 | 0.9997 |
|  | 2011-26 | diff<0.015 | 0.4670 | -0.2442 | 0.4670 | -0.5229 | 0.9998 |

Run 1 used all the slope/offset found for 2011-26 data.

Run 2 used the slope/offset result from run 1 and rejected differences >0.05

Run 3 used the slope/offset result from run 2 and various choices of data.

The fit with all data is reasonably good and removing points has little effect on the slope and offset. From Run 3 excluding 26 or 36 points have very similar slope, offset and R2 values. So the choice was made of using differences <0.02 with SOC and Voffset 0.4622 and -0.4935; R2=0.9989. The R2 values are a little lower than for the other 3 cruises but more than a third of the points need to be rejected to achieve higher values and the range of DO in the fit is reduced significantly.

During 2011-26 it was found that the downcast SBE DO data seemed a little low early in the cruise and high later. A time-dependent offset was applied that while not ideal for the deepest data worked pretty well. A plot was made of the titrated DO values and differences from the 2011-09 fit versus event number, and there is no evidence of time-dependence though the DO values are higher later. When outliers are excluded some of the later casts look better than the early ones. At least one early cast has small differences.  Plots against pressure show the greatest differences at the surface which is also where DO values are high. There were only 9 differences >0.05 and all came from above 10db.

The configuration files were updated with the new parameters Soc and Voffset and saved with names 2011-09-ctd-new.con.

##### Hysteresis Study

There was no deep sampling so no study is necessary or possible.

##### BOTTLE FILE PREPARATION

The station names and water depths will not appear in the IOS headers because a colon is missing in those entries in the hex files. The headers were edited to fix this. At the same time some station names were changed to standard format (Sta 58 to 58) to match those generally used for this program, and bottom depths were corrected to match log entries in a few cases where the reading from the previous cast had been accidentally repeated.

The ROS files were recreated with the new configuration parameters, 2011-09-ctd-new.con.

The files were then converted to IOS format.

They IOS files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files and a few casts with possible outliers were identified and examined in CTDEDIT. Only cast #13 required any editing; a few points in secondary salinity were cleaned. The output file was copied to the BOT file.

A preliminary header check turned up no problems with off-scale fluorescence, but the minimum values are slightly negative, as was found for 2011-26 and 2011-44. For those cruises the data were recalibrated by adding 0.062mL/L and 0.036mL/L. A similar correction will be needed for this cruise too.

The addsamp.csv file prepared in the DO calibration step was sorted on Event_Number and Sample_Number and then converted to CST files.

It was also used to add sample numbers to the BOT files, creating SAM files.

The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-09-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2011-09chl.xls. The file included comments and flags. A simplified version of the spreadsheet was prepared in which some columns were removed, an event number column was added and the information completed based on the log book. The file was saved as 2011-09chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-09oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2011-09oxy.csv. One set of duplicates had not been averaged because they were outliers in the precision study. One value was found to be an outlier in the comparison with bottles, so was selected instead of the average. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was done at IOS and received in file 2011-09.xls. This was renamed as 2011-09-sal.xls.  There was a comments column with no entries though there were comments on the analysis log sheet. There were no entries for some samples that were rejected. They were entered with flag “5” and pad value. There was a column with sample number and station name together and no event number. Separate columns were created for event number, station name and sample number. The event numbers were completed based on the log book. Some columns were removed. The file was then saved as 2011-09-sal.csv. There were no duplicates or loop samples.  The file was converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-09nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers, header names were changed to standard format and the file was saved as 2011-09-nuts.csv. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. In a few cases the comment appeared to affect all sampling, so “ALL” was placed in front instead of “Nuts”. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so in order to merge them, the MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### 6) Compare

Salinity

Compare was run with pressure as reference channel. The scatter in the fit is high with 1 extreme outlier that is associated with a very high standard deviation in the CTD data. There are only 11 bottles available and 2 of those are above 150m. To achieve a reasonably flat fit only 4 bottles are included. The outliers are the same for both CTD sensor pairs and except for the extreme outlier the standard deviations in the CTD data are low. So the problem is unlikely to be with the CTD data. The 4 bottles that remain are from the most reliable casts, 2 near the mouth of Juan de Fuca Strait and two in the middle of the Strait of Georgia. The worst outliers are from Saanich Inlet, Haro Strait and near Lasqueti Island, areas where local gradients might account for a significant difference between bottle and CTD. It must also be considered that since many bottles were rejected because they had no liners, there may also have been some poor seals on other samples, though the analyst did not note that.

Based on just 4 samples the primary sensor is low by 0.0013 and the secondary is low by 0.0009 with standard deviations of ~0.001 for both. If more data are included the CTD would appear to be reading even lower.

These data were analyzed 65 days after the end of the cruise. These sensors were calibrated in the spring and have been used for 3 other cruises before this one. The sensor pairs are in good agreement as they were on the earlier cruises. There is some evidence that when bottles sit for many months before analysis, the scatter increases with a tendency towards high values. After only a 2-month delay in analysis we would not expect so much scatter. That may be due to poor sampling protocols.

There is no suggestion of time-dependence or salinity-dependence.

The outliers in the comparison excluding those flagged previously included:

- Cast #1, sample #2– The bottle value looks like the CTD value just as the CTD stopped. The Niskin may contain water from the shed wake and not have flushed during the stop. So while it does not match the CTD well, there is likely no problem with analysis or sampling. A “3” flag should be added.

- Cast #3, sample #11 – This is in Haro Strait and there was a lot of variability at this level. The CTD standard deviation is a little high, and local gradients probably explain the differences.

- Cast #22, sample 87. There is a minor outlier and from only 100m in waters with fairly high gradients, so not worthy of a flag.

- Cast #30, sample #112 – The standard deviation in the CTD data was very high – no flag added.

- Cast #49, sample #171 – This two CTD channel pairs are not in agreement. The upcast primary salinity is very different from the downcast and from the secondary downcast and upcast. There are spikes and a lot of variability. No flag is justified.

- Cast #53, sample #187 – This cast is south of Lasqueti Island. The CTD data is quiet and the local gradient doesn’t explain a 0.013 difference- a “3” flag will be added.

- Cast #58, sample #218 – This cast is north of Lasqueti Island and there is a fairly high gradient at this level during the downcast. No flag is justified.

The salinity spreadsheet was updated.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected the fit of differences against DO concentration is quite flat. The major outliers are all at high values of DO. A plot against file pair # shows some time variation, but the differences are more likely due to the DO range than calibration drift. Most of the cases where the CTD DO looked higher than the titrated samples were at the end of the cruise, but they were also higher early in the cruise, corresponding to casts with higher DO values. The CTD DO readings were lower than the titrated values for casts #13 and 15, and the DO values were relatively low for those casts. The plot against pressure shows SBE DO values too high right at the surface. Values around 250m seem slightly high.

Checks were made in COMPARE of bottles that were flagged by the analyst:

- Cast #3, sample #11 – changed  comment to “OXY: Stopper loose/No good; outlier in comparison with CTD”

- Cast #15, sample #69 – Poor dups; one value rejected, flag removed, add to comment.

- Cast #41, sample #139 – Cap not sealed, slightly out of line in comparison, add to comment.

- Cast #46, sample #169 – Titration stopped early, slightly out of line in comparison, add to comment.

- Cast #53, sample #186 – Leaky Niskin, ok in comparison, add to comment.

- Cast #57, sample #201 – Leaky Niskin, ok in comparison, add to comment

- Cast #57, sample #202 – Poor dups; one value rejected, flag removed, add to comment.

- Cast #57, sample #214 – Stopper not sealed, chain too short; outlier in comparison, add to comment.

Outliers with differences >0.15mL/L were investigated. Some had a high standard deviation in the SBE DO that likely explains the differences. Most were in the top 5m where the sensor and bottle often don’t match well and where values are high so that 0.15mL/L is not as significant an error as it would be at lower DO values. Other outliers were:

- Cast #1, sample #2 – This is out of line as was the salinity sample. This could be a problem of closing too early, but more likely the Niskin may not have flushed well. The draw temperature seems ok for sample #2 but looks a little off for the bottom sample – we don’t know how reliable this measurement is. A flag should be added.

- Cast #7, sample #82 – The local gradient is quite high. No flag was assigned.

- Cast #37, sample #135 – Local gradient very high. No flag assigned.

- Cast #49, sample #185 – Standard deviation in CTD data very high. No flag assigned.

- Cast #49, samples #182 and 184 – temperature very noisy – assume DO could not equilibrate. No flags.

- Cast #53, sample #199 – unusual conditions, lot of variability – likely just mismatch in depth accounts for the difference. No flag assigned.

- Cast #57, sample #215 – high variability. No flag assigned.

- Cast #58, sample #219 – This sample looks out of line – looks like this bottle closed late. There is a note on the rosette sheet that it “closed nicely for once”. There had been many problems with this particular bottle closing poorly with no water, or leaking. It is likely it fired ok, but didn’t close fully until later. The nutrients are consistent with a mistrip, but could be ok. Flag DO “3”.

- Cast #58, samples #228 and 229 – high variability. No flags assigned.

- Cast #59, samples #239, 240, 241 – high variability. SBE DO didn’t equilibrate. No flags.

- Cast #62, samples #256, 257 - high variability. SBE DO didn’t equilibrate. No flags.

- Cast #63, sample #259, 260 - high variability. SBE DO didn’t equilibrate. No flags.

(See 2011-09-dox-comp1.xls.)

Fluorescence

COMPARE was run using the Wet Labs ECO CTD Fluorescence and the Extracted Chlorophyll from bottles. There is a lot of scatter, but for CHL values <0.4ug/L the fluorescence is generally higher than the extracted chlorophyll, while it is mostly lower for higher CHL. During 2011-26 the sensor showed little response for CHL<1 but that does not seem to be the case here, but there were not so many low CHL values as for the earlier cruise. (See 2011-09-chl-fluor-comp1.xls.)

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

The MRGCLN2 files were exported to a spreadsheet for a check that all samples are present. A few problems were found:

- Some missing chlorophyll data had been saved with the wrong file name, so were renamed and the merge process rerun.

- The 4 bottles of cast #28 with no sampling, were removed from the MRG file with a note of explanation.

- There was a bottle fired during cast #29, but the sampling was not anything intended for the archive, so this file will not be processed further.

- There was a surface bottle fired during cast #35, but no indication anywhere that it was intended or sampled, so it will not be processed further.

##### 7. Conversion of Full Files from Raw Data

All files were converted using 2011-09-ctd550-new.con.

A few casts were examined and all expected channels are present. There are frequent spikes in pressure at the surface.

The two temperature channels are fairly close during the downcasts but there are some odd excursions in the primary channel. The upcast temperature traces differ much more with what looks like the CTD was swinging so that alignment varied. The conductivity channels are much like the temperature with primary looking a little odd and upcasts very noisy.

As noted on other cruises the ECO fluorometer sensor traces are offset vertically. When compared with temperature traces, it appears that the fluorometer response time is about 1.5 to 4s.

The DO voltage looks as expected with a vertical offset.

Transmissivity looks normal.

PAR looks fine but Surface PAR does not. The maximum values are 2 where 2000 might be expected. There are also many small spikes in some records which is odd. The serial number was not recorded, but there is only one surface PAR instrument available. A few test conversions using different order of magnitude Conversion Factor and Ratio Factor – the 2nd had no effect while the 1st did increase the values, but the traces look very odd with large steps and spikes. During a November cruise it was found that the Surface PAR data was bad until a cable was replaced, so it is likely that the problem during this cruise was due to the same cable.

The altimetry looks ok near the bottom, but there are some spikes at the bottom.

##### 8. WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### 9. ALIGN DO

Tests were done on 3 casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge but by trying to match the depth of notable features in downcasts and upcasts, a 4s offset looks reasonable. That was found appropriate for 3 other recent cruises for which this sensor was used. ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### 10. CELLTM

These casts are too noisy to run satisfactory tests for choosing settings for CELLTM. The goal is to make upcasts look closer to downcasts on a T-S surface. The choice (α = 0.02, β=7) for the primary and (α = 0.03, β=9) looked best overall during 2011-16 and 2011-26 using the same equipment. Two casts were tested with those choices and there was a small improvement.

CELLTM was run on all casts using those settings.

##### 11. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on 2 casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 12. Test Plots and Channel Check

Two casts were plotted to check for agreement between the pairs of T and C sensors. The casts are shallow and the differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen. Results of 2 other cruises using the same equipment are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-16-0034 | 800 | -0.0003 | +0.00004 | +0.0008 | Extremely noisy |
| 2011-16-0111 | 800 | -0.0001 | +0.00010 | +0.0014 | Noisy |
| 2011-26-0017 | 800 | -0.0001 | +0.00012 | +0.0016 | Noisy,high |
| 2011-26-0055 | 800 | ~0 | +0.00007 | +0.0009 | VNoisy, VHigh |
| 2011-26-0082 | 800 | ~0 | +0.00006 | +0.0007 | VNoisy, VHigh |
| 2011-09-0056 | 330 | -0.0005 | +0.00004 | +0.0008 | High, Steady |
| 2011-09-0065 | 330 | -0.0002 | +0.00004 | +0.0005 | High, Steady |

The differences are small and show little change with time. The pressure dependence is slight.

##### 13. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### 14. Checking Headers

The header check was run.  There are pressure spikes, at least most of which occur during the initial soak period. Fluorescence does not go off-scale at the high end but has some negative values; again, most seem to be in the initial soak period, but checks should be made after DELETE to see if any remain. From other cruises we expect some negative dark values from this sensor, and checks of a few casts did find some readings of -0.031ug/L, so it is likely a tiny correction will be needed later to avoid negative values in the record.

Surface check was run and shows an average surface pressure for the cruise was 2.1db which looks low for the Tully, but given the work was in inland waters, this is likely accurate. Many of the values were very low, but the associated salinity values are also very low, so they are likely very close to the surface. One cast (#30) was found during which there are slightly negative values of pressure at the end of the cast with pumps running. In that case it looks like the CTD sensors were bobbing in and out of the water as one salinity channel looked like it was in water and the other out; pressure was on the order of -0.4db. With water splashing around the equipment it is not clear, but -0.4db is probably just out of the water. There is no need to recalibrate pressure.

The cross-reference check was compared with the log book and the only problems were 4 casts with the missing or incorrect station names; those were corrected in both bottle files and full profile files.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet, together with water depths and maximum pressure in the file. A few problems were found:

- For cast 1 the CTD did not get within 15m of the bottom but a spike was interpreted as a signal, so the header was removed from the CLN file.

- For cast #2 the altimeter reading is fine, but the bottom depth is wrong; that was replaced with the value from the log book.

- For cast #63 the bottle file should not contain an altimeter header since the deepest bottle was not within 15m, so it was removed from the MRGCLN2 and SAMAVG files.

Overall, the algorithm worked very well, ignoring spikes at the bottom.

##### 15. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset.

For the ECO sensor the estimates give highly variable results, from about 1s to 5s, with 2s looking about average. On previous cruises a +48 records (2 seconds) was found most appropriate.

SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records. (Output: SHFFL1)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.8s was applied to the primary and a shift of -0.5s to the secondary conductivity. During 2011-16 and 2011-26 the best results were with -0.5 and +0.5s.

SHIFT was run twice on all casts using -0.8s and -0.5s.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 16. DELETE

Almost all files contain data from an initial drop to 10m. This needs to be removed since DELETE is likely to join the initial drop to 10 to the full cast from 10 to the bottom, rather than the full cast from 2 to the bottom. It is easier to check them all first than to examine the DELETE output and then go back and fix them. Each file was opened and initial records removed if appropriate.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: Warnings were found for 2 casts. Cast #2 contained many pressure spikes around 13.5db – 15db. The SHFC1 file was edited to replace suspect pressure values with pad values. This leaves little data between 14 and 15db. DELETE was rerun for that cast; there remain warnings but they concern points with pad values in the pressure, so are not significant. For cast #52 the warnings all pertain to spikes at the end of the file, during the upcast.

##### 17. Other Comparisons

Previous experience with these sensors –

1. Conductivity:

The sensors were both recalibrated in late March 2011 and were used for 3 other cruises. There were doubts about the quality of the comparison with bottles. No corrections were applied.

2. Dissolved Oxygen

The DO sensor was repaired and recalibrated in April 2011. It was used for 3 other cruises. The correction was time-dependent for the most recent, 2011-26, but not for the other 2.

3. Pressure

The sensor was recalibrated in April 2011 and was used for 2011-44, 2011-16 and 2011-26. No further offset was applied for those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity was slightly low for a few metres during cast #14 and high below 140m for cast #30. Temperatures west of 124° longitude were near or below the local minima around 100db, with those to the north being the coldest. These excursions look real, not indicative of instrument malfunction.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### 18. DETAILED EDITING

A few casts were opened in CTDEDIT to compare the salinity channels to see whether primary or secondary should be chosen for archiving, and hence, editing. There are fewer spikes in the secondary salinity and it was closer to the bottles than the primary (though that comparison is not trusted.) For 2011-16 and 2011-26 the secondary channels were selected. So the secondary were selected again.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as they appear to be corrupted by either ship effects or too short a wait at the top. Two casts had unusual problems:

- Cast 1 – The secondary salinity looks poor in the top 8db; the primary looks more reliable there, but below that level the primary has spikes where the secondary looks fine. The upcast is not useful because there were many stops for bottles, so it is heavily corrupted by shed wakes. Two approaches were tried – using the secondary but removing the top 8db, or using the primary with more editing needed. The resulting profiles are very close below 8db, and the file with primary channels edited looks satisfactory in the top 8db, so that will be used. Since the secondary will be used for most casts, an adjustment should be made later to bring the primary into line with the secondary based on the results of section 11.

- Cast 2 – There were pressure spikes during the downcast when the CTD was at about 15db. Given the location in Haro Strait it is hard to judge whether data are reliable. There were no stops for bottles for this cast, so the upcast data can be used. The SHFC1 file was put through REVERSE and then DELETE to produce a file with upcast data. That was edited and looks better.

In cases where more than 1db of data were removed from the bottom, the altimetry header was adjusted to show that the deepest values are further from the bottom than in the unedited files. (Casts #26, 35, 52)

All EDU files were copied to EDT.

##### 19. Initial Recalibration

Based on the studies described in section 12, the primary salinity appears to be lower than the secondary by ~0.0008. Since the primary was chosen for one cast and secondary for others, an adjustment will be made by adding 0.0008 to the primary salinity. The comparison with bottles described in section 6 does not justify recalibration – the 4 bottles that produced a flat fit of differences against pressure suggest that the secondary was within 0.001 of bottles.

The pressure and dissolved oxygen do not require recalibration.

The fluorescence will be adjusted by adding 0.031 to avoid negative values in the final files.

CALIBRATE was run using file 2011-09-recal1.ccf to apply the corrections to the Salinity:T0:C0 and Fluorescence channels.

##### 20. Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences are small. Plots of differences have fairly flat trendlines, though there is some suggestion of errors being higher at depth than at the surface. As noted earlier, the plot of differences against file pair number shows variability but no systematic trend. No further recalibration is justified.

##### 21. Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD. (Since the ECO fluorometer was used, there was no need to prepare a  second set that had been filtered.) The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

The ECO fluorometer data is quite smooth and does not need filtering.

##### 22. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### 23. Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except cast #1 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag

For cast #1 only, the following casts were removed:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1.0 mL/L from 0 to 10db

±0.4 mL/L from 10db to 200db

±0.1 mL/L below 200db

For details on the processing see processing report: 2011-09-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

Profile plots and T-S plots were examined and no problems were found.

The track plot looks ok.

##### 24. Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from ~70% to ~170%. Values were very low in Haro Strait and Juan de Fuca Strait. In Georgia Strait most values were >100% with very high values at stations 1, 11, 12, 14, 24, 25, 26 and in Saanich Inlet. A check was made of 3 of these casts for which titrated samples were available at 5m. In each case the SBE DO was a little higher at the surface than the titrated value, so the high saturation may be due, at least partly, to the calibration not working as well at the surface as it does below 10m. The fit against bottles worked very well for the casts with lower DO values, but it is likely that many of the high DO bottles were rejected as outliers. It might be better in future to do separate fits when different parts of the cruise have very different DO ranges, though at the high end there maybe be too much variability to come up with a fit that works any better than this. This problem may just be the inability of the instrument to work well in very high gradients.

##### 25. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts except cast #1 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag

For cast #1 only, the following casts were removed (there was no problem with the secondary salinity during the upcast, but since the primary channels were selected for the CTD file, the same was done for the bottle file just for consistency):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run on the final files and no problems were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no errors were found.

##### 26. Thermosalinograph Data

Data were provided in 1 hex file. There was no loop sampling.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong scale factor. After that correction the CON file was saved as 2011-09-tsg.con. The remote temperature calibration parameters could not be checked. They are not listed in the configuration file nor in the header records. This is because it is connected serially with the new TSG system and details can only be checked through SEATERM when it is connected to a computer directly. This is awkward and we have no way of being sure it was done correctly. This should be checked at the beginning of every cruise to be sure no-one has changed it. If we did find an error at the end, there does not appear to be any easy way to correct it.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format. The file name was changed to standard format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

A track plot shows no data beyond Juan de Fuca Strait, yet the times in the file cover the whole cruise. It was discovered that the position data did not update, but times and other data seem ok. A plot of GPS positions and times was obtained from the Chief Scientist. These data are irregularly spaced in time, roughly every 10s but with large gaps. The TSG file was opened in EXCEL and the GPS data was added to it. The GPS data was then matched for time as closely as possible to the TSG file times. So approximate positions could be inserted in the TSG file, or pad values where there was no match within 30s of the TSG time. A column with “1” entered for all rows was added to serve as the file break indicator.

The file thus created was saved as a 2011-09-0001 with gps3.csv. It was converted to file 2011-09-0001.IOS. Most header information was added in the conversion, but the time interval was added to the IOS file using a text editor. Then the file was put through CLEAN and ADD TIME CHANNEL.

Time-series plots were produced. Overall the records look good, though it is sometimes very noisy in the Strait of Georgia.

The flow rate was not converted, so a test conversion was run to ensure that was ok. File 2011-09-0001.flow contains that channel. The flow rate was very steady with minimum and maximum values of 0.958 and 1.040.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2011-09-ctd-tsg-comp.xls.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 62 matches since there was no match for cast #1 and no data from cast #14 at 4m. Those data were added to 2011-01-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.00032° and in longitude <0.00034°. The median differences were 0.00005° and 0.00002°. This shows both the times and positions are reliable for both systems, despite the method used to get positions after the TSG positions froze, though there were only 9 points of comparison for data added in that way

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. Using all 9847 records, the lab temperature was higher than the intake by an average of 0.23Cº, with a median difference of 0.27Cº and standard deviation of 0.47Cº. Picking out a section where the differences look fairly steady, the average of 500 points showed the lab temperature to be higher than the intake by an average of 0.26Cº, a median difference of 0.28Cº and standard deviation of 0.05Cº.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. The differences were very noisy for the casts in the Strait of Georgia, presumably because the surface gradients were large and variability was likely high, greatly increasing the effect of small mismatches in time and in depth. The lowest standard deviations in intake temperature are all in Juan de Fuca Strait.

When all data were included the TSG intake temperature was higher than the CTD by a median value of 0.32Cº while the median difference was 0.10Cº. When only data from the 14 Juan de Fuca casts with the lowest standard deviation in the intake temperature were included, the average and median differences were both 0.017Cº. Using the same points, the TSG lab temperature is higher than the intake temperature by an average and median of 0.28Cº and the TSG lab temperature was higher than the intake temperature by 0.30 Cº.

Using the 14 Juan de Fuca casts with lowest standard deviation in the TSG salinity, the TSG values are lower than the CTD by an average of 0.22 and a median of 0.21.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.5 to 11.4, with a median of 1.6. When restricted to the 14 cases with lowest standard deviation in the TSG fluorescence, the ratio has an average and median of 1.7 and 1.6.

For more details see 2011-09-ctd-tsg-comp.xls.

- Loop Bottle - TSG Comparisons

There were no loop bottles taken during the cruise.

- Loop Bottle - Rosette Comparisons

There were no loop bottles taken during the cruise.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and were used for 2011-44, 2011-16 and 2011-26, but there are not yet any data available for 2011-16.

For 2011-44 the intake temperature looked unbelievable being higher than the CTD by 0.41Cº and higher than the lab temperature by 0.2 Cº. The lab temperature was higher than the CTD by 0.25Cº which is about the amount of heating we would expect to find in the loop at that time of year. The TSG Salinity was lower than the CTD salinity by ~0.007 during offshore casts though it was lower by much more in inlets with high near-surface gradients. The TSG fluorescence was higher than the CTD fluorescence by a median factor of 2.

For 2011-26 the TSG intake temperature was within 0.005Cº of the CTD, the salinity was lower than the CTD by ~0.029 and the ratio of TSG to CTD fluorescence was ~2.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was very steady.

3. The temperature in the loop increases by an average of 0.30Cº.

4. The TSG intake temperature appears to be higher than the CTD by ~0.02Cº. This is a little higher than found for 2011-26, but given mismatches in depth and time, and it is a good result and suggests the calibration of both temperature sensors is good.

5. The TSG Salinity is lower than the CTD salinity by a median value of ~0.21. This is much higher than the correction found for 2011-44 and 2011-26. This may indicate conductivity sensor calibration drift, or may be the result of a systematic mismatch in depths in the presence of higher salinity gradients. However, the near-surface CTD salinity was very noisy for many of the casts in Juan de Fuca Strait, so the comparison is not very reliable. In the absence of loop sampling, the results of 2011-26 are the best information available.

6. The TSG fluorometer had values higher than the CTD fluorometer by a median factor of 1.6. This is reasonably close to the results of other spring 2011 cruises. The sensor needs recalibration.

f.) Editing

No editing was applied to these data. There is a section at the beginning that looks a little odd, but given the location near Satellite Channel and Haro Strait, this may be fine. There is a lot of noise in the Strait of Georgia section, but nothing that appears to be due to instrumental problems. Some very low salinity values were recorded near the mouth of the Fraser River, but the CTD data shows the same thing; the spring run-off was very high in spring 2011.

g.) Recalibration

File 2011-01-tsg-recal1.ccf was prepared to adjust salinity by adding 0.02. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary and Flag.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files. The End Time was added to the headers using a text editor.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data and all looked fine.

The cruise plot was added to the end of this report.

##### 27. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (Notes from log book and sampling notes)

50. Archiving started early.

62. Header error –should be stn 12, not 62.

Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3396 | 1Apr2011 | Factory |  |  |
| Conductivity | 2374 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2754 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2668 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug2010 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2216 | ? |  |  |  |
| Surface PAR | ? | ? |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2011-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 23Mar11 | Factory |  |  |
| Conductivity | 3363 | 23Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 603 | 03Mar11 | “ |  |  |