## comparison-1176
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
46 | 49.104 | 9.0419 | 9.0394 | 29.3827 | 29.3844 | 6.6868 | 2.041 | 11 | 11 | 163 | 6.6 | 0.71384 | 4.7114 | 3.125e-07 | 0.457 | -0.2189 |  | Summary of Soc Voffset including the original values in the factory calibration
46 | 369.718 | 8.6746 | 8.6746 | 30.8281 | 30.8286 | 6.6793 | 1.5427 | 1 | 1 | 153 | 6.8708 | 0.48625 | 3.3409 | 0.00011967
3 | 75.192 | 8.9681 | 8.9684 | 31.0748 | 31.0751 | 6.625 | 1.8354 | 7 | 7 | 16 | 6.5613 | 0.61973 | 4.0662 | 0.00012534 | [5 empty] | m | b | Soc | Voffset | R2
62 | 300.672 | 8.589 | 8.5887 | 30.5561 | 30.5572 | 6.7039 | 1.5972 | 2 | 2 | 243 | 6.8363 | 0.51116 | 3.4944 | 0.00014712 |  |  |  | # of points | factory setting | 0.4453 | -0.2321 | 0.4453 | -0.5212
46 | 1.444 | 13.091 | 13.0973 | 11.5121 | 11.4643 | 6.8497 | 2.9102 | 17 | 17 | 169 | 6.7029 | 1.1109 | 7.4463 | 0.00017035 |  |  | run 1 | 251 | all | 0.4418 | -0.1921 | 0.4418 | -0.43481 | 0.9946
3 | 151.218 | 8.2993 | 8.3 | 32.1536 | 32.1539 | 6.6781 | 1.6003 | 4 | 4 | 13 | 6.6834 | 0.51225 | 3.4236 | 0.00021167 |  |  | run 2 | 199 | <.05 | 0.457 | -0.2189 | 0.457 | -0.47899 | 0.9978 | used results from all
15 | 173.432 | 6.8083 | 6.8085 | 33.8681 | 33.8684 | 6.8316 | 1.1138 | 1 | 1 | 64 | 6.8705 | 0.28988 | 1.9916 | 0.00023138 |  |  | run 3 | 225 | <.02 | 0.4622 | -0.2281 | 0.4622 | -0.49351 | 0.9989 | used results from <0.05
62 | 250.575 | 8.4925 | 8.4915 | 30.5105 | 30.5106 | 6.7205 | 1.6563 | 3 | 3 | 244 | 6.8102 | 0.53827 | 3.6657 | 0.00023342 |  |  | run 3 | 215 | diff<0.015 | 0.4614 | -0.2264 | 0.4614 | -0.49068 | 0.9989 | used results from <0.05
46 | 125.564 | 8.9321 | 8.9327 | 30.2327 | 30.233 | 6.6665 | 1.9209 | 8 | 8 | 160 | 6.6452 | 0.6592 | 4.3806 | 0.00024911 |  |  | run 3 | 189 | diff<0.01 | 0.4599 | -0.2237 | 0.4599 | -0.48641 | 0.9993 | used results from <0.05
53 | 40.602 | 9.2407 | 9.2439 | 29.3596 | 29.3592 | 6.6582 | 2.0162 | 10 | 10 | 195 | 6.5634 | 0.70279 | 4.6127 | 0.00029009 |  |  | run 3 | 172 | diff<0.008 | 0.4597 | -0.2232 | 0.4597 | -0.48553 | 0.9995 | used results from <0.05
53 | 175.396 | 8.3671 | 8.3673 | 30.4282 | 30.4286 | 6.7431 | 1.8303 | 4 | 4 | 189 | 6.7688 | 0.61785 | 4.1821 | 0.00029288 |  |  |  | The results of 3 previous uses in time order
13 | 183.158 | 6.7629 | 6.7632 | 33.8805 | 33.8807 | 6.8381 | 1.0982 | 1 | 1 | 51 | 6.8862 | 0.28266 | 1.9465 | 0.00032149 |  |  |  | 2011-44 | diff<0.015 | 0.4604 | -0.2352 | 0.4604 | -0.51086 | 0.9997
49 | 149.97 | 8.3729 | 8.3733 | 30.4063 | 30.4334 | 6.7432 | 1.8693 | 6 | 6 | 175 | 6.7469 | 0.63504 | 4.2845 | 0.00033242 |  |  |  | 2011-16 | diff<0.02 | 0.4661 | -0.2471 | 0.4661 | -0.53014 | 0.9997
57 | 30.72 | 9.5349 | 9.5337 | 29.1913 | 29.1931 | 6.6219 | 2.0499 | 12 | 12 | 212 | 6.5177 | 0.71829 | 4.6816 | 0.00040168 |  |  |  | 2011-26 | diff<0.015 | 0.467 | -0.2442 | 0.467 | -0.52291 | 0.9998 | time-dependence found in COMPARE
53 | 273.03 | 8.9054 | 8.9001 | 30.8204 | 30.8151 | 6.6452 | 1.4627 | 1 | 1 | 186 | 6.75 | 0.44996 | 3.0372 | 0.00041744
41 | 200.738 | 8.5192 | 8.5191 | 30.6156 | 30.616 | 6.7119 | 1.7254 | 4 | 4 | 140 | 6.7581 | 0.56917 | 3.8465 | 0.00042331 |  |  | run 1 used m/b from 2011-26 and all data
59 | 164.526 | 8.6187 | 8.6201 | 30.5959 | 30.5971 | 6.6977 | 1.4607 | 1 | 1 | 230 | 6.7119 | 0.44909 | 3.0143 | 0.00043188 |  |  | run 2 used m/b from run 1 rejecting differences >0.05
59 | 40.21 | 9.5463 | 9.5205 | 29.1797 | 29.1897 | 6.6208 | 1.99 | 7 | 7 | 236 | 6.5243 | 0.69007 | 4.5022 | 0.00043754 |  |  | run 3 used m/b from run 2 rejecting various differences
7 | 50.932 | 7.8064 | 7.8055 | 32.859 | 32.86 | 6.7219 | 1.4776 | 9 | 9 | 32 | 6.6454 | 0.45588 | 3.0295 | 0.00050375
46 | 149.251 | 8.6884 | 8.6931 | 30.3266 | 30.3276 | 6.6989 | 1.9051 | 7 | 7 | 159 | 6.6995 | 0.65118 | 4.3626 | 0.00056299
53 | 150.577 | 8.3248 | 8.3248 | 30.3425 | 30.3427 | 6.7533 | 1.8934 | 5 | 5 | 190 | 6.7579 | 0.64581 | 4.3643 | 0.00057367
25 | 10.513 | 9.9498 | 9.9502 | 30.338 | 30.3378 | 6.5135 | 2.2438 | 13 | 13 | 108 | 6.3922 | 0.80721 | 5.1599 | 0.000688
18 | 2.005 | 9.6987 | 9.6985 | 30.7598 | 30.7602 | 6.5318 | 2.1742 | 10 | 10 | 85 | 6.4046 | 0.77402 | 4.9573 | 0.00069746
13 | 175.581 | 6.7615 | 6.7617 | 33.8806 | 33.8808 | 6.8384 | 1.0963 | 2 | 2 | 52 | 6.8797 | 0.28138 | 1.9358 | 0.00073545
7 | 39.963 | 8.2094 | 8.2088 | 32.367 | 32.368 | 6.6823 | 1.6251 | 10 | 10 | 33 | 6.5938 | 0.52453 | 3.4587 | 0.00074211
62 | 75.762 | 8.2552 | 8.2561 | 29.8889 | 29.8889 | 6.7839 | 1.8565 | 9 | 9 | 250 | 6.7244 | 0.62875 | 4.228 | 0.00079175
53 | 100.845 | 8.1816 | 8.1813 | 30.1028 | 30.1065 | 6.7858 | 1.9707 | 7 | 7 | 192 | 6.7485 | 0.68252 | 4.606 | 0.00081078
59 | 100.108 | 8.1795 | 8.1801 | 30.194 | 30.1958 | 6.7821 | 1.877 | 4 | 4 | 233 | 6.7442 | 0.63806 | 4.3032 | 0.0008292
49 | 100.907 | 8.2036 | 8.204 | 30.0981 | 30.124 | 6.7826 | 1.9746 | 8 | 8 | 177 | 6.7452 | 0.68261 | 4.6043 | 0.00090318
46 | 74.712 | 9.0141 | 9.0176 | 29.6896 | 29.6908 | 6.6777 | 1.9892 | 10 | 10 | 162 | 6.6128 | 0.68925 | 4.5578 | 0.00091062
30 | 175.405 | 8.6064 | 8.6067 | 31.6419 | 31.642 | 6.6543 | 1.6862 | 3 | 3 | 113 | 6.6778 | 0.55061 | 3.6768 | 0.0010988
3 | 20.196 | 9.5322 | 9.5328 | 30.2533 | 30.2536 | 6.5773 | 2.0504 | 11 | 11 | 20 | 6.465 | 0.71925 | 4.65 | 0.001112
57 | 125.67 | 8.2253 | 8.2252 | 30.1981 | 30.1981 | 6.7749 | 1.9471 | 7 | 7 | 207 | 6.7587 | 0.67206 | 4.5423 | 0.0011177
62 | 150.917 | 8.2405 | 8.2405 | 30.2646 | 30.2948 | 6.7696 | 1.8728 | 6 | 6 | 247 | 6.7752 | 0.63827 | 4.3244 | 0.0012744
46 | 30.326 | 9.3208 | 9.3202 | 29.1447 | 29.1456 | 6.6555 | 2.1889 | 13 | 13 | 165 | 6.5517 | 0.78272 | 5.1281 | 0.0012994
57 | 40.894 | 9.1217 | 9.123 | 29.4429 | 29.4434 | 6.6723 | 1.9699 | 11 | 11 | 211 | 6.5782 | 0.68268 | 4.4908 | 0.0013468
25 | 75.471 | 8.7745 | 8.775 | 31.5387 | 31.5383 | 6.6337 | 1.7505 | 8 | 8 | 103 | 6.5715 | 0.58247 | 3.8277 | 0.0013696
46 | 100.321 | 8.8617 | 8.8603 | 29.9603 | 29.9598 | 6.6888 | 1.9694 | 9 | 9 | 161 | 6.6465 | 0.67969 | 4.5175 | 0.0014274
3 | 30.416 | 9.3504 | 9.3513 | 30.5138 | 30.5133 | 6.5927 | 1.9784 | 10 | 10 | 19 | 6.4897 | 0.68668 | 4.4564 | 0.001446
22 | 50.486 | 8.944 | 8.9443 | 31.2898 | 31.2899 | 6.6193 | 1.8146 | 4 | 4 | 89 | 6.5352 | 0.61181 | 3.9983 | 0.001455
41 | 301.004 | 8.6508 | 8.6509 | 30.7753 | 30.7759 | 6.6851 | 1.5845 | 2 | 2 | 138 | 6.8169 | 0.50374 | 3.434 | 0.0014976
37 | 19.914 | 9.7021 | 9.7015 | 28.8594 | 28.8589 | 6.6116 | 2.2606 | 9 | 9 | 133 | 6.4975 | 0.81267 | 5.2804 | 0.0015216
1 | 4.88 | 11.1368 | 11.1128 | 28.4283 | 28.4314 | 6.4251 | 2.6247 | 8 | 8 | 8 | 6.2958 | 0.97905 | 6.1638 | 0.0015225
58 | 75.062 | 8.5045 | 8.5047 | 29.7719 | 29.7723 | 6.7509 | 1.8514 | 6 | 6 | 222 | 6.6892 | 0.62882 | 4.2063 | 0.0016131
10 | 51.321 | 8.0086 | 8.0101 | 32.5666 | 32.5652 | 6.704 | 1.5423 | 6 | 6 | 44 | 6.6264 | 0.48428 | 3.209 | 0.0016551
22 | 40.569 | 9.0184 | 9.0213 | 31.1763 | 31.1731 | 6.6132 | 1.8385 | 5 | 5 | 90 | 6.5204 | 0.62295 | 4.0619 | 0.0016558
7 | 31.227 | 8.4759 | 8.4692 | 32.1691 | 32.18 | 6.651 | 1.7296 | 11 | 11 | 34 | 6.5536 | 0.56985 | 3.7346 | 0.0017004
58 | 175.117 | 8.3963 | 8.3964 | 30.4654 | 30.466 | 6.737 | 1.7933 | 2 | 2 | 218 | 6.7622 | 0.5989 | 4.0499 | 0.0017193
58 | 99.928 | 8.1448 | 8.1452 | 29.9953 | 29.9953 | 6.7962 | 1.962 | 5 | 5 | 221 | 6.7584 | 0.676 | 4.5687 | 0.0017402
41 | 30.769 | 8.8092 | 8.8095 | 29.2349 | 29.2355 | 6.7282 | 2.0863 | 12 | 12 | 148 | 6.6269 | 0.73634 | 4.8797 | 0.0018232
3 | 175.878 | 8.2142 | 8.2156 | 32.2819 | 32.2808 | 6.6853 | 1.5601 | 3 | 3 | 12 | 6.7125 | 0.49592 | 3.3288 | 0.0018551
13 | 126.292 | 6.7651 | 6.7654 | 33.8784 | 33.8784 | 6.8379 | 1.0989 | 4 | 4 | 54 | 6.8358 | 0.2814 | 1.9236 | 0.0019065
30 | 5.188 | 10.3765 | 10.3766 | 28.3726 | 28.3733 | 6.5345 | 2.4403 | 13 | 13 | 123 | 6.4064 | 0.89825 | 5.7546 | 0.0019305
57 | 100.553 | 8.1646 | 8.1645 | 30.0599 | 30.0595 | 6.7903 | 1.9679 | 8 | 8 | 208 | 6.7529 | 0.68236 | 4.6079 | 0.0019491
59 | 75.204 | 8.247 | 8.2475 | 29.9447 | 29.9451 | 6.7827 | 1.8637 | 5 | 5 | 234 | 6.7228 | 0.63083 | 4.2409 | 0.0019843
62 | 20.506 | 9.9242 | 9.9289 | 28.8687 | 28.8739 | 6.5788 | 2.1151 | 13 | 13 | 254 | 6.4647 | 0.74567 | 4.8205 | 0.0020156
58 | 124.712 | 8.2393 | 8.2396 | 30.2972 | 30.2977 | 6.7684 | 1.8994 | 4 | 4 | 220 | 6.7513 | 0.65113 | 4.396 | 0.0020163
53 | 30.803 | 9.5805 | 9.5804 | 29.1733 | 29.1732 | 6.616 | 2.0906 | 11 | 11 | 196 | 6.5116 | 0.73857 | 4.8093 | 0.0020618
41 | 40.822 | 8.5285 | 8.5288 | 29.4522 | 29.4524 | 6.7613 | 2.0161 | 11 | 11 | 147 | 6.6701 | 0.70452 | 4.6992 | 0.0020757
37 | 29.996 | 9.7056 | 9.7048 | 29.2467 | 29.2451 | 6.5946 | 2.1416 | 8 | 8 | 132 | 6.4892 | 0.76191 | 4.9442 | 0.0021126
49 | 40.323 | 8.2871 | 8.2886 | 29.5165 | 29.5405 | 6.7955 | 1.9522 | 11 | 11 | 180 | 6.7052 | 0.67538 | 4.5285 | 0.0021448
62 | 101.311 | 8.1606 | 8.1623 | 30.0245 | 30.0245 | 6.7925 | 1.8949 | 8 | 8 | 249 | 6.7557 | 0.64923 | 4.3861 | 0.0021738
46 | 40.197 | 9.1702 | 9.1684 | 29.3122 | 29.3135 | 6.6707 | 2.0758 | 12 | 12 | 164 | 6.5758 | 0.73192 | 4.813 | 0.002179
46 | 350.203 | 8.6723 | 8.6724 | 30.827 | 30.8273 | 6.6796 | 1.5483 | 2 | 2 | 154 | 6.8541 | 0.48649 | 3.3344 | 0.0021947
46 | 300.16 | 8.6672 | 8.6673 | 30.7763 | 30.7767 | 6.6826 | 1.5963 | 3 | 3 | 155 | 6.8135 | 0.50836 | 3.4637 | 0.0022418
41 | 20.828 | 9.3373 | 9.3349 | 28.762 | 28.7654 | 6.6695 | 2.2022 | 13 | 13 | 149 | 6.5574 | 0.78976 | 5.1788 | 0.0022424
7 | 19.885 | 8.6306 | 8.629 | 31.9725 | 31.9737 | 6.6364 | 1.7848 | 12 | 12 | 35 | 6.5287 | 0.59901 | 3.9108 | 0.0022697
3 | 1.381 | 9.7178 | 9.719 | 29.9179 | 29.9168 | 6.5645 | 2.1473 | 14 | 14 | 23 | 6.436 | 0.76476 | 4.922 | 0.0023602
41 | 50.831 | 8.3823 | 8.3817 | 29.5748 | 29.5757 | 6.7783 | 1.9932 | 10 | 10 | 146 | 6.6965 | 0.69435 | 4.6497 | 0.0023664
53 | 75.569 | 8.1789 | 8.1789 | 29.8437 | 29.8452 | 6.7977 | 2.0001 | 8 | 8 | 193 | 6.7385 | 0.69755 | 4.7005 | 0.0024262
3 | 5.12 | 9.664 | 9.6644 | 30.0011 | 30.0015 | 6.5688 | 2.121 | 13 | 13 | 22 | 6.4436 | 0.75288 | 4.8512 | 0.0024956
57 | 150.516 | 8.2745 | 8.275 | 30.3398 | 30.3403 | 6.7611 | 1.8911 | 6 | 6 | 206 | 6.7661 | 0.64783 | 4.3832 | 0.002509
30 | 75.947 | 9.2047 | 9.2055 | 30.3475 | 30.3466 | 6.6212 | 1.9421 | 7 | 7 | 117 | 6.5566 | 0.67118 | 4.4007 | 0.002525
59 | 29.807 | 10.0571 | 10.0615 | 28.844 | 28.8413 | 6.5606 | 2.0943 | 8 | 8 | 237 | 6.4537 | 0.73568 | 4.7479 | 0.0025287
41 | 100.809 | 8.2659 | 8.266 | 30.0399 | 30.0405 | 6.7756 | 2.0195 | 8 | 8 | 144 | 6.7377 | 0.70655 | 4.7605 | 0.0025566
3 | 201.122 | 8.211 | 8.2113 | 32.2903 | 32.291 | 6.6854 | 1.5554 | 2 | 2 | 11 | 6.7343 | 0.49448 | 3.33 | 0.002586
3 | 10.618 | 9.5821 | 9.5838 | 30.0436 | 30.037 | 6.5789 | 2.0957 | 12 | 12 | 21 | 6.4585 | 0.74142 | 4.7884 | 0.0026072
41 | 5.49 | 11.0153 | 11.0173 | 25.2165 | 25.2101 | 6.5747 | 2.4869 | 15 | 15 | 151 | 6.4433 | 0.91496 | 5.8954 | 0.0026572
62 | 125.832 | 8.181 | 8.1814 | 30.171 | 30.1967 | 6.7829 | 1.8891 | 7 | 7 | 248 | 6.7672 | 0.64709 | 4.379 | 0.0026611
30 | 200.613 | 8.5109 | 8.5187 | 31.8098 | 31.7965 | 6.6613 | 1.645 | 2 | 2 | 112 | 6.7071 | 0.53554 | 3.5919 | 0.0026658
13 | 100.315 | 7.1307 | 7.1287 | 33.6179 | 33.619 | 6.7923 | 1.2089 | 5 | 5 | 55 | 6.7639 | 0.33638 | 2.2752 | 0.002808
22 | 10.485 | 9.8853 | 9.8865 | 30.5803 | 30.5802 | 6.5126 | 2.2222 | 8 | 8 | 93 | 6.3916 | 0.79945 | 5.1098 | 0.0028251
41 | 174.914 | 8.4427 | 8.4426 | 30.5354 | 30.536 | 6.7269 | 1.8134 | 5 | 5 | 141 | 6.7515 | 0.60688 | 4.0974 | 0.0029207
15 | 99.695 | 7.2203 | 7.223 | 33.5467 | 33.5445 | 6.7816 | 1.256 | 4 | 4 | 67 | 6.7518 | 0.35801 | 2.4172 | 0.0029209
3 | 100.96 | 8.8084 | 8.8058 | 31.3636 | 31.3685 | 6.6362 | 1.7737 | 6 | 6 | 15 | 6.5952 | 0.59464 | 3.9217 | 0.0029361
30 | 40.444 | 9.5171 | 9.518 | 29.6911 | 29.6903 | 6.6033 | 2.0899 | 9 | 9 | 119 | 6.5075 | 0.73912 | 4.8098 | 0.0029442
41 | 150.898 | 8.4225 | 8.4224 | 30.4344 | 30.4347 | 6.7344 | 1.8412 | 6 | 6 | 142 | 6.7385 | 0.62551 | 4.215 | 0.0030091
37 | 150.359 | 9.0426 | 9.0412 | 30.5743 | 30.58 | 6.6354 | 1.8683 | 2 | 2 | 126 | 6.6344 | 0.63794 | 4.2324 | 0.0030134
49 | 75.301 | 8.047 | 8.0482 | 29.8616 | 29.8877 | 6.8173 | 1.9785 | 9 | 9 | 178 | 6.7588 | 0.68834 | 4.6524 | 0.003056
41 | 75.276 | 8.1923 | 8.1923 | 29.8036 | 29.8043 | 6.7974 | 1.996 | 9 | 9 | 145 | 6.7379 | 0.69635 | 4.6919 | 0.0030828
37 | 175.041 | 9.0139 | 9.0143 | 30.6512 | 30.6514 | 6.6364 | 1.8511 | 1 | 1 | 125 | 6.6565 | 0.63021 | 4.195 | 0.0031602
3 | 50.94 | 9.1815 | 9.1826 | 30.7368 | 30.7367 | 6.608 | 1.9175 | 8 | 8 | 17 | 6.5229 | 0.66057 | 4.3088 | 0.0031963
10 | 20.841 | 8.827 | 8.8254 | 31.8418 | 31.8429 | 6.613 | 1.8586 | 9 | 9 | 47 | 6.505 | 0.63373 | 4.1224 | 0.0032559
57 | 315.07 | 9.044 | 9.0439 | 30.8674 | 30.8679 | 6.6227 | 1.3649 | 1 | 1 | 201 | 6.7623 | 0.40819 | 2.7603 | 0.0033353
49 | 49.777 | 8.2 | 8.201 | 29.6614 | 29.6857 | 6.8025 | 1.9665 | 10 | 10 | 179 | 6.7209 | 0.68314 | 4.5913 | 0.0033492
49 | 200.624 | 8.5052 | 8.5008 | 30.6091 | 30.633 | 6.7143 | 1.7523 | 4 | 4 | 173 | 6.7605 | 0.57857 | 3.9114 | 0.0033524
49 | 342.43 | 8.7177 | 8.7184 | 30.8205 | 30.8572 | 6.6731 | 1.5568 | 1 | 1 | 170 | 6.8402 | 0.4892 | 3.3462 | 0.0033612
62 | 50.809 | 8.7459 | 8.7477 | 29.5744 | 29.5746 | 6.7229 | 1.9308 | 10 | 10 | 251 | 6.6392 | 0.66692 | 4.4278 | 0.0034373
57 | 50.148 | 8.8639 | 8.8696 | 29.5659 | 29.5641 | 6.7055 | 1.9362 | 10 | 10 | 210 | 6.6206 | 0.66248 | 4.386 | 0.0034395
1 | 19.606 | 9.8952 | 9.9042 | 29.002 | 28.9991 | 6.5774 | 2.2157 | 6 | 6 | 6 | 6.4627 | 0.79712 | 5.1515 | 0.0034585
30 | 1.736 | 10.4698 | 10.468 | 28.224 | 28.2278 | 6.5274 | 2.5398 | 14 | 14 | 124 | 6.3962 | 0.93828 | 6.0014 | 0.0035144
25 | 100.568 | 8.1678 | 8.1715 | 32.3658 | 32.3609 | 6.6887 | 1.5507 | 7 | 7 | 102 | 6.6518 | 0.49335 | 3.2817 | 0.00359
30 | 150.409 | 8.6481 | 8.6487 | 31.557 | 31.5565 | 6.6517 | 1.7065 | 4 | 4 | 114 | 6.6536 | 0.55736 | 3.7084 | 0.0036014
37 | 9.644 | 10.0092 | 10.0296 | 28.4139 | 28.3805 | 6.5857 | 2.3802 | 10 | 10 | 134 | 6.462 | 0.87253 | 5.6383 | 0.0036647
25 | 49.706 | 9.3745 | 9.376 | 30.7571 | 30.7552 | 6.5789 | 1.9565 | 9 | 9 | 104 | 6.4919 | 0.67888 | 4.4072 | 0.0036662
25 | 5.489 | 9.9589 | 9.9588 | 30.2944 | 30.2953 | 6.514 | 2.2637 | 14 | 14 | 109 | 6.3886 | 0.8193 | 5.2342 | 0.0036724
49 | 174.907 | 8.446 | 8.4469 | 30.5402 | 30.567 | 6.7262 | 1.8143 | 5 | 5 | 174 | 6.7508 | 0.60653 | 4.0945 | 0.0037086
15 | 1.502 | 9.4119 | 9.4113 | 31.3992 | 31.3997 | 6.5463 | 2.0602 | 13 | 13 | 75 | 6.42 | 0.72635 | 4.6632 | 0.0037297
37 | 40.002 | 9.5746 | 9.5749 | 29.3932 | 29.3927 | 6.6075 | 2.0706 | 7 | 7 | 131 | 6.5109 | 0.73111 | 4.7602 | 0.0037482
57 | 175.486 | 8.3963 | 8.3955 | 30.4623 | 30.4619 | 6.7372 | 1.8086 | 5 | 5 | 205 | 6.7627 | 0.60385 | 4.0836 | 0.0037592
25 | 40.18 | 9.6386 | 9.6389 | 30.5713 | 30.5714 | 6.5484 | 2.0557 | 10 | 10 | 105 | 6.4525 | 0.72433 | 4.6737 | 0.003764
37 | 74.976 | 9.2315 | 9.2301 | 29.8618 | 29.8616 | 6.638 | 1.9967 | 5 | 5 | 129 | 6.5723 | 0.69738 | 4.5834 | 0.003798
30 | 100.794 | 9.1024 | 9.1025 | 30.5776 | 30.5783 | 6.6264 | 1.8939 | 6 | 6 | 116 | 6.5833 | 0.65043 | 4.282 | 0.0038228
30 | 30.966 | 9.6749 | 9.6743 | 29.39 | 29.3916 | 6.593 | 2.1639 | 10 | 10 | 120 | 6.4886 | 0.77382 | 5.021 | 0.0038275
3 | 125.142 | 8.5877 | 8.5898 | 31.7174 | 31.7155 | 6.6538 | 1.7006 | 5 | 5 | 14 | 6.6347 | 0.56211 | 3.7294 | 0.0038416
46 | 250.333 | 8.6361 | 8.6359 | 30.7116 | 30.7115 | 6.6901 | 1.614 | 4 | 4 | 156 | 6.7781 | 0.5148 | 3.4893 | 0.0039076
62 | 175.175 | 8.3425 | 8.3429 | 30.4034 | 30.4043 | 6.7479 | 1.7748 | 5 | 5 | 246 | 6.7737 | 0.59607 | 4.0376 | 0.0039078
3 | 227.407 | 8.214 | 8.2142 | 32.2899 | 32.2909 | 6.685 | 1.554 | 1 | 1 | 10 | 6.7565 | 0.48726 | 3.2922 | 0.0040106
53 | 50.864 | 8.9384 | 8.9453 | 29.5386 | 29.5365 | 6.6955 | 1.9359 | 9 | 9 | 194 | 6.6108 | 0.66176 | 4.3748 | 0.0040409
62 | 40.801 | 9.0667 | 9.0674 | 29.4537 | 29.4592 | 6.68 | 2.0191 | 11 | 11 | 252 | 6.5862 | 0.70789 | 4.6623 | 0.0040796
37 | 125.272 | 9.0916 | 9.0918 | 30.4329 | 30.4331 | 6.6342 | 1.8997 | 3 | 3 | 127 | 6.6117 | 0.65348 | 4.3206 | 0.0042311
18 | 51.337 | 7.5887 | 7.5847 | 33.0742 | 33.0798 | 6.7457 | 1.3842 | 4 | 4 | 79 | 6.6712 | 0.41794 | 2.7882 | 0.0042425
13 | 50.875 | 8.3849 | 8.3855 | 32.2261 | 32.2258 | 6.6621 | 1.6561 | 7 | 7 | 57 | 6.5817 | 0.5422 | 3.5686 | 0.0042867
18 | 101.096 | 7.1571 | 7.1573 | 33.6027 | 33.6028 | 6.7889 | 1.2322 | 2 | 2 | 77 | 6.7609 | 0.33987 | 2.2978 | 0.0043308
46 | 200.292 | 8.5302 | 8.5295 | 30.5869 | 30.5867 | 6.7115 | 1.7418 | 5 | 5 | 157 | 6.7572 | 0.57276 | 3.8703 | 0.0043379
58 | 30.081 | 9.9606 | 9.9633 | 28.9794 | 28.9773 | 6.5689 | 2.0719 | 9 | 9 | 225 | 6.4626 | 0.73229 | 4.7325 | 0.0043457
3 | 41.022 | 9.2239 | 9.2241 | 30.6902 | 30.6908 | 6.6037 | 1.9313 | 9 | 9 | 18 | 6.5102 | 0.66808 | 4.3493 | 0.0043576
37 | 100.239 | 9.1851 | 9.1835 | 30.2065 | 30.2076 | 6.6301 | 1.9426 | 4 | 4 | 128 | 6.5859 | 0.67324 | 4.4339 | 0.004362
22 | 20.849 | 9.7727 | 9.7723 | 30.6144 | 30.6156 | 6.5273 | 2.1421 | 7 | 7 | 92 | 6.4151 | 0.76441 | 4.9038 | 0.0043741
30 | 20.065 | 9.8129 | 9.8128 | 29.1852 | 29.1858 | 6.5816 | 2.2147 | 11 | 11 | 121 | 6.4676 | 0.79768 | 5.1591 | 0.0044803
30 | 124.528 | 8.9496 | 8.9533 | 30.9183 | 30.9092 | 6.6344 | 1.8297 | 5 | 5 | 115 | 6.6123 | 0.62193 | 4.1124 | 0.0046776
15 | 125.874 | 6.8195 | 6.8196 | 33.8591 | 33.8592 | 6.8302 | 1.1154 | 3 | 3 | 66 | 6.8271 | 0.29562 | 2.0182 | 0.0047886
59 | 20.067 | 10.9076 | 10.9069 | 28.3184 | 28.3234 | 6.4616 | 2.2239 | 9 | 9 | 238 | 6.3446 | 0.79257 | 5.0285 | 0.0048583
13 | 39.866 | 8.5936 | 8.5936 | 31.9793 | 31.9795 | 6.6416 | 1.7219 | 8 | 8 | 58 | 6.5508 | 0.57298 | 3.7535 | 0.0049895
10 | 30.82 | 8.7982 | 8.7991 | 32.0895 | 32.0888 | 6.6066 | 1.8679 | 8 | 8 | 46 | 6.5073 | 0.63972 | 4.1628 | 0.0050035
30 | 50.6 | 9.438 | 9.4379 | 29.8377 | 29.8381 | 6.6086 | 2.0487 | 8 | 8 | 118 | 6.5216 | 0.72239 | 4.7111 | 0.0050111
53 | 10.793 | 11.9409 | 11.9324 | 26.204 | 26.2134 | 6.4044 | 2.3846 | 13 | 13 | 198 | 6.2775 | 0.86559 | 5.4337 | 0.0052645
10 | 5.772 | 9.7217 | 9.7209 | 31.762 | 31.7622 | 6.4866 | 2.2219 | 11 | 11 | 49 | 6.3632 | 0.80182 | 5.1022 | 0.0053134
37 | 50.061 | 9.4709 | 9.4707 | 29.5464 | 29.5466 | 6.6162 | 2.0487 | 6 | 6 | 130 | 6.5285 | 0.7227 | 4.7181 | 0.0053399
18 | 76.128 | 7.3077 | 7.3115 | 33.4346 | 33.4303 | 6.773 | 1.2863 | 3 | 3 | 78 | 6.7221 | 0.36354 | 2.4437 | 0.0054173
15 | 76.135 | 7.8423 | 7.844 | 32.9673 | 32.9623 | 6.7117 | 1.4633 | 5 | 5 | 68 | 6.6565 | 0.45529 | 3.0306 | 0.0054677
57 | 75.015 | 8.2186 | 8.2141 | 29.7905 | 29.7938 | 6.7939 | 1.9291 | 9 | 9 | 209 | 6.734 | 0.66829 | 4.5003 | 0.005573
57 | 300.351 | 9.0666 | 9.0664 | 30.8621 | 30.8625 | 6.6195 | 1.3754 | 2 | 2 | 202 | 6.7462 | 0.40408 | 2.726 | 0.0055769
18 | 40.259 | 7.9755 | 7.9708 | 32.6095 | 32.6154 | 6.7071 | 1.5031 | 5 | 5 | 80 | 6.6204 | 0.46243 | 3.0614 | 0.0056044
25 | 29.9 | 9.8048 | 9.8057 | 30.476 | 30.476 | 6.5285 | 2.1262 | 11 | 11 | 106 | 6.4235 | 0.75841 | 4.8716 | 0.0056224
49 | 30.335 | 8.5112 | 8.5104 | 29.3039 | 29.3293 | 6.7705 | 2.0004 | 12 | 12 | 181 | 6.6703 | 0.68957 | 4.5996 | 0.0057087
13 | 9.963 | 9.4981 | 9.5003 | 31.5421 | 31.5416 | 6.5278 | 2.0381 | 11 | 11 | 61 | 6.4083 | 0.71827 | 4.6029 | 0.0057824
7 | 150.88 | 6.9855 | 6.9849 | 33.7892 | 33.79 | 6.8073 | 1.1881 | 5 | 5 | 28 | 6.8245 | 0.31828 | 2.1721 | 0.005791
41 | 1.328 | 13.5112 | 13.5477 | 7.6875 | 7.4685 | 6.9506 | 2.852 | 16 | 16 | 152 | 6.8007 | 1.0903 | 7.4146 | 0.0058016
22 | 5.35 | 9.8965 | 9.8967 | 30.5768 | 30.577 | 6.5111 | 2.2275 | 9 | 9 | 94 | 6.386 | 0.80489 | 5.14 | 0.0058312
49 | 300.369 | 8.7711 | 8.7709 | 30.7811 | 30.8122 | 6.6669 | 1.5541 | 2 | 2 | 171 | 6.7968 | 0.48532 | 3.2986 | 0.005999
22 | 75.617 | 8.5126 | 8.5117 | 31.9018 | 31.9036 | 6.6571 | 1.6778 | 3 | 3 | 88 | 6.5966 | 0.54183 | 3.5743 | 0.0060047
7 | 125.074 | 7.1047 | 7.105 | 33.7234 | 33.7236 | 6.7917 | 1.2584 | 6 | 6 | 29 | 6.785 | 0.35016 | 2.3758 | 0.0060069
7 | 75.513 | 7.4348 | 7.4349 | 33.4729 | 33.4732 | 6.7517 | 1.414 | 8 | 8 | 31 | 6.6992 | 0.42124 | 2.822 | 0.006042
41 | 125.077 | 8.4756 | 8.4753 | 30.2813 | 30.281 | 6.733 | 1.9534 | 7 | 7 | 143 | 6.7145 | 0.67994 | 4.5655 | 0.0061472
62 | 30.721 | 9.3036 | 9.3051 | 29.2773 | 29.2814 | 6.6524 | 2.0497 | 12 | 12 | 253 | 6.549 | 0.72399 | 4.7414 | 0.00617
58 | 40.04 | 9.7916 | 9.7922 | 29.1132 | 29.1092 | 6.5877 | 2.0565 | 8 | 8 | 224 | 6.4903 | 0.72718 | 4.7196 | 0.0062427
62 | 10.277 | 12.6543 | 12.6489 | 26.9616 | 26.9771 | 6.2782 | 2.4983 | 14 | 14 | 255 | 6.1514 | 0.91652 | 5.6378 | 0.0063105
53 | 124.726 | 8.2081 | 8.2087 | 30.2231 | 30.2231 | 6.7764 | 1.9456 | 6 | 6 | 191 | 6.7596 | 0.66377 | 4.4868 | 0.0064749
10 | 75.714 | 7.3751 | 7.3751 | 33.4843 | 33.4845 | 6.7604 | 1.3439 | 5 | 5 | 43 | 6.7086 | 0.3888 | 2.6083 | 0.0064848
49 | 125.246 | 8.2048 | 8.2034 | 30.222 | 30.2457 | 6.777 | 1.9519 | 7 | 7 | 176 | 6.7606 | 0.66654 | 4.5062 | 0.0065785
46 | 174.884 | 8.592 | 8.5754 | 30.4834 | 30.4898 | 6.7066 | 1.8199 | 6 | 6 | 158 | 6.73 | 0.6061 | 4.079 | 0.0066796
57 | 10.569 | 13.8554 | 13.8552 | 26.2497 | 26.2502 | 6.1498 | 3.124 | 14 | 14 | 214 | 6.0235 | 1.2156 | 7.3222 | 0.0068326
18 | 30.932 | 8.2297 | 8.2287 | 32.2979 | 32.2992 | 6.6823 | 1.5894 | 6 | 6 | 81 | 6.586 | 0.50064 | 3.2972 | 0.0068351
22 | 2.437 | 9.8949 | 9.8948 | 30.5769 | 30.5772 | 6.5113 | 2.2259 | 10 | 10 | 95 | 6.3838 | 0.80535 | 5.1412 | 0.0070024
1 | 10.051 | 10.265 | 10.2622 | 28.7558 | 28.7604 | 6.5344 | 2.2955 | 7 | 7 | 7 | 6.4108 | 0.823 | 5.2761 | 0.0071701
15 | 30.637 | 8.6997 | 8.7016 | 31.8677 | 31.8663 | 6.6307 | 1.7651 | 9 | 9 | 71 | 6.5315 | 0.59495 | 3.8859 | 0.0071956
25 | 125.405 | 7.6072 | 7.5946 | 33.0909 | 33.1071 | 6.7421 | 1.3591 | 6 | 6 | 101 | 6.7311 | 0.39498 | 2.6587 | 0.007217
15 | 5.078 | 9.4293 | 9.4286 | 31.3906 | 31.3911 | 6.5442 | 2.0581 | 12 | 12 | 74 | 6.4207 | 0.72899 | 4.6806 | 0.0073385
15 | 40.617 | 8.6268 | 8.628 | 31.9393 | 31.938 | 6.6384 | 1.7391 | 8 | 8 | 70 | 6.548 | 0.58322 | 3.8189 | 0.0073392
62 | 349.956 | 8.7967 | 8.7969 | 30.609 | 30.6097 | 6.6705 | 1.431 | 1 | 1 | 242 | 6.8434 | 0.44247 | 3.028 | 0.007385
15 | 9.932 | 9.4061 | 9.4071 | 31.4056 | 31.405 | 6.5469 | 2.0411 | 11 | 11 | 73 | 6.4275 | 0.7213 | 4.6361 | 0.0074098
18 | 116.696 | 7.1044 | 7.1046 | 33.6647 | 33.6649 | 6.7943 | 1.217 | 1 | 1 | 76 | 6.7804 | 0.34496 | 2.339 | 0.0076746
10 | 99.841 | 7.2616 | 7.2609 | 33.5857 | 33.5865 | 6.7735 | 1.2982 | 4 | 4 | 42 | 6.7434 | 0.36668 | 2.4727 | 0.0076822
25 | 1.289 | 9.9611 | 9.961 | 30.2834 | 30.2851 | 6.5141 | 2.2714 | 15 | 15 | 110 | 6.3853 | 0.82691 | 5.2801 | 0.0077795
25 | 148.064 | 7.4648 | 7.4649 | 33.2758 | 33.2761 | 6.7558 | 1.3195 | 5 | 5 | 100 | 6.7657 | 0.3762 | 2.5453 | 0.0078972
10 | 2.215 | 9.726 | 9.7257 | 31.7609 | 31.7611 | 6.4861 | 2.2223 | 12 | 12 | 50 | 6.3597 | 0.80464 | 5.1173 | 0.0079455
10 | 126.161 | 7.0647 | 7.0655 | 33.7609 | 33.7608 | 6.7962 | 1.2277 | 3 | 3 | 41 | 6.7909 | 0.33413 | 2.269 | 0.008031
57 | 20.47 | 11.0967 | 11.0873 | 28.2481 | 28.2538 | 6.438 | 2.2286 | 13 | 13 | 213 | 6.321 | 0.80762 | 5.105 | 0.0080427
25 | 20.098 | 9.8521 | 9.8523 | 30.4485 | 30.4488 | 6.5228 | 2.1562 | 12 | 12 | 107 | 6.4097 | 0.77458 | 4.9648 | 0.0081108
53 | 250.404 | 8.8295 | 8.8292 | 30.757 | 30.7575 | 6.6592 | 1.5139 | 2 | 2 | 187 | 6.7453 | 0.46456 | 3.1336 | 0.0084173
30 | 200.611 | 8.5016 | 8.5003 | 31.8263 | 31.8287 | 6.662 | 1.6438 | 1 | 1 | 111 | 6.7079 | 0.54073 | 3.6272 | 0.0084262
46 | 10.054 | 10.7695 | 10.7872 | 28.3348 | 28.3176 | 6.4803 | 2.3847 | 15 | 15 | 167 | 6.3554 | 0.86246 | 5.4813 | 0.0084473
41 | 314.549 | 8.6573 | 8.6571 | 30.786 | 30.7862 | 6.6837 | 1.565 | 1 | 1 | 137 | 6.8272 | 0.50521 | 3.4492 | 0.0088994
41 | 10.39 | 10.1778 | 10.1775 | 27.5045 | 27.5064 | 6.5996 | 2.3629 | 14 | 14 | 150 | 6.4754 | 0.852 | 5.5171 | 0.0089452
58 | 20.033 | 10.4568 | 10.4458 | 28.5731 | 28.5816 | 6.5147 | 2.1455 | 10 | 10 | 226 | 6.3986 | 0.75261 | 4.8157 | 0.0089619
59 | 49.956 | 9.0524 | 9.052 | 29.4484 | 29.4486 | 6.6824 | 1.9398 | 6 | 6 | 235 | 6.5963 | 0.65836 | 4.3428 | 0.0092156
57 | 250.587 | 8.8878 | 8.8874 | 30.7752 | 30.7755 | 6.6497 | 1.4742 | 3 | 3 | 203 | 6.7355 | 0.4455 | 3.0006 | 0.0092893
13 | 2.007 | 9.5179 | 9.5175 | 31.5406 | 31.5414 | 6.5251 | 2.0434 | 13 | 13 | 63 | 6.399 | 0.72445 | 4.6357 | 0.0094902
10 | 215.414 | 6.5931 | 6.5928 | 33.9171 | 33.9178 | 6.8633 | 1.0697 | 1 | 1 | 39 | 6.9422 | 0.26027 | 1.8068 | 0.0096652
7 | 5.547 | 9.8523 | 9.8436 | 31.7031 | 31.704 | 6.4705 | 2.2229 | 14 | 14 | 37 | 6.3465 | 0.8067 | 5.1198 | 0.0097248
22 | 123.938 | 7.2327 | 7.2327 | 33.4834 | 33.4837 | 6.7825 | 1.2514 | 1 | 1 | 86 | 6.7736 | 0.34317 | 2.3245 | 0.0098228
46 | 20.075 | 9.7648 | 9.7644 | 28.9651 | 28.9655 | 6.5979 | 2.3234 | 14 | 14 | 166 | 6.4839 | 0.83299 | 5.401 | 0.0099036
15 | 149.521 | 6.8053 | 6.8055 | 33.8682 | 33.8684 | 6.832 | 1.1139 | 2 | 2 | 65 | 6.8499 | 0.30047 | 2.0582 | 0.010339
30 | 10.07 | 9.9106 | 9.9108 | 29.016 | 29.0159 | 6.5746 | 2.2553 | 12 | 12 | 122 | 6.452 | 0.82215 | 5.3045 | 0.010357
57 | 200.456 | 8.6083 | 8.6083 | 30.6226 | 30.6234 | 6.6981 | 1.657 | 4 | 4 | 204 | 6.7433 | 0.52794 | 3.5601 | 0.010396
15 | 50.755 | 8.6012 | 8.6013 | 32.005 | 32.0043 | 6.6394 | 1.7205 | 6 | 6 | 69 | 6.5576 | 0.57799 | 3.7902 | 0.010606
13 | 20.133 | 9.331 | 9.3355 | 31.5761 | 31.5747 | 6.5506 | 1.9781 | 10 | 10 | 60 | 6.4399 | 0.69589 | 4.4815 | 0.010807
10 | 150.173 | 6.989 | 6.9892 | 33.7997 | 33.8001 | 6.8063 | 1.1989 | 2 | 2 | 40 | 6.8228 | 0.31818 | 2.1709 | 0.010829
7 | 2.455 | 9.8544 | 9.8541 | 31.7036 | 31.7039 | 6.4702 | 2.216 | 15 | 15 | 38 | 6.3437 | 0.80463 | 5.1043 | 0.010832
13 | 30.419 | 8.8908 | 8.8895 | 31.7205 | 31.7215 | 6.6087 | 1.8173 | 9 | 9 | 59 | 6.5084 | 0.62284 | 4.0537 | 0.011224
7 | 201.053 | 6.6059 | 6.6062 | 33.9078 | 33.9081 | 6.8617 | 1.0468 | 3 | 3 | 26 | 6.9276 | 0.24819 | 1.7194 | 0.011281
10 | 39.977 | 8.5319 | 8.5315 | 32.2147 | 32.215 | 6.6407 | 1.7414 | 7 | 7 | 45 | 6.5503 | 0.58824 | 3.8532 | 0.011297
18 | 9.861 | 9.6598 | 9.664 | 30.7889 | 30.7856 | 6.5362 | 2.1163 | 8 | 8 | 83 | 6.4155 | 0.7599 | 4.8752 | 0.011646
22 | 100.409 | 7.238 | 7.2387 | 33.4751 | 33.4748 | 6.782 | 1.2515 | 2 | 2 | 87 | 6.7527 | 0.34116 | 2.3038 | 0.011897
58 | 188.513 | 8.4518 | 8.4506 | 30.5117 | 30.5116 | 6.7266 | 1.7402 | 1 | 1 | 217 | 6.7629 | 0.58829 | 3.9785 | 0.011929
22 | 30.314 | 9.4224 | 9.4372 | 30.7609 | 30.7562 | 6.5717 | 1.9834 | 6 | 6 | 91 | 6.4686 | 0.69959 | 4.5253 | 0.012079
7 | 249.672 | 6.4942 | 6.4944 | 33.94 | 33.9404 | 6.878 | 1.0455 | 1 | 1 | 24 | 6.9889 | 0.24664 | 1.7237 | 0.012271
1 | 30.103 | 9.2095 | 9.2091 | 29.2737 | 29.275 | 6.6665 | 2.1681 | 5 | 5 | 5 | 6.563 | 0.75942 | 4.984 | 0.012513
15 | 20.751 | 8.9023 | 8.9003 | 31.7204 | 31.7218 | 6.6071 | 1.8426 | 10 | 10 | 72 | 6.4987 | 0.63585 | 4.1321 | 0.012659
49 | 10.696 | 11.7375 | 11.7309 | 27.0317 | 27.0612 | 6.3989 | 2.447 | 14 | 14 | 183 | 6.2726 | 0.88647 | 5.5604 | 0.012932
18 | 6.247 | 9.7117 | 9.7116 | 30.7478 | 30.7481 | 6.5305 | 2.1539 | 9 | 9 | 84 | 6.4066 | 0.77941 | 4.9934 | 0.013974
59 | 125.215 | 8.2371 | 8.2363 | 30.3076 | 30.3067 | 6.7683 | 1.8041 | 3 | 3 | 232 | 6.7516 | 0.59158 | 3.9941 | 0.014001
7 | 175.007 | 6.7308 | 6.731 | 33.8521 | 33.8524 | 6.8445 | 1.0217 | 4 | 4 | 27 | 6.8857 | 0.2338 | 1.6099 | 0.014243
13 | 5.121 | 9.5002 | 9.4995 | 31.542 | 31.5423 | 6.5275 | 2.0394 | 12 | 12 | 62 | 6.404 | 0.72759 | 4.6595 | 0.014486
53 | 200.576 | 8.5187 | 8.5186 | 30.5545 | 30.5552 | 6.7146 | 1.7086 | 3 | 3 | 188 | 6.7607 | 0.54717 | 3.6993 | 0.014733
37 | 1.418 | 13.5964 | 13.5287 | 23.8174 | 23.9353 | 6.2771 | 3.7162 | 12 | 12 | 136 | 6.1416 | 1.4943 | 9.1774 | 0.014902
1 | 39.768 | 8.9342 | 8.9396 | 29.5067 | 29.5112 | 6.6975 | 2.0494 | 4 | 4 | 4 | 6.6035 | 0.70267 | 4.6401 | 0.01501
7 | 249.892 | 6.495 | 6.4952 | 33.9398 | 33.9402 | 6.8779 | 1.0454 | 2 | 2 | 25 | 6.989 | 0.24369 | 1.7032 | 0.015171
13 | 74.78 | 7.7868 | 7.7792 | 32.8944 | 32.9055 | 6.7233 | 1.4578 | 6 | 6 | 56 | 6.6674 | 0.46323 | 3.0885 | 0.015912
53 | 2.715 | 15.8103 | 15.7907 | 17.3848 | 17.4185 | 6.2383 | 3.5829 | 15 | 15 | 200 | 6.1024 | 1.4349 | 8.7566 | 0.016439
46 | 5.039 | 11.7519 | 11.8362 | 27.163 | 27.1114 | 6.3917 | 2.5761 | 16 | 16 | 168 | 6.261 | 0.94174 | 5.8962 | 0.016636
41 | 249.652 | 8.6001 | 8.6001 | 30.7182 | 30.7187 | 6.6952 | 1.6771 | 3 | 3 | 139 | 6.7829 | 0.53075 | 3.6001 | 0.016769
58 | 50.293 | 9.1962 | 9.1973 | 29.3941 | 29.3922 | 6.6633 | 1.9889 | 7 | 7 | 223 | 6.5768 | 0.67252 | 4.423 | 0.017512
57 | 2.537 | 15.3286 | 15.3231 | 16.4833 | 16.6714 | 6.3351 | 3.5099 | 16 | 16 | 216 | 6.1973 | 1.4026 | 8.6925 | 0.017524
53 | 20.157 | 10.6885 | 10.6881 | 28.5346 | 28.5362 | 6.4835 | 2.2467 | 12 | 12 | 197 | 6.3671 | 0.78829 | 5.0191 | 0.019566
1 | 50.163 | 9.1496 | 9.1496 | 29.8368 | 29.8382 | 6.6512 | 1.965 | 3 | 3 | 3 | 6.5651 | 0.65928 | 4.3282 | 0.01982
59 | 149.883 | 8.4029 | 8.4052 | 30.4418 | 30.4449 | 6.7371 | 1.5688 | 2 | 2 | 231 | 6.7404 | 0.47675 | 3.2135 | 0.021284
7 | 10.555 | 9.0745 | 9.1021 | 31.7756 | 31.7709 | 6.5795 | 1.9339 | 13 | 13 | 36 | 6.462 | 0.68619 | 4.4342 | 0.021303
1 | 99.943 | 8.5669 | 8.5667 | 30.6614 | 30.6625 | 6.7027 | 1.1668 | 1 | 1 | 1 | 6.6621 | 0.33563 | 2.236 | 0.021305
59 | 5.101 | 14.4738 | 14.4741 | 24.6475 | 24.6505 | 6.1323 | 3.1279 | 11 | 11 | 240 | 6.0015 | 1.1881 | 7.1304 | 0.022442
10 | 10.689 | 9.4011 | 9.4017 | 31.7628 | 31.7626 | 6.5326 | 2.0552 | 10 | 10 | 48 | 6.4141 | 0.74372 | 4.7703 | 0.023384
18 | 21.559 | 8.5438 | 8.5373 | 31.9032 | 31.9127 | 6.6523 | 1.7207 | 7 | 7 | 82 | 6.5463 | 0.59254 | 3.879 | 0.025083
49 | 19.761 | 9.5475 | 9.54 | 28.7119 | 28.7423 | 6.6405 | 2.2743 | 13 | 13 | 182 | 6.5268 | 0.79503 | 5.189 | 0.025432
63 | 1.24 | 15.683 | 15.6847 | 24.0168 | 24.0108 | 6.0063 | 3.7618 | 3 | 3 | 260 | 5.8745 | 1.4702 | 8.6365 | 0.030054
7 | 100.782 | 7.2895 | 7.2862 | 33.589 | 33.5918 | 6.769 | 1.3537 | 7 | 7 | 30 | 6.7395 | 0.43044 | 2.901 | 0.030711
58 | 1.309 | 14.9341 | 14.9719 | 19.9841 | 19.3653 | 6.2513 | 3.4929 | 13 | 13 | 229 | 6.1147 | 1.3455 | 8.2272 | 0.031861
58 | 150.282 | 8.2942 | 8.2945 | 30.3558 | 30.3563 | 6.7574 | 1.8848 | 3 | 3 | 219 | 6.762 | 0.67488 | 4.5636 | 0.03243
37 | 4.637 | 11.2178 | 11.22 | 27.0612 | 27.063 | 6.4697 | 2.7826 | 11 | 11 | 135 | 6.339 | 1.014 | 6.428 | 0.038713
62 | 1.713 | 17.1707 | 17.1724 | 18.8475 | 18.8445 | 6.0157 | 4.2912 | 16 | 16 | 257 | 5.8837 | 1.6972 | 9.9855 | 0.045004
63 | 9.757 | 12.9494 | 12.949 | 26.7171 | 26.7173 | 6.2488 | 2.657 | 2 | 2 | 259 | 6.1216 | 0.9462 | 5.7922 | 0.049158
1 | 74.892 | 8.6383 | 8.6383 | 30.4251 | 30.4267 | 6.7022 | 1.5845 | 2 | 2 | 2 | 6.6398 | 0.45549 | 3.0244 | 0.049708
59 | 1.349 | 15.4318 | 15.415 | 21.4345 | 21.4765 | 6.1331 | 3.4852 | 12 | 12 | 241 | 5.9987 | 1.3194 | 7.9144 | 0.054463
57 | 5.254 | 14.0558 | 13.9552 | 22.5958 | 22.6038 | 6.2645 | 2.9214 | 15 | 15 | 215 | 6.1315 | 1.171 | 7.1798 | 0.054775
58 | 4.879 | 14.2394 | 14.2536 | 25.6331 | 25.647 | 6.1245 | 3.1717 | 12 | 12 | 228 | 5.9941 | 1.1703 | 7.015 | 0.06026
59 | 9.861 | 12.9235 | 12.9232 | 26.9475 | 26.9482 | 6.2432 | 2.8436 | 10 | 10 | 239 | 6.1162 | 1.02 | 6.2387 | 0.060627
49 | 5.143 | 13.4116 | 13.4144 | 24.2925 | 24.3103 | 6.2828 | 2.8946 | 15 | 15 | 184 | 6.1504 | 1.0252 | 6.3054 | 0.078719
53 | 5.506 | 15.2884 | 15.3116 | 18.7332 | 18.6744 | 6.2537 | 3.4503 | 14 | 14 | 199 | 6.1199 | 1.2765 | 7.8118 | 0.08145
1 | 1.128 | 16.4589 | 16.4644 | 28.077 | 28.0872 | 5.7697 | 3.4231 | 9 | 9 | 9 | 5.6427 | 1.253 | 7.0705 | 0.092407
49 | 2.001 | 14.8123 | 14.8216 | 14.7042 | 14.7849 | 6.4738 | 3.7594 | 16 | 16 | 185 | 6.333 | 1.3829 | 8.7579 | 0.11625
62 | 5.5 | 16.7816 | 16.7775 | 19.8642 | 19.8758 | 6.0255 | 3.8645 | 15 | 15 | 256 | 5.896 | 1.4029 | 8.2715 | 0.14428

## bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Flag | Sample_Number | Number_of_bin_records | Sample_Number | Oxygen:Dissolved
1 | 99.943 | 8.5669 | 8.5667 | 30.6614 | 30.6625 | 6.7027 | 1.1668 | 1 | 1 | 0 | 1 | 241 | 1 | 2.236
1 | 74.892 | 8.6383 | 8.6383 | 30.4251 | 30.4267 | 6.7022 | 1.5845 | 2 | 2 | 0 | 2 | 241 | 2 | 3.0244
1 | 50.163 | 9.1496 | 9.1496 | 29.8368 | 29.8382 | 6.6512 | 1.965 | 3 | 3 | 0 | 3 | 241 | 3 | 4.3282
1 | 39.768 | 8.9342 | 8.9396 | 29.5067 | 29.5112 | 6.6975 | 2.0494 | 4 | 4 | 0 | 4 | 241 | 4 | 4.6401
1 | 30.103 | 9.2095 | 9.2091 | 29.2737 | 29.275 | 6.6665 | 2.1681 | 5 | 5 | 0 | 5 | 241 | 5 | 4.984
1 | 19.606 | 9.8952 | 9.9042 | 29.002 | 28.9991 | 6.5774 | 2.2157 | 6 | 6 | 0 | 6 | 241 | 6 | 5.1515
1 | 10.051 | 10.265 | 10.2622 | 28.7558 | 28.7604 | 6.5344 | 2.2955 | 7 | 7 | 0 | 7 | 241 | 7 | 5.2761
1 | 4.88 | 11.1368 | 11.1128 | 28.4283 | 28.4314 | 6.4251 | 2.6247 | 8 | 8 | 0 | 8 | 241 | 8 | 6.1638
1 | 1.128 | 16.4589 | 16.4644 | 28.077 | 28.0872 | 5.7697 | 3.4231 | 9 | 9 | 0 | 9 | 241 | 9 | 7.0705
3 | 227.407 | 8.214 | 8.2142 | 32.2899 | 32.2909 | 6.685 | 1.554 | 1 | 1 | 0 | 10 | 241 | 10 | 3.2922
3 | 201.122 | 8.211 | 8.2113 | 32.2903 | 32.291 | 6.6854 | 1.5554 | 2 | 2 | 0 | 11 | 241 | 11 | 3.33
3 | 175.878 | 8.2142 | 8.2156 | 32.2819 | 32.2808 | 6.6853 | 1.5601 | 3 | 3 | 0 | 12 | 241 | 12 | 3.3288
3 | 151.218 | 8.2993 | 8.3 | 32.1536 | 32.1539 | 6.6781 | 1.6003 | 4 | 4 | 0 | 13 | 241 | 13 | 3.4236
3 | 125.142 | 8.5877 | 8.5898 | 31.7174 | 31.7155 | 6.6538 | 1.7006 | 5 | 5 | 0 | 14 | 241 | 14 | 3.7294
3 | 100.96 | 8.8084 | 8.8058 | 31.3636 | 31.3685 | 6.6362 | 1.7737 | 6 | 6 | 0 | 15 | 241 | 15 | 3.9217
3 | 75.192 | 8.9681 | 8.9684 | 31.0748 | 31.0751 | 6.625 | 1.8354 | 7 | 7 | 0 | 16 | 241 | 16 | 4.0662
3 | 50.94 | 9.1815 | 9.1826 | 30.7368 | 30.7367 | 6.608 | 1.9175 | 8 | 8 | 0 | 17 | 241 | 17 | 4.3088
3 | 41.022 | 9.2239 | 9.2241 | 30.6902 | 30.6908 | 6.6037 | 1.9313 | 9 | 9 | 0 | 18 | 241 | 18 | 4.3493
3 | 30.416 | 9.3504 | 9.3513 | 30.5138 | 30.5133 | 6.5927 | 1.9784 | 10 | 10 | 0 | 19 | 241 | 19 | 4.4564
3 | 20.196 | 9.5322 | 9.5328 | 30.2533 | 30.2536 | 6.5773 | 2.0504 | 11 | 11 | 0 | 20 | 241 | 20 | 4.65
3 | 10.618 | 9.5821 | 9.5838 | 30.0436 | 30.037 | 6.5789 | 2.0957 | 12 | 12 | 0 | 21 | 241 | 21 | 4.7884
3 | 5.12 | 9.664 | 9.6644 | 30.0011 | 30.0015 | 6.5688 | 2.121 | 13 | 13 | 0 | 22 | 241 | 22 | 4.8512
3 | 1.381 | 9.7178 | 9.719 | 29.9179 | 29.9168 | 6.5645 | 2.1473 | 14 | 14 | 0 | 23 | 241 | 23 | 4.922
7 | 249.672 | 6.4942 | 6.4944 | 33.94 | 33.9404 | 6.878 | 1.0455 | 1 | 1 | 0 | 24 | 106 | 24 | 1.7237
7 | 249.892 | 6.495 | 6.4952 | 33.9398 | 33.9402 | 6.8779 | 1.0454 | 2 | 2 | 0 | 25 | 241 | 25 | 1.7032
7 | 201.053 | 6.6059 | 6.6062 | 33.9078 | 33.9081 | 6.8617 | 1.0468 | 3 | 3 | 0 | 26 | 241 | 26 | 1.7194
7 | 175.007 | 6.7308 | 6.731 | 33.8521 | 33.8524 | 6.8445 | 1.0217 | 4 | 4 | 0 | 27 | 241 | 27 | 1.6099
7 | 150.88 | 6.9855 | 6.9849 | 33.7892 | 33.79 | 6.8073 | 1.1881 | 5 | 5 | 0 | 28 | 241 | 28 | 2.1721
7 | 125.074 | 7.1047 | 7.105 | 33.7234 | 33.7236 | 6.7917 | 1.2584 | 6 | 6 | 0 | 29 | 241 | 29 | 2.3758
7 | 100.782 | 7.2895 | 7.2862 | 33.589 | 33.5918 | 6.769 | 1.3537 | 7 | 7 | 0 | 30 | 241 | 30 | 2.901
7 | 75.513 | 7.4348 | 7.4349 | 33.4729 | 33.4732 | 6.7517 | 1.414 | 8 | 8 | 0 | 31 | 241 | 31 | 2.822
7 | 50.932 | 7.8064 | 7.8055 | 32.859 | 32.86 | 6.7219 | 1.4776 | 9 | 9 | 0 | 32 | 241 | 32 | 3.0295
7 | 39.963 | 8.2094 | 8.2088 | 32.367 | 32.368 | 6.6823 | 1.6251 | 10 | 10 | 0 | 33 | 241 | 33 | 3.4587
7 | 31.227 | 8.4759 | 8.4692 | 32.1691 | 32.18 | 6.651 | 1.7296 | 11 | 11 | 0 | 34 | 241 | 34 | 3.7346
7 | 19.885 | 8.6306 | 8.629 | 31.9725 | 31.9737 | 6.6364 | 1.7848 | 12 | 12 | 0 | 35 | 241 | 35 | 3.9108
7 | 10.555 | 9.0745 | 9.1021 | 31.7756 | 31.7709 | 6.5795 | 1.9339 | 13 | 13 | 0 | 36 | 241 | 36 | 4.4342
7 | 5.547 | 9.8523 | 9.8436 | 31.7031 | 31.704 | 6.4705 | 2.2229 | 14 | 14 | 0 | 37 | 241 | 37 | 5.1198
7 | 2.455 | 9.8544 | 9.8541 | 31.7036 | 31.7039 | 6.4702 | 2.216 | 15 | 15 | 0 | 38 | 241 | 38 | 5.1043
10 | 215.414 | 6.5931 | 6.5928 | 33.9171 | 33.9178 | 6.8633 | 1.0697 | 1 | 1 | 0 | 39 | 241 | 39 | 1.8068
10 | 150.173 | 6.989 | 6.9892 | 33.7997 | 33.8001 | 6.8063 | 1.1989 | 2 | 2 | 0 | 40 | 241 | 40 | 2.1709
10 | 126.161 | 7.0647 | 7.0655 | 33.7609 | 33.7608 | 6.7962 | 1.2277 | 3 | 3 | 0 | 41 | 241 | 41 | 2.269
10 | 99.841 | 7.2616 | 7.2609 | 33.5857 | 33.5865 | 6.7735 | 1.2982 | 4 | 4 | 0 | 42 | 241 | 42 | 2.4727
10 | 75.714 | 7.3751 | 7.3751 | 33.4843 | 33.4845 | 6.7604 | 1.3439 | 5 | 5 | 0 | 43 | 241 | 43 | 2.6083
10 | 51.321 | 8.0086 | 8.0101 | 32.5666 | 32.5652 | 6.704 | 1.5423 | 6 | 6 | 0 | 44 | 241 | 44 | 3.209
10 | 39.977 | 8.5319 | 8.5315 | 32.2147 | 32.215 | 6.6407 | 1.7414 | 7 | 7 | 0 | 45 | 241 | 45 | 3.8532
10 | 30.82 | 8.7982 | 8.7991 | 32.0895 | 32.0888 | 6.6066 | 1.8679 | 8 | 8 | 0 | 46 | 241 | 46 | 4.1628
10 | 20.841 | 8.827 | 8.8254 | 31.8418 | 31.8429 | 6.613 | 1.8586 | 9 | 9 | 0 | 47 | 241 | 47 | 4.1224
10 | 10.689 | 9.4011 | 9.4017 | 31.7628 | 31.7626 | 6.5326 | 2.0552 | 10 | 10 | 0 | 48 | 241 | 48 | 4.7703
10 | 5.772 | 9.7217 | 9.7209 | 31.762 | 31.7622 | 6.4866 | 2.2219 | 11 | 11 | 0 | 49 | 241 | 49 | 5.1022
10 | 2.215 | 9.726 | 9.7257 | 31.7609 | 31.7611 | 6.4861 | 2.2223 | 12 | 12 | 0 | 50 | 191 | 50 | 5.1173
13 | 183.158 | 6.7629 | 6.7632 | 33.8805 | 33.8807 | 6.8381 | 1.0982 | 1 | 1 | 0 | 51 | 241 | 51 | 1.9465
13 | 175.581 | 6.7615 | 6.7617 | 33.8806 | 33.8808 | 6.8384 | 1.0963 | 2 | 2 | 0 | 52 | 241 | 52 | 1.9358
13 | 126.292 | 6.7651 | 6.7654 | 33.8784 | 33.8784 | 6.8379 | 1.0989 | 4 | 4 | 0 | 54 | 241 | 54 | 1.9236
13 | 100.315 | 7.1307 | 7.1287 | 33.6179 | 33.619 | 6.7923 | 1.2089 | 5 | 5 | 0 | 55 | 241 | 55 | 2.2752
13 | 74.78 | 7.7868 | 7.7792 | 32.8944 | 32.9055 | 6.7233 | 1.4578 | 6 | 6 | 0 | 56 | 241 | 56 | 3.0885
13 | 50.875 | 8.3849 | 8.3855 | 32.2261 | 32.2258 | 6.6621 | 1.6561 | 7 | 7 | 0 | 57 | 241 | 57 | 3.5686
13 | 39.866 | 8.5936 | 8.5936 | 31.9793 | 31.9795 | 6.6416 | 1.7219 | 8 | 8 | 0 | 58 | 241 | 58 | 3.7535
13 | 30.419 | 8.8908 | 8.8895 | 31.7205 | 31.7215 | 6.6087 | 1.8173 | 9 | 9 | 0 | 59 | 241 | 59 | 4.0537
13 | 20.133 | 9.331 | 9.3355 | 31.5761 | 31.5747 | 6.5506 | 1.9781 | 10 | 10 | 0 | 60 | 241 | 60 | 4.4815
13 | 9.963 | 9.4981 | 9.5003 | 31.5421 | 31.5416 | 6.5278 | 2.0381 | 11 | 11 | 0 | 61 | 241 | 61 | 4.6029
13 | 5.121 | 9.5002 | 9.4995 | 31.542 | 31.5423 | 6.5275 | 2.0394 | 12 | 12 | 0 | 62 | 241 | 62 | 4.6595
13 | 2.007 | 9.5179 | 9.5175 | 31.5406 | 31.5414 | 6.5251 | 2.0434 | 13 | 13 | 0 | 63 | 241 | 63 | 4.6357
15 | 173.432 | 6.8083 | 6.8085 | 33.8681 | 33.8684 | 6.8316 | 1.1138 | 1 | 1 | 0 | 64 | 241 | 64 | 1.9916
15 | 149.521 | 6.8053 | 6.8055 | 33.8682 | 33.8684 | 6.832 | 1.1139 | 2 | 2 | 0 | 65 | 241 | 65 | 2.0582
15 | 125.874 | 6.8195 | 6.8196 | 33.8591 | 33.8592 | 6.8302 | 1.1154 | 3 | 3 | 0 | 66 | 241 | 66 | 2.0182
15 | 99.695 | 7.2203 | 7.223 | 33.5467 | 33.5445 | 6.7816 | 1.256 | 4 | 4 | 0 | 67 | 241 | 67 | 2.4172
15 | 76.135 | 7.8423 | 7.844 | 32.9673 | 32.9623 | 6.7117 | 1.4633 | 5 | 5 | 0 | 68 | 241 | 68 | 3.0306
15 | 50.755 | 8.6012 | 8.6013 | 32.005 | 32.0043 | 6.6394 | 1.7205 | 6 | 6 | 0 | 69 | 241 | 69 | 3.7902
15 | 40.617 | 8.6268 | 8.628 | 31.9393 | 31.938 | 6.6384 | 1.7391 | 8 | 8 | 0 | 70 | 241 | 70 | 3.8189
15 | 30.637 | 8.6997 | 8.7016 | 31.8677 | 31.8663 | 6.6307 | 1.7651 | 9 | 9 | 0 | 71 | 241 | 71 | 3.8859
15 | 20.751 | 8.9023 | 8.9003 | 31.7204 | 31.7218 | 6.6071 | 1.8426 | 10 | 10 | 0 | 72 | 241 | 72 | 4.1321
15 | 9.932 | 9.4061 | 9.4071 | 31.4056 | 31.405 | 6.5469 | 2.0411 | 11 | 11 | 0 | 73 | 241 | 73 | 4.6361
15 | 5.078 | 9.4293 | 9.4286 | 31.3906 | 31.3911 | 6.5442 | 2.0581 | 12 | 12 | 0 | 74 | 241 | 74 | 4.6806
15 | 1.502 | 9.4119 | 9.4113 | 31.3992 | 31.3997 | 6.5463 | 2.0602 | 13 | 13 | 0 | 75 | 241 | 75 | 4.6632
18 | 116.696 | 7.1044 | 7.1046 | 33.6647 | 33.6649 | 6.7943 | 1.217 | 1 | 1 | 0 | 76 | 241 | 76 | 2.339
18 | 101.096 | 7.1571 | 7.1573 | 33.6027 | 33.6028 | 6.7889 | 1.2322 | 2 | 2 | 0 | 77 | 241 | 77 | 2.2978
18 | 76.128 | 7.3077 | 7.3115 | 33.4346 | 33.4303 | 6.773 | 1.2863 | 3 | 3 | 0 | 78 | 241 | 78 | 2.4437
18 | 51.337 | 7.5887 | 7.5847 | 33.0742 | 33.0798 | 6.7457 | 1.3842 | 4 | 4 | 0 | 79 | 241 | 79 | 2.7882
18 | 40.259 | 7.9755 | 7.9708 | 32.6095 | 32.6154 | 6.7071 | 1.5031 | 5 | 5 | 0 | 80 | 241 | 80 | 3.0614
18 | 30.932 | 8.2297 | 8.2287 | 32.2979 | 32.2992 | 6.6823 | 1.5894 | 6 | 6 | 0 | 81 | 241 | 81 | 3.2972
18 | 21.559 | 8.5438 | 8.5373 | 31.9032 | 31.9127 | 6.6523 | 1.7207 | 7 | 7 | 0 | 82 | 241 | 82 | 3.879
18 | 9.861 | 9.6598 | 9.664 | 30.7889 | 30.7856 | 6.5362 | 2.1163 | 8 | 8 | 0 | 83 | 241 | 83 | 4.8752
18 | 6.247 | 9.7117 | 9.7116 | 30.7478 | 30.7481 | 6.5305 | 2.1539 | 9 | 9 | 0 | 84 | 241 | 84 | 4.9934
18 | 2.005 | 9.6987 | 9.6985 | 30.7598 | 30.7602 | 6.5318 | 2.1742 | 10 | 10 | 0 | 85 | 241 | 85 | 4.9573
22 | 123.938 | 7.2327 | 7.2327 | 33.4834 | 33.4837 | 6.7825 | 1.2514 | 1 | 1 | 0 | 86 | 241 | 86 | 2.3245
22 | 100.409 | 7.238 | 7.2387 | 33.4751 | 33.4748 | 6.782 | 1.2515 | 2 | 2 | 0 | 87 | 241 | 87 | 2.3038
22 | 75.617 | 8.5126 | 8.5117 | 31.9018 | 31.9036 | 6.6571 | 1.6778 | 3 | 3 | 0 | 88 | 241 | 88 | 3.5743
22 | 50.486 | 8.944 | 8.9443 | 31.2898 | 31.2899 | 6.6193 | 1.8146 | 4 | 4 | 0 | 89 | 241 | 89 | 3.9983
22 | 40.569 | 9.0184 | 9.0213 | 31.1763 | 31.1731 | 6.6132 | 1.8385 | 5 | 5 | 0 | 90 | 241 | 90 | 4.0619
22 | 30.314 | 9.4224 | 9.4372 | 30.7609 | 30.7562 | 6.5717 | 1.9834 | 6 | 6 | 0 | 91 | 241 | 91 | 4.5253
22 | 20.849 | 9.7727 | 9.7723 | 30.6144 | 30.6156 | 6.5273 | 2.1421 | 7 | 7 | 0 | 92 | 241 | 92 | 4.9038
22 | 10.485 | 9.8853 | 9.8865 | 30.5803 | 30.5802 | 6.5126 | 2.2222 | 8 | 8 | 0 | 93 | 241 | 93 | 5.1098
22 | 5.35 | 9.8965 | 9.8967 | 30.5768 | 30.577 | 6.5111 | 2.2275 | 9 | 9 | 0 | 94 | 241 | 94 | 5.14
22 | 2.437 | 9.8949 | 9.8948 | 30.5769 | 30.5772 | 6.5113 | 2.2259 | 10 | 10 | 0 | 95 | 241 | 95 | 5.1412
25 | 148.064 | 7.4648 | 7.4649 | 33.2758 | 33.2761 | 6.7558 | 1.3195 | 5 | 5 | 0 | 100 | 241 | 100 | 2.5453
25 | 125.405 | 7.6072 | 7.5946 | 33.0909 | 33.1071 | 6.7421 | 1.3591 | 6 | 6 | 0 | 101 | 241 | 101 | 2.6587
25 | 100.568 | 8.1678 | 8.1715 | 32.3658 | 32.3609 | 6.6887 | 1.5507 | 7 | 7 | 0 | 102 | 241 | 102 | 3.2817
25 | 75.471 | 8.7745 | 8.775 | 31.5387 | 31.5383 | 6.6337 | 1.7505 | 8 | 8 | 0 | 103 | 241 | 103 | 3.8277
25 | 49.706 | 9.3745 | 9.376 | 30.7571 | 30.7552 | 6.5789 | 1.9565 | 9 | 9 | 0 | 104 | 241 | 104 | 4.4072
25 | 40.18 | 9.6386 | 9.6389 | 30.5713 | 30.5714 | 6.5484 | 2.0557 | 10 | 10 | 0 | 105 | 241 | 105 | 4.6737
25 | 29.9 | 9.8048 | 9.8057 | 30.476 | 30.476 | 6.5285 | 2.1262 | 11 | 11 | 0 | 106 | 241 | 106 | 4.8716
25 | 20.098 | 9.8521 | 9.8523 | 30.4485 | 30.4488 | 6.5228 | 2.1562 | 12 | 12 | 0 | 107 | 241 | 107 | 4.9648
25 | 10.513 | 9.9498 | 9.9502 | 30.338 | 30.3378 | 6.5135 | 2.2438 | 13 | 13 | 0 | 108 | 241 | 108 | 5.1599
25 | 5.489 | 9.9589 | 9.9588 | 30.2944 | 30.2953 | 6.514 | 2.2637 | 14 | 14 | 0 | 109 | 241 | 109 | 5.2342
25 | 1.289 | 9.9611 | 9.961 | 30.2834 | 30.2851 | 6.5141 | 2.2714 | 15 | 15 | 0 | 110 | 211 | 110 | 5.2801
30 | 200.611 | 8.5016 | 8.5003 | 31.8263 | 31.8287 | 6.662 | 1.6438 | 1 | 1 | 0 | 111 | 66 | 111 | 3.6272
30 | 200.613 | 8.5109 | 8.5187 | 31.8098 | 31.7965 | 6.6613 | 1.645 | 2 | 2 | 0 | 112 | 241 | 112 | 3.5919
30 | 175.405 | 8.6064 | 8.6067 | 31.6419 | 31.642 | 6.6543 | 1.6862 | 3 | 3 | 0 | 113 | 241 | 113 | 3.6768
30 | 150.409 | 8.6481 | 8.6487 | 31.557 | 31.5565 | 6.6517 | 1.7065 | 4 | 4 | 0 | 114 | 241 | 114 | 3.7084
30 | 124.528 | 8.9496 | 8.9533 | 30.9183 | 30.9092 | 6.6344 | 1.8297 | 5 | 5 | 0 | 115 | 241 | 115 | 4.1124
30 | 100.794 | 9.1024 | 9.1025 | 30.5776 | 30.5783 | 6.6264 | 1.8939 | 6 | 6 | 0 | 116 | 241 | 116 | 4.282
30 | 75.947 | 9.2047 | 9.2055 | 30.3475 | 30.3466 | 6.6212 | 1.9421 | 7 | 7 | 0 | 117 | 241 | 117 | 4.4007
30 | 50.6 | 9.438 | 9.4379 | 29.8377 | 29.8381 | 6.6086 | 2.0487 | 8 | 8 | 0 | 118 | 241 | 118 | 4.7111
30 | 40.444 | 9.5171 | 9.518 | 29.6911 | 29.6903 | 6.6033 | 2.0899 | 9 | 9 | 0 | 119 | 241 | 119 | 4.8098
30 | 30.966 | 9.6749 | 9.6743 | 29.39 | 29.3916 | 6.593 | 2.1639 | 10 | 10 | 0 | 120 | 241 | 120 | 5.021
30 | 20.065 | 9.8129 | 9.8128 | 29.1852 | 29.1858 | 6.5816 | 2.2147 | 11 | 11 | 0 | 121 | 241 | 121 | 5.1591
30 | 10.07 | 9.9106 | 9.9108 | 29.016 | 29.0159 | 6.5746 | 2.2553 | 12 | 12 | 0 | 122 | 241 | 122 | 5.3045
30 | 5.188 | 10.3765 | 10.3766 | 28.3726 | 28.3733 | 6.5345 | 2.4403 | 13 | 13 | 0 | 123 | 241 | 123 | 5.7546
30 | 1.736 | 10.4698 | 10.468 | 28.224 | 28.2278 | 6.5274 | 2.5398 | 14 | 14 | 0 | 124 | 241 | 124 | 6.0014
37 | 175.041 | 9.0139 | 9.0143 | 30.6512 | 30.6514 | 6.6364 | 1.8511 | 1 | 1 | 0 | 125 | 241 | 125 | 4.195
37 | 150.359 | 9.0426 | 9.0412 | 30.5743 | 30.58 | 6.6354 | 1.8683 | 2 | 2 | 0 | 126 | 241 | 126 | 4.2324
37 | 125.272 | 9.0916 | 9.0918 | 30.4329 | 30.4331 | 6.6342 | 1.8997 | 3 | 3 | 0 | 127 | 241 | 127 | 4.3206
37 | 100.239 | 9.1851 | 9.1835 | 30.2065 | 30.2076 | 6.6301 | 1.9426 | 4 | 4 | 0 | 128 | 241 | 128 | 4.4339
37 | 74.976 | 9.2315 | 9.2301 | 29.8618 | 29.8616 | 6.638 | 1.9967 | 5 | 5 | 0 | 129 | 241 | 129 | 4.5834
37 | 50.061 | 9.4709 | 9.4707 | 29.5464 | 29.5466 | 6.6162 | 2.0487 | 6 | 6 | 0 | 130 | 241 | 130 | 4.7181
37 | 40.002 | 9.5746 | 9.5749 | 29.3932 | 29.3927 | 6.6075 | 2.0706 | 7 | 7 | 0 | 131 | 241 | 131 | 4.7602
37 | 29.996 | 9.7056 | 9.7048 | 29.2467 | 29.2451 | 6.5946 | 2.1416 | 8 | 8 | 0 | 132 | 241 | 132 | 4.9442
37 | 19.914 | 9.7021 | 9.7015 | 28.8594 | 28.8589 | 6.6116 | 2.2606 | 9 | 9 | 0 | 133 | 241 | 133 | 5.2804
37 | 9.644 | 10.0092 | 10.0296 | 28.4139 | 28.3805 | 6.5857 | 2.3802 | 10 | 10 | 0 | 134 | 241 | 134 | 5.6383
37 | 4.637 | 11.2178 | 11.22 | 27.0612 | 27.063 | 6.4697 | 2.7826 | 11 | 11 | 0 | 135 | 241 | 135 | 6.428
37 | 1.418 | 13.5964 | 13.5287 | 23.8174 | 23.9353 | 6.2771 | 3.7162 | 12 | 12 | 0 | 136 | 241 | 136 | 9.1774
41 | 314.549 | 8.6573 | 8.6571 | 30.786 | 30.7862 | 6.6837 | 1.565 | 1 | 1 | 0 | 137 | 241 | 137 | 3.4492
41 | 301.004 | 8.6508 | 8.6509 | 30.7753 | 30.7759 | 6.6851 | 1.5845 | 2 | 2 | 0 | 138 | 241 | 138 | 3.434
41 | 249.652 | 8.6001 | 8.6001 | 30.7182 | 30.7187 | 6.6952 | 1.6771 | 3 | 3 | 0 | 139 | 241 | 139 | 3.6001
41 | 200.738 | 8.5192 | 8.5191 | 30.6156 | 30.616 | 6.7119 | 1.7254 | 4 | 4 | 0 | 140 | 241 | 140 | 3.8465
41 | 174.914 | 8.4427 | 8.4426 | 30.5354 | 30.536 | 6.7269 | 1.8134 | 5 | 5 | 0 | 141 | 241 | 141 | 4.0974
41 | 150.898 | 8.4225 | 8.4224 | 30.4344 | 30.4347 | 6.7344 | 1.8412 | 6 | 6 | 0 | 142 | 241 | 142 | 4.215
41 | 125.077 | 8.4756 | 8.4753 | 30.2813 | 30.281 | 6.733 | 1.9534 | 7 | 7 | 0 | 143 | 241 | 143 | 4.5655
41 | 100.809 | 8.2659 | 8.266 | 30.0399 | 30.0405 | 6.7756 | 2.0195 | 8 | 8 | 0 | 144 | 241 | 144 | 4.7605
41 | 75.276 | 8.1923 | 8.1923 | 29.8036 | 29.8043 | 6.7974 | 1.996 | 9 | 9 | 0 | 145 | 241 | 145 | 4.6919
41 | 50.831 | 8.3823 | 8.3817 | 29.5748 | 29.5757 | 6.7783 | 1.9932 | 10 | 10 | 0 | 146 | 241 | 146 | 4.6497
41 | 40.822 | 8.5285 | 8.5288 | 29.4522 | 29.4524 | 6.7613 | 2.0161 | 11 | 11 | 0 | 147 | 241 | 147 | 4.6992
41 | 30.769 | 8.8092 | 8.8095 | 29.2349 | 29.2355 | 6.7282 | 2.0863 | 12 | 12 | 0 | 148 | 241 | 148 | 4.8797
41 | 20.828 | 9.3373 | 9.3349 | 28.762 | 28.7654 | 6.6695 | 2.2022 | 13 | 13 | 0 | 149 | 241 | 149 | 5.1788
41 | 10.39 | 10.1778 | 10.1775 | 27.5045 | 27.5064 | 6.5996 | 2.3629 | 14 | 14 | 0 | 150 | 241 | 150 | 5.5171
41 | 5.49 | 11.0153 | 11.0173 | 25.2165 | 25.2101 | 6.5747 | 2.4869 | 15 | 15 | 0 | 151 | 241 | 151 | 5.8954
41 | 1.328 | 13.5112 | 13.5477 | 7.6875 | 7.4685 | 6.9506 | 2.852 | 16 | 16 | 0 | 152 | 241 | 152 | 7.4146
46 | 369.718 | 8.6746 | 8.6746 | 30.8281 | 30.8286 | 6.6793 | 1.5427 | 1 | 1 | 0 | 153 | 241 | 153 | 3.3409
46 | 350.203 | 8.6723 | 8.6724 | 30.827 | 30.8273 | 6.6796 | 1.5483 | 2 | 2 | 0 | 154 | 241 | 154 | 3.3344
46 | 300.16 | 8.6672 | 8.6673 | 30.7763 | 30.7767 | 6.6826 | 1.5963 | 3 | 3 | 0 | 155 | 241 | 155 | 3.4637
46 | 250.333 | 8.6361 | 8.6359 | 30.7116 | 30.7115 | 6.6901 | 1.614 | 4 | 4 | 0 | 156 | 241 | 156 | 3.4893
46 | 200.292 | 8.5302 | 8.5295 | 30.5869 | 30.5867 | 6.7115 | 1.7418 | 5 | 5 | 0 | 157 | 241 | 157 | 3.8703
46 | 174.884 | 8.592 | 8.5754 | 30.4834 | 30.4898 | 6.7066 | 1.8199 | 6 | 6 | 0 | 158 | 241 | 158 | 4.079
46 | 149.251 | 8.6884 | 8.6931 | 30.3266 | 30.3276 | 6.6989 | 1.9051 | 7 | 7 | 0 | 159 | 241 | 159 | 4.3626
46 | 125.564 | 8.9321 | 8.9327 | 30.2327 | 30.233 | 6.6665 | 1.9209 | 8 | 8 | 0 | 160 | 241 | 160 | 4.3806
46 | 100.321 | 8.8617 | 8.8603 | 29.9603 | 29.9598 | 6.6888 | 1.9694 | 9 | 9 | 0 | 161 | 241 | 161 | 4.5175
46 | 74.712 | 9.0141 | 9.0176 | 29.6896 | 29.6908 | 6.6777 | 1.9892 | 10 | 10 | 0 | 162 | 241 | 162 | 4.5578
46 | 49.104 | 9.0419 | 9.0394 | 29.3827 | 29.3844 | 6.6868 | 2.041 | 11 | 11 | 0 | 163 | 241 | 163 | 4.7114
46 | 40.197 | 9.1702 | 9.1684 | 29.3122 | 29.3135 | 6.6707 | 2.0758 | 12 | 12 | 0 | 164 | 241 | 164 | 4.813
46 | 30.326 | 9.3208 | 9.3202 | 29.1447 | 29.1456 | 6.6555 | 2.1889 | 13 | 13 | 0 | 165 | 241 | 165 | 5.1281
46 | 20.075 | 9.7648 | 9.7644 | 28.9651 | 28.9655 | 6.5979 | 2.3234 | 14 | 14 | 0 | 166 | 241 | 166 | 5.401
46 | 10.054 | 10.7695 | 10.7872 | 28.3348 | 28.3176 | 6.4803 | 2.3847 | 15 | 15 | 0 | 167 | 241 | 167 | 5.4813
46 | 5.039 | 11.7519 | 11.8362 | 27.163 | 27.1114 | 6.3917 | 2.5761 | 16 | 16 | 0 | 168 | 241 | 168 | 5.8962
46 | 1.444 | 13.091 | 13.0973 | 11.5121 | 11.4643 | 6.8497 | 2.9102 | 17 | 17 | 0 | 169 | 241 | 169 | 7.4463
49 | 342.43 | 8.7177 | 8.7184 | 30.8205 | 30.8572 | 6.6731 | 1.5568 | 1 | 1 | 0 | 170 | 241 | 170 | 3.3462
49 | 300.369 | 8.7711 | 8.7709 | 30.7811 | 30.8122 | 6.6669 | 1.5541 | 2 | 2 | 0 | 171 | 241 | 171 | 3.2986
49 | 200.624 | 8.5052 | 8.5008 | 30.6091 | 30.633 | 6.7143 | 1.7523 | 4 | 4 | 0 | 173 | 241 | 173 | 3.9114
49 | 174.907 | 8.446 | 8.4469 | 30.5402 | 30.567 | 6.7262 | 1.8143 | 5 | 5 | 0 | 174 | 241 | 174 | 4.0945
49 | 149.97 | 8.3729 | 8.3733 | 30.4063 | 30.4334 | 6.7432 | 1.8693 | 6 | 6 | 0 | 175 | 241 | 175 | 4.2845
49 | 125.246 | 8.2048 | 8.2034 | 30.222 | 30.2457 | 6.777 | 1.9519 | 7 | 7 | 0 | 176 | 241 | 176 | 4.5062
49 | 100.907 | 8.2036 | 8.204 | 30.0981 | 30.124 | 6.7826 | 1.9746 | 8 | 8 | 0 | 177 | 241 | 177 | 4.6043
49 | 75.301 | 8.047 | 8.0482 | 29.8616 | 29.8877 | 6.8173 | 1.9785 | 9 | 9 | 0 | 178 | 241 | 178 | 4.6524
49 | 49.777 | 8.2 | 8.201 | 29.6614 | 29.6857 | 6.8025 | 1.9665 | 10 | 10 | 0 | 179 | 241 | 179 | 4.5913
49 | 40.323 | 8.2871 | 8.2886 | 29.5165 | 29.5405 | 6.7955 | 1.9522 | 11 | 11 | 0 | 180 | 241 | 180 | 4.5285
49 | 30.335 | 8.5112 | 8.5104 | 29.3039 | 29.3293 | 6.7705 | 2.0004 | 12 | 12 | 0 | 181 | 241 | 181 | 4.5996
49 | 19.761 | 9.5475 | 9.54 | 28.7119 | 28.7423 | 6.6405 | 2.2743 | 13 | 13 | 0 | 182 | 241 | 182 | 5.189
49 | 10.696 | 11.7375 | 11.7309 | 27.0317 | 27.0612 | 6.3989 | 2.447 | 14 | 14 | 0 | 183 | 241 | 183 | 5.5604
49 | 5.143 | 13.4116 | 13.4144 | 24.2925 | 24.3103 | 6.2828 | 2.8946 | 15 | 15 | 0 | 184 | 241 | 184 | 6.3054
49 | 2.001 | 14.8123 | 14.8216 | 14.7042 | 14.7849 | 6.4738 | 3.7594 | 16 | 16 | 0 | 185 | 241 | 185 | 8.7579
53 | 273.03 | 8.9054 | 8.9001 | 30.8204 | 30.8151 | 6.6452 | 1.4627 | 1 | 1 | 0 | 186 | 241 | 186 | 3.0372
53 | 250.404 | 8.8295 | 8.8292 | 30.757 | 30.7575 | 6.6592 | 1.5139 | 2 | 2 | 0 | 187 | 241 | 187 | 3.1336
53 | 200.576 | 8.5187 | 8.5186 | 30.5545 | 30.5552 | 6.7146 | 1.7086 | 3 | 3 | 0 | 188 | 241 | 188 | 3.6993
53 | 175.396 | 8.3671 | 8.3673 | 30.4282 | 30.4286 | 6.7431 | 1.8303 | 4 | 4 | 0 | 189 | 241 | 189 | 4.1821
53 | 150.577 | 8.3248 | 8.3248 | 30.3425 | 30.3427 | 6.7533 | 1.8934 | 5 | 5 | 0 | 190 | 241 | 190 | 4.3643
53 | 124.726 | 8.2081 | 8.2087 | 30.2231 | 30.2231 | 6.7764 | 1.9456 | 6 | 6 | 0 | 191 | 241 | 191 | 4.4868
53 | 100.845 | 8.1816 | 8.1813 | 30.1028 | 30.1065 | 6.7858 | 1.9707 | 7 | 7 | 0 | 192 | 241 | 192 | 4.606
53 | 75.569 | 8.1789 | 8.1789 | 29.8437 | 29.8452 | 6.7977 | 2.0001 | 8 | 8 | 0 | 193 | 241 | 193 | 4.7005
53 | 50.864 | 8.9384 | 8.9453 | 29.5386 | 29.5365 | 6.6955 | 1.9359 | 9 | 9 | 0 | 194 | 241 | 194 | 4.3748
53 | 40.602 | 9.2407 | 9.2439 | 29.3596 | 29.3592 | 6.6582 | 2.0162 | 10 | 10 | 0 | 195 | 241 | 195 | 4.6127
53 | 30.803 | 9.5805 | 9.5804 | 29.1733 | 29.1732 | 6.616 | 2.0906 | 11 | 11 | 0 | 196 | 241 | 196 | 4.8093
53 | 20.157 | 10.6885 | 10.6881 | 28.5346 | 28.5362 | 6.4835 | 2.2467 | 12 | 12 | 0 | 197 | 241 | 197 | 5.0191
53 | 10.793 | 11.9409 | 11.9324 | 26.204 | 26.2134 | 6.4044 | 2.3846 | 13 | 13 | 0 | 198 | 241 | 198 | 5.4337
53 | 5.506 | 15.2884 | 15.3116 | 18.7332 | 18.6744 | 6.2537 | 3.4503 | 14 | 14 | 0 | 199 | 241 | 199 | 7.8118
53 | 2.715 | 15.8103 | 15.7907 | 17.3848 | 17.4185 | 6.2383 | 3.5829 | 15 | 15 | 0 | 200 | 241 | 200 | 8.7566
57 | 315.07 | 9.044 | 9.0439 | 30.8674 | 30.8679 | 6.6227 | 1.3649 | 1 | 1 | 0 | 201 | 241 | 201 | 2.7603 | dups
57 | 300.351 | 9.0666 | 9.0664 | 30.8621 | 30.8625 | 6.6195 | 1.3754 | 2 | 2 | 0 | 202 | 241 | 202 | 2.684 | 2.726 | 2.643
57 | 250.587 | 8.8878 | 8.8874 | 30.7752 | 30.7755 | 6.6497 | 1.4742 | 3 | 3 | 0 | 203 | 241 | 203 | 3.0006
57 | 200.456 | 8.6083 | 8.6083 | 30.6226 | 30.6234 | 6.6981 | 1.657 | 4 | 4 | 0 | 204 | 241 | 204 | 3.5601
57 | 175.486 | 8.3963 | 8.3955 | 30.4623 | 30.4619 | 6.7372 | 1.8086 | 5 | 5 | 0 | 205 | 241 | 205 | 4.0836
57 | 150.516 | 8.2745 | 8.275 | 30.3398 | 30.3403 | 6.7611 | 1.8911 | 6 | 6 | 0 | 206 | 241 | 206 | 4.3832
57 | 125.67 | 8.2253 | 8.2252 | 30.1981 | 30.1981 | 6.7749 | 1.9471 | 7 | 7 | 0 | 207 | 241 | 207 | 4.5423
57 | 100.553 | 8.1646 | 8.1645 | 30.0599 | 30.0595 | 6.7903 | 1.9679 | 8 | 8 | 0 | 208 | 241 | 208 | 4.6079
57 | 75.015 | 8.2186 | 8.2141 | 29.7905 | 29.7938 | 6.7939 | 1.9291 | 9 | 9 | 0 | 209 | 241 | 209 | 4.5003
57 | 50.148 | 8.8639 | 8.8696 | 29.5659 | 29.5641 | 6.7055 | 1.9362 | 10 | 10 | 0 | 210 | 241 | 210 | 4.386
57 | 40.894 | 9.1217 | 9.123 | 29.4429 | 29.4434 | 6.6723 | 1.9699 | 11 | 11 | 0 | 211 | 241 | 211 | 4.4908
57 | 30.72 | 9.5349 | 9.5337 | 29.1913 | 29.1931 | 6.6219 | 2.0499 | 12 | 12 | 0 | 212 | 241 | 212 | 4.6816
57 | 20.47 | 11.0967 | 11.0873 | 28.2481 | 28.2538 | 6.438 | 2.2286 | 13 | 13 | 0 | 213 | 241 | 213 | 5.105
57 | 10.569 | 13.8554 | 13.8552 | 26.2497 | 26.2502 | 6.1498 | 3.124 | 14 | 14 | 0 | 214 | 241 | 214 | 7.3222
57 | 5.254 | 14.0558 | 13.9552 | 22.5958 | 22.6038 | 6.2645 | 2.9214 | 15 | 15 | 0 | 215 | 241 | 215 | 7.1798
57 | 2.537 | 15.3286 | 15.3231 | 16.4833 | 16.6714 | 6.3351 | 3.5099 | 16 | 16 | 0 | 216 | 241 | 216 | 8.6925
58 | 188.513 | 8.4518 | 8.4506 | 30.5117 | 30.5116 | 6.7266 | 1.7402 | 1 | 1 | 0 | 217 | 241 | 217 | 3.9785
58 | 175.117 | 8.3963 | 8.3964 | 30.4654 | 30.466 | 6.737 | 1.7933 | 2 | 2 | 0 | 218 | 241 | 218 | 4.0499
58 | 150.282 | 8.2942 | 8.2945 | 30.3558 | 30.3563 | 6.7574 | 1.8848 | 3 | 3 | 0 | 219 | 241 | 219 | 4.5636
58 | 124.712 | 8.2393 | 8.2396 | 30.2972 | 30.2977 | 6.7684 | 1.8994 | 4 | 4 | 0 | 220 | 241 | 220 | 4.396
58 | 99.928 | 8.1448 | 8.1452 | 29.9953 | 29.9953 | 6.7962 | 1.962 | 5 | 5 | 0 | 221 | 241 | 221 | 4.5687
58 | 75.062 | 8.5045 | 8.5047 | 29.7719 | 29.7723 | 6.7509 | 1.8514 | 6 | 6 | 0 | 222 | 241 | 222 | 4.2063
58 | 50.293 | 9.1962 | 9.1973 | 29.3941 | 29.3922 | 6.6633 | 1.9889 | 7 | 7 | 0 | 223 | 241 | 223 | 4.423
58 | 40.04 | 9.7916 | 9.7922 | 29.1132 | 29.1092 | 6.5877 | 2.0565 | 8 | 8 | 0 | 224 | 241 | 224 | 4.7196
58 | 30.081 | 9.9606 | 9.9633 | 28.9794 | 28.9773 | 6.5689 | 2.0719 | 9 | 9 | 0 | 225 | 241 | 225 | 4.7325
58 | 20.033 | 10.4568 | 10.4458 | 28.5731 | 28.5816 | 6.5147 | 2.1455 | 10 | 10 | 0 | 226 | 241 | 226 | 4.8157
58 | 4.879 | 14.2394 | 14.2536 | 25.6331 | 25.647 | 6.1245 | 3.1717 | 12 | 12 | 0 | 228 | 241 | 228 | 7.015
58 | 1.309 | 14.9341 | 14.9719 | 19.9841 | 19.3653 | 6.2513 | 3.4929 | 13 | 13 | 0 | 229 | 241 | 229 | 8.2272
59 | 164.526 | 8.6187 | 8.6201 | 30.5959 | 30.5971 | 6.6977 | 1.4607 | 1 | 1 | 0 | 230 | 241 | 230 | 3.0143
59 | 149.883 | 8.4029 | 8.4052 | 30.4418 | 30.4449 | 6.7371 | 1.5688 | 2 | 2 | 0 | 231 | 241 | 231 | 3.2135
59 | 125.215 | 8.2371 | 8.2363 | 30.3076 | 30.3067 | 6.7683 | 1.8041 | 3 | 3 | 0 | 232 | 241 | 232 | 3.9941
59 | 100.108 | 8.1795 | 8.1801 | 30.194 | 30.1958 | 6.7821 | 1.877 | 4 | 4 | 0 | 233 | 241 | 233 | 4.3032
59 | 75.204 | 8.247 | 8.2475 | 29.9447 | 29.9451 | 6.7827 | 1.8637 | 5 | 5 | 0 | 234 | 241 | 234 | 4.2409
59 | 49.956 | 9.0524 | 9.052 | 29.4484 | 29.4486 | 6.6824 | 1.9398 | 6 | 6 | 0 | 235 | 241 | 235 | 4.3428
59 | 40.21 | 9.5463 | 9.5205 | 29.1797 | 29.1897 | 6.6208 | 1.99 | 7 | 7 | 0 | 236 | 241 | 236 | 4.5022
59 | 29.807 | 10.0571 | 10.0615 | 28.844 | 28.8413 | 6.5606 | 2.0943 | 8 | 8 | 0 | 237 | 241 | 237 | 4.7479
59 | 20.067 | 10.9076 | 10.9069 | 28.3184 | 28.3234 | 6.4616 | 2.2239 | 9 | 9 | 0 | 238 | 241 | 238 | 5.0285
59 | 9.861 | 12.9235 | 12.9232 | 26.9475 | 26.9482 | 6.2432 | 2.8436 | 10 | 10 | 0 | 239 | 241 | 239 | 6.2387
59 | 5.101 | 14.4738 | 14.4741 | 24.6475 | 24.6505 | 6.1323 | 3.1279 | 11 | 11 | 0 | 240 | 241 | 240 | 7.1304
59 | 1.349 | 15.4318 | 15.415 | 21.4345 | 21.4765 | 6.1331 | 3.4852 | 12 | 12 | 0 | 241 | 241 | 241 | 7.9144
62 | 349.956 | 8.7967 | 8.7969 | 30.609 | 30.6097 | 6.6705 | 1.431 | 1 | 1 | 0 | 242 | 241 | 242 | 3.028
62 | 300.672 | 8.589 | 8.5887 | 30.5561 | 30.5572 | 6.7039 | 1.5972 | 2 | 2 | 0 | 243 | 241 | 243 | 3.4944
62 | 250.575 | 8.4925 | 8.4915 | 30.5105 | 30.5106 | 6.7205 | 1.6563 | 3 | 3 | 0 | 244 | 241 | 244 | 3.6657
62 | 175.175 | 8.3425 | 8.3429 | 30.4034 | 30.4043 | 6.7479 | 1.7748 | 5 | 5 | 0 | 246 | 241 | 246 | 4.0376
62 | 150.917 | 8.2405 | 8.2405 | 30.2646 | 30.2948 | 6.7696 | 1.8728 | 6 | 6 | 0 | 247 | 241 | 247 | 4.3244
62 | 125.832 | 8.181 | 8.1814 | 30.171 | 30.1967 | 6.7829 | 1.8891 | 7 | 7 | 0 | 248 | 241 | 248 | 4.379
62 | 101.311 | 8.1606 | 8.1623 | 30.0245 | 30.0245 | 6.7925 | 1.8949 | 8 | 8 | 0 | 249 | 241 | 249 | 4.3861
62 | 75.762 | 8.2552 | 8.2561 | 29.8889 | 29.8889 | 6.7839 | 1.8565 | 9 | 9 | 0 | 250 | 241 | 250 | 4.228
62 | 50.809 | 8.7459 | 8.7477 | 29.5744 | 29.5746 | 6.7229 | 1.9308 | 10 | 10 | 0 | 251 | 241 | 251 | 4.4278
62 | 40.801 | 9.0667 | 9.0674 | 29.4537 | 29.4592 | 6.68 | 2.0191 | 11 | 11 | 0 | 252 | 241 | 252 | 4.6623
62 | 30.721 | 9.3036 | 9.3051 | 29.2773 | 29.2814 | 6.6524 | 2.0497 | 12 | 12 | 0 | 253 | 241 | 253 | 4.7414
62 | 20.506 | 9.9242 | 9.9289 | 28.8687 | 28.8739 | 6.5788 | 2.1151 | 13 | 13 | 0 | 254 | 241 | 254 | 4.8205
62 | 10.277 | 12.6543 | 12.6489 | 26.9616 | 26.9771 | 6.2782 | 2.4983 | 14 | 14 | 0 | 255 | 241 | 255 | 5.6378
62 | 5.5 | 16.7816 | 16.7775 | 19.8642 | 19.8758 | 6.0255 | 3.8645 | 15 | 15 | 0 | 256 | 241 | 256 | 8.2715
62 | 1.713 | 17.1707 | 17.1724 | 18.8475 | 18.8445 | 6.0157 | 4.2912 | 16 | 16 | 0 | 257 | 241 | 257 | 9.9855
63 | 9.757 | 12.9494 | 12.949 | 26.7171 | 26.7173 | 6.2488 | 2.657 | 2 | 2 | 0 | 259 | 241 | 259 | 5.7922
63 | 1.24 | 15.683 | 15.6847 | 24.0168 | 24.0108 | 6.0063 | 3.7618 | 3 | 3 | 0 | 260 | 241 | 260 | 8.6365

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers ( diff from fit = slope *Oxy Volt + Offset - Phi/DO)
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers