## comparison-0997
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
7 | 30.22 | 8.356 | 8.3608 | 32.77 | 32.766 | 6.6429 | 1.4433 | 10 | 10 | 33 | 6.7581 | 0.44056 | 2.9773 |  | 8.8117e-05 | 0.4378 | -0.1914 |  |  |  | Summary of Soc Voffset including the original values in the factory calibration
51 | 10.208 | 10.9504 | 10.9472 | 27.7424 | 27.7436 | 6.4792 | 1.849 | 14 | 14 | 178 | 6.6038 | 0.61799 | 4.0811 |  | 0.00010483 | [7 empty] | m | b | Soc | Voffset | R2
51 | 149.41 | 8.9349 | 8.936 | 30.6531 | 30.6541 | 6.648 | 1.5439 | 6 | 6 | 170 | 6.8736 | 0.48441 | 3.3296 |  | 0.00011681 | [5 empty] | Bottles used | Original | 0.4348 | -0.21101 | 0.4348 | -0.4853 |  | 9.1691 |  | 2.956
55 | 39.971 | 9.4426 | 9.4429 | 29.1606 | 29.1616 | 6.6368 | 1.7669 | 10 | 10 | 190 | 6.7727 | 0.58198 | 3.9416 |  | 0.00015429 | [5 empty] | 258 | all unflagged bottles | 0.4378 | -0.1914 | 0.4378 | -0.43719 | 0.9965 | 9.3588 |  | 3.1175
70 | 29.066 | 9.989 | 9.9913 | 29.0548 | 29.0545 | 6.5616 | 1.8486 | 7 | 7 | 274 | 6.6928 | 0.61807 | 4.1366 |  | 0.00015932 | [5 empty] | 256 | excl. outliers diff>0.05 | 0.441 | -0.1962 | 0.441 | -0.4449 | 0.9977 | 9.4068 |  | 3.1175
39 | 172.975 | 9.1688 | 9.1695 | 31.4228 | 31.4237 | 6.5807 | 1.4505 | 1 | 1 | 121 | 6.8271 | 0.44383 | 3.0301 |  | 0.00020347 | [5 empty] | 253 | excl. outliers diff>0.04 | 0.4418 | -0.1975 | 0.4418 | -0.44703 | 0.998 | 9.4182 |  | 3.1168
48 | 124.611 | 9.0798 | 9.0795 | 30.3717 | 30.373 | 6.6386 | 1.5936 | 8 | 8 | 155 | 6.8438 | 0.50607 | 3.4634 |  | 0.00020739 | [5 empty] | 249 | excl. outliers diff>0.03 | 0.4415 | -0.1973 | 0.4415 | -0.44689 | 0.9984 | 9.4122 |  | 3.1152
43 | 49.92 | 9.6835 | 9.6824 | 29.4774 | 29.4782 | 6.5881 | 1.6702 | 10 | 10 | 142 | 6.7342 | 0.53961 | 3.6338 |  | 0.00022438 | [5 empty] | 241 | excl. outliers diff>0.02 | 0.4417 | -0.1979 | 0.4417 | -0.44804 | 0.9989 | 9.4134 |  | 3.1132
1 | 19.644 | 11.7827 | 11.7868 | 28.5391 | 28.5379 | 6.3322 | 1.6508 | 6 | 6 | 6 | 6.4708 | 0.53157 | 3.4396 | 6 | 0.00024501 | [5 empty] | 214 | excl. outliers diff>0.01 | 0.4406 | -0.1963 | 0.4406 | -0.44553 | 0.9992 | 9.3966 |  | 3.1128
63 | 39.712 | 9.6856 | 9.6848 | 28.9314 | 28.9318 | 6.6109 | 1.7904 | 7 | 7 | 231 | 6.7488 | 0.59219 | 3.9966 |  | 0.00024718
19 | 11.111 | 9.6891 | 9.6903 | 31.246 | 31.2469 | 6.5128 | 1.6388 | 7 | 7 | 83 | 6.6246 | 0.52581 | 3.4832 |  | 0.00026537 | [6 empty] | 0.015
13 | 52.477 | 8.4739 | 8.4739 | 32.6047 | 32.6054 | 6.6325 | 1.4382 | 7 | 7 | 58 | 6.7681 | 0.43796 | 2.9642 |  | 0.00027275
39 | 9.918 | 14.8044 | 14.8305 | 24.8454 | 24.8175 | 6.0835 | 2.3536 | 10 | 10 | 130 | 6.2404 | 0.8393 | 5.2376 |  | 0.00028981 | [5 empty] | 19 |  |  |  | V+Voffset | Soc *(V+Voffset) |  | DO/AA | Phi
19 | 30.108 | 9.1176 | 9.1223 | 31.7841 | 31.7797 | 6.5728 | 1.5149 | 5 | 5 | 81 | 6.6952 | 0.47151 | 3.1569 |  | 0.00030905 | [8 empty] | old | 1.1814 | 0.51367 |  | 7.0726
7 | 48.117 | 7.9339 | 7.9339 | 33.299 | 33.3009 | 6.6833 | 1.4448 | 8 | 8 | 31 | 6.8099 | 0.44083 | 3.002 |  | 0.00032437 | [8 empty] | new | 1.1447 | 0.54121 |  | 6.7127
7 | 124.195 | 7.217 | 7.2169 | 33.7433 | 33.7442 | 6.7734 | 1.2298 | 5 | 5 | 28 | 6.9611 | 0.34737 | 2.4181 |  | 0.00035063
48 | 199.644 | 8.8199 | 8.8203 | 30.8321 | 30.8329 | 6.6574 | 1.4795 | 5 | 5 | 152 | 6.9263 | 0.45597 | 3.1582 |  | 0.00036239
43 | 251.714 | 8.9271 | 8.9276 | 30.9606 | 30.9614 | 6.6359 | 1.445 | 3 | 3 | 135 | 6.9512 | 0.4416 | 3.0697 |  | 0.00037064
55 | 199.67 | 8.8236 | 8.8253 | 30.803 | 30.8038 | 6.6581 | 1.4934 | 3 | 3 | 183 | 6.9271 | 0.46283 | 3.2061 |  | 0.00041223
19 | 4.916 | 10.4921 | 10.4941 | 30.8267 | 30.8268 | 6.4168 | 1.8887 | 8 | 8 | 84 | 6.5306 | 0.635 | 4.147 |  | 0.00045135
26 | 39.674 | 9.8018 | 9.8031 | 30.9892 | 30.9896 | 6.5074 | 1.6327 | 6 | 6 | 101 | 6.6444 | 0.52292 | 3.4745 |  | 0.00048265
31 | 50.048 | 10.447 | 10.4455 | 29.6331 | 29.6372 | 6.4722 | 1.7281 | 8 | 8 | 114 | 6.6243 | 0.56567 | 3.7472 |  | 0.00050278
10 | 127.01 | 7.0643 | 7.0649 | 33.8429 | 33.8435 | 6.7926 | 1.2005 | 5 | 5 | 42 | 6.9817 | 0.33367 | 2.3296 |  | 0.00050995
48 | 40.012 | 9.4532 | 9.4547 | 29.0448 | 29.0447 | 6.6402 | 1.7605 | 12 | 12 | 159 | 6.7763 | 0.57999 | 3.9302 |  | 0.00066369
43 | 19.743 | 10.7306 | 10.7458 | 28.1879 | 28.1833 | 6.4918 | 1.8289 | 12 | 12 | 144 | 6.6221 | 0.60999 | 4.0395 | 6 | 0.00071557 | [5 empty] | Summary of Soc Voffset including the original values in the factory calibration
48 | 174.43 | 8.8532 | 8.8537 | 30.768 | 30.7689 | 6.6552 | 1.5007 | 6 | 6 | 153 | 6.9022 | 0.46487 | 3.2086 |  | 0.00071678 | [7 empty] | m | b | Soc | Voffset | R2
55 | 73.771 | 8.9226 | 8.9231 | 29.888 | 29.8882 | 6.6828 | 1.6809 | 8 | 8 | 188 | 6.8431 | 0.54525 | 3.7312 |  | 0.00074014
70 | 124.847 | 8.5776 | 8.578 | 30.4889 | 30.4897 | 6.7085 | 1.5751 | 2 | 2 | 269 | 6.9105 | 0.49743 | 3.4375 |  | 0.00075133
48 | 99.851 | 9.0463 | 9.0466 | 30.1634 | 30.1639 | 6.6525 | 1.6336 | 9 | 9 | 156 | 6.8362 | 0.52453 | 3.5858 |  | 0.00076027
55 | 20.12 | 10.2651 | 10.2678 | 28.4652 | 28.4645 | 6.5466 | 1.8308 | 12 | 12 | 192 | 6.6731 | 0.60935 | 4.0662 | 6 | 0.00076984
48 | 150.303 | 8.9621 | 8.9625 | 30.6335 | 30.6341 | 6.6448 | 1.5388 | 7 | 7 | 154 | 6.8714 | 0.48147 | 3.3083 | 6 | 0.00080633
55 | 124.103 | 8.7834 | 8.7816 | 30.402 | 30.4022 | 6.6814 | 1.6506 | 6 | 6 | 186 | 6.8842 | 0.53209 | 3.663 |  | 0.00083216
51 | 100.165 | 9.2273 | 9.2276 | 30.2986 | 30.2994 | 6.6199 | 1.587 | 8 | 8 | 172 | 6.805 | 0.50425 | 3.4315 |  | 0.00085222
26 | 99.808 | 8.13 | 8.1304 | 32.7029 | 32.7036 | 6.6797 | 1.278 | 3 | 3 | 98 | 6.8537 | 0.36715 | 2.5163 |  | 0.00094344
31 | 146.371 | 9.1396 | 9.1402 | 31.4611 | 31.4618 | 6.5833 | 1.4524 | 1 | 1 | 107 | 6.8064 | 0.44541 | 3.0316 |  | 0.00096671
51 | 174.815 | 8.9079 | 8.9086 | 30.7841 | 30.7846 | 6.6464 | 1.4989 | 5 | 5 | 169 | 6.8939 | 0.46583 | 3.2114 |  | 0.0010004
70 | 74.475 | 8.7702 | 8.7707 | 29.9072 | 29.9078 | 6.7048 | 1.6756 | 4 | 4 | 271 | 6.8645 | 0.54117 | 3.7148 |  | 0.0010163
39 | 29.882 | 10.2584 | 10.2559 | 29.2472 | 29.2463 | 6.5149 | 1.6959 | 8 | 8 | 128 | 6.6489 | 0.55001 | 3.657 |  | 0.0010379
58 | 40.522 | 9.2088 | 9.2099 | 29.2819 | 29.2826 | 6.6662 | 1.732 | 11 | 11 | 206 | 6.8005 | 0.56579 | 3.8477 | 6 | 0.001055
70 | 100.123 | 8.5525 | 8.5526 | 30.2439 | 30.2432 | 6.723 | 1.6224 | 3 | 3 | 270 | 6.9032 | 0.51778 | 3.5744 |  | 0.0010874
7 | 39.245 | 7.9689 | 7.9715 | 33.1273 | 33.124 | 6.6855 | 1.4065 | 9 | 9 | 32 | 6.8048 | 0.42553 | 2.8957 |  | 0.0011425
23 | 2.128 | 10.9752 | 10.9752 | 30.204 | 30.2053 | 6.3753 | 2.015 | 10 | 10 | 95 | 6.4914 | 0.6896 | 4.4765 |  | 0.0011714
43 | 302.588 | 8.9279 | 8.929 | 31.0445 | 31.0457 | 6.6322 | 1.4171 | 2 | 2 | 134 | 6.9926 | 0.42784 | 2.9917 |  | 0.0011888
68 | 9.833 | 11.7254 | 11.7256 | 27.9063 | 27.9075 | 6.3653 | 2.1368 | 13 | 13 | 265 | 6.4959 | 0.74286 | 4.8256 |  | 0.0012055
55 | 49.566 | 9.2371 | 9.2375 | 29.4387 | 29.4397 | 6.6553 | 1.729 | 9 | 9 | 189 | 6.7975 | 0.56684 | 3.8531 |  | 0.0012984
55 | 174.973 | 8.7697 | 8.77 | 30.6976 | 30.6985 | 6.6707 | 1.5508 | 4 | 4 | 184 | 6.9178 | 0.48885 | 3.3817 | 6 | 0.0013116
51 | 200.21 | 8.8132 | 8.8136 | 30.842 | 30.8421 | 6.6579 | 1.4732 | 4 | 4 | 168 | 6.9273 | 0.45226 | 3.1329 | 6 | 0.0013195
51 | 39.629 | 9.3258 | 9.3255 | 29.2093 | 29.2106 | 6.652 | 1.7516 | 11 | 11 | 175 | 6.7865 | 0.5768 | 3.9145 |  | 0.0013533
60 | 188.384 | 8.7872 | 8.7876 | 30.7999 | 30.801 | 6.6636 | 1.4348 | 1 | 1 | 212 | 6.9225 | 0.43538 | 3.0139 |  | 0.001386
68 | 39.678 | 10.054 | 10.0537 | 28.9572 | 28.9583 | 6.5563 | 1.8631 | 10 | 10 | 262 | 6.6972 | 0.62278 | 4.1709 | 6 | 0.0015032
10 | 102.106 | 7.3492 | 7.3488 | 33.7165 | 33.7176 | 6.7541 | 1.305 | 6 | 6 | 43 | 6.9232 | 0.37841 | 2.6198 |  | 0.0015349
55 | 4.929 | 16.9705 | 17.0113 | 25.2125 | 25.2122 | 5.8118 | 2.869 | 14 | 14 | 194 | 5.9789 | 1.0662 | 6.3746 |  | 0.0015386
51 | 346.355 | 8.9195 | 8.9201 | 31.0745 | 31.0753 | 6.6322 | 1.4294 | 1 | 1 | 165 | 7.0316 | 0.43594 | 3.0654 |  | 0.001567
10 | 177.187 | 6.8199 | 6.8204 | 33.9196 | 33.9201 | 6.8274 | 1.0977 | 3 | 3 | 40 | 7.0601 | 0.2876 | 2.0305 | 6 | 0.0015752
58 | 20.723 | 10.3656 | 10.3646 | 28.3303 | 28.3297 | 6.5378 | 1.8131 | 13 | 13 | 208 | 6.6658 | 0.60395 | 4.0258 |  | 0.0015769
43 | 175.631 | 8.8708 | 8.8707 | 30.7532 | 30.7539 | 6.6532 | 1.5102 | 5 | 5 | 137 | 6.9014 | 0.46818 | 3.2311 | 6 | 0.0015803
51 | 300.086 | 8.9094 | 8.9099 | 31.0178 | 31.0185 | 6.6361 | 1.4446 | 2 | 2 | 166 | 6.9943 | 0.4427 | 3.0963 |  | 0.0016516
43 | 75.567 | 9.6155 | 9.6155 | 29.7962 | 29.7954 | 6.5845 | 1.6052 | 9 | 9 | 141 | 6.7518 | 0.50967 | 3.4412 |  | 0.0016901
16 | 29.957 | 9.2593 | 9.2603 | 31.6706 | 31.6709 | 6.557 | 1.5564 | 8 | 8 | 72 | 6.6806 | 0.4917 | 3.2849 |  | 0.0017129
66 | 122.893 | 8.8498 | 8.8504 | 30.3499 | 30.351 | 6.6737 | 1.6163 | 7 | 7 | 243 | 6.8759 | 0.51451 | 3.5377 |  | 0.0017191
43 | 101.115 | 9.7405 | 9.7408 | 30.1481 | 30.1493 | 6.5515 | 1.5419 | 8 | 8 | 140 | 6.7412 | 0.48191 | 3.2486 |  | 0.0017518
48 | 9.947 | 12.5761 | 12.6325 | 26.6196 | 26.6042 | 6.302 | 2.0474 | 15 | 15 | 162 | 6.4407 | 0.70319 | 4.529 |  | 0.0017843
39 | 124.506 | 9.4554 | 9.4556 | 31.0082 | 31.0096 | 6.5565 | 1.5125 | 3 | 3 | 123 | 6.7633 | 0.46896 | 3.1717 |  | 0.0017994
16 | 40.461 | 8.8222 | 8.822 | 32.1617 | 32.1625 | 6.6 | 1.4768 | 7 | 7 | 71 | 6.7285 | 0.45332 | 3.0501 |  | 0.0018332
7 | 150.051 | 6.9337 | 6.9345 | 33.8673 | 33.8678 | 6.8119 | 1.1203 | 4 | 4 | 27 | 7.0208 | 0.29718 | 2.0864 | 6 | 0.0019141
60 | 150.841 | 8.7367 | 8.7371 | 30.5768 | 30.5775 | 6.6808 | 1.5615 | 3 | 3 | 214 | 6.9066 | 0.49027 | 3.3861 |  | 0.0019799
43 | 150.567 | 8.9481 | 8.9586 | 30.6376 | 30.6285 | 6.6467 | 1.531 | 6 | 6 | 138 | 6.8734 | 0.47689 | 3.2779 |  | 0.0019975
10 | 77.085 | 7.5928 | 7.5956 | 33.5294 | 33.5286 | 6.725 | 1.3872 | 7 | 7 | 44 | 6.874 | 0.41392 | 2.8453 |  | 0.0020048
63 | 49.854 | 9.5734 | 9.5729 | 29.1423 | 29.1435 | 6.6184 | 1.7848 | 6 | 6 | 230 | 6.7639 | 0.5879 | 3.9765 |  | 0.002086
3 | 75.428 | 9.5414 | 9.5413 | 31.1202 | 31.1228 | 6.5394 | 1.5637 | 7 | 7 | 16 | 6.7046 | 0.49531 | 3.3209 |  | 0.0021125
39 | 102.815 | 9.715 | 9.7164 | 30.4745 | 30.4762 | 6.5415 | 1.5635 | 4 | 4 | 124 | 6.7321 | 0.49095 | 3.3051 |  | 0.0021431
26 | 49.894 | 9.6975 | 9.698 | 31.0847 | 31.0854 | 6.5184 | 1.6018 | 5 | 5 | 100 | 6.6631 | 0.50767 | 3.3827 |  | 0.0021779
70 | 10.018 | 10.4517 | 10.4536 | 28.773 | 28.7737 | 6.5071 | 1.9244 | 9 | 9 | 276 | 6.6264 | 0.65339 | 4.3296 |  | 0.0022953
66 | 40.374 | 9.3682 | 9.3715 | 29.1414 | 29.1404 | 6.6486 | 1.7469 | 11 | 11 | 247 | 6.7843 | 0.5711 | 3.8745 |  | 0.0023018
19 | 19.177 | 9.6244 | 9.6248 | 31.3008 | 31.3018 | 6.5198 | 1.6234 | 6 | 6 | 82 | 6.6377 | 0.51698 | 3.4316 |  | 0.0023225
55 | 273.211 | 8.8383 | 8.8388 | 30.9992 | 31.0003 | 6.6475 | 1.4299 | 1 | 1 | 181 | 6.9815 | 0.43692 | 3.0503 |  | 0.002333
23 | 51.716 | 8.7929 | 8.794 | 32.1047 | 32.1052 | 6.6067 | 1.4324 | 4 | 4 | 89 | 6.7447 | 0.43336 | 2.9229 |  | 0.0023344
3 | 151.857 | 8.7993 | 8.8 | 31.9901 | 31.9913 | 6.6107 | 1.4126 | 4 | 4 | 13 | 6.8357 | 0.42461 | 2.9025 | 6 | 0.0024168
58 | 314.138 | 8.8599 | 8.8607 | 31.0699 | 31.0706 | 6.6412 | 1.404 | 1 | 1 | 196 | 7.0117 | 0.42569 | 2.9848 |  | 0.0024194
60 | 100.837 | 8.8302 | 8.8291 | 30.2021 | 30.2024 | 6.683 | 1.6294 | 5 | 5 | 216 | 6.866 | 0.51954 | 3.5671 |  | 0.0024279
55 | 30.429 | 9.5906 | 9.59 | 29.0184 | 29.02 | 6.6211 | 1.7798 | 11 | 11 | 191 | 6.7502 | 0.59027 | 3.9844 |  | 0.0024607
48 | 349.674 | 8.8672 | 8.8677 | 31.1904 | 31.191 | 6.6349 | 1.3827 | 2 | 2 | 149 | 7.037 | 0.41644 | 2.9305 |  | 0.0024794
26 | 9.268 | 10.4011 | 10.4013 | 30.4376 | 30.4389 | 6.4454 | 1.8012 | 9 | 9 | 104 | 6.5624 | 0.59964 | 3.9351 |  | 0.0024808
19 | 2.745 | 10.5393 | 10.5452 | 30.7832 | 30.7822 | 6.412 | 1.8932 | 9 | 9 | 85 | 6.5245 | 0.63995 | 4.1753 |  | 0.0024988
3 | 203.081 | 8.5362 | 8.5363 | 32.3035 | 32.3052 | 6.6362 | 1.3568 | 2 | 2 | 11 | 6.9042 | 0.40009 | 2.7623 |  | 0.0025411
68 | 30.04 | 10.0916 | 10.0921 | 28.9421 | 28.943 | 6.5515 | 1.8721 | 11 | 11 | 263 | 6.6846 | 0.62565 | 4.1822 |  | 0.0025766
48 | 50.353 | 9.2819 | 9.2827 | 29.2896 | 29.2897 | 6.6551 | 1.7503 | 11 | 11 | 158 | 6.7985 | 0.57747 | 3.9259 |  | 0.0026065
55 | 149.899 | 8.7899 | 8.7918 | 30.5622 | 30.563 | 6.6735 | 1.5922 | 5 | 5 | 185 | 6.8988 | 0.50302 | 3.4703 |  | 0.0026277
31 | 1.864 | 16.0922 | 16.2503 | 21.979 | 21.8011 | 6.0315 | 2.6679 | 14 | 14 | 120 | 6.1939 | 0.97926 | 6.0655 |  | 0.0026362
10 | 202.889 | 6.7751 | 6.7762 | 33.9252 | 33.9259 | 6.8342 | 1.0829 | 2 | 2 | 39 | 7.09 | 0.28004 | 1.9855 |  | 0.0026509
66 | 150.148 | 8.7411 | 8.7412 | 30.4408 | 30.4417 | 6.6861 | 1.6176 | 6 | 6 | 242 | 6.9114 | 0.51413 | 3.5534 | 6 | 0.0026674
48 | 30.204 | 9.9264 | 9.9269 | 28.6638 | 28.6646 | 6.5871 | 1.7774 | 13 | 13 | 160 | 6.7191 | 0.58945 | 3.9606 |  | 0.0026983
51 | 49.992 | 9.114 | 9.1154 | 29.3744 | 29.3744 | 6.6764 | 1.7357 | 10 | 10 | 174 | 6.818 | 0.5712 | 3.8944 |  | 0.0027029
55 | 249.878 | 8.8206 | 8.8202 | 30.9275 | 30.9302 | 6.6532 | 1.4475 | 2 | 2 | 182 | 6.9665 | 0.43961 | 3.0626 |  | 0.0027233
13 | 175.84 | 6.9351 | 6.9353 | 33.8746 | 33.8755 | 6.8114 | 1.0768 | 2 | 2 | 53 | 7.0436 | 0.27725 | 1.9528 |  | 0.002777
70 | 49.763 | 9.1413 | 9.1434 | 29.621 | 29.6217 | 6.6617 | 1.729 | 5 | 5 | 272 | 6.8031 | 0.56274 | 3.8284 |  | 0.0028083
66 | 174.763 | 8.6223 | 8.6214 | 30.5295 | 30.5302 | 6.7001 | 1.5818 | 5 | 5 | 241 | 6.9464 | 0.49826 | 3.4611 |  | 0.0028522
55 | 99.944 | 8.8265 | 8.8268 | 30.2729 | 30.2736 | 6.6805 | 1.675 | 7 | 7 | 187 | 6.8626 | 0.54485 | 3.7391 |  | 0.0029177
7 | 75.097 | 7.4032 | 7.4047 | 33.6575 | 33.6575 | 6.7484 | 1.2872 | 7 | 7 | 30 | 6.894 | 0.37507 | 2.5858 |  | 0.0029347
66 | 354.123 | 8.7586 | 8.7592 | 30.6877 | 30.6884 | 6.6727 | 1.4363 | 1 | 1 | 237 | 7.0799 | 0.43443 | 3.0757 |  | 0.0030015
60 | 75.278 | 8.7686 | 8.7697 | 29.835 | 29.8341 | 6.7082 | 1.6571 | 6 | 6 | 217 | 6.8687 | 0.53098 | 3.6471 | 6 | 0.003097
39 | 5.611 | 17.8737 | 17.8677 | 21.964 | 21.9135 | 5.8226 | 2.9636 | 11 | 11 | 131 | 5.9991 | 1.1092 | 6.654 |  | 0.0031323
60 | 50.207 | 9.5076 | 9.507 | 29.1678 | 29.1687 | 6.627 | 1.7809 | 7 | 7 | 218 | 6.7722 | 0.58515 | 3.9628 |  | 0.0031373
70 | 39.528 | 9.6474 | 9.6517 | 29.2973 | 29.2953 | 6.601 | 1.7952 | 6 | 6 | 273 | 6.738 | 0.5914 | 3.9849 |  | 0.003143
51 | 75.204 | 8.972 | 8.9724 | 29.9023 | 29.9027 | 6.6748 | 1.6832 | 9 | 9 | 173 | 6.8367 | 0.54868 | 3.7512 |  | 0.0031771
1 | 49.865 | 9.9211 | 9.9208 | 29.8363 | 29.8378 | 6.5386 | 1.3971 | 3 | 3 | 3 | 6.6862 | 0.417 | 2.7881 |  | 0.0032348
39 | 151.327 | 9.2023 | 9.2028 | 31.374 | 31.3751 | 6.5778 | 1.4623 | 2 | 2 | 122 | 6.8057 | 0.44554 | 3.0322 |  | 0.0032389
26 | 74.775 | 8.9727 | 8.9733 | 31.8481 | 31.8485 | 6.5913 | 1.4491 | 4 | 4 | 99 | 6.7508 | 0.43973 | 2.9685 | 6 | 0.0032722
68 | 310.112 | 8.7381 | 8.7385 | 30.6643 | 30.665 | 6.6768 | 1.4786 | 1 | 1 | 253 | 7.0443 | 0.45259 | 3.1882 |  | 0.0033397
39 | 75.587 | 9.7677 | 9.7673 | 30.1087 | 30.1102 | 6.5492 | 1.5848 | 5 | 5 | 125 | 6.7174 | 0.49907 | 3.3525 |  | 0.0033429
26 | 4.72 | 10.4247 | 10.4245 | 30.4086 | 30.4105 | 6.4433 | 1.8044 | 10 | 10 | 105 | 6.5567 | 0.60192 | 3.9466 |  | 0.0033722
60 | 30.034 | 9.8989 | 9.8988 | 28.8978 | 28.8992 | 6.5813 | 1.8522 | 9 | 9 | 220 | 6.7127 | 0.61607 | 4.1355 |  | 0.0034085
60 | 20.176 | 10.0597 | 10.0724 | 28.4226 | 28.4195 | 6.578 | 1.7966 | 10 | 10 | 221 | 6.7028 | 0.5986 | 4.0123 |  | 0.0034352
3 | 39.909 | 11.0502 | 11.0506 | 29.6401 | 29.6417 | 6.3878 | 1.8946 | 9 | 9 | 18 | 6.5362 | 0.6415 | 4.193 |  | 0.0034403
3 | 100.778 | 9.3556 | 9.36 | 31.358 | 31.3554 | 6.5562 | 1.5275 | 6 | 6 | 15 | 6.7415 | 0.48079 | 3.2412 |  | 0.0034442
51 | 30.513 | 9.6681 | 9.6683 | 28.8308 | 28.8319 | 6.6177 | 1.7644 | 12 | 12 | 176 | 6.7477 | 0.58448 | 3.9439 |  | 0.003447
68 | 199.91 | 8.5651 | 8.5656 | 30.5052 | 30.5062 | 6.7097 | 1.5596 | 3 | 3 | 255 | 6.9782 | 0.48795 | 3.405 |  | 0.0034471
68 | 249.624 | 8.6823 | 8.6825 | 30.6078 | 30.6085 | 6.6876 | 1.5291 | 2 | 2 | 254 | 7.0008 | 0.47457 | 3.3223 |  | 0.0034567
58 | 74.146 | 8.9691 | 8.9689 | 30.0073 | 30.008 | 6.6707 | 1.6578 | 9 | 9 | 204 | 6.8315 | 0.53087 | 3.6266 |  | 0.0035271
16 | 50.28 | 7.8453 | 7.8437 | 33.0407 | 33.0433 | 6.708 | 1.2411 | 6 | 6 | 70 | 6.836 | 0.34844 | 2.3819 | 6 | 0.0035358
26 | 19.37 | 10.399 | 10.3993 | 30.4406 | 30.4417 | 6.4456 | 1.8036 | 8 | 8 | 103 | 6.571 | 0.60176 | 3.9542 |  | 0.0035479
68 | 99.606 | 8.6513 | 8.6508 | 29.9887 | 29.9897 | 6.7192 | 1.6675 | 7 | 7 | 259 | 6.9 | 0.53507 | 3.692 |  | 0.0035672
66 | 250.34 | 8.7173 | 8.7178 | 30.6314 | 30.632 | 6.6814 | 1.5218 | 3 | 3 | 239 | 6.9953 | 0.47127 | 3.2967 |  | 0.0035698
3 | 51.149 | 10.6374 | 10.6385 | 29.9979 | 29.9984 | 6.4304 | 1.8003 | 8 | 8 | 17 | 6.5847 | 0.60038 | 3.9533 |  | 0.0035907
68 | 125.05 | 8.5582 | 8.5581 | 30.2277 | 30.2283 | 6.7229 | 1.6401 | 6 | 6 | 258 | 6.9252 | 0.52302 | 3.622 | 6 | 0.0036051
58 | 123.893 | 8.8653 | 8.8659 | 30.4865 | 30.4872 | 6.6655 | 1.6044 | 7 | 7 | 202 | 6.8686 | 0.50735 | 3.4848 |  | 0.0036472
48 | 20.241 | 10.3219 | 10.3332 | 28.3726 | 28.3655 | 6.5423 | 1.8077 | 14 | 14 | 161 | 6.6694 | 0.60368 | 4.0262 | 6 | 0.0036633
10 | 31.847 | 9.6347 | 9.6333 | 31.4376 | 31.437 | 6.5126 | 1.6929 | 10 | 10 | 47 | 6.6412 | 0.55346 | 3.6756 | 6 | 0.0037055
7 | 199.54 | 6.7021 | 6.7027 | 33.9202 | 33.9207 | 6.846 | 1.0329 | 2 | 2 | 25 | 7.0983 | 0.25707 | 1.8248 |  | 0.0037337
39 | 40.215 | 10.1362 | 10.1342 | 29.6049 | 29.6075 | 6.5174 | 1.6531 | 7 | 7 | 127 | 6.6589 | 0.5286 | 3.5198 |  | 0.0037531
23 | 29.605 | 9.9445 | 9.9468 | 31.232 | 31.2304 | 6.477 | 1.7106 | 6 | 6 | 91 | 6.6065 | 0.56129 | 3.7082 |  | 0.0037873
68 | 150.113 | 8.4942 | 8.4945 | 30.3917 | 30.3923 | 6.7254 | 1.6065 | 5 | 5 | 257 | 6.9493 | 0.50812 | 3.5311 |  | 0.0038192
13 | 75.71 | 7.2472 | 7.2473 | 33.7397 | 33.7409 | 6.7688 | 1.2541 | 6 | 6 | 57 | 6.9137 | 0.36147 | 2.4991 |  | 0.0038272
66 | 49.812 | 9.2143 | 9.2144 | 29.4029 | 29.4037 | 6.6602 | 1.7322 | 10 | 10 | 246 | 6.8025 | 0.56313 | 3.8307 |  | 0.0038282
70 | 2.444 | 11.0443 | 11.0428 | 28.2071 | 28.2105 | 6.447 | 2.025 | 11 | 11 | 278 | 6.5655 | 0.69908 | 4.5898 |  | 0.0039561
51 | 250.907 | 8.8853 | 8.8853 | 30.9484 | 30.9491 | 6.6427 | 1.4617 | 3 | 3 | 167 | 6.9571 | 0.45248 | 3.148 |  | 0.00397
31 | 39.277 | 10.555 | 10.5548 | 29.4179 | 29.4203 | 6.4658 | 1.7509 | 9 | 9 | 115 | 6.61 | 0.5791 | 3.8279 |  | 0.0039724
68 | 19.955 | 10.4122 | 10.4126 | 28.7135 | 28.7106 | 6.5152 | 1.9226 | 12 | 12 | 264 | 6.6426 | 0.64628 | 4.293 |  | 0.0040361
23 | 123.257 | 7.4888 | 7.491 | 33.3628 | 33.3624 | 6.7483 | 1.175 | 1 | 1 | 86 | 6.9376 | 0.32709 | 2.2692 |  | 0.0040491
7 | 11.201 | 9.8276 | 9.8237 | 31.5369 | 31.5316 | 6.4809 | 1.7879 | 12 | 12 | 35 | 6.5937 | 0.59537 | 3.9257 |  | 0.0040491
31 | 124.822 | 9.2828 | 9.2856 | 31.2611 | 31.2588 | 6.5709 | 1.4791 | 5 | 5 | 111 | 6.7765 | 0.46029 | 3.1192 | 6 | 0.0041313
60 | 39.855 | 9.7677 | 9.7656 | 29.0212 | 29.0223 | 6.5951 | 1.8323 | 8 | 8 | 219 | 6.7337 | 0.60666 | 4.0851 |  | 0.0041362
26 | 2.199 | 10.437 | 10.4372 | 30.3915 | 30.3926 | 6.4423 | 1.8085 | 11 | 11 | 106 | 6.5537 | 0.60449 | 3.9617 |  | 0.0041364
31 | 100.124 | 9.9255 | 9.9268 | 30.3774 | 30.3769 | 6.5153 | 1.6043 | 6 | 6 | 112 | 6.7052 | 0.51522 | 3.4546 |  | 0.0042309
66 | 300.562 | 8.7466 | 8.7471 | 30.673 | 30.6737 | 6.6752 | 1.4849 | 2 | 2 | 238 | 7.0341 | 0.4544 | 3.1963 |  | 0.0042774
68 | 74.817 | 8.844 | 8.8433 | 29.7282 | 29.7289 | 6.7015 | 1.704 | 8 | 8 | 260 | 6.8622 | 0.55034 | 3.7766 |  | 0.0042897
13 | 186.599 | 6.9361 | 6.9364 | 33.8748 | 33.8757 | 6.8112 | 1.0765 | 1 | 1 | 52 | 7.0532 | 0.27557 | 1.9437 |  | 0.0043066
39 | 51.069 | 9.9759 | 9.9764 | 29.8712 | 29.8727 | 6.5292 | 1.6233 | 6 | 6 | 126 | 6.6783 | 0.52365 | 3.4971 | 6 | 0.0043544
3 | 30.238 | 11.315 | 11.3174 | 29.2856 | 29.284 | 6.3658 | 1.9823 | 10 | 10 | 19 | 6.5087 | 0.68082 | 4.4312 |  | 0.0043681
43 | 312.485 | 8.9349 | 8.9349 | 31.0599 | 31.0605 | 6.6305 | 1.4105 | 1 | 1 | 133 | 6.9997 | 0.43056 | 3.0138 |  | 0.0044356
16 | 5.267 | 11.9419 | 11.9291 | 30.4699 | 30.4736 | 6.2344 | 2.249 | 11 | 11 | 75 | 6.361 | 0.79782 | 5.0749 |  | 0.0046192
13 | 20.222 | 9.6863 | 9.6882 | 31.2957 | 31.2952 | 6.5112 | 1.6625 | 10 | 10 | 61 | 6.6305 | 0.54106 | 3.5875 |  | 0.0046287
19 | 40.089 | 8.8197 | 8.8205 | 32.1086 | 32.1094 | 6.6026 | 1.4544 | 4 | 4 | 80 | 6.7308 | 0.44062 | 2.9657 | 6 | 0.0047218
26 | 124.915 | 7.7647 | 7.7648 | 33.066 | 33.067 | 6.7192 | 1.2087 | 2 | 2 | 97 | 6.9122 | 0.33303 | 2.302 |  | 0.0047345
43 | 39.874 | 9.8431 | 9.8499 | 29.2835 | 29.2823 | 6.5731 | 1.6951 | 11 | 11 | 143 | 6.7121 | 0.54599 | 3.6648 |  | 0.004744
13 | 30.816 | 9.3601 | 9.3621 | 31.629 | 31.6285 | 6.5441 | 1.5941 | 9 | 9 | 60 | 6.6694 | 0.51128 | 3.4099 |  | 0.0047666
48 | 299.458 | 8.9219 | 8.9224 | 31.0765 | 31.0772 | 6.6317 | 1.4279 | 3 | 3 | 150 | 6.9892 | 0.42892 | 2.9978 |  | 0.0048187
31 | 76.317 | 10.1635 | 10.1625 | 30.0674 | 30.07 | 6.4943 | 1.6618 | 7 | 7 | 113 | 6.666 | 0.54095 | 3.606 |  | 0.004822
58 | 51.232 | 8.9412 | 8.9417 | 29.5335 | 29.5347 | 6.6953 | 1.6916 | 10 | 10 | 205 | 6.8364 | 0.54423 | 3.7206 |  | 0.0049436
51 | 124.158 | 9.1557 | 9.1562 | 30.5451 | 30.5461 | 6.62 | 1.5324 | 7 | 7 | 171 | 6.8251 | 0.48443 | 3.3063 |  | 0.0049641
13 | 151.004 | 6.9356 | 6.9359 | 33.8747 | 33.8755 | 6.8113 | 1.0805 | 3 | 3 | 54 | 7.0211 | 0.27661 | 1.9421 |  | 0.0050291
7 | 174.664 | 6.778 | 6.7785 | 33.8989 | 33.8995 | 6.835 | 1.0462 | 3 | 3 | 26 | 7.0651 | 0.26139 | 1.8468 |  | 0.0052163
63 | 74.572 | 8.7961 | 8.7967 | 29.8893 | 29.8901 | 6.7017 | 1.6456 | 5 | 5 | 229 | 6.8617 | 0.52377 | 3.594 |  | 0.0052486
19 | 114.171 | 7.3014 | 7.3016 | 33.5454 | 33.5466 | 6.7691 | 1.1491 | 1 | 1 | 77 | 6.9488 | 0.30645 | 2.1294 |  | 0.0052518
58 | 174.692 | 8.768 | 8.7685 | 30.736 | 30.7365 | 6.6693 | 1.5028 | 5 | 5 | 200 | 6.916 | 0.46122 | 3.1898 | 6 | 0.0052966
19 | 51.635 | 8.707 | 8.707 | 32.1947 | 32.1959 | 6.6156 | 1.4268 | 3 | 3 | 79 | 6.7527 | 0.42796 | 2.8899 |  | 0.0053166
66 | 19.457 | 10.2923 | 10.2927 | 28.316 | 28.3139 | 6.5489 | 1.8276 | 13 | 13 | 249 | 6.6752 | 0.60331 | 4.0272 |  | 0.0054299
43 | 125.449 | 9.4818 | 9.4589 | 30.3921 | 30.3922 | 6.5787 | 1.5399 | 7 | 7 | 139 | 6.7873 | 0.48821 | 3.3136 |  | 0.0054426
23 | 42.017 | 9.2163 | 9.2187 | 31.7379 | 31.7371 | 6.5604 | 1.5266 | 5 | 5 | 90 | 6.6939 | 0.48249 | 3.2297 |  | 0.0055292
58 | 100.47 | 8.9217 | 8.9221 | 30.3224 | 30.3233 | 6.6642 | 1.6194 | 8 | 8 | 203 | 6.8473 | 0.51202 | 3.5059 |  | 0.005545
58 | 150.255 | 8.8106 | 8.8106 | 30.6655 | 30.6657 | 6.6659 | 1.549 | 6 | 6 | 201 | 6.8915 | 0.48115 | 3.3158 |  | 0.0056139
3 | 19.758 | 11.413 | 11.4131 | 29.1006 | 29.1022 | 6.3599 | 2.0081 | 11 | 11 | 20 | 6.4951 | 0.69344 | 4.504 | 6 | 0.0056789
48 | 377.653 | 8.8629 | 8.8636 | 31.2166 | 31.2169 | 6.6345 | 1.3742 | 1 | 1 | 148 | 7.0616 | 0.4045 | 2.8564 |  | 0.0057335
48 | 74.291 | 9.225 | 9.2251 | 29.8924 | 29.8932 | 6.6376 | 1.6412 | 10 | 10 | 157 | 6.8007 | 0.53285 | 3.6238 |  | 0.0057541
23 | 19.765 | 10.3389 | 10.339 | 30.604 | 30.6043 | 6.4473 | 1.7801 | 7 | 7 | 92 | 6.5724 | 0.59371 | 3.9021 |  | 0.0057845
63 | 165.711 | 8.7835 | 8.7841 | 30.7697 | 30.7709 | 6.6655 | 1.3915 | 1 | 1 | 225 | 6.9044 | 0.412 | 2.8446 |  | 0.0057902
3 | 174.091 | 8.6293 | 8.6307 | 32.1958 | 32.196 | 6.627 | 1.3768 | 3 | 3 | 12 | 6.8701 | 0.40552 | 2.786 |  | 0.005862
23 | 75.425 | 8.1912 | 8.1915 | 32.7002 | 32.701 | 6.6706 | 1.316 | 3 | 3 | 88 | 6.8237 | 0.37881 | 2.5849 | 6 | 0.0059224
58 | 250.352 | 8.7898 | 8.7902 | 30.9396 | 30.9397 | 6.6572 | 1.4408 | 3 | 3 | 198 | 6.9708 | 0.43326 | 3.0202 |  | 0.0061079
16 | 169.361 | 7.0108 | 7.0112 | 33.8401 | 33.841 | 6.8011 | 1.1023 | 1 | 1 | 65 | 7.028 | 0.28507 | 2.0035 |  | 0.0061103
19 | 75.908 | 7.6214 | 7.6407 | 33.2162 | 33.1987 | 6.7344 | 1.2124 | 2 | 2 | 78 | 6.8829 | 0.34572 | 2.3795 |  | 0.006327
26 | 144.918 | 7.7457 | 7.7461 | 33.086 | 33.0868 | 6.7212 | 1.2007 | 1 | 1 | 96 | 6.9318 | 0.32794 | 2.2732 |  | 0.0063475
63 | 125.14 | 8.572 | 8.5724 | 30.4132 | 30.414 | 6.7127 | 1.6039 | 3 | 3 | 227 | 6.9149 | 0.50437 | 3.4877 |  | 0.006436
16 | 20.355 | 9.7723 | 9.771 | 31.2407 | 31.2426 | 6.5012 | 1.6822 | 9 | 9 | 73 | 6.6214 | 0.55154 | 3.652 |  | 0.0064859
58 | 300.365 | 8.8511 | 8.8516 | 31.0467 | 31.0472 | 6.6435 | 1.4294 | 2 | 2 | 197 | 7.0017 | 0.42762 | 2.9941 |  | 0.0067514
3 | 221.734 | 8.5109 | 8.5112 | 32.3323 | 32.3339 | 6.6387 | 1.3504 | 1 | 1 | 10 | 6.923 | 0.40659 | 2.8148 |  | 0.0067623
10 | 217.116 | 6.7429 | 6.7435 | 33.9298 | 33.9306 | 6.8391 | 1.0719 | 1 | 1 | 38 | 7.1077 | 0.28473 | 2.0238 |  | 0.0068603
3 | 10.398 | 11.4026 | 11.4039 | 29.0645 | 29.0651 | 6.3628 | 2.0071 | 12 | 12 | 21 | 6.4902 | 0.69421 | 4.5056 |  | 0.0068893
26 | 29.683 | 10.2055 | 10.205 | 30.6213 | 30.6227 | 6.4654 | 1.7306 | 7 | 7 | 102 | 6.5976 | 0.57328 | 3.7823 |  | 0.0070413
51 | 20.166 | 10.1415 | 10.1416 | 28.4508 | 28.4518 | 6.565 | 1.7897 | 13 | 13 | 177 | 6.6904 | 0.59927 | 4.0094 | 6 | 0.0071165
70 | 19.527 | 10.149 | 10.1521 | 28.9406 | 28.9386 | 6.5434 | 1.8755 | 8 | 8 | 275 | 6.6679 | 0.6368 | 4.2462 | 6 | 0.0071172
10 | 6.66 | 10.1258 | 10.1347 | 31.0502 | 31.0461 | 6.4589 | 1.8452 | 13 | 13 | 50 | 6.5709 | 0.62369 | 4.0982 |  | 0.0072786
60 | 5.104 | 17.3344 | 17.3323 | 24.8717 | 24.8739 | 5.7823 | 2.9852 | 12 | 12 | 223 | 5.9522 | 1.1082 | 6.5962 |  | 0.007301
13 | 125.492 | 6.9489 | 6.9494 | 33.8734 | 33.8741 | 6.8093 | 1.0907 | 4 | 4 | 55 | 6.9962 | 0.27878 | 1.9504 | 6 | 0.007311
66 | 75.838 | 8.8114 | 8.812 | 29.7841 | 29.7845 | 6.704 | 1.7064 | 9 | 9 | 245 | 6.8653 | 0.54833 | 3.7645 |  | 0.0073262
16 | 98.72 | 7.0058 | 7.0061 | 33.8385 | 33.8393 | 6.8019 | 1.1065 | 4 | 4 | 68 | 6.9653 | 0.28565 | 1.9896 |  | 0.0073898
1 | 29.955 | 10.6329 | 10.6397 | 28.9905 | 28.9911 | 6.4725 | 1.5252 | 5 | 5 | 5 | 6.6099 | 0.46888 | 3.0993 |  | 0.0074567
13 | 39.974 | 8.8642 | 8.8673 | 32.1989 | 32.1978 | 6.5922 | 1.5138 | 8 | 8 | 59 | 6.7206 | 0.479 | 3.2192 |  | 0.0076414
63 | 29.858 | 9.9426 | 9.9431 | 28.7879 | 28.7882 | 6.5796 | 1.8324 | 8 | 8 | 232 | 6.7113 | 0.60309 | 4.0475 |  | 0.0077527
68 | 49.98 | 9.7767 | 9.7699 | 29.0657 | 29.0668 | 6.5919 | 1.8093 | 9 | 9 | 261 | 6.7393 | 0.59295 | 3.996 |  | 0.0077667
16 | 75.077 | 7.3739 | 7.3741 | 33.545 | 33.5463 | 6.7579 | 1.2281 | 5 | 5 | 69 | 6.9034 | 0.33825 | 2.3351 |  | 0.0079949
63 | 2.263 | 16.8657 | 16.9069 | 25.456 | 25.4456 | 5.8152 | 2.7795 | 12 | 12 | 236 | 5.9795 | 1.0174 | 6.0838 |  | 0.0080376
16 | 124.25 | 7.0044 | 7.0048 | 33.8406 | 33.8413 | 6.8021 | 1.1028 | 3 | 3 | 67 | 6.9883 | 0.28302 | 1.9778 |  | 0.0083992
16 | 150.178 | 7.0078 | 7.0083 | 33.8403 | 33.8411 | 6.8015 | 1.1024 | 2 | 2 | 66 | 7.0111 | 0.28271 | 1.9821 |  | 0.0085168
55 | 2.02 | 17.6449 | 17.6447 | 24.9506 | 24.9463 | 5.7446 | 2.9761 | 15 | 15 | 195 | 5.914 | 1.1201 | 6.6241 |  | 0.0085559
1 | 40.002 | 10.2113 | 10.2127 | 29.5419 | 29.5432 | 6.5093 | 1.4687 | 4 | 4 | 4 | 6.6513 | 0.44299 | 2.9464 |  | 0.0086011
13 | 4.952 | 11.3567 | 11.3544 | 30.454 | 30.4561 | 6.3132 | 2.0858 | 12 | 12 | 63 | 6.4347 | 0.73041 | 4.7 |  | 0.0086376
1 | 99.874 | 9.0497 | 9.0508 | 31.0725 | 31.0745 | 6.613 | 0.76771 | 1 | 1 | 1 | 6.7957 | 0.15338 | 1.0423 |  | 0.0086734
10 | 11.993 | 9.9316 | 9.9343 | 31.2109 | 31.2108 | 6.4797 | 1.7941 | 12 | 12 | 49 | 6.5943 | 0.60278 | 3.9749 |  | 0.0087195
60 | 174.448 | 8.7466 | 8.747 | 30.7065 | 30.7076 | 6.6737 | 1.497 | 2 | 2 | 213 | 6.9202 | 0.45496 | 3.1484 | 6 | 0.0090311
3 | 124.152 | 9.1001 | 9.1023 | 31.6351 | 31.6347 | 6.5817 | 1.4714 | 5 | 5 | 14 | 6.785 | 0.44337 | 3.0083 |  | 0.0093968
31 | 30.626 | 10.6227 | 10.6216 | 29.2954 | 29.2986 | 6.4614 | 1.7636 | 10 | 10 | 116 | 6.599 | 0.59016 | 3.8944 |  | 0.0094524
7 | 248.365 | 6.611 | 6.6117 | 33.9444 | 33.9451 | 6.8593 | 1.0281 | 1 | 1 | 24 | 7.1559 | 0.26818 | 1.9191 |  | 0.0094656
58 | 29.582 | 9.4434 | 9.4453 | 29.0685 | 29.0683 | 6.6406 | 1.749 | 12 | 12 | 207 | 6.7676 | 0.58415 | 3.9533 |  | 0.0098128
7 | 3.071 | 11.2614 | 11.3416 | 30.623 | 30.5823 | 6.3194 | 2.1604 | 14 | 14 | 37 | 6.4385 | 0.76439 | 4.9215 |  | 0.0099849
58 | 5.326 | 16.6704 | 16.707 | 25.1353 | 25.1286 | 5.8491 | 2.8771 | 15 | 15 | 210 | 6.0147 | 1.0578 | 6.3622 |  | 0.01042
23 | 4.799 | 10.8333 | 10.8344 | 30.2105 | 30.2118 | 6.3945 | 2.0034 | 9 | 9 | 94 | 6.5117 | 0.67518 | 4.3966 |  | 0.010487
16 | 1.677 | 11.9098 | 11.8999 | 30.4768 | 30.4801 | 6.2384 | 2.2454 | 12 | 12 | 76 | 6.3618 | 0.80258 | 5.1059 |  | 0.01093
23 | 9.307 | 10.63 | 10.6308 | 30.2786 | 30.279 | 6.42 | 1.9103 | 8 | 8 | 93 | 6.5391 | 0.63402 | 4.1459 |  | 0.010935
1 | 1.758 | 15.3085 | 15.3959 | 27.5131 | 27.515 | 5.9238 | 3.1839 | 9 | 9 | 9 | 6.0755 | 1.2136 | 7.3733 |  | 0.011115
70 | 5.031 | 10.8644 | 10.8644 | 28.365 | 28.3543 | 6.4657 | 1.9873 | 10 | 10 | 277 | 6.5847 | 0.69009 | 4.544 |  | 0.011427
63 | 9.948 | 11.9794 | 11.974 | 27.69 | 27.6991 | 6.3396 | 2.1659 | 10 | 10 | 234 | 6.4725 | 0.74518 | 4.8232 |  | 0.011662
58 | 2.293 | 16.8583 | 16.8637 | 25.1132 | 25.1142 | 5.8282 | 2.9005 | 16 | 16 | 211 | 5.9928 | 1.0911 | 6.5386 |  | 0.012623
31 | 19.203 | 12.0971 | 12.1312 | 27.3994 | 27.358 | 6.3353 | 2.0082 | 11 | 11 | 117 | 6.477 | 0.70059 | 4.5377 |  | 0.012804
13 | 1.867 | 11.5207 | 11.5181 | 30.4391 | 30.4405 | 6.2917 | 2.1119 | 13 | 13 | 64 | 6.4121 | 0.74617 | 4.7845 |  | 0.01296
66 | 1.822 | 18.0802 | 18.0731 | 24.9738 | 24.9786 | 5.6955 | 2.9248 | 16 | 16 | 252 | 5.8672 | 1.1025 | 6.4687 |  | 0.013449
10 | 22.007 | 9.8128 | 9.8133 | 31.3097 | 31.3111 | 6.4925 | 1.7533 | 11 | 11 | 48 | 6.6144 | 0.58963 | 3.9001 |  | 0.013455
13 | 10.173 | 10.9836 | 11.0075 | 30.5259 | 30.5219 | 6.3611 | 2.0229 | 11 | 11 | 62 | 6.4837 | 0.70796 | 4.5902 | 6 | 0.013728
3 | 5.298 | 11.4766 | 11.4829 | 28.9575 | 28.9533 | 6.357 | 2.0182 | 13 | 13 | 22 | 6.4809 | 0.70605 | 4.5759 |  | 0.013868
7 | 19.388 | 8.9216 | 8.9364 | 32.2162 | 32.2037 | 6.5831 | 1.5475 | 11 | 11 | 34 | 6.6943 | 0.50047 | 3.3503 | 6 | 0.014363
63 | 5.096 | 15.8741 | 16.2171 | 25.844 | 25.7037 | 5.917 | 2.6481 | 11 | 11 | 235 | 6.0766 | 0.95329 | 5.7928 |  | 0.014644
68 | 4.74 | 13.0243 | 13.0006 | 27.3536 | 27.3626 | 6.2142 | 2.311 | 14 | 14 | 266 | 6.3516 | 0.80536 | 5.1153 |  | 0.015003
70 | 151.125 | 8.6454 | 8.646 | 30.5921 | 30.5933 | 6.6939 | 1.5284 | 1 | 1 | 268 | 6.9193 | 0.49288 | 3.4104 |  | 0.015137
51 | 1.937 | 16.0186 | 16.012 | 24.3624 | 24.3678 | 5.9534 | 2.9151 | 16 | 16 | 180 | 6.113 | 1.1001 | 6.7246 |  | 0.01524
39 | 1.937 | 17.8392 | 17.8374 | 21.1989 | 21.2032 | 5.8534 | 3.0633 | 12 | 12 | 132 | 6.0277 | 1.1652 | 7.0236 |  | 0.015486
66 | 10.198 | 12.7353 | 12.725 | 27.0895 | 27.0857 | 6.2624 | 2.202 | 14 | 14 | 250 | 6.4021 | 0.75698 | 4.8463 |  | 0.015631
60 | 1.883 | 18.8345 | 18.8372 | 24.5968 | 24.5987 | 5.6261 | 2.9772 | 13 | 13 | 224 | 5.8025 | 1.1282 | 6.5461 |  | 0.016161
1 | 74.552 | 8.9253 | 8.925 | 30.8447 | 30.8462 | 6.6412 | 0.7762 | 2 | 2 | 2 | 6.8012 | 0.13136 | 0.89338 |  | 0.017064
43 | 2.381 | 17.4508 | 17.4526 | 23.4521 | 23.4528 | 5.8187 | 2.8572 | 15 | 15 | 147 | 5.9887 | 1.0408 | 6.2331 |  | 0.018686
31 | 11.246 | 12.2727 | 12.2701 | 27.1744 | 27.1798 | 6.3206 | 2.0312 | 12 | 12 | 118 | 6.4574 | 0.71668 | 4.6279 |  | 0.0188
43 | 9.758 | 13.9092 | 13.8961 | 26.1406 | 26.1514 | 6.1471 | 2.1656 | 13 | 13 | 145 | 6.2962 | 0.77643 | 4.8886 |  | 0.019725
43 | 4.993 | 17.3132 | 17.308 | 23.4842 | 23.4866 | 5.8333 | 2.7982 | 14 | 14 | 146 | 6.0044 | 1.0134 | 6.0849 |  | 0.020239
55 | 9.511 | 12.981 | 12.9697 | 26.5625 | 26.5735 | 6.2507 | 2.1208 | 13 | 13 | 193 | 6.3922 | 0.71592 | 4.5764 |  | 0.021184
16 | 9.995 | 10.5863 | 10.588 | 30.6945 | 30.6952 | 6.409 | 2.0354 | 10 | 10 | 74 | 6.528 | 0.72122 | 4.7081 | 6 | 0.021529
39 | 19.981 | 11.4729 | 11.4802 | 27.6267 | 27.6227 | 6.4112 | 1.9018 | 9 | 9 | 129 | 6.5484 | 0.61845 | 4.0499 |  | 0.022749
7 | 6.192 | 10.045 | 10.0597 | 31.204 | 31.1969 | 6.4639 | 1.7899 | 13 | 13 | 36 | 6.5747 | 0.61621 | 4.0514 |  | 0.023971
3 | 2.182 | 11.8339 | 11.8273 | 28.6693 | 28.6761 | 6.3201 | 2.1039 | 14 | 14 | 23 | 6.4447 | 0.7565 | 4.8754 |  | 0.026807
68 | 2.443 | 13.732 | 13.676 | 27.086 | 27.1057 | 6.1337 | 2.3796 | 15 | 15 | 267 | 6.2749 | 0.87844 | 5.5121 |  | 0.028053
10 | 41.847 | 9.2075 | 9.2069 | 31.8442 | 31.8459 | 6.5572 | 1.5691 | 9 | 9 | 46 | 6.6903 | 0.52432 | 3.5079 |  | 0.028776
48 | 5.392 | 16.3579 | 16.3455 | 24.0807 | 24.0901 | 5.9233 | 3.055 | 16 | 16 | 163 | 6.088 | 1.1136 | 6.7795 |  | 0.032523
60 | 10.046 | 12.6884 | 12.6903 | 27.0125 | 27.0132 | 6.2716 | 2.1728 | 11 | 11 | 222 | 6.4109 | 0.79415 | 5.0912 |  | 0.034316
10 | 2.88 | 10.3632 | 10.3658 | 30.9184 | 30.9179 | 6.431 | 1.9214 | 14 | 14 | 51 | 6.542 | 0.6862 | 4.4891 |  | 0.036405
48 | 2.106 | 18.055 | 17.9826 | 23.1719 | 23.2061 | 5.7602 | 2.8983 | 17 | 17 | 164 | 5.9338 | 1.1158 | 6.6209 |  | 0.038328
31 | 4.553 | 13.5399 | 13.4994 | 25.4901 | 25.5458 | 6.2195 | 2.2256 | 13 | 13 | 119 | 6.3623 | 0.82503 | 5.2491 |  | 0.042064
66 | 5.026 | 17.4526 | 17.3044 | 25.2677 | 25.3183 | 5.7552 | 2.8623 | 15 | 15 | 251 | 5.9253 | 1.0172 | 6.0275 |  | 0.044456
1 | 9.941 | 13.2718 | 13.2816 | 27.7346 | 27.7376 | 6.1675 | 2.4337 | 7 | 7 | 7 | 6.3106 | 0.82959 | 5.2352 |  | 0.04446
51 | 5.037 | 13.447 | 13.5184 | 26.0748 | 26.0423 | 6.2089 | 2.2588 | 15 | 15 | 179 | 6.3509 | 0.73728 | 4.6824 |  | 0.060239
1 | 5.123 | 14.4689 | 14.48 | 27.5053 | 27.5089 | 6.0257 | 2.9509 | 8 | 8 | 8 | 6.1741 | 1.0252 | 6.3296 |  | 0.075311
63 | 150.147 | 8.6905 | 8.6907 | 30.6116 | 30.6122 | 6.6862 | 1.454 | 2 | 2 | 226 | 6.9111 | 0.43843 | 3.03 | 4 | 0.0067348
58 | 200.582 | 8.8024 | 8.8031 | 30.8193 | 30.8202 | 6.6605 | 1.4384 | 4 | 4 | 199 | 6.9302 | 0.43023 | 2.9816 | 3 | 0.0080858
43 | 201.863 | 8.9053 | 8.9089 | 30.8334 | 30.836 | 6.6446 | 1.4814 | 4 | 4 | 136 | 6.916 | 0.45297 | 3.1327 | 3 | 0.0041805
10 | 152.22 | 6.9083 | 6.9088 | 33.9015 | 33.9021 | 6.8144 | 1.1288 | 4 | 4 | 41 | 7.025 | 0.29983 | 2.1063 | 3 | 0.00298
66 | 103.164 | 8.7523 | 8.7535 | 30.2171 | 30.2178 | 6.6941 | 1.6526 | 8 | 8 | 244 | 6.8785 | 0.52878 | 3.6372 | 3 | 0.0033277
48 | 249.14 | 8.9221 | 8.9225 | 30.9638 | 30.9646 | 6.6365 | 1.4575 | 4 | 4 | 151 | 6.9495 | 0.44583 | 3.0983 | 3 | 0.00088525
68 | 175.23 | 8.5543 | 8.5543 | 30.4545 | 30.4545 | 6.7136 | 1.5767 | 4 | 4 | 256 | 6.96 | 0.4977 | 3.464 | 3 | 0.0011751
60 | 126.432 | 8.7798 | 8.7805 | 30.381 | 30.3814 | 6.6828 | 1.6033 | 4 | 4 | 215 | 6.8877 | 0.50953 | 3.5095 | 3 | 0.00099229
10 | 51.947 | 8.7869 | 8.7906 | 32.2356 | 32.2328 | 6.602 | 1.4766 | 8 | 8 | 45 | 6.74 | 0.46085 | 3.1062 | 3 | 0.0057948
58 | 10.324 | 12.2835 | 12.2899 | 27.2171 | 27.2172 | 6.3175 | 2.1071 | 14 | 14 | 209 | 6.4536 | 0.73885 | 4.7682 | 3 | 0.0077578
13 | 100.25 | 7.0156 | 7.0167 | 33.8516 | 33.8522 | 6.7998 | 1.1257 | 5 | 5 | 56 | 6.9646 | 0.32186 | 2.2416 | 3 | 0.020433
23 | 101.686 | 7.712 | 7.7126 | 33.1459 | 33.1466 | 6.7237 | 1.2206 | 2 | 2 | 87 | 6.8957 | 0.48516 | 3.3456 | 3 | 0.1422
63 | 100.461 | 8.6452 | 8.6457 | 30.2741 | 30.2749 | 6.7077 | 1.6473 | 4 | 4 | 228 | 6.8889 | 0.52349 | 3.6063 | 36 | 0.0062979
66 | 29.848 | 9.5816 | 9.5734 | 28.8606 | 28.8693 | 6.6292 | 1.7465 | 12 | 12 | 248 | 6.7577 | 0.58125 | 3.928 | 36 | 0.0080457

## ctd + bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
1 | 99.874 | 9.0497 | 9.0508 | 31.0725 | 31.0745 | 6.613 | 0.76771 | 1 | 1 | 1 | 1.0423 |  | 9.2
1 | 74.552 | 8.9253 | 8.925 | 30.8447 | 30.8462 | 6.6412 | 0.7762 | 2 | 2 | 2 | 0.89338 |  | 9.7
1 | 49.865 | 9.9211 | 9.9208 | 29.8363 | 29.8378 | 6.5386 | 1.3971 | 3 | 3 | 3 | 2.7881 |  | 10.5
1 | 40.002 | 10.2113 | 10.2127 | 29.5419 | 29.5432 | 6.5093 | 1.4687 | 4 | 4 | 4 | 2.9464 |  | 10.5
1 | 29.955 | 10.6329 | 10.6397 | 28.9905 | 28.9911 | 6.4725 | 1.5252 | 5 | 5 | 5 | 3.0993 |  | 11
1 | 19.644 | 11.7827 | 11.7868 | 28.5391 | 28.5379 | 6.3322 | 1.6508 | 6 | 6 | 6 | 3.4396 | 6 | 12.3
1 | 9.941 | 13.2718 | 13.2816 | 27.7346 | 27.7376 | 6.1675 | 2.4337 | 7 | 7 | 7 | 5.2352 |  | 13.6
1 | 5.123 | 14.4689 | 14.48 | 27.5053 | 27.5089 | 6.0257 | 2.9509 | 8 | 8 | 8 | 6.3296 |  | 14.5
1 | 1.758 | 15.3085 | 15.3959 | 27.5131 | 27.515 | 5.9238 | 3.1839 | 9 | 9 | 9 | 7.3733 |  | 16.1
3 | 221.734 | 8.5109 | 8.5112 | 32.3323 | 32.3339 | 6.6387 | 1.3504 | 1 | 1 | 10 | 2.8148 |  | 8.9
3 | 203.081 | 8.5362 | 8.5363 | 32.3035 | 32.3052 | 6.6362 | 1.3568 | 2 | 2 | 11 | 2.7623 |  | 8.7
3 | 174.091 | 8.6293 | 8.6307 | 32.1958 | 32.196 | 6.627 | 1.3768 | 3 | 3 | 12 | 2.786 |  | 8.8
3 | 151.857 | 8.7993 | 8.8 | 31.9901 | 31.9913 | 6.6107 | 1.4126 | 4 | 4 | 13 | 2.9025 | 6 | 8.9
3 | 124.152 | 9.1001 | 9.1023 | 31.6351 | 31.6347 | 6.5817 | 1.4714 | 5 | 5 | 14 | 3.0083 |  | 9.1
3 | 100.778 | 9.3556 | 9.36 | 31.358 | 31.3554 | 6.5562 | 1.5275 | 6 | 6 | 15 | 3.2412 |  | 9.6
3 | 75.428 | 9.5414 | 9.5413 | 31.1202 | 31.1228 | 6.5394 | 1.5637 | 7 | 7 | 16 | 3.3209 |  | 9.6
3 | 51.149 | 10.6374 | 10.6385 | 29.9979 | 29.9984 | 6.4304 | 1.8003 | 8 | 8 | 17 | 3.9533 |  | 10.7
3 | 39.909 | 11.0502 | 11.0506 | 29.6401 | 29.6417 | 6.3878 | 1.8946 | 9 | 9 | 18 | 4.193 |  | 11
3 | 30.238 | 11.315 | 11.3174 | 29.2856 | 29.284 | 6.3658 | 1.9823 | 10 | 10 | 19 | 4.4312 |  | 11.3
3 | 19.758 | 11.413 | 11.4131 | 29.1006 | 29.1022 | 6.3599 | 2.0081 | 11 | 11 | 20 | 4.504 | 6 | 11.4
3 | 10.398 | 11.4026 | 11.4039 | 29.0645 | 29.0651 | 6.3628 | 2.0071 | 12 | 12 | 21 | 4.5056 |  | 11.4
3 | 5.298 | 11.4766 | 11.4829 | 28.9575 | 28.9533 | 6.357 | 2.0182 | 13 | 13 | 22 | 4.5759 |  | 11.5
3 | 2.182 | 11.8339 | 11.8273 | 28.6693 | 28.6761 | 6.3201 | 2.1039 | 14 | 14 | 23 | 4.8754 |  | 11.9
7 | 248.365 | 6.611 | 6.6117 | 33.9444 | 33.9451 | 6.8593 | 1.0281 | 1 | 1 | 24 | 1.9191 |  | 7.1
7 | 199.54 | 6.7021 | 6.7027 | 33.9202 | 33.9207 | 6.846 | 1.0329 | 2 | 2 | 25 | 1.8248 |  | 7.1
7 | 174.664 | 6.778 | 6.7785 | 33.8989 | 33.8995 | 6.835 | 1.0462 | 3 | 3 | 26 | 1.8468 |  | 7.3
7 | 150.051 | 6.9337 | 6.9345 | 33.8673 | 33.8678 | 6.8119 | 1.1203 | 4 | 4 | 27 | 2.0864 | 6 | 7.4
7 | 124.195 | 7.217 | 7.2169 | 33.7433 | 33.7442 | 6.7734 | 1.2298 | 5 | 5 | 28 | 2.4181 |  | 7.6
7 | 99.381 | 7.2517 | 7.2526 | 33.7267 | 33.7271 | 6.7687 | 1.2368 | 6 | 6 | 29 | -99 | 5 | 7.7 | OXY: Stopper not sealed. Sample not analyzed
7 | 75.097 | 7.4032 | 7.4047 | 33.6575 | 33.6575 | 6.7484 | 1.2872 | 7 | 7 | 30 | 2.5858 |  | 7.8
7 | 48.117 | 7.9339 | 7.9339 | 33.299 | 33.3009 | 6.6833 | 1.4448 | 8 | 8 | 31 | 3.002 |  | 8.2
7 | 39.245 | 7.9689 | 7.9715 | 33.1273 | 33.124 | 6.6855 | 1.4065 | 9 | 9 | 32 | 2.8957 |  | 8.3
7 | 30.22 | 8.356 | 8.3608 | 32.77 | 32.766 | 6.6429 | 1.4433 | 10 | 10 | 33 | 2.9773 |  | 8.7
7 | 19.388 | 8.9216 | 8.9364 | 32.2162 | 32.2037 | 6.5831 | 1.5475 | 11 | 11 | 34 | 3.3503 | 6 | 9.2
7 | 11.201 | 9.8276 | 9.8237 | 31.5369 | 31.5316 | 6.4809 | 1.7879 | 12 | 12 | 35 | 3.9257 |  | 9.9
7 | 6.192 | 10.045 | 10.0597 | 31.204 | 31.1969 | 6.4639 | 1.7899 | 13 | 13 | 36 | 4.0514 |  | 10.3
7 | 3.071 | 11.2614 | 11.3416 | 30.623 | 30.5823 | 6.3194 | 2.1604 | 14 | 14 | 37 | 4.9215 |  | 11.4
10 | 217.116 | 6.7429 | 6.7435 | 33.9298 | 33.9306 | 6.8391 | 1.0719 | 1 | 1 | 38 | 2.0238 |  | 7.6
10 | 202.889 | 6.7751 | 6.7762 | 33.9252 | 33.9259 | 6.8342 | 1.0829 | 2 | 2 | 39 | 1.9855 |  | 7.5
10 | 177.187 | 6.8199 | 6.8204 | 33.9196 | 33.9201 | 6.8274 | 1.0977 | 3 | 3 | 40 | 2.0305 | 6 | 7.4
10 | 152.22 | 6.9083 | 6.9088 | 33.9015 | 33.9021 | 6.8144 | 1.1288 | 4 | 4 | 41 | 2.1063 | 3 | 7.5 | OXY: Niskin leaked, not air tight
10 | 127.01 | 7.0643 | 7.0649 | 33.8429 | 33.8435 | 6.7926 | 1.2005 | 5 | 5 | 42 | 2.3296 |  | 7.7
10 | 102.106 | 7.3492 | 7.3488 | 33.7165 | 33.7176 | 6.7541 | 1.305 | 6 | 6 | 43 | 2.6198 |  | 8
10 | 77.085 | 7.5928 | 7.5956 | 33.5294 | 33.5286 | 6.725 | 1.3872 | 7 | 7 | 44 | 2.8453 |  | 8.2
10 | 51.947 | 8.7869 | 8.7906 | 32.2356 | 32.2328 | 6.602 | 1.4766 | 8 | 8 | 45 | 3.1062 | 3 | 9.2 | OXY: Only two points on the ascending slope. Difficult to assess validity of end point
10 | 41.847 | 9.2075 | 9.2069 | 31.8442 | 31.8459 | 6.5572 | 1.5691 | 9 | 9 | 46 | 3.5079 |  | 9.6
10 | 31.847 | 9.6347 | 9.6333 | 31.4376 | 31.437 | 6.5126 | 1.6929 | 10 | 10 | 47 | 3.6756 | 6 | 9.9
10 | 22.007 | 9.8128 | 9.8133 | 31.3097 | 31.3111 | 6.4925 | 1.7533 | 11 | 11 | 48 | 3.9001 |  | 10.3
10 | 11.993 | 9.9316 | 9.9343 | 31.2109 | 31.2108 | 6.4797 | 1.7941 | 12 | 12 | 49 | 3.9749 |  | 10.3
10 | 6.66 | 10.1258 | 10.1347 | 31.0502 | 31.0461 | 6.4589 | 1.8452 | 13 | 13 | 50 | 4.0982 |  | 10.5
10 | 2.88 | 10.3632 | 10.3658 | 30.9184 | 30.9179 | 6.431 | 1.9214 | 14 | 14 | 51 | 4.4891 |  | 10.8
13 | 186.599 | 6.9361 | 6.9364 | 33.8748 | 33.8757 | 6.8112 | 1.0765 | 1 | 1 | 52 | 1.9437 |  | 7.6
13 | 175.84 | 6.9351 | 6.9353 | 33.8746 | 33.8755 | 6.8114 | 1.0768 | 2 | 2 | 53 | 1.9528 |  | 7.7
13 | 151.004 | 6.9356 | 6.9359 | 33.8747 | 33.8755 | 6.8113 | 1.0805 | 3 | 3 | 54 | 1.9421 |  | 7.5
13 | 125.492 | 6.9489 | 6.9494 | 33.8734 | 33.8741 | 6.8093 | 1.0907 | 4 | 4 | 55 | 1.9504 | 6 | 7.5
13 | 100.25 | 7.0156 | 7.0167 | 33.8516 | 33.8522 | 6.7998 | 1.1257 | 5 | 5 | 56 | 2.2416 | 3 | 7.6 | OXY: No good, bubble on lens
13 | 75.71 | 7.2472 | 7.2473 | 33.7397 | 33.7409 | 6.7688 | 1.2541 | 6 | 6 | 57 | 2.4991 |  | 7.8
13 | 52.477 | 8.4739 | 8.4739 | 32.6047 | 32.6054 | 6.6325 | 1.4382 | 7 | 7 | 58 | 2.9642 |  | 8.9
13 | 39.974 | 8.8642 | 8.8673 | 32.1989 | 32.1978 | 6.5922 | 1.5138 | 8 | 8 | 59 | 3.2192 |  | 9.3
13 | 30.816 | 9.3601 | 9.3621 | 31.629 | 31.6285 | 6.5441 | 1.5941 | 9 | 9 | 60 | 3.4099 |  | 9.7
13 | 20.222 | 9.6863 | 9.6882 | 31.2957 | 31.2952 | 6.5112 | 1.6625 | 10 | 10 | 61 | 3.5875 |  | 10
13 | 10.173 | 10.9836 | 11.0075 | 30.5259 | 30.5219 | 6.3611 | 2.0229 | 11 | 11 | 62 | 4.5902 | 6 | 11.4
13 | 4.952 | 11.3567 | 11.3544 | 30.454 | 30.4561 | 6.3132 | 2.0858 | 12 | 12 | 63 | 4.7 |  | 11.7
13 | 1.867 | 11.5207 | 11.5181 | 30.4391 | 30.4405 | 6.2917 | 2.1119 | 13 | 13 | 64 | 4.7845 |  | 11.8
16 | 169.361 | 7.0108 | 7.0112 | 33.8401 | 33.841 | 6.8011 | 1.1023 | 1 | 1 | 65 | 2.0035 |  | 7.6
16 | 150.178 | 7.0078 | 7.0083 | 33.8403 | 33.8411 | 6.8015 | 1.1024 | 2 | 2 | 66 | 1.9821 |  | 7.6
16 | 124.25 | 7.0044 | 7.0048 | 33.8406 | 33.8413 | 6.8021 | 1.1028 | 3 | 3 | 67 | 1.9778 |  | 7.5
16 | 98.72 | 7.0058 | 7.0061 | 33.8385 | 33.8393 | 6.8019 | 1.1065 | 4 | 4 | 68 | 1.9896 |  | 7.6
16 | 75.077 | 7.3739 | 7.3741 | 33.545 | 33.5463 | 6.7579 | 1.2281 | 5 | 5 | 69 | 2.3351 |  | 7.9
16 | 50.28 | 7.8453 | 7.8437 | 33.0407 | 33.0433 | 6.708 | 1.2411 | 6 | 6 | 70 | 2.3819 | 6 | 8.3
16 | 40.461 | 8.8222 | 8.822 | 32.1617 | 32.1625 | 6.6 | 1.4768 | 7 | 7 | 71 | 3.0501 |  | 9.1
16 | 29.957 | 9.2593 | 9.2603 | 31.6706 | 31.6709 | 6.557 | 1.5564 | 8 | 8 | 72 | 3.2849 |  | 9.5
16 | 20.355 | 9.7723 | 9.771 | 31.2407 | 31.2426 | 6.5012 | 1.6822 | 9 | 9 | 73 | 3.652 |  | 10
16 | 9.995 | 10.5863 | 10.588 | 30.6945 | 30.6952 | 6.409 | 2.0354 | 10 | 10 | 74 | 4.7081 | 6 | 10.7
16 | 5.267 | 11.9419 | 11.9291 | 30.4699 | 30.4736 | 6.2344 | 2.249 | 11 | 11 | 75 | 5.0749 |  | 12
16 | 1.677 | 11.9098 | 11.8999 | 30.4768 | 30.4801 | 6.2384 | 2.2454 | 12 | 12 | 76 | 5.1059 |  | 12
19 | 114.171 | 7.3014 | 7.3016 | 33.5454 | 33.5466 | 6.7691 | 1.1491 | 1 | 1 | 77 | 2.1294 |  | 8
19 | 75.908 | 7.6214 | 7.6407 | 33.2162 | 33.1987 | 6.7344 | 1.2124 | 2 | 2 | 78 | 2.3795 |  | 8.3
19 | 51.635 | 8.707 | 8.707 | 32.1947 | 32.1959 | 6.6156 | 1.4268 | 3 | 3 | 79 | 2.8899 |  | 9
19 | 40.089 | 8.8197 | 8.8205 | 32.1086 | 32.1094 | 6.6026 | 1.4544 | 4 | 4 | 80 | 2.9657 | 6 | 9.1
19 | 30.108 | 9.1176 | 9.1223 | 31.7841 | 31.7797 | 6.5728 | 1.5149 | 5 | 5 | 81 | 3.1569 |  | 9.4
19 | 19.177 | 9.6244 | 9.6248 | 31.3008 | 31.3018 | 6.5198 | 1.6234 | 6 | 6 | 82 | 3.4316 |  | 9.8
19 | 11.111 | 9.6891 | 9.6903 | 31.246 | 31.2469 | 6.5128 | 1.6388 | 7 | 7 | 83 | 3.4832 |  | 9.9
19 | 4.916 | 10.4921 | 10.4941 | 30.8267 | 30.8268 | 6.4168 | 1.8887 | 8 | 8 | 84 | 4.147 |  | 10.6
19 | 2.745 | 10.5393 | 10.5452 | 30.7832 | 30.7822 | 6.412 | 1.8932 | 9 | 9 | 85 | 4.1753 |  | 10.6
23 | 123.257 | 7.4888 | 7.491 | 33.3628 | 33.3624 | 6.7483 | 1.175 | 1 | 1 | 86 | 2.2692 |  | 7.9
23 | 101.686 | 7.712 | 7.7126 | 33.1459 | 33.1466 | 6.7237 | 1.2206 | 2 | 2 | 87 | 3.3456 | 3 | 8.1 | OXY: Bubble on lens. Endpoint at 84%
23 | 75.425 | 8.1912 | 8.1915 | 32.7002 | 32.701 | 6.6706 | 1.316 | 3 | 3 | 88 | 2.5849 | 6 | 8.5
23 | 51.716 | 8.7929 | 8.794 | 32.1047 | 32.1052 | 6.6067 | 1.4324 | 4 | 4 | 89 | 2.9229 |  | 9
23 | 42.017 | 9.2163 | 9.2187 | 31.7379 | 31.7371 | 6.5604 | 1.5266 | 5 | 5 | 90 | 3.2297 |  | 9.5
23 | 29.605 | 9.9445 | 9.9468 | 31.232 | 31.2304 | 6.477 | 1.7106 | 6 | 6 | 91 | 3.7082 |  | 10.1
23 | 19.765 | 10.3389 | 10.339 | 30.604 | 30.6043 | 6.4473 | 1.7801 | 7 | 7 | 92 | 3.9021 |  | 10.4
23 | 9.307 | 10.63 | 10.6308 | 30.2786 | 30.279 | 6.42 | 1.9103 | 8 | 8 | 93 | 4.1459 |  | 10.7
23 | 4.799 | 10.8333 | 10.8344 | 30.2105 | 30.2118 | 6.3945 | 2.0034 | 9 | 9 | 94 | 4.3966 |  | 10.9
23 | 2.128 | 10.9752 | 10.9752 | 30.204 | 30.2053 | 6.3753 | 2.015 | 10 | 10 | 95 | 4.4765 |  | 11
26 | 144.918 | 7.7457 | 7.7461 | 33.086 | 33.0868 | 6.7212 | 1.2007 | 1 | 1 | 96 | 2.2732 |  | 8.1
26 | 124.915 | 7.7647 | 7.7648 | 33.066 | 33.067 | 6.7192 | 1.2087 | 2 | 2 | 97 | 2.302 |  | 8.1
26 | 99.808 | 8.13 | 8.1304 | 32.7029 | 32.7036 | 6.6797 | 1.278 | 3 | 3 | 98 | 2.5163 |  | 8.5
26 | 74.775 | 8.9727 | 8.9733 | 31.8481 | 31.8485 | 6.5913 | 1.4491 | 4 | 4 | 99 | 2.9685 | 6 | 9.3
26 | 49.894 | 9.6975 | 9.698 | 31.0847 | 31.0854 | 6.5184 | 1.6018 | 5 | 5 | 100 | 3.3827 |  | 9.9
26 | 39.674 | 9.8018 | 9.8031 | 30.9892 | 30.9896 | 6.5074 | 1.6327 | 6 | 6 | 101 | 3.4745 |  | 10.1
26 | 29.683 | 10.2055 | 10.205 | 30.6213 | 30.6227 | 6.4654 | 1.7306 | 7 | 7 | 102 | 3.7823 |  | 10.4
26 | 19.37 | 10.399 | 10.3993 | 30.4406 | 30.4417 | 6.4456 | 1.8036 | 8 | 8 | 103 | 3.9542 |  | 10.5
26 | 9.268 | 10.4011 | 10.4013 | 30.4376 | 30.4389 | 6.4454 | 1.8012 | 9 | 9 | 104 | 3.9351 |  | 10.6
26 | 4.72 | 10.4247 | 10.4245 | 30.4086 | 30.4105 | 6.4433 | 1.8044 | 10 | 10 | 105 | 3.9466 |  | 10.6
26 | 2.199 | 10.437 | 10.4372 | 30.3915 | 30.3926 | 6.4423 | 1.8085 | 11 | 11 | 106 | 3.9617 |  | 10.6
31 | 146.371 | 9.1396 | 9.1402 | 31.4611 | 31.4618 | 6.5833 | 1.4524 | 1 | 1 | 107 | 3.0316 |  | 9.7
31 | 124.725 | 9.2323 | 9.2344 | 31.3307 | 31.3293 | 6.5753 | 1.4714 | 2 | 2 | 108 | -99 | 9 |  | OXY: Redundant
31 | 124.322 | 9.271 | 9.2684 | 31.2762 | 31.2809 | 6.572 | 1.4721 | 3 | 3 | 109 | -99 | 9 |  | OXY: Redundant
31 | 124.035 | 9.2815 | 9.281 | 31.2623 | 31.2644 | 6.571 | 1.4743 | 4 | 4 | 110 | -99 | 9 |  | OXY: Redundant
31 | 124.822 | 9.2828 | 9.2856 | 31.2611 | 31.2588 | 6.5709 | 1.4791 | 5 | 5 | 111 | 3.1192 | 6 | 9.8
31 | 100.124 | 9.9255 | 9.9268 | 30.3774 | 30.3769 | 6.5153 | 1.6043 | 6 | 6 | 112 | 3.4546 |  | 10.4
31 | 76.317 | 10.1635 | 10.1625 | 30.0674 | 30.07 | 6.4943 | 1.6618 | 7 | 7 | 113 | 3.606 |  | 10.6
31 | 50.048 | 10.447 | 10.4455 | 29.6331 | 29.6372 | 6.4722 | 1.7281 | 8 | 8 | 114 | 3.7472 |  | 11.1
31 | 39.277 | 10.555 | 10.5548 | 29.4179 | 29.4203 | 6.4658 | 1.7509 | 9 | 9 | 115 | 3.8279 |  | 11.2
31 | 30.626 | 10.6227 | 10.6216 | 29.2954 | 29.2986 | 6.4614 | 1.7636 | 10 | 10 | 116 | 3.8944 |  | 11.2
31 | 19.203 | 12.0971 | 12.1312 | 27.3994 | 27.358 | 6.3353 | 2.0082 | 11 | 11 | 117 | 4.5377 |  | 12.7
31 | 11.246 | 12.2727 | 12.2701 | 27.1744 | 27.1798 | 6.3206 | 2.0312 | 12 | 12 | 118 | 4.6279 |  | 13
31 | 4.553 | 13.5399 | 13.4994 | 25.4901 | 25.5458 | 6.2195 | 2.2256 | 13 | 13 | 119 | 5.2491 |  | 14.6
31 | 1.864 | 16.0922 | 16.2503 | 21.979 | 21.8011 | 6.0315 | 2.6679 | 14 | 14 | 120 | 6.0655 |  | 16.7
39 | 172.975 | 9.1688 | 9.1695 | 31.4228 | 31.4237 | 6.5807 | 1.4505 | 1 | 1 | 121 | 3.0301 |  | 9.7
39 | 151.327 | 9.2023 | 9.2028 | 31.374 | 31.3751 | 6.5778 | 1.4623 | 2 | 2 | 122 | 3.0322 |  | 9.7
39 | 124.506 | 9.4554 | 9.4556 | 31.0082 | 31.0096 | 6.5565 | 1.5125 | 3 | 3 | 123 | 3.1717 |  | 9.9
39 | 102.815 | 9.715 | 9.7164 | 30.4745 | 30.4762 | 6.5415 | 1.5635 | 4 | 4 | 124 | 3.3051 |  | 10.6
39 | 75.587 | 9.7677 | 9.7673 | 30.1087 | 30.1102 | 6.5492 | 1.5848 | 5 | 5 | 125 | 3.3525 |  | 10.3
39 | 51.069 | 9.9759 | 9.9764 | 29.8712 | 29.8727 | 6.5292 | 1.6233 | 6 | 6 | 126 | 3.4971 | 6 | 10.4 | OXY: Poor replication but no evidence to support rejection
39 | 40.215 | 10.1362 | 10.1342 | 29.6049 | 29.6075 | 6.5174 | 1.6531 | 7 | 7 | 127 | 3.5198 |  | 10.6
39 | 29.882 | 10.2584 | 10.2559 | 29.2472 | 29.2463 | 6.5149 | 1.6959 | 8 | 8 | 128 | 3.657 |  | 10.7
39 | 19.981 | 11.4729 | 11.4802 | 27.6267 | 27.6227 | 6.4112 | 1.9018 | 9 | 9 | 129 | 4.0499 |  | 11.4
39 | 9.918 | 14.8044 | 14.8305 | 24.8454 | 24.8175 | 6.0835 | 2.3536 | 10 | 10 | 130 | 5.2376 |  | 15.2
39 | 5.611 | 17.8737 | 17.8677 | 21.964 | 21.9135 | 5.8226 | 2.9636 | 11 | 11 | 131 | 6.654 |  | 18.4
39 | 1.937 | 17.8392 | 17.8374 | 21.1989 | 21.2032 | 5.8534 | 3.0633 | 12 | 12 | 132 | 7.0236 |  | 18.5
43 | 312.485 | 8.9349 | 8.9349 | 31.0599 | 31.0605 | 6.6305 | 1.4105 | 1 | 1 | 133 | 3.0138 |  | 9.7
43 | 302.588 | 8.9279 | 8.929 | 31.0445 | 31.0457 | 6.6322 | 1.4171 | 2 | 2 | 134 | 2.9917 |  | 9.5
43 | 251.714 | 8.9271 | 8.9276 | 30.9606 | 30.9614 | 6.6359 | 1.445 | 3 | 3 | 135 | 3.0697 |  | 9.5
43 | 201.863 | 8.9053 | 8.9089 | 30.8334 | 30.836 | 6.6446 | 1.4814 | 4 | 4 | 136 | 3.1327 | 3 | 9.4 | OXY: Spigot dripping
43 | 175.631 | 8.8708 | 8.8707 | 30.7532 | 30.7539 | 6.6532 | 1.5102 | 5 | 5 | 137 | 3.2311 | 6 | 9.5
43 | 150.567 | 8.9481 | 8.9586 | 30.6376 | 30.6285 | 6.6467 | 1.531 | 6 | 6 | 138 | 3.2779 |  | 9.5
43 | 125.449 | 9.4818 | 9.4589 | 30.3921 | 30.3922 | 6.5787 | 1.5399 | 7 | 7 | 139 | 3.3136 |  | 9.9
43 | 101.115 | 9.7405 | 9.7408 | 30.1481 | 30.1493 | 6.5515 | 1.5419 | 8 | 8 | 140 | 3.2486 |  | 10.3
43 | 75.567 | 9.6155 | 9.6155 | 29.7962 | 29.7954 | 6.5845 | 1.6052 | 9 | 9 | 141 | 3.4412 |  | 10.1
43 | 49.92 | 9.6835 | 9.6824 | 29.4774 | 29.4782 | 6.5881 | 1.6702 | 10 | 10 | 142 | 3.6338 |  | 10.2
43 | 39.874 | 9.8431 | 9.8499 | 29.2835 | 29.2823 | 6.5731 | 1.6951 | 11 | 11 | 143 | 3.6648 |  | 10.4
43 | 19.743 | 10.7306 | 10.7458 | 28.1879 | 28.1833 | 6.4918 | 1.8289 | 12 | 12 | 144 | 4.0395 | 6 | 11.1
43 | 9.758 | 13.9092 | 13.8961 | 26.1406 | 26.1514 | 6.1471 | 2.1656 | 13 | 13 | 145 | 4.8886 |  | 14.5
43 | 4.993 | 17.3132 | 17.308 | 23.4842 | 23.4866 | 5.8333 | 2.7982 | 14 | 14 | 146 | 6.0849 |  | 17.7
43 | 2.381 | 17.4508 | 17.4526 | 23.4521 | 23.4528 | 5.8187 | 2.8572 | 15 | 15 | 147 | 6.2331 |  | 17.9
48 | 377.653 | 8.8629 | 8.8636 | 31.2166 | 31.2169 | 6.6345 | 1.3742 | 1 | 1 | 148 | 2.8564 |  | 9.3
48 | 349.674 | 8.8672 | 8.8677 | 31.1904 | 31.191 | 6.6349 | 1.3827 | 2 | 2 | 149 | 2.9305 |  | 9.3
48 | 299.458 | 8.9219 | 8.9224 | 31.0765 | 31.0772 | 6.6317 | 1.4279 | 3 | 3 | 150 | 2.9978 |  | 9.4
48 | 249.14 | 8.9221 | 8.9225 | 30.9638 | 30.9646 | 6.6365 | 1.4575 | 4 | 4 | 151 | 3.0983 | 3 | 9.4 | OXY: Wavy curve. Endpoint seems reasonable based on the regression analysis
48 | 199.644 | 8.8199 | 8.8203 | 30.8321 | 30.8329 | 6.6574 | 1.4795 | 5 | 5 | 152 | 3.1582 |  | 9.4
48 | 174.43 | 8.8532 | 8.8537 | 30.768 | 30.7689 | 6.6552 | 1.5007 | 6 | 6 | 153 | 3.2086 |  | 9.4
48 | 150.303 | 8.9621 | 8.9625 | 30.6335 | 30.6341 | 6.6448 | 1.5388 | 7 | 7 | 154 | 3.3083 | 6 | 9.5
48 | 124.611 | 9.0798 | 9.0795 | 30.3717 | 30.373 | 6.6386 | 1.5936 | 8 | 8 | 155 | 3.4634 |  | 9.6
48 | 99.851 | 9.0463 | 9.0466 | 30.1634 | 30.1639 | 6.6525 | 1.6336 | 9 | 9 | 156 | 3.5858 |  | 9.7
48 | 74.291 | 9.225 | 9.2251 | 29.8924 | 29.8932 | 6.6376 | 1.6412 | 10 | 10 | 157 | 3.6238 |  | 9.8
48 | 50.353 | 9.2819 | 9.2827 | 29.2896 | 29.2897 | 6.6551 | 1.7503 | 11 | 11 | 158 | 3.9259 |  | 9.9
48 | 40.012 | 9.4532 | 9.4547 | 29.0448 | 29.0447 | 6.6402 | 1.7605 | 12 | 12 | 159 | 3.9302 |  | 10.1
48 | 30.204 | 9.9264 | 9.9269 | 28.6638 | 28.6646 | 6.5871 | 1.7774 | 13 | 13 | 160 | 3.9606 |  | 10.5
48 | 20.241 | 10.3219 | 10.3332 | 28.3726 | 28.3655 | 6.5423 | 1.8077 | 14 | 14 | 161 | 4.0262 | 6 | 11.1
48 | 9.947 | 12.5761 | 12.6325 | 26.6196 | 26.6042 | 6.302 | 2.0474 | 15 | 15 | 162 | 4.529 |  | 13.1
48 | 5.392 | 16.3579 | 16.3455 | 24.0807 | 24.0901 | 5.9233 | 3.055 | 16 | 16 | 163 | 6.7795 |  | 16.9
48 | 2.106 | 18.055 | 17.9826 | 23.1719 | 23.2061 | 5.7602 | 2.8983 | 17 | 17 | 164 | 6.6209 |  | 17.9
51 | 346.355 | 8.9195 | 8.9201 | 31.0745 | 31.0753 | 6.6322 | 1.4294 | 1 | 1 | 165 | 3.0654 |  | 9.5
51 | 300.086 | 8.9094 | 8.9099 | 31.0178 | 31.0185 | 6.6361 | 1.4446 | 2 | 2 | 166 | 3.0963 |  | 9.5
51 | 250.907 | 8.8853 | 8.8853 | 30.9484 | 30.9491 | 6.6427 | 1.4617 | 3 | 3 | 167 | 3.148 |  | 9.5
51 | 200.21 | 8.8132 | 8.8136 | 30.842 | 30.8421 | 6.6579 | 1.4732 | 4 | 4 | 168 | 3.1329 | 6 | 9.3
51 | 174.815 | 8.9079 | 8.9086 | 30.7841 | 30.7846 | 6.6464 | 1.4989 | 5 | 5 | 169 | 3.2114 |  | 9.5
51 | 149.41 | 8.9349 | 8.936 | 30.6531 | 30.6541 | 6.648 | 1.5439 | 6 | 6 | 170 | 3.3296 |  | 9.4
51 | 124.158 | 9.1557 | 9.1562 | 30.5451 | 30.5461 | 6.62 | 1.5324 | 7 | 7 | 171 | 3.3063 |  | 9.7
51 | 100.165 | 9.2273 | 9.2276 | 30.2986 | 30.2994 | 6.6199 | 1.587 | 8 | 8 | 172 | 3.4315 |  | 9.8
51 | 75.204 | 8.972 | 8.9724 | 29.9023 | 29.9027 | 6.6748 | 1.6832 | 9 | 9 | 173 | 3.7512 |  | 9.5
51 | 49.992 | 9.114 | 9.1154 | 29.3744 | 29.3744 | 6.6764 | 1.7357 | 10 | 10 | 174 | 3.8944 |  | 9.7
51 | 39.629 | 9.3258 | 9.3255 | 29.2093 | 29.2106 | 6.652 | 1.7516 | 11 | 11 | 175 | 3.9145 |  | 9.9
51 | 30.513 | 9.6681 | 9.6683 | 28.8308 | 28.8319 | 6.6177 | 1.7644 | 12 | 12 | 176 | 3.9439 |  | 10.1
51 | 20.166 | 10.1415 | 10.1416 | 28.4508 | 28.4518 | 6.565 | 1.7897 | 13 | 13 | 177 | 4.0094 | 6 | 10.6
51 | 10.208 | 10.9504 | 10.9472 | 27.7424 | 27.7436 | 6.4792 | 1.849 | 14 | 14 | 178 | 4.0811 |  | 11.3
51 | 5.037 | 13.447 | 13.5184 | 26.0748 | 26.0423 | 6.2089 | 2.2588 | 15 | 15 | 179 | 4.6824 |  | 13.1
51 | 1.937 | 16.0186 | 16.012 | 24.3624 | 24.3678 | 5.9534 | 2.9151 | 16 | 16 | 180 | 6.7246 |  | 16.5
55 | 273.211 | 8.8383 | 8.8388 | 30.9992 | 31.0003 | 6.6475 | 1.4299 | 1 | 1 | 181 | 3.0503 |  | 9.5
55 | 249.878 | 8.8206 | 8.8202 | 30.9275 | 30.9302 | 6.6532 | 1.4475 | 2 | 2 | 182 | 3.0626 |  | 9.5
55 | 199.67 | 8.8236 | 8.8253 | 30.803 | 30.8038 | 6.6581 | 1.4934 | 3 | 3 | 183 | 3.2061 |  | 9.5
55 | 174.973 | 8.7697 | 8.77 | 30.6976 | 30.6985 | 6.6707 | 1.5508 | 4 | 4 | 184 | 3.3817 | 6 | 9.5
55 | 149.899 | 8.7899 | 8.7918 | 30.5622 | 30.563 | 6.6735 | 1.5922 | 5 | 5 | 185 | 3.4703 |  | 9.5
55 | 124.103 | 8.7834 | 8.7816 | 30.402 | 30.4022 | 6.6814 | 1.6506 | 6 | 6 | 186 | 3.663 |  | 9.4
55 | 99.944 | 8.8265 | 8.8268 | 30.2729 | 30.2736 | 6.6805 | 1.675 | 7 | 7 | 187 | 3.7391 |  | 9.5
55 | 73.771 | 8.9226 | 8.9231 | 29.888 | 29.8882 | 6.6828 | 1.6809 | 8 | 8 | 188 | 3.7312 |  | 9.5
55 | 49.566 | 9.2371 | 9.2375 | 29.4387 | 29.4397 | 6.6553 | 1.729 | 9 | 9 | 189 | 3.8531 |  | 9.9
55 | 39.971 | 9.4426 | 9.4429 | 29.1606 | 29.1616 | 6.6368 | 1.7669 | 10 | 10 | 190 | 3.9416 |  | 10
55 | 30.429 | 9.5906 | 9.59 | 29.0184 | 29.02 | 6.6211 | 1.7798 | 11 | 11 | 191 | 3.9844 |  | 10.1
55 | 20.12 | 10.2651 | 10.2678 | 28.4652 | 28.4645 | 6.5466 | 1.8308 | 12 | 12 | 192 | 4.0662 | 6 | 10.7
55 | 9.511 | 12.981 | 12.9697 | 26.5625 | 26.5735 | 6.2507 | 2.1208 | 13 | 13 | 193 | 4.5764 |  | 12.9
55 | 4.929 | 16.9705 | 17.0113 | 25.2125 | 25.2122 | 5.8118 | 2.869 | 14 | 14 | 194 | 6.3746 |  | 17.5
55 | 2.02 | 17.6449 | 17.6447 | 24.9506 | 24.9463 | 5.7446 | 2.9761 | 15 | 15 | 195 | 6.6241 |  | 18.1
58 | 314.138 | 8.8599 | 8.8607 | 31.0699 | 31.0706 | 6.6412 | 1.404 | 1 | 1 | 196 | 2.9848 |  | 9.4
58 | 300.365 | 8.8511 | 8.8516 | 31.0467 | 31.0472 | 6.6435 | 1.4294 | 2 | 2 | 197 | 2.9941 |  | 9.4
58 | 250.352 | 8.7898 | 8.7902 | 30.9396 | 30.9397 | 6.6572 | 1.4408 | 3 | 3 | 198 | 3.0202 |  | 9.3
58 | 200.582 | 8.8024 | 8.8031 | 30.8193 | 30.8202 | 6.6605 | 1.4384 | 4 | 4 | 199 | 2.9816 | 3 | 9.5 | OXY: Leaker, partial seal
58 | 174.692 | 8.768 | 8.7685 | 30.736 | 30.7365 | 6.6693 | 1.5028 | 5 | 5 | 200 | 3.1898 | 6 | 9.5
58 | 150.255 | 8.8106 | 8.8106 | 30.6655 | 30.6657 | 6.6659 | 1.549 | 6 | 6 | 201 | 3.3158 |  | 9.5
58 | 123.893 | 8.8653 | 8.8659 | 30.4865 | 30.4872 | 6.6655 | 1.6044 | 7 | 7 | 202 | 3.4848 |  | 9.6
58 | 100.47 | 8.9217 | 8.9221 | 30.3224 | 30.3233 | 6.6642 | 1.6194 | 8 | 8 | 203 | 3.5059 |  | 9.6
58 | 74.146 | 8.9691 | 8.9689 | 30.0073 | 30.008 | 6.6707 | 1.6578 | 9 | 9 | 204 | 3.6266 |  | 9.7
58 | 51.232 | 8.9412 | 8.9417 | 29.5335 | 29.5347 | 6.6953 | 1.6916 | 10 | 10 | 205 | 3.7206 |  | 9.7
58 | 40.522 | 9.2088 | 9.2099 | 29.2819 | 29.2826 | 6.6662 | 1.732 | 11 | 11 | 206 | 3.8477 | 6 | 10
58 | 29.582 | 9.4434 | 9.4453 | 29.0685 | 29.0683 | 6.6406 | 1.749 | 12 | 12 | 207 | 3.9533 |  | 10.3
58 | 20.723 | 10.3656 | 10.3646 | 28.3303 | 28.3297 | 6.5378 | 1.8131 | 13 | 13 | 208 | 4.0258 |  | 11.1
58 | 10.324 | 12.2835 | 12.2899 | 27.2171 | 27.2172 | 6.3175 | 2.1071 | 14 | 14 | 209 | 4.7682 | 3 | 12.9 | OXY: Bad curve. End point ambiguous
58 | 5.326 | 16.6704 | 16.707 | 25.1353 | 25.1286 | 5.8491 | 2.8771 | 15 | 15 | 210 | 6.3622 |  | 17.4
58 | 2.293 | 16.8583 | 16.8637 | 25.1132 | 25.1142 | 5.8282 | 2.9005 | 16 | 16 | 211 | 6.5386 |  | 17.8
60 | 188.384 | 8.7872 | 8.7876 | 30.7999 | 30.801 | 6.6636 | 1.4348 | 1 | 1 | 212 | 3.0139 |  | 9.5
60 | 174.448 | 8.7466 | 8.747 | 30.7065 | 30.7076 | 6.6737 | 1.497 | 2 | 2 | 213 | 3.1484 | 6 | 9.5
60 | 150.841 | 8.7367 | 8.7371 | 30.5768 | 30.5775 | 6.6808 | 1.5615 | 3 | 3 | 214 | 3.3861 |  | 9.4
60 | 126.432 | 8.7798 | 8.7805 | 30.381 | 30.3814 | 6.6828 | 1.6033 | 4 | 4 | 215 | 3.5095 | 3 | 9.5 | OXY: Leaky Niskin
60 | 100.837 | 8.8302 | 8.8291 | 30.2021 | 30.2024 | 6.683 | 1.6294 | 5 | 5 | 216 | 3.5671 |  | 9.6
60 | 75.278 | 8.7686 | 8.7697 | 29.835 | 29.8341 | 6.7082 | 1.6571 | 6 | 6 | 217 | 3.6471 | 6 | 9.5
60 | 50.207 | 9.5076 | 9.507 | 29.1678 | 29.1687 | 6.627 | 1.7809 | 7 | 7 | 218 | 3.9628 |  | 10.2
60 | 39.855 | 9.7677 | 9.7656 | 29.0212 | 29.0223 | 6.5951 | 1.8323 | 8 | 8 | 219 | 4.0851 |  | 10.3
60 | 30.034 | 9.8989 | 9.8988 | 28.8978 | 28.8992 | 6.5813 | 1.8522 | 9 | 9 | 220 | 4.1355 |  | 10.5
60 | 20.176 | 10.0597 | 10.0724 | 28.4226 | 28.4195 | 6.578 | 1.7966 | 10 | 10 | 221 | 4.0123 |  | 10.8
60 | 10.046 | 12.6884 | 12.6903 | 27.0125 | 27.0132 | 6.2716 | 2.1728 | 11 | 11 | 222 | 5.0912 |  | 13.7
60 | 5.104 | 17.3344 | 17.3323 | 24.8717 | 24.8739 | 5.7823 | 2.9852 | 12 | 12 | 223 | 6.5962 |  | 19
60 | 1.883 | 18.8345 | 18.8372 | 24.5968 | 24.5987 | 5.6261 | 2.9772 | 13 | 13 | 224 | 6.5461 |  | 19.4
63 | 165.711 | 8.7835 | 8.7841 | 30.7697 | 30.7709 | 6.6655 | 1.3915 | 1 | 1 | 225 | 2.8446 |  | 9.5
63 | 150.147 | 8.6905 | 8.6907 | 30.6116 | 30.6122 | 6.6862 | 1.454 | 2 | 2 | 226 | 3.03 | 4 | 9.3 | OXY: Very messy curve. No obvious flattening of curve or endpoint.
63 | 125.14 | 8.572 | 8.5724 | 30.4132 | 30.414 | 6.7127 | 1.6039 | 3 | 3 | 227 | 3.4877 |  | 9.2
63 | 100.461 | 8.6452 | 8.6457 | 30.2741 | 30.2749 | 6.7077 | 1.6473 | 4 | 4 | 228 | 3.6063 | 36 | 9.3 | OXY: Leaky Niskin. Reps ok
63 | 74.572 | 8.7961 | 8.7967 | 29.8893 | 29.8901 | 6.7017 | 1.6456 | 5 | 5 | 229 | 3.594 |  | 9.4
63 | 49.854 | 9.5734 | 9.5729 | 29.1423 | 29.1435 | 6.6184 | 1.7848 | 6 | 6 | 230 | 3.9765 |  | 10.1
63 | 39.712 | 9.6856 | 9.6848 | 28.9314 | 28.9318 | 6.6109 | 1.7904 | 7 | 7 | 231 | 3.9966 |  | 10.3
63 | 29.858 | 9.9426 | 9.9431 | 28.7879 | 28.7882 | 6.5796 | 1.8324 | 8 | 8 | 232 | 4.0475 |  | 10.4
63 | 19.777 | 10.5693 | 10.5684 | 28.4937 | 28.4948 | 6.502 | 1.929 | 9 | 9 | 233 | -99 | 5 | 11 | OXY: Chain too short. Air leak therefore sample not analyzed.
63 | 9.948 | 11.9794 | 11.974 | 27.69 | 27.6991 | 6.3396 | 2.1659 | 10 | 10 | 234 | 4.8232 |  | 12.3
63 | 5.096 | 15.8741 | 16.2171 | 25.844 | 25.7037 | 5.917 | 2.6481 | 11 | 11 | 235 | 5.7928 |  | 15.9
63 | 2.263 | 16.8657 | 16.9069 | 25.456 | 25.4456 | 5.8152 | 2.7795 | 12 | 12 | 236 | 6.0838 |  | 17.2
66 | 354.123 | 8.7586 | 8.7592 | 30.6877 | 30.6884 | 6.6727 | 1.4363 | 1 | 1 | 237 | 3.0757 |  | 9.3
66 | 300.562 | 8.7466 | 8.7471 | 30.673 | 30.6737 | 6.6752 | 1.4849 | 2 | 2 | 238 | 3.1963 |  | 9.4
66 | 250.34 | 8.7173 | 8.7178 | 30.6314 | 30.632 | 6.6814 | 1.5218 | 3 | 3 | 239 | 3.2967 |  | 9.3
66 | 200.213 | 8.6517 | 8.6522 | 30.5713 | 30.5721 | 6.6938 | 1.5562 | 4 | 4 | 240 | -99 | 5 | 9.3 | OXY: Chain too short. Stopper pulled out therefore not analyzed
66 | 174.763 | 8.6223 | 8.6214 | 30.5295 | 30.5302 | 6.7001 | 1.5818 | 5 | 5 | 241 | 3.4611 |  | 9.2
66 | 150.148 | 8.7411 | 8.7412 | 30.4408 | 30.4417 | 6.6861 | 1.6176 | 6 | 6 | 242 | 3.5534 | 6 | 9.3
66 | 122.893 | 8.8498 | 8.8504 | 30.3499 | 30.351 | 6.6737 | 1.6163 | 7 | 7 | 243 | 3.5377 |  | 9.5
66 | 103.164 | 8.7523 | 8.7535 | 30.2171 | 30.2178 | 6.6941 | 1.6526 | 8 | 8 | 244 | 3.6372 | 3 | 9.4 | OXY: Messy curve. Endpoint ambiguous but is in most reasonable position.
66 | 75.838 | 8.8114 | 8.812 | 29.7841 | 29.7845 | 6.704 | 1.7064 | 9 | 9 | 245 | 3.7645 |  | 9.5
66 | 49.812 | 9.2143 | 9.2144 | 29.4029 | 29.4037 | 6.6602 | 1.7322 | 10 | 10 | 246 | 3.8307 |  | 9.8
66 | 40.374 | 9.3682 | 9.3715 | 29.1414 | 29.1404 | 6.6486 | 1.7469 | 11 | 11 | 247 | 3.8745 |  | 10
66 | 29.848 | 9.5816 | 9.5734 | 28.8606 | 28.8693 | 6.6292 | 1.7465 | 12 | 12 | 248 | 3.928 | 36 | 10.1 | OXY: Replicate outliers
66 | 19.457 | 10.2923 | 10.2927 | 28.316 | 28.3139 | 6.5489 | 1.8276 | 13 | 13 | 249 | 4.0272 |  | 10.7
66 | 10.198 | 12.7353 | 12.725 | 27.0895 | 27.0857 | 6.2624 | 2.202 | 14 | 14 | 250 | 4.8463 |  | 13.1
66 | 5.026 | 17.4526 | 17.3044 | 25.2677 | 25.3183 | 5.7552 | 2.8623 | 15 | 15 | 251 | 6.0275 |  | 17.1
66 | 1.822 | 18.0802 | 18.0731 | 24.9738 | 24.9786 | 5.6955 | 2.9248 | 16 | 16 | 252 | 6.4687 |  | 18.5
68 | 310.112 | 8.7381 | 8.7385 | 30.6643 | 30.665 | 6.6768 | 1.4786 | 1 | 1 | 253 | 3.1882 |  | 9.3
68 | 249.624 | 8.6823 | 8.6825 | 30.6078 | 30.6085 | 6.6876 | 1.5291 | 2 | 2 | 254 | 3.3223 |  | 9.3
68 | 199.91 | 8.5651 | 8.5656 | 30.5052 | 30.5062 | 6.7097 | 1.5596 | 3 | 3 | 255 | 3.405 |  | 9.2
68 | 175.23 | 8.5543 | 8.5543 | 30.4545 | 30.4545 | 6.7136 | 1.5767 | 4 | 4 | 256 | 3.464 | 3 | 9.1 | OXY: Leaky Niskin
68 | 150.113 | 8.4942 | 8.4945 | 30.3917 | 30.3923 | 6.7254 | 1.6065 | 5 | 5 | 257 | 3.5311 |  | 9
68 | 125.05 | 8.5582 | 8.5581 | 30.2277 | 30.2283 | 6.7229 | 1.6401 | 6 | 6 | 258 | 3.622 | 6 | 9.1
68 | 99.606 | 8.6513 | 8.6508 | 29.9887 | 29.9897 | 6.7192 | 1.6675 | 7 | 7 | 259 | 3.692 |  | 9.1
68 | 74.817 | 8.844 | 8.8433 | 29.7282 | 29.7289 | 6.7015 | 1.704 | 8 | 8 | 260 | 3.7766 |  | 9.3
68 | 49.98 | 9.7767 | 9.7699 | 29.0657 | 29.0668 | 6.5919 | 1.8093 | 9 | 9 | 261 | 3.996 |  | 10
68 | 39.678 | 10.054 | 10.0537 | 28.9572 | 28.9583 | 6.5563 | 1.8631 | 10 | 10 | 262 | 4.1709 | 6 | 10.4
68 | 30.04 | 10.0916 | 10.0921 | 28.9421 | 28.943 | 6.5515 | 1.8721 | 11 | 11 | 263 | 4.1822 |  | 10.5
68 | 19.955 | 10.4122 | 10.4126 | 28.7135 | 28.7106 | 6.5152 | 1.9226 | 12 | 12 | 264 | 4.293 |  | 10.7
68 | 9.833 | 11.7254 | 11.7256 | 27.9063 | 27.9075 | 6.3653 | 2.1368 | 13 | 13 | 265 | 4.8256 |  | 12
68 | 4.74 | 13.0243 | 13.0006 | 27.3536 | 27.3626 | 6.2142 | 2.311 | 14 | 14 | 266 | 5.1153 |  | 13
68 | 2.443 | 13.732 | 13.676 | 27.086 | 27.1057 | 6.1337 | 2.3796 | 15 | 15 | 267 | 5.5121 |  | 14.3
70 | 151.125 | 8.6454 | 8.646 | 30.5921 | 30.5933 | 6.6939 | 1.5284 | 1 | 1 | 268 | 3.4104 |  | 9.2
70 | 124.847 | 8.5776 | 8.578 | 30.4889 | 30.4897 | 6.7085 | 1.5751 | 2 | 2 | 269 | 3.4375 |  | 9.1
70 | 100.123 | 8.5525 | 8.5526 | 30.2439 | 30.2432 | 6.723 | 1.6224 | 3 | 3 | 270 | 3.5744 |  | 9
70 | 74.475 | 8.7702 | 8.7707 | 29.9072 | 29.9078 | 6.7048 | 1.6756 | 4 | 4 | 271 | 3.7148 |  | 9.2
70 | 49.763 | 9.1413 | 9.1434 | 29.621 | 29.6217 | 6.6617 | 1.729 | 5 | 5 | 272 | 3.8284 |  | 9.5
70 | 39.528 | 9.6474 | 9.6517 | 29.2973 | 29.2953 | 6.601 | 1.7952 | 6 | 6 | 273 | 3.9849 |  | 10
70 | 29.066 | 9.989 | 9.9913 | 29.0548 | 29.0545 | 6.5616 | 1.8486 | 7 | 7 | 274 | 4.1366 |  | 10.5
70 | 19.527 | 10.149 | 10.1521 | 28.9406 | 28.9386 | 6.5434 | 1.8755 | 8 | 8 | 275 | 4.2462 | 6 | 10.6
70 | 10.018 | 10.4517 | 10.4536 | 28.773 | 28.7737 | 6.5071 | 1.9244 | 9 | 9 | 276 | 4.3296 |  | 10.9
70 | 5.031 | 10.8644 | 10.8644 | 28.365 | 28.3543 | 6.4657 | 1.9873 | 10 | 10 | 277 | 4.544 |  | 11.4
70 | 2.444 | 11.0443 | 11.0428 | 28.2071 | 28.2105 | 6.447 | 2.025 | 11 | 11 | 278 | 4.5898 |  | 11.5

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers