## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 30 Nov 2020 | ### Added HPLC data. S.H. |
| 21 March 2014 | Corrections made to Extracted Chlorophyll & Phaeo-Pigment data in analysis spreadsheet and in CHE files. GG |

## PROCESSING NOTES

Cruise: 2011-11

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: 28 November 2011 – 5 December 2011

Processed by: Germaine Gatien

Date of Processing: 9 March 2012 – 16 March 2012

Number of original HEX files: 81

Number of CTD files:  81

Number of CHE files: 21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0997), a Wet Labs Eco-AFL/FL Fluorometer (#2215), a PAR sensor (#4601), Surface PAR (#16504), a pH sensor (#0692) and an altimeter.

The deck unit was an SBE 11+ (#0425). Seasave V7 was used.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The Rosette and CTD logs were in excellent order with a list of all equipment and clear notes about problems encountered.

The event number in the log should be used for naming files even if CTD casts do not have consecutive numbers. It does not matter if some event numbers are used for activities other than CTD casts. The file names for the final 4 CTD casts were changed to 2011-11-0080, -0081, -0084 and -0085. There is a note of explanation in the file headers.

The SBE data are considered of lower quality than normal because 24 scans were averaged in acquisition due to an improper setting in the configuration file. These averages may contain spikes that would normally be removed in processing and might be systematic in nature. Fine-tuning of alignment settings and CELLTM correction are also less sensitive with smoothed data. Unfortunately the same error was made on at least 4 other cruises using the same equipment, so we can not rely on settings determined from those data sets. Given that sampling was in an area that is not usually affected much by shed wakes and the CTD is believed to have worked well with few spikes, it is likely the errors caused by this are small.

The salinity samples were analyzed within 39 days of collection. The comparison of CTD salinity and bottles suggests that both the primary and secondary channels were very close to the bottles.

The Tau correction of the SBE DO data could not be applied leading to values that are too high in regions of high SBE DO gradient. This error is likely up to 0.5mL/L in the strongest gradients.

The tests for accuracy of Oxygen:Dissolved:SBE data are less reliable than usual, but the data are considered, very roughly, to be:

high in regions of high DO gradient

±0.5mL/L from 0 to 20db

±0.2mL/L from 20 to 150db

±0.06mL/L below 150db

Data from the SBE pH sensor were collected but are not to be archived at this time. Files were provided to the chief scientist with the pH data included.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names had a syntax error which was fixed.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. The DO sensor is listed as #97 but is believed to be #997. There was a cruise report available; no problems are noted concerning CTD data.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. In future references to these files, the * stands for the date.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. Since the last factory calibration of most of the sensors there have been 3 cruises that have been processed, 2011-60, 2011-63 and 2011-10.

The calibration constants were checked for all instruments. There were 2 errors in the con files:

- The dissolved oxygen sensor serial number was changed from #0097 to #0997.

- The calibration for the Wetlabs Eco fluorometer could not be confirmed. The parameters entered are close to those on the characterization sheet from the factory, but a field test was done just before cruise 2011-60 in August and no record is available for that, but it looks reasonable

- The most serious error is that “Scans to average” was set to 24. This should be set to 1. There is no way to fix this and acquiring metre-averaged data only.  Single bad points will be included in the average since there is no opportunity to remove them first and fine-tuning of settings will be limited or impossible. (This problem also occurred during 2011-60, 2011-63 and 2011-10.)

A configuration file used at sea was corrected and saved as 2011-11-ctd.xmlcon.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

There were several occasions when bottles were fired but no sample number assigned. Those were removed from the ADDSAMP file.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-11-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. A calculation was made of Ф using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees. Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.4769 and B = -0.2546 with an R2 value of 0.9971. The initial estimates of M and B were used to calculate a difference between each point and the fit:
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. The fit with all unflagged bottles was used to update the value of M and B. When removing a little more data has little effect on the fit, it is judged that a reasonable value has been found unless, in so doing, a whole class of points has been removed such as all high values, or all values from late in the cruise. The following table shows that results:

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4348 | -0.2110 | 0.4348 | -0.4853 |  |
| 211 | all unflagged bottles | 0.4785 | -0.2317 | 0.4785 | -0.4842 | 0.9975 |
| 207 | excl. outliers diff>0.03 | 0.4790 | -0.2317 | 0.4790 | -0.4837 | 0.9989 |
| 191 | excl. outliers diff>0.01 | 0.4794 | -0.2318 | 0.4794 | -0.4835 | 0.9997 |
| 185 | excl. outliers diff>0.008 | 0.4794 | -0.2318 | 0.4794 | -0.4835 | 0.9997 |
| 175 | excl. outliers diff>0.006 | 0.4797 | -0.2324 | 0.4797 | -0.4849 | 0.9998 |
| 160 | excl. outliers diff>0.005 | 0.4795 | -0.2325 | 0.4795 | -0.4845 | 0.9998 |
| 147 | excl. outliers diff>0.004 | 0.4792 | -0.2322 | 0.4792 | -0.4849 | 0.9999 |
| 112 | excl. outliers diff>0.003 | 0.4794 | -0.2327 | 0.4794 | -0.4849 | 0.9999 |

The fit excluding differences >0.006 still has a lot of bottles and the higher end of the range is still well represented. The single anoxic value has been removed, but there are concerns about including that anyway because the SBE sensor is slow to equilibrate to those conditions. Removing more data has little effect on the fit. To see if there is any time-dependence or range-dependence the cruise was divided into two sections including casts 2-39 (Juan de Fuca and Haro Straits and the southern Strait of Georgia) and cast #43-75 in the central and northern Strait of Georgia.) The following table shows a small variation, but given the limited bottles no attempt will be made to do a time-dependent calibration. It is more likely a matter of range and equilibration than temporal drift.

| excluding diff from fit >0.006 | m | b | Soc | Voffset | R2 |
| --- | --- | --- | --- | --- | --- |
| Southern SoG +JdF (not SI) | 0.4835 | -0.2394 | 0.4835 | -0.4951 | 0.9997 |
| Northern SoG | 0.4780 | -0.2301 | 0.4780 | -0.4814 | 0.9998 |
| All areas | 0.4797 | -0.2324 | 0.4797 | -0.4845 | 0.9998 |

A choice of 0.006 as a cutoff difference looks reasonable. These values are quite different from those selected for this sensor for the summer cruises, but similar to that seen for 2011-76.

| Bottles Used | Cruise | m | b | Soc | Voffset | R2 | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 21 | 2011-60 | 0.4493 | -0.2214 | 0.4493 | -0.4928 | 0.9991 | few bottles |
| 93 | 2011-63 | 0.4474 | -0.2213 | 0.4474 | -0.4946 | 0.9991 | problems with bottles |
| 241 | 2011-10 | 0.4417 | -0.1979 | 0.4417 | -0.4480 | 0.9989 |  |
| 17 | 2011-76 | 0.4803 | -0.2259 | 0.4803 | -0.4703 | 0.9990 | Highly variable results |
| 175 | 2011-11 | 0.4797 | -0.2324 | 0.4797 | -0.4845 | 0.9998 |  |

While it is unclear if the quality of the SBE data justifies choosing a fit so different from those seen in September, the COMPARE stage will give a 2nd chance to assess how well it works. The summer fits may be the suspect ones given that high surface gradients mean more data are rejected and the average data are more likely to have spikes in those conditions. The temperature ranges for many of these casts are low which may also affect the DO fits.

The configuration files were updated with the new parameters Soc and Voffset and saved with name 2011-11-ctd-new.xmlcon.

##### Hysteresis Study

This sensor has been recalibrated since the last hysteresis checks were done. There was no deep sampling during this cruise, so hysteresis is not going to be a problem and there is not enough deep data to enable reliable tests to be done to fine-tune H1, H3 and E.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters.

The files were put through CLEAN to create BOT files.

A track plot and header check turned up no problems.

There were no negative or off-scale fluorescence values.

Temperature and salinity were plotted for all BOT files. No outliers were noted which is not as meaningful as usual given the averaging on acquisition.

The addsamp.csv file prepared in the DO calibration step is already ordered on sample #s. It was converted to CST files to be used as a framework for the bottle files. It was also used to create SAM files with bottle #s and bottle positions. The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets was examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-11-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2011-11chl*.xls which includes an analysis of variability. A simplified version of the spreadsheet was prepared and saved as 2011-11chl.csv which was then converted to individual CHL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-11nuts*.xls which included a report on precision.  The file was simplified, reordered on sample numbers and the file was saved as 2011-11-nuts.csv. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-11oxy*.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified; the file was then saved as 2011-11oxy.csv. That file was converted into individual *.ADD files.

SALINITY

Salinity analysis was done at IOS and saved in file 2011-11sal*.xls. The file was simplified, and saved as 2011-11sal.csv. There were no duplicate samples. The file was then converted to individual SAL files.

Note: The rosette cast at station ADCP was named #18 on the rosette log, but should be #19. This was fixed in the nutrient, DO and salinity files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps. After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, and SORT to get the files in bottle number order.

##### Compare

Salinity

The salinity samples were analyzed on January 11, 2012, 39 days after the last sample was collected.

Compare was run with pressure as reference channel. There were 2 major outliers (differences >0.04):

- Cast #1 – sample #2– The bottle value is higher than the CTD salinity by 0.06. There is no evidence of a problem with the sample or with the CTD data. A “3” flag was attached with a comment indicating it may be a problem.

- Cast #75 – sample #273– This bottle is higher than the CTD by 0.05, but the standard deviation is high in the CTD salinity, so the bottle is likely fine. A note was added to the header about this.

Next, two samples that had been flagged by the analyst (due to salt being seen on the cap) were examined:

- Cast #66 – sample #242 – salt on cap; higher than CTD by 0.005 – comparison comment was added but no change to flag.

- Cast #70 – sample #258 – salt on cap; bottle value looks ok in comparison to CTD salinity– comparison comment was added but no change made to flag.

When the 2 major outliers plus 3 minor ones (differences >0.004) are excluded the primary salinity is found to be low by an average of 0.0005 and the secondary by 0.0003.

There is some evidence of time dependence but it looks the same for both channels. This is often noted from this project, probably because the different parts of the region vary in bottom depth and salinity gradients. The most reliable casts for calibration purposes are likely those in the central basin of the Strait of Georgia and that is where the differences are closest to zero.

For details see file 2011-11-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. There was a lot of scatter in the fits, but the outliers were not severe. The largest outliers were all near the surface except for a sample from just above the hypoxic layer in Saanich Inlet; that is likely due to the SBE:DO sensor being a little slow to equilibrate after sampling very low DO water.  The samples that had been flagged by the analyst were examined in COMPARE. The analyst was consulted and some changes made to flags/comments:

Cast #1 – Sample #3 – duplicates were outilers. The 1st value looks better in COMPARE. The 2nd sample was rejected.

Cast #7 – Sample #35 – looks slightly high, but at the surface the SBE DO is not that reliable

Cast #13 – Sample #54 – looks ok in COMPARE

Cast #13 – Sample #58 – probably ok, hard to judge

Cast #19 – Sample #79 – looks ok

Cast #32 – Sample #109 – looks ok

Cast #48 – Sample #155 – probably ok, hard to judge

Cast #48 – Sample #163 – looks a bit off, but hard to judge

Cast #48 – Sample #165 – outlier in COMPARE – changed flag to “4” and amended comment.

Cast #60 – Sample #225 – probably ok

Cast #75 – Sample #277 – looks slightly low in COMPARE; added that to comment

For details see file 2011-11-dox-comp1*.xls.

Extracted Chlorophyll

COMPARE was run with pressure as the reference channel.

For CHL<0.4ug/L the fluorescence was higher than the CHL with the ratio of CTD fluorescence to extracted chlorophyll samples ranging from ~15 at the lowest values and being about 1 above 0.4, but there were very few points >0.4.

The MERGE steps were rerun so that added flags and comments were captured.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-11-ctd-new.con.

A few casts were examined and all expected channels are present. The conductivity channels track reasonably well during downcasts; while a little further apart during upcasts, they are not very far apart. There appear to be a few more spikes in the primary channel. The two temperature traces track well especially during the downcasts. Upcast and downcast traces are very different near the mouth of Juan de Fuca Strait with a large vertical offset.

The fluorescence looks reasonable with the usual vertical offset but values are often too low to distinguish from the noise level.

The DO voltage looks as expected with a vertical offset; the pH trace is similar to the DO voltage..

PAR, SPAR and transmissivity look normal. The altimetry is very noisy near the mouth of Juan de Fuca, presumably due to the weights being on; elsewhere the signal looks useful.

##### WILDEDIT

Program WILDEDIT is usually run to remove spikes from the pressure, conductivity & temperature channels. Because the data are already averaged, this step will be skipped.

##### ALIGN DO

Tests to determine the optimal alignment of the DO voltage do not offer clear results. It is very hard to judge in these data because of the lack of detail, but a setting of +4.5s looks reasonable.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

On 3 previous cruises using these sensors and for which the 24 scans were averaged on acquisition, tests to determine the best settings for CELLTM were inconclusive. The goal is to make upcasts look closer to downcasts on a T-S surface. There is little detail in the data to aid the judgment, but all choices looked better than no application, with the choice of (α = 0.0245, β=9.5) seeming best overall for both channels on those cruises and looks reasonable for this cruise.

CELLTM was run on all casts using setting (α = 0.0245, β=9.5) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The Tau correction feature was turned off since it caused some bad data that prevented conversion to IOS headers, and is unlikely to have any good effect on averaged data. The first attempt to convert these files failed due to off-scale dissolved oxygen values. When the TAU correction was turned off, the conversion worked fine. This did not happen for other cruises using these sensors and with the 24 scans averaged, but perhaps the small temperature range leads to over-interpretation of spikes. Whatever the reason, DERIVE was rerun with the TAU correction turned off.

on 2 casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The two deep casts were plotted to check for agreement between the pairs of T and C sensors. They were compared with 2 casts from 2 earlier cruises when the same sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-60-0013 | 320 | +0.0005 | -0.00004 | -0.0008 | High, F. Steady |
| 2011-60-0015 | 300 | +0.0007 | -0.00003 | -0.0009 | High, F. Steady |
| 2011-10-0047 | 290 | +0.0010 | +0.0001 | +0.0006 | High, Steady |
| 2011-10-0053 | 380 | +0.0009 | +0.00012 | +0.0003 | High, Steady |
| 2011-11-0048 | 360 | +0.0007 | +0.00009 | +0.0003 | High, Steady |
| 2011-11-0053 | 360 | +0.00075 | +0.00008 | +0.0002 | High, Steady |

The differences are all small. The differences in salinity are in agreement with the results of COMPARE when the 2 channels differed by an average of 0.0002.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Header Check was run and the minimum fluorescence value was -0.061mg/m^3, so CALIBRATE will be used to add that amount so that there are no negative values. This is the same value used for 2011-10. The bottle files had no negative values but the correction will be applied to those as well. There was an off-scale fluorescence value for cast #37 but it occurred at the surface after the pumps were turned off and dissolved oxygen values were bad there too.

Surface check was run and shows an average surface pressure for the cruise was 2.5db which is slightly deeper than usual for the Vector, but given that the data were averaged in acquisition it looks about right.

The cross-reference check was compared with the log book. 4 casts had been saved at sea with file names that were not consistent with the event number in the log book. It is usual to use event number, not consecutive CTD #, so the file names were changed to match the log event number.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. The values for a selection of casts were checked and the only problems found were casts near the mouth of Juan de Fuca Strait when the weights were on. Any readings of ~6.6 were checked and the header entries were removed for casts 6, 7, 8 and 11 where the entry is likely wrong. The header was also removed from the bottle file for cast #7.

The Water Depth headers have some readings that differ from the log book. In some cases it had not been changed since the previous cast, and for others there was a change made to the log entry, but not the header. Each case was examined carefully and changes were made to the header entry in casts #4, 11, 35, 42, 49, 50 and 80. The bottle file depths were ok.

##### 13. Shift

Fluorescence

Fluorometers usually require alignment, either to remove the effects of pumping for SeaPoint sensors or to correct for slow response time in ECO sensors. A number of cruises using ECO sensor (#2216) have been processed and it was found that an advance of ~2s suited those data. The 2011-11 data are from a different sensor (#2215) which has only been used for 3 previous cruises and just as for this cruise, the data had been averaged in acquisition.

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. Usually an advance of +48 records (2s) is found appropriate. Given the averaging 2s is equivalent to +2 records. Tests using that setting show the resulting offset in fluorescence sometimes seemed too high, sometimes too low, but overall, the results look reasonable. Low fluorescence values make the judgment even more difficult. SHIFT was applied with a +2 records setting.

Conductivity

The pre-averaging makes the alignment more difficult to assess. During 2011-60 tests were run using values from -0.1 records to +0.1 and none looked any better than the original. For 2011-63, 2011-10 and this cruise it was hard to find suitable features to aid this assessment. No shift was applied.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked and no further alignment was found necessary. SHIFT was not run on DO.

pH

During 2011-63 a setting of +4 records was used for this data. As for fluorescence it is difficult to determine the best choice. SHIFT was applied with a +4 records setting.

##### 14. DELETE

The following DELETE parameters were used (adjustments were made to the usual settings to allow for the fact the data are already averaged):

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

**Pressure filtered over 15 points**NOT APPLIED

Swells deleted. Warning message if pressure difference of 2.00

**Drop rates <   0.30m/s (calculated over 11 points) **NOT APPLIED.

**Drop rate applies in the range:  10db to 10db less than the maximum pressure

**NOT APPLIED

Sample interval:  taken from header

COMMENTS ON WARNINGS: All warnings pertained to upcast sections of casts except for cast #6 which had 1 warning that there were differences of >2db between successive records during the downcast section. This is not too surprising since the descent rate was high at that level and the data have been averaged over 24 scans, but it is surprising that the average descent rate was >1.5m/s. None of these warnings indicate any problem in how DELETE worked, but might suggest there were pressure spikes in the average.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  During 2011-60 in August there were only 5 bottles with a lot of scatter, but both CTD salinity channels appeared to be close to the bottles. There were no bottles from 2011-63 and there were many flagged bottles and outliers for 2011-10.

2. Dissolved Oxygen: The DO sensor was repaired and recalibrated in April 2011. During 2011-60 and -2011-63 there were too many problems with the comparison with bottles to justify a 2nd recalibration. For 2011-10 the average offset was 0.003mL/L and no 2nd recalibration was applied.

3. Pressure: The sensor was recalibrated in April 2011. Because all data were averaged on acquisition surface tests are not very useful, but no problem has been detected.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data in the central part of the Strait of Georgia fell within the ranges. At the northern end of the Strait, near Texada Island, salinity was low at mid-depths in areas of very low salinity gradient, so this may be a result of strong vertical mixing. In the most southerly part of the Strait there were also observations of low temperatures at mid-depths and low salinity near the bottom; again these casts showed strong vertical mixing. In Juan de Fuca Strait there was low salinity near the bottom of many casts and very low salinity near the mouth. Excursions of this sort are not unexpected in these waters where the 3-standard deviation climatology is too rigid a standard. There is no evidence of CTD calibration problems.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

Both T/C sensor pairs produce salinity that is very close to the bottles. A decision was made to select the secondary T and S for archiving because they have fewer spikes than the primary channels.

CTDEDIT was used to edit the secondary salinity very lightly, remove surface and bottom records and a few records obviously corrupted by shed wakes (mostly near the bottom). No casts needed heavy editing and 19 casts needed no editing.

##### Initial Recalibration

File 2011-11-recal1.ccf was prepared to add 0.061mg/m^3 to the fluorescence to avoid negative values in the files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response-time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are compared to bottle data from the same pressure. For this cruise, the problem of data smoothed during acquisition means the comparison is unlikely to be sensitive enough to justify further recalibration, but is still worth doing to get some idea of the accuracy of the SBE DO data.

The downcast files are usually averaged in 0.5db bins, but in this case the full files were thinned to the levels at which bottles are usually fired. COMPARE was then run to compare the CTD and bottle data. The comparison has many outliers, mostly ones with the CTD higher than the bottles. When differences are plotted against pressure and cases where differences are >0.1mL/L are excluded the CTD DO is high by an average of 0.03mL/L with a standard deviation of 0.06. Plots against file pair number show many more outliers in the Strait of Georgia, the most severe in the southern part. Examination of the profiles from each bottle cast show that the outliers are all from depths with high DO gradient. In the southern Strait of Georgia these are in the top 10 to 50m, but in the northern SoG there are higher gradients around 125m. This problem is likely due to having to turn off the Tau correction when the DO data were derived.

While it does appear that the SBE DO values are too high where DO gradients are high, any recalibration scheme will improve some areas but make others worse. Below 150db no correction appears appropriate as the bottles and CTD are very close. For other casts a pressure-dependent correction looks possible, but the depth and extent of the high-gradient zones varies significantly from cast to cast, so that some surface data would be made worse, while other data would be improved. There are only bottles for a third of the casts, so individual choices would depend on the SBE profile. There was also active mixing going on at many casts, so there could be significant differences that are real. Given the reduced quality of the SBE data, a cast-by-cast recalibration is not justified. A note will be added to the headers to suggest that the DO values are too high in high-gradient regions.

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD and saved. Since there was only an ECO fluorometer which does not need filtering, the usual second set was not prepared. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files. Special note was made that the data is of lower quality and that most bins are either empty or contain only 1 point.

No filtering is required for the ECO fluorescence.

##### BIN AVERAGE of CTD files

While the data were average over 24 scans, the descent rate is often less than 1m/s so Bin Average was still run on the COR1 files using the following settings:

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, and to add the following comments:

Data Processing Notes:

----------------------

The SBE data are considered of lower quality than normal because 24 scans

were averaged in acquisition. These averages may contain spikes that

would normally be removed in processing and might be systematic in nature.

Fine-tuning of alignment settings and CELLTM correction are also less

sensitive with smoothed data. Unfortunately the same error was made on

several other cruises using the same equipment, so we can not rely on

settings determined from those data sets. Given the area is in protected

waters and the CTD is believed to have worked well with few spikes, it is

likely the errors caused by this problem are small.

Fluorescence, PAR and Transmissivity data are nominal and unedited.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Tau correction of the SBE DO data could not be applied leading to values

that are too high in regions of high SBE DO gradient. This error is likely

up to 0.5mL/L in the strongest gradients, but much lower for most casts.

The tests for accuracy of Oxygen:Dissolved:SBE data are less reliable than

usual, but the data are considered, very roughly, to be:

high in regions of high DO gradient

±0.5mL/L from 0 to 20db

±0.2mL/L from 20 to 150db

±0.06mL/L below 150db

For details on the processing see processing report: 2011-11-proc.doc.

A second set of files was produced with pH:SBE channel included. These will not be archived.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

Profile and T-S plots were examined and no problems found.

The track plot looks ok.

Special files were prepared that included the pH channel (except for cast #72 when the cap was left on) and were given extension "ctdph

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values in most of the Strait of Georgia and the eastern part of Juan de Fuca Strait were between 80% and 100%, mostly close to 90%. Values between 90% and 100% were found in the western half of Juan de Fuca Strait. The only values <70% were seen in Nodales Channel and Frederick Arm. Those values look reasonable given the season and the strong vertical mixing seen at many casts.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, SBE:pH, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data. Those files were named *.CHE.

A header check was run on the final files and no problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Data from the CHE files were exported to spreadsheet 2011-11-bottles-final.xls; no errors were found.

Standards check was run repeatedly on all files until all errors were found and fixed except for the format for Temperature:Draw that was deliberately given a non-standard format.

Special files were prepared that included the pH channel; they were named with extension “cheph”.

##### 26. Producing final files

The CHE and CHEPH files for cast #1 were edited to remove bottles for which there was no sampling. CLEAN was run to fix the headers.

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (Notes from log book and sampling notes)

1. No samples from bottles #10 – 18.

6. Weights on – interfered with altimetry.

12. Weights off

72. Cap left on pH sensor.

80-85 – File names used at sea used consecutive CTD # but should use event #.

CRUISE SUMMARY - CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 4Aug2011 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 23Apr2011 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2215 | 4Aug2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |
| pH | 392 | 29Dec0210 | Factory |  |  |
| Altimeter | 2013380 |  |  |  |  |