## comparison-0997
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp # | [6 empty] | 0.4348 | -0.21101 | 0.4348 | -0.4853
48 | 353.272 | 9.0303 | 9.0307 | 31.1644 | 31.1649 | 6.612 | 1.2361 | 2 | 2 | 153 | 7.0176 | 0.35981 | 2.525 |  | 1.6807e-05 | 0.4785 | -0.2317 |  |  |  | Summary of Soc Voffset including the original values in the factory calibration
43 | 100.438 | 9.2773 | 9.2777 | 30.3511 | 30.3514 | 6.6104 | 1.5583 | 8 | 8 | 143 | 6.796 | 0.51398 | 3.493 |  | 1.8813e-05 | [7 empty] | m | b | Soc | Voffset | R2
43 | 50.274 | 9.3636 | 9.364 | 29.9248 | 29.9251 | 6.6159 | 1.5455 | 10 | 10 | 145 | 6.7593 | 0.50775 | 3.432 |  | 5.9581e-05 | [5 empty] | Bottles used | Original | 0.4348 | -0.21101 | 0.4348 | -0.4853 |  |  |  | 2.956
66 | 174.996 | 9.1319 | 9.1318 | 30.7001 | 30.7015 | 6.6169 | 1.3506 | 5 | 5 | 245 | 6.866 | 0.41451 | 2.846 |  | 6.5143e-05 | [5 empty] | 211 | all unflagged bottles | 0.4785 | -0.2317 | 0.4785 | -0.48422 | 0.9975 |  |  | 3.2565
51 | 2.691 | 7.5448 | 7.5494 | 27.1139 | 27.1221 | 7.021 | 2.3852 | 16 | 16 | 184 | 7.1079 | 0.9097 | 6.466 |  | 7.8111e-05 | [5 empty] | 207 | excl. outliers diff>0.03 | 0.479 | -0.2317 | 0.479 | -0.48372 | 0.9989 |  |  | 3.2616
66 | 125.028 | 9.1872 | 9.187 | 30.529 | 30.5309 | 6.616 | 1.3694 | 7 | 7 | 247 | 6.8221 | 0.42347 | 2.889 |  | 8.831e-05 | [5 empty] | 191 | excl. outliers diff>0.01 | 0.4794 | -0.2318 | 0.4794 | -0.48352 | 0.9997 |  |  | 3.2649
32 | 216.77 | 8.4627 | 8.4634 | 30.7827 | 30.7817 | 6.7131 | 1.9042 | 1 | 1 | 108 | 6.9956 | 0.67957 | 4.754 |  | 0.00010283 | [5 empty] | 185 | excl. outliers diff>0.008 | 0.4794 | -0.2318 | 0.4794 | -0.48352 | 0.9997 |  |  | 3.2649
75 | 101.363 | 8.5096 | 8.51 | 29.3416 | 29.3425 | 6.769 | 1.9376 | 3 | 3 | 274 | 6.9511 | 0.69557 | 4.835 |  | 0.00013668 | [5 empty] | 175 | excl. outliers diff>0.006 | 0.4797 | -0.2324 | 0.4797 | -0.48488 | 0.9998 |  |  | 3.2626
43 | 9.636 | 8.2296 | 8.2297 | 27.0374 | 27.0369 | 6.9149 | 2.2137 | 14 | 14 | 149 | 7.0148 | 0.82739 | 5.804 |  | 0.00014751 | [5 empty] | 160 | excl. outliers diff>0.005 | 0.4795 | -0.2325 | 0.4795 | -0.48447 | 0.9998 |  |  | 3.2625
60 | 174.484 | 9.0335 | 9.0338 | 30.9076 | 30.9083 | 6.6225 | 1.333 | 2 | 2 | 217 | 6.8703 | 0.40595 | 2.789 |  | 0.00017471 | [5 empty] | 147 | excl. outliers diff>0.004 | 0.4792 | -0.2322 | 0.4792 | -0.48488 | 0.9999 |  |  | 3.2592
32 | 31.839 | 8.9183 | 8.92 | 30.2717 | 30.2719 | 6.6669 | 1.7692 | 10 | 10 | 117 | 6.7903 | 0.6147 | 4.174 |  | 0.00018519 | [5 empty] | 112 | excl. outliers diff>0.003 | 0.4794 | -0.2327 | 0.4794 | -0.48488 | 0.9999 |  |  | 3.2605
66 | 250.077 | 9.0942 | 9.0941 | 30.7787 | 30.7801 | 6.619 | 1.324 | 3 | 3 | 243 | 6.9339 | 0.40165 | 2.785 | 6 | 0.00020521
70 | 4.821 | 8.1232 | 8.1233 | 28.6736 | 28.6738 | 6.8583 | 2.3281 | 14 | 14 | 270 | 6.9518 | 0.8825 | 6.135 |  | 0.00021949
66 | 149.819 | 9.1584 | 9.1584 | 30.6398 | 30.6412 | 6.6155 | 1.3594 | 6 | 6 | 246 | 6.8429 | 0.41854 | 2.864 |  | 0.00025579 | [7 empty] | m | b | Soc | Voffset | R2
43 | 174.572 | 9.1844 | 9.1848 | 30.7917 | 30.7924 | 6.6052 | 1.4319 | 5 | 5 | 140 | 6.8541 | 0.45374 | 3.11 |  | 0.00029113 | [5 empty] | 21 | 2011-60 | 0.4493 | -0.2214 | 0.4493 | -0.49277 | 0.9991 | few bottles |  | 3.0321
43 | 30.495 | 9.153 | 9.1523 | 29.3623 | 29.3574 | 6.6711 | 1.7937 | 12 | 12 | 147 | 6.7961 | 0.62624 | 4.256 |  | 0.00032966 | [5 empty] | 93 | 2011-63 | 0.4474 | -0.2213 | 0.4474 | -0.49464 | 0.9991 | problems with bottles |  | 3.0137
60 | 189.595 | 8.9982 | 8.9986 | 30.9633 | 30.9638 | 6.6253 | 1.3184 | 1 | 1 | 216 | 6.8861 | 0.39878 | 2.746 |  | 0.00036487 | [5 empty] | 241 | 2011-10 | 0.4417 | -0.1979 | 0.4417 | -0.44804 | 0.9989 |  |  | 3.1132
3 | 9.813 | 8.4729 | 8.4747 | 30.5372 | 30.5314 | 6.7223 | 1.922 | 12 | 12 | 21 | 6.8224 | 0.68759 | 4.691 |  | 0.00037942 | [5 empty] | 17 | 2011-76 | 0.4803 | -0.2259 | 0.4803 | -0.47033 | 0.999
66 | 99.935 | 9.1651 | 9.1652 | 30.3829 | 30.3847 | 6.6255 | 1.3804 | 8 | 8 | 248 | 6.8098 | 0.42923 | 2.923 |  | 0.0004051 | [5 empty] | 175 | 2011-11 | 0.4797 | -0.2324 | 0.4797 | -0.48447 | 0.9998 |  |  | 3.2639
55 | 40.2 | 8.5819 | 8.5865 | 28.7138 | 28.7203 | 6.7856 | 2.0993 | 10 | 10 | 194 | 6.9146 | 0.77242 | 5.341 |  | 0.00041547 | [14 empty] | 0
55 | 250.74 | 9.0604 | 9.0611 | 31.0039 | 31.0039 | 6.6144 | 1.3257 | 2 | 2 | 186 | 6.9293 | 0.40221 | 2.787 |  | 0.00042716 | [5 empty] | excluding diff from fit >0.006 |  | m | b | Soc | Voffset | R2
48 | 100.638 | 9.3588 | 9.3586 | 30.3202 | 30.3213 | 6.5997 | 1.5366 | 9 | 9 | 160 | 6.7861 | 0.50308 | 3.414 |  | 0.00045495 | [5 empty] | southern SoG +JdF (not SI) |  | 0.4835 | -0.2394 | 0.4835 | -0.49514 | 0.9997 |  |  | 3.2552
32 | 149.739 | 8.4766 | 8.477 | 30.6973 | 30.6979 | 6.7147 | 1.9135 | 4 | 4 | 111 | 6.9377 | 0.68437 | 4.748 | 6 | 0.00045526 | [5 empty] | northern SoG |  | 0.478 | -0.2301 | 0.478 | -0.48138 | 0.9998 |  |  | 3.2622
66 | 2.392 | 8.0791 | 8.0783 | 28.2963 | 28.2982 | 6.8821 | 2.3544 | 16 | 16 | 256 | 6.9733 | 0.89442 | 6.237 |  | 0.0004792 | [5 empty] | all |  | 0.4797 | -0.2324 | 0.4797 | -0.48447 | 0.9998 |  |  | 3.2639
26 | 125.728 | 8.5056 | 8.5062 | 31.1024 | 31.1028 | 6.6927 | 2.038 | 2 | 2 | 98 | 6.8941 | 0.74397 | 5.129 |  | 0.00049311
66 | 199.666 | 9.1267 | 9.1267 | 30.7202 | 30.7216 | 6.6167 | 1.3461 | 4 | 4 | 244 | 6.8875 | 0.41292 | 2.844 |  | 0.00052034 | [5 empty] | Summary of Soc Voffset including the original values in the factory calibration
39 | 100.563 | 9.3462 | 9.3465 | 30.3338 | 30.3344 | 6.601 | 1.5486 | 4 | 4 | 126 | 6.7872 | 0.50875 | 3.453 |  | 0.00054669 | [7 empty] | m | b | Soc | Voffset | R2
48 | 378.184 | 9.0143 | 9.0147 | 31.1859 | 31.1864 | 6.6134 | 1.2283 | 1 | 1 | 152 | 7.0413 | 0.35547 | 2.503 |  | 0.00056776
60 | 2.383 | 8.0808 | 8.0819 | 28.1306 | 28.1312 | 6.8893 | 2.3599 | 13 | 13 | 228 | 6.9805 | 0.89807 | 6.269 |  | 0.00057083
39 | 29.768 | 8.9617 | 8.9608 | 29.5863 | 29.5865 | 6.69 | 1.7908 | 8 | 8 | 130 | 6.8125 | 0.62576 | 4.263 |  | 0.0005848
39 | 149.733 | 8.6202 | 8.6207 | 30.4585 | 30.4589 | 6.7035 | 1.8512 | 2 | 2 | 124 | 6.9277 | 0.65346 | 4.527 |  | 0.00061117
43 | 249.978 | 9.0847 | 9.0852 | 31.0333 | 31.0338 | 6.6096 | 1.3289 | 3 | 3 | 138 | 6.9238 | 0.40354 | 2.794 | 6 | 0.00064362
51 | 5.535 | 7.8229 | 7.8182 | 27.8336 | 27.821 | 6.9435 | 2.3654 | 15 | 15 | 183 | 7.0352 | 0.89947 | 6.328 |  | 0.00068564
66 | 75.702 | 8.9979 | 8.9979 | 30.0555 | 30.0569 | 6.6643 | 1.4996 | 9 | 9 | 249 | 6.8267 | 0.48515 | 3.312 |  | 0.00071564
43 | 299.799 | 9.0466 | 9.047 | 31.1231 | 31.1235 | 6.6113 | 1.2642 | 2 | 2 | 137 | 6.9694 | 0.37248 | 2.596 |  | 0.00072045
70 | 40.441 | 8.4847 | 8.4857 | 29.1082 | 29.1086 | 6.7831 | 2.0808 | 10 | 10 | 266 | 6.9112 | 0.7647 | 5.285 |  | 0.00072622
75 | 152.406 | 8.8889 | 8.8897 | 30.471 | 30.4729 | 6.6627 | 1.4391 | 1 | 1 | 272 | 6.8909 | 0.4577 | 3.154 |  | 0.00077632
32 | 174.486 | 8.4653 | 8.4658 | 30.7471 | 30.7474 | 6.7143 | 1.9111 | 3 | 3 | 110 | 6.9591 | 0.68357 | 4.757 |  | 0.00079566
26 | 146.042 | 8.4948 | 8.4951 | 31.1083 | 31.109 | 6.6941 | 2.0293 | 1 | 1 | 97 | 6.9133 | 0.74016 | 5.117 |  | 0.00084277
60 | 99.941 | 9.148 | 9.1483 | 30.3127 | 30.3133 | 6.631 | 1.3781 | 5 | 5 | 220 | 6.8153 | 0.42683 | 2.909 | 6 | 0.00087073
60 | 150.202 | 9.0831 | 9.0834 | 30.8452 | 30.8458 | 6.6178 | 1.3483 | 3 | 3 | 218 | 6.8448 | 0.41258 | 2.824 |  | 0.00089975
3 | 151.683 | 8.3715 | 8.3717 | 31.0364 | 31.0379 | 6.7158 | 1.9314 | 4 | 4 | 13 | 6.9394 | 0.69156 | 4.799 |  | 0.00090734
39 | 74.288 | 9.3702 | 9.3714 | 30.1716 | 30.1715 | 6.6044 | 1.5551 | 5 | 5 | 127 | 6.7683 | 0.5115 | 3.462 |  | 0.00092651
16 | 76.231 | 8.4871 | 8.4875 | 31.4321 | 31.4325 | 6.6812 | 1.9476 | 5 | 5 | 69 | 6.8386 | 0.70116 | 4.795 |  | 0.00093177
43 | 75.187 | 9.386 | 9.3863 | 30.2218 | 30.222 | 6.5999 | 1.5206 | 9 | 9 | 144 | 6.7647 | 0.49684 | 3.361 |  | 0.00094715
66 | 29.919 | 8.3976 | 8.3975 | 28.6928 | 28.6939 | 6.8148 | 2.2649 | 12 | 12 | 252 | 6.9332 | 0.853 | 5.914 |  | 0.00095102
16 | 101.392 | 8.5119 | 8.5125 | 31.4929 | 31.4931 | 6.6748 | 1.9432 | 4 | 4 | 68 | 6.8544 | 0.69911 | 4.792 |  | 0.00096616
39 | 40.69 | 8.8025 | 8.801 | 29.7103 | 29.7108 | 6.7085 | 1.8379 | 7 | 7 | 129 | 6.8391 | 0.64877 | 4.437 |  | 0.0010207
60 | 124.916 | 9.2275 | 9.2278 | 30.6297 | 30.6306 | 6.6058 | 1.3861 | 4 | 4 | 219 | 6.812 | 0.43262 | 2.947 |  | 0.0010491
32 | 100.025 | 8.509 | 8.5092 | 30.5884 | 30.5897 | 6.7146 | 1.9129 | 6 | 6 | 113 | 6.894 | 0.68466 | 4.72 |  | 0.001052
32 | 76.071 | 8.5282 | 8.5286 | 30.5106 | 30.5105 | 6.7151 | 1.9112 | 7 | 7 | 114 | 6.8736 | 0.68392 | 4.701 |  | 0.0010941
51 | 299.965 | 9.0707 | 9.0711 | 31.064 | 31.0645 | 6.6103 | 1.3168 | 2 | 2 | 170 | 6.9687 | 0.3995 | 2.784 |  | 0.0011146
70 | 149.788 | 9.0558 | 9.0554 | 30.5577 | 30.559 | 6.6342 | 1.3603 | 5 | 5 | 261 | 6.861 | 0.42035 | 2.884 | 6 | 0.0011381
16 | 150.202 | 7.9316 | 7.9321 | 32.6716 | 32.6716 | 6.7111 | 1.4923 | 2 | 2 | 66 | 6.9282 | 0.48353 | 3.35 |  | 0.001156
66 | 299.961 | 9.0923 | 9.0922 | 30.79 | 30.7913 | 6.6188 | 1.317 | 2 | 2 | 242 | 6.978 | 0.39725 | 2.772 |  | 0.0012475
70 | 249.78 | 8.9796 | 8.98 | 30.7572 | 30.7578 | 6.6369 | 1.3162 | 2 | 2 | 258 | 6.9511 | 0.39936 | 2.776 |  | 0.0012594
43 | 316.402 | 9.0419 | 9.0422 | 31.1362 | 31.1367 | 6.6115 | 1.253 | 1 | 1 | 136 | 6.9843 | 0.36654 | 2.56 |  | 0.0013467
60 | 29.981 | 8.4366 | 8.4372 | 28.7804 | 28.7807 | 6.805 | 2.2424 | 9 | 9 | 224 | 6.9237 | 0.84262 | 5.834 | 6 | 0.0013526
60 | 5.131 | 8.0819 | 8.0828 | 28.1306 | 28.1314 | 6.8891 | 2.3575 | 12 | 12 | 227 | 6.9828 | 0.89777 | 6.269 |  | 0.0013869
3 | 75.406 | 8.3915 | 8.3919 | 30.9184 | 30.9193 | 6.718 | 1.9202 | 7 | 7 | 16 | 6.8745 | 0.68849 | 4.733 |  | 0.0013945
16 | 125.503 | 8.2224 | 8.2223 | 32.0488 | 32.0497 | 6.6943 | 1.7322 | 3 | 3 | 67 | 6.8923 | 0.59574 | 4.106 |  | 0.0014005
60 | 9.518 | 8.1618 | 8.1483 | 28.1737 | 28.1662 | 6.8746 | 2.3253 | 11 | 11 | 226 | 6.973 | 0.87953 | 6.133 |  | 0.0014461
48 | 2.777 | 7.5686 | 7.5548 | 27.1257 | 27.0961 | 7.0166 | 2.4251 | 17 | 17 | 168 | 7.1038 | 0.92711 | 6.586 |  | 0.001606
32 | 123.134 | 8.4859 | 8.4865 | 30.6611 | 30.661 | 6.7149 | 1.9132 | 5 | 5 | 112 | 6.9144 | 0.68538 | 4.739 |  | 0.0016151
3 | 201.317 | 8.3675 | 8.3678 | 31.1301 | 31.1305 | 6.7124 | 1.9644 | 2 | 2 | 11 | 6.9799 | 0.70989 | 4.955 |  | 0.0016362
66 | 351.312 | 9.069 | 9.0683 | 30.7952 | 30.7966 | 6.622 | 1.2997 | 1 | 1 | 241 | 7.027 | 0.3885 | 2.73 |  | 0.0016984
75 | 9.464 | 8.0434 | 8.043 | 28.4896 | 28.4898 | 6.879 | 2.3554 | 9 | 9 | 280 | 6.9761 | 0.89707 | 6.258 |  | 0.0017043
3 | 177.799 | 8.3662 | 8.3665 | 31.0884 | 31.0886 | 6.7144 | 1.9508 | 3 | 3 | 12 | 6.961 | 0.70363 | 4.898 |  | 0.0018484
70 | 75.045 | 8.5994 | 8.5999 | 29.4441 | 29.4447 | 6.7508 | 1.8633 | 8 | 8 | 264 | 6.9102 | 0.66178 | 4.573 |  | 0.0018803
32 | 10.201 | 8.5741 | 8.5767 | 28.8795 | 28.8876 | 6.7795 | 1.9695 | 12 | 12 | 119 | 6.8819 | 0.71259 | 4.904 |  | 0.001901
48 | 302.913 | 9.0616 | 9.0621 | 31.112 | 31.1123 | 6.6096 | 1.2767 | 3 | 3 | 154 | 6.9705 | 0.3773 | 2.63 | 6 | 0.0019114
70 | 175.523 | 9.0113 | 9.0118 | 30.6574 | 30.658 | 6.6365 | 1.3389 | 4 | 4 | 260 | 6.8855 | 0.41087 | 2.829 |  | 0.0019256
70 | 199.755 | 9.0312 | 9.0316 | 30.7178 | 30.7184 | 6.6309 | 1.3347 | 3 | 3 | 259 | 6.9013 | 0.40891 | 2.822 |  | 0.0019524
3 | 222.97 | 8.3697 | 8.37 | 31.1327 | 31.133 | 6.7119 | 1.9597 | 1 | 1 | 10 | 6.9988 | 0.70797 | 4.955 |  | 0.0019576
75 | 50.292 | 8.3891 | 8.39 | 29.1252 | 29.1261 | 6.797 | 2.0512 | 5 | 5 | 276 | 6.933 | 0.75176 | 5.212 | 6 | 0.0019695
66 | 19.853 | 8.3542 | 8.3536 | 28.5597 | 28.5602 | 6.8274 | 2.2734 | 13 | 13 | 253 | 6.9366 | 0.8582 | 5.953 |  | 0.002085
60 | 39.788 | 8.4309 | 8.4306 | 28.8298 | 28.8296 | 6.8037 | 2.2196 | 8 | 8 | 223 | 6.931 | 0.8325 | 5.77 |  | 0.0021284
66 | 39.895 | 8.4353 | 8.4352 | 28.7499 | 28.7505 | 6.8065 | 2.2306 | 11 | 11 | 251 | 6.934 | 0.83775 | 5.809 | 6 | 0.0021364
55 | 30.038 | 8.4892 | 8.4902 | 28.6295 | 28.6305 | 6.8035 | 2.1363 | 11 | 11 | 195 | 6.9229 | 0.7884 | 5.458 |  | 0.0021369
43 | 124.794 | 9.1582 | 9.1595 | 30.5001 | 30.5003 | 6.6215 | 1.575 | 7 | 7 | 142 | 6.8273 | 0.52407 | 3.578 |  | 0.0021407
3 | 2.741 | 8.5009 | 8.5012 | 30.3864 | 30.3871 | 6.7246 | 1.9178 | 14 | 14 | 23 | 6.8189 | 0.68809 | 4.692 |  | 0.002142
70 | 310.494 | 8.9756 | 8.976 | 30.7812 | 30.7817 | 6.6364 | 1.3028 | 1 | 1 | 257 | 7.0047 | 0.39388 | 2.759 |  | 0.0021912
55 | 175.532 | 9.1268 | 9.1276 | 30.8677 | 30.8677 | 6.6104 | 1.3616 | 4 | 4 | 188 | 6.8597 | 0.42203 | 2.895 |  | 0.0021921
55 | 199.893 | 9.0953 | 9.0957 | 30.9319 | 30.9323 | 6.6123 | 1.3514 | 3 | 3 | 187 | 6.8827 | 0.41713 | 2.871 |  | 0.0021957
43 | 150.057 | 9.1575 | 9.158 | 30.7099 | 30.7106 | 6.6127 | 1.4884 | 6 | 6 | 141 | 6.8402 | 0.48274 | 3.302 |  | 0.0022381
13 | 5.306 | 9.2922 | 9.293 | 30.2508 | 30.2516 | 6.6124 | 2.3722 | 12 | 12 | 63 | 6.7164 | 0.9057 | 6.083 |  | 0.0023108
75 | 4.987 | 8.0285 | 8.0291 | 28.4896 | 28.4901 | 6.8814 | 2.3584 | 10 | 10 | 281 | 6.9743 | 0.89916 | 6.271 |  | 0.0023454
70 | 124.964 | 9.1032 | 9.1036 | 30.453 | 30.4536 | 6.6316 | 1.355 | 6 | 6 | 262 | 6.8373 | 0.41903 | 2.865 |  | 0.0023743
55 | 76.002 | 9.2052 | 9.2051 | 30.1128 | 30.1132 | 6.6311 | 1.4869 | 8 | 8 | 192 | 6.7953 | 0.47739 | 3.244 |  | 0.0023807
13 | 2.586 | 9.269 | 9.2691 | 30.2411 | 30.2414 | 6.6163 | 2.3782 | 13 | 13 | 64 | 6.7176 | 0.90865 | 6.104 |  | 0.0023927
13 | 10.059 | 9.3062 | 9.3064 | 30.3485 | 30.3496 | 6.6062 | 2.3455 | 11 | 11 | 62 | 6.7143 | 0.89302 | 5.996 |  | 0.0024019
16 | 40.648 | 8.9377 | 8.9376 | 30.9683 | 30.969 | 6.634 | 2.1249 | 7 | 7 | 71 | 6.7647 | 0.78747 | 5.327 | 6 | 0.0024237
13 | 175.581 | 7.7141 | 7.7147 | 33.3 | 33.3004 | 6.7166 | 1.2794 | 2 | 2 | 53 | 6.954 | 0.37805 | 2.629 |  | 0.0024294
16 | 174.404 | 7.9033 | 7.9037 | 32.7487 | 32.749 | 6.712 | 1.4625 | 1 | 1 | 65 | 6.9503 | 0.46559 | 3.236 |  | 0.0024938
3 | 4.804 | 8.5006 | 8.5011 | 30.3882 | 30.3882 | 6.7246 | 1.9179 | 13 | 13 | 22 | 6.8206 | 0.6885 | 4.696 |  | 0.0024978
70 | 29.989 | 8.4599 | 8.4603 | 29.0789 | 29.0795 | 6.7882 | 2.1174 | 11 | 11 | 267 | 6.9069 | 0.784 | 5.415 |  | 0.0025007
19 | 102.344 | 8.4965 | 8.4968 | 31.2469 | 31.2476 | 6.6878 | 2.0146 | 2 | 2 | 78 | 6.8684 | 0.73481 | 5.047 |  | 0.0025362
3 | 126.227 | 8.3812 | 8.3817 | 30.9859 | 30.9864 | 6.7166 | 1.9266 | 5 | 5 | 14 | 6.9177 | 0.69271 | 4.792 | 6 | 0.0025364
48 | 76.2 | 9.2776 | 9.2783 | 30.1279 | 30.1285 | 6.6198 | 1.4566 | 10 | 10 | 161 | 6.7847 | 0.46781 | 3.174 |  | 0.0025415
16 | 49.954 | 8.7989 | 8.7982 | 31.0443 | 31.0455 | 6.6514 | 2.0894 | 6 | 6 | 70 | 6.7888 | 0.77068 | 5.232 |  | 0.0026254
39 | 177.901 | 8.6129 | 8.6133 | 30.5086 | 30.5091 | 6.7024 | 1.8334 | 1 | 1 | 123 | 6.9515 | 0.64821 | 4.506 |  | 0.0026503
1 | 29.971 | 8.9799 | 8.9803 | 29.8673 | 29.8685 | 6.6751 | 1.7346 | 5 | 5 | 5 | 6.7978 | 0.59564 | 4.049 |  | 0.0026516
75 | 75.393 | 8.4542 | 8.4546 | 29.2047 | 29.2051 | 6.7835 | 1.9943 | 4 | 4 | 275 | 6.9423 | 0.71994 | 4.998 |  | 0.0026526
3 | 20.095 | 8.4293 | 8.4297 | 30.7941 | 30.7948 | 6.7177 | 1.9191 | 11 | 11 | 20 | 6.8261 | 0.68926 | 4.705 |  | 0.0026875
55 | 150.578 | 9.2067 | 9.2072 | 30.7787 | 30.779 | 6.6025 | 1.3992 | 5 | 5 | 189 | 6.8306 | 0.44052 | 3.009 | 6 | 0.0027075
48 | 151.041 | 8.9749 | 8.9764 | 30.7185 | 30.719 | 6.6392 | 1.559 | 7 | 7 | 158 | 6.8665 | 0.51701 | 3.55 |  | 0.002725
70 | 99.815 | 8.8711 | 8.8707 | 29.9825 | 29.9825 | 6.6864 | 1.579 | 7 | 7 | 263 | 6.869 | 0.52657 | 3.617 |  | 0.0027308
32 | 2.571 | 8.392 | 8.3919 | 28.2703 | 28.2691 | 6.8344 | 2.0651 | 14 | 14 | 121 | 6.9288 | 0.75367 | 5.222 |  | 0.0027847
26 | 31.385 | 8.6103 | 8.6085 | 30.9881 | 30.9905 | 6.682 | 2.0982 | 7 | 7 | 103 | 6.8017 | 0.7751 | 5.272 |  | 0.0027924
3 | 50.563 | 8.393 | 8.3938 | 30.9047 | 30.9047 | 6.7183 | 1.9223 | 8 | 8 | 17 | 6.853 | 0.69093 | 4.735 |  | 0.002794
43 | 39.997 | 9.2045 | 9.2047 | 29.6861 | 29.6866 | 6.6495 | 1.702 | 11 | 11 | 146 | 6.7829 | 0.57984 | 3.933 | 6 | 0.002866
13 | 38.866 | 8.8697 | 8.8696 | 31.0254 | 31.0264 | 6.6417 | 2.108 | 8 | 8 | 59 | 6.7701 | 0.7799 | 5.28 |  | 0.0029229
16 | 30.343 | 9.0498 | 9.0504 | 30.9051 | 30.9054 | 6.6202 | 2.1535 | 8 | 8 | 72 | 6.7429 | 0.80173 | 5.406 |  | 0.0029825
19 | 116.354 | 8.4966 | 8.4968 | 31.2445 | 31.245 | 6.6879 | 2.0106 | 1 | 1 | 77 | 6.8808 | 0.73334 | 5.046 |  | 0.0029908
48 | 176.432 | 9.1046 | 9.105 | 30.8697 | 30.8703 | 6.6136 | 1.42 | 6 | 6 | 157 | 6.8636 | 0.45078 | 3.094 |  | 0.0030141
3 | 100.975 | 8.3933 | 8.3936 | 30.9245 | 30.925 | 6.7174 | 1.9173 | 6 | 6 | 15 | 6.8964 | 0.68876 | 4.75 |  | 0.0030195
26 | 50.694 | 8.6329 | 8.6313 | 30.974 | 30.9757 | 6.6792 | 2.1023 | 5 | 5 | 101 | 6.816 | 0.77729 | 5.298 |  | 0.0030229
55 | 9.79 | 8.1516 | 8.1474 | 28.0734 | 28.0723 | 6.8807 | 2.3008 | 13 | 13 | 197 | 6.9793 | 0.86613 | 6.045 |  | 0.0031171
43 | 199.575 | 9.1478 | 9.1483 | 30.8869 | 30.8873 | 6.6065 | 1.3997 | 4 | 4 | 139 | 6.877 | 0.43493 | 2.991 |  | 0.0031425
66 | 5.357 | 8.0748 | 8.0745 | 28.2962 | 28.2977 | 6.8828 | 2.3534 | 15 | 15 | 255 | 6.9766 | 0.89127 | 6.218 |  | 0.0031453
26 | 100.565 | 8.5159 | 8.5156 | 31.0911 | 31.0929 | 6.6917 | 2.0492 | 3 | 3 | 99 | 6.871 | 0.752 | 5.167 |  | 0.003175
51 | 249.566 | 9.0835 | 9.084 | 31.0268 | 31.0273 | 6.61 | 1.3374 | 3 | 3 | 171 | 6.9239 | 0.41147 | 2.849 |  | 0.003208
16 | 10.174 | 9.9065 | 9.9086 | 30.088 | 30.0919 | 6.5301 | 2.3831 | 10 | 10 | 74 | 6.6438 | 0.91183 | 6.058 |  | 0.0032268
26 | 1.975 | 8.7153 | 8.7159 | 30.9297 | 30.9304 | 6.6688 | 2.1607 | 11 | 11 | 107 | 6.7641 | 0.80543 | 5.448 |  | 0.0032297
13 | 125.092 | 8.1511 | 8.1513 | 32.1327 | 32.1334 | 6.7014 | 1.6962 | 4 | 4 | 55 | 6.8984 | 0.58318 | 4.023 |  | 0.0032483
3 | 40.238 | 8.4056 | 8.4066 | 30.8756 | 30.8747 | 6.7177 | 1.9183 | 9 | 9 | 18 | 6.8435 | 0.68971 | 4.72 |  | 0.0034945
51 | 125.512 | 9.2019 | 9.2103 | 30.6163 | 30.6171 | 6.6101 | 1.5067 | 7 | 7 | 175 | 6.8166 | 0.49276 | 3.359 |  | 0.0035089
55 | 50.679 | 8.6402 | 8.6407 | 28.9467 | 28.9474 | 6.7664 | 2.1178 | 9 | 9 | 193 | 6.9051 | 0.77813 | 5.373 | 6 | 0.0035176
55 | 2.507 | 7.7222 | 7.7279 | 27.5439 | 27.5497 | 6.9727 | 2.3627 | 15 | 15 | 199 | 7.0609 | 0.89535 | 6.322 |  | 0.0035226
13 | 75.96 | 8.4619 | 8.4624 | 31.4604 | 31.4608 | 6.6838 | 1.938 | 6 | 6 | 57 | 6.8407 | 0.69919 | 4.783 |  | 0.0035467
26 | 74.786 | 8.5581 | 8.5582 | 31.041 | 31.0419 | 6.6875 | 2.0706 | 4 | 4 | 100 | 6.8446 | 0.76264 | 5.22 | 6 | 0.0035691
48 | 126.49 | 9.0748 | 9.0755 | 30.5708 | 30.5712 | 6.6308 | 1.5755 | 8 | 8 | 159 | 6.8374 | 0.52579 | 3.595 |  | 0.0035898
48 | 29.948 | 8.9624 | 8.9634 | 29.3351 | 29.3367 | 6.7007 | 1.8455 | 13 | 13 | 164 | 6.8237 | 0.64775 | 4.42 |  | 0.00362
19 | 10.659 | 8.5459 | 8.5467 | 31.0852 | 31.0859 | 6.6874 | 2.0777 | 8 | 8 | 84 | 6.7885 | 0.76614 | 5.201 |  | 0.0036783
66 | 10.29 | 8.2667 | 8.2676 | 28.4458 | 28.4487 | 6.8461 | 2.3118 | 14 | 14 | 254 | 6.946 | 0.8782 | 6.1 |  | 0.0036897
16 | 4.58 | 9.8633 | 9.864 | 29.8337 | 29.8346 | 6.547 | 2.4264 | 11 | 11 | 75 | 6.6558 | 0.93303 | 6.21 |  | 0.0037085
43 | 19.805 | 9.0088 | 9.0092 | 28.9232 | 28.9222 | 6.7116 | 1.8927 | 13 | 13 | 148 | 6.8264 | 0.67019 | 4.575 |  | 0.0037637
16 | 2.673 | 9.804 | 9.8052 | 29.7865 | 29.7878 | 6.5575 | 2.4309 | 12 | 12 | 76 | 6.6642 | 0.93545 | 6.234 |  | 0.0039547
43 | 5.382 | 7.2204 | 7.2195 | 22.882 | 22.8785 | 7.273 | 2.4665 | 15 | 15 | 150 | 7.3615 | 0.94451 | 6.953 |  | 0.0040265
55 | 4.877 | 8.0614 | 8.0851 | 27.907 | 27.9138 | 6.9024 | 2.3159 | 14 | 14 | 198 | 6.9958 | 0.88053 | 6.16 |  | 0.0040638
51 | 100.961 | 9.3461 | 9.3466 | 30.5373 | 30.5378 | 6.5923 | 1.4514 | 8 | 8 | 176 | 6.7787 | 0.4669 | 3.165 |  | 0.0040898
48 | 49.975 | 9.1213 | 9.1218 | 29.6865 | 29.6875 | 6.6619 | 1.6697 | 11 | 11 | 162 | 6.8032 | 0.56312 | 3.831 |  | 0.0041123
48 | 4.958 | 7.6891 | 7.6918 | 27.3991 | 27.4025 | 6.9846 | 2.4213 | 16 | 16 | 167 | 7.0748 | 0.92271 | 6.528 |  | 0.0041933
55 | 275.596 | 9.0085 | 9.0089 | 31.0813 | 31.0817 | 6.6187 | 1.2951 | 1 | 1 | 185 | 6.9553 | 0.39222 | 2.728 |  | 0.0042342
7 | 2.601 | 10.0555 | 10.0558 | 29.9743 | 29.9754 | 6.5136 | 2.4512 | 14 | 14 | 37 | 6.6222 | 0.94545 | 6.261 |  | 0.0042742
7 | 5.008 | 10.0616 | 10.0619 | 29.9751 | 29.9762 | 6.5126 | 2.4489 | 13 | 13 | 36 | 6.6234 | 0.94438 | 6.255 |  | 0.0042971
75 | 30.152 | 8.3239 | 8.3239 | 28.9301 | 28.9298 | 6.8157 | 2.1688 | 7 | 7 | 278 | 6.9334 | 0.81042 | 5.619 |  | 0.0043244
7 | 40.211 | 10.3284 | 10.3286 | 30.3621 | 30.362 | 6.4588 | 2.4042 | 9 | 9 | 32 | 6.6011 | 0.92303 | 6.093 |  | 0.0043344
19 | 31.659 | 8.5093 | 8.5095 | 31.1612 | 31.1619 | 6.6896 | 2.0546 | 6 | 6 | 82 | 6.8086 | 0.75581 | 5.146 |  | 0.0043961
70 | 2.811 | 8.1311 | 8.1282 | 28.6829 | 28.6815 | 6.8567 | 2.3223 | 15 | 15 | 271 | 6.9485 | 0.88393 | 6.142 |  | 0.0044129
70 | 19.562 | 8.4199 | 8.4212 | 29.048 | 29.049 | 6.7957 | 2.1696 | 12 | 12 | 268 | 6.9049 | 0.81088 | 5.599 |  | 0.0044297
75 | 2.611 | 8.036 | 8.0365 | 28.4912 | 28.4918 | 6.8801 | 2.3581 | 11 | 11 | 282 | 6.9709 | 0.90117 | 6.282 |  | 0.0045225
51 | 75.318 | 9.4023 | 9.4028 | 30.3553 | 30.3558 | 6.5919 | 1.4761 | 9 | 9 | 177 | 6.7567 | 0.47923 | 3.238 |  | 0.0046172
16 | 19.99 | 9.1345 | 9.1428 | 30.8183 | 30.8066 | 6.6114 | 2.1797 | 9 | 9 | 73 | 6.7261 | 0.81593 | 5.488 |  | 0.0046306
26 | 5.534 | 8.7199 | 8.7217 | 30.9299 | 30.9298 | 6.6681 | 2.1696 | 10 | 10 | 106 | 6.7665 | 0.81121 | 5.489 |  | 0.0047473
70 | 49.788 | 8.5089 | 8.5095 | 29.1677 | 29.1674 | 6.7768 | 2.0142 | 9 | 9 | 265 | 6.9133 | 0.73684 | 5.094 | 6 | 0.004749
51 | 201.255 | 9.0049 | 9.0044 | 30.8799 | 30.8798 | 6.6279 | 1.4887 | 4 | 4 | 172 | 6.8991 | 0.48543 | 3.349 | 6 | 0.0047677
51 | 175.851 | 9.0525 | 9.0528 | 30.8077 | 30.8083 | 6.6239 | 1.4983 | 5 | 5 | 173 | 6.8732 | 0.49016 | 3.369 |  | 0.0049215
13 | 99.481 | 8.3933 | 8.3933 | 31.6083 | 31.6094 | 6.6876 | 1.8841 | 5 | 5 | 56 | 6.8645 | 0.67477 | 4.632 |  | 0.0049247
60 | 75.34 | 8.9386 | 8.9394 | 29.7919 | 29.7943 | 6.6845 | 1.6438 | 6 | 6 | 221 | 6.8464 | 0.54992 | 3.765 |  | 0.0049469
19 | 40.648 | 8.5225 | 8.5225 | 31.172 | 31.1719 | 6.6872 | 2.0521 | 5 | 5 | 81 | 6.8141 | 0.7552 | 5.146 | 6 | 0.0049902
26 | 20.514 | 8.6142 | 8.6142 | 30.9813 | 30.9822 | 6.6817 | 2.105 | 8 | 8 | 104 | 6.792 | 0.78062 | 5.302 |  | 0.0050769
19 | 4.961 | 8.5964 | 8.5985 | 31.0652 | 31.0645 | 6.6807 | 2.0869 | 9 | 9 | 85 | 6.7774 | 0.77198 | 5.232 |  | 0.0050885
26 | 40.391 | 8.6067 | 8.6088 | 30.9905 | 30.9902 | 6.6824 | 2.1036 | 6 | 6 | 102 | 6.81 | 0.78003 | 5.312 |  | 0.0051744
39 | 20.205 | 9.0485 | 9.0453 | 29.5435 | 29.5457 | 6.6788 | 1.7673 | 9 | 9 | 131 | 6.7939 | 0.61923 | 4.207 |  | 0.0052549
19 | 50.783 | 8.5193 | 8.5199 | 31.2083 | 31.2091 | 6.6861 | 2.0401 | 4 | 4 | 80 | 6.8218 | 0.7498 | 5.115 |  | 0.0053266
7 | 76.691 | 9.1035 | 9.1052 | 31.2832 | 31.2832 | 6.5961 | 2.138 | 7 | 7 | 30 | 6.7589 | 0.79672 | 5.385 |  | 0.0053918
7 | 129.571 | 8.0357 | 8.0334 | 32.3334 | 32.3422 | 6.7101 | 1.6038 | 5 | 5 | 28 | 6.9101 | 0.54124 | 3.74 | 6 | 0.0054982
51 | 20.551 | 8.8022 | 8.8005 | 28.9864 | 28.9834 | 6.7401 | 1.982 | 13 | 13 | 181 | 6.8536 | 0.72225 | 4.95 |  | 0.0055613
13 | 31.456 | 8.96 | 8.9607 | 30.9429 | 30.9421 | 6.6318 | 2.1364 | 9 | 9 | 60 | 6.7547 | 0.79618 | 5.378 | 6 | 0.005595
51 | 150.283 | 9.1226 | 9.1239 | 30.7345 | 30.7349 | 6.6167 | 1.5052 | 6 | 6 | 174 | 6.8442 | 0.49414 | 3.382 |  | 0.0056192
32 | 5.158 | 8.4054 | 8.4054 | 28.3303 | 28.3292 | 6.8297 | 2.0524 | 13 | 13 | 120 | 6.9265 | 0.74468 | 5.158 |  | 0.0056779
48 | 201.965 | 9.1059 | 9.1062 | 30.9194 | 30.92 | 6.6113 | 1.3904 | 5 | 5 | 156 | 6.8836 | 0.43931 | 3.024 |  | 0.005714
13 | 20.394 | 9.1448 | 9.1439 | 30.7339 | 30.7356 | 6.6135 | 2.2261 | 10 | 10 | 61 | 6.7287 | 0.83924 | 5.647 |  | 0.005749
26 | 10.437 | 8.6613 | 8.6602 | 30.9577 | 30.9591 | 6.6757 | 2.1262 | 9 | 9 | 105 | 6.7777 | 0.79156 | 5.365 |  | 0.0058807
1 | 20.013 | 9.0455 | 9.0457 | 29.7891 | 29.7904 | 6.6687 | 1.7357 | 6 | 6 | 6 | 6.7834 | 0.60501 | 4.104 | 6 | 0.0061704
66 | 50.154 | 8.5068 | 8.5066 | 28.9105 | 28.9121 | 6.7884 | 2.1623 | 10 | 10 | 250 | 6.9255 | 0.79677 | 5.518 |  | 0.006189
19 | 21.399 | 8.476 | 8.4766 | 31.1482 | 31.149 | 6.6952 | 2.0579 | 7 | 7 | 83 | 6.805 | 0.7593 | 5.167 |  | 0.0062683
7 | 20.166 | 10.1363 | 10.1415 | 30.0176 | 30.0204 | 6.5002 | 2.4296 | 11 | 11 | 34 | 6.6244 | 0.93745 | 6.21 |  | 0.0065818
43 | 2.643 | 7.0746 | 7.0782 | 22.1161 | 22.1326 | 7.3348 | 2.4658 | 16 | 16 | 151 | 7.4196 | 0.94102 | 6.982 |  | 0.0071483
39 | 51.306 | 9.0466 | 9.0455 | 29.8977 | 29.8978 | 6.6639 | 1.7226 | 6 | 6 | 128 | 6.8056 | 0.59995 | 4.083 |  | 0.0073753
7 | 50.712 | 10.259 | 10.2603 | 30.6628 | 30.663 | 6.4561 | 2.3708 | 8 | 8 | 31 | 6.6064 | 0.91033 | 6.014 |  | 0.0076224
39 | 6.199 | 8.2981 | 8.2991 | 27.4514 | 27.4517 | 6.8856 | 2.1948 | 11 | 11 | 133 | 6.9828 | 0.82618 | 5.769 |  | 0.0076604
51 | 342.907 | 9.0151 | 9.0156 | 31.1658 | 31.1661 | 6.6142 | 1.2565 | 1 | 1 | 169 | 7.0105 | 0.37729 | 2.645 |  | 0.0077427
51 | 40.545 | 9.1527 | 9.1437 | 29.7781 | 29.7703 | 6.6533 | 1.5928 | 11 | 11 | 179 | 6.7866 | 0.53826 | 3.653 |  | 0.007789
19 | 2.344 | 8.6702 | 8.6639 | 30.9835 | 30.9931 | 6.6732 | 2.1145 | 10 | 10 | 86 | 6.7683 | 0.78808 | 5.334 |  | 0.0079692
7 | 30.226 | 10.1942 | 10.1956 | 30.0594 | 30.0586 | 6.4903 | 2.4167 | 10 | 10 | 33 | 6.6233 | 0.93306 | 6.18 |  | 0.008377
51 | 10.101 | 8.1691 | 8.163 | 28.1925 | 28.19 | 6.8726 | 2.2816 | 14 | 14 | 182 | 6.9716 | 0.85102 | 5.933 |  | 0.0090262
7 | 100.739 | 8.2737 | 8.2749 | 31.6435 | 31.6419 | 6.7042 | 1.8948 | 6 | 6 | 29 | 6.8812 | 0.68418 | 4.708 |  | 0.0092277
32 | 48.746 | 8.8031 | 8.8037 | 30.295 | 30.2956 | 6.6831 | 1.8371 | 8 | 8 | 115 | 6.8202 | 0.65717 | 4.482 |  | 0.0098373
75 | 19.982 | 8.1779 | 8.1791 | 28.6439 | 28.645 | 6.8511 | 2.3049 | 8 | 8 | 279 | 6.9587 | 0.86137 | 5.994 |  | 0.0098466
60 | 49.548 | 8.5082 | 8.5091 | 28.9868 | 28.988 | 6.7848 | 2.138 | 7 | 7 | 222 | 6.9213 | 0.77976 | 5.397 |  | 0.011569
32 | 19.823 | 8.9964 | 8.9965 | 30.1637 | 30.1645 | 6.6599 | 1.7401 | 11 | 11 | 118 | 6.7737 | 0.61296 | 4.152 | 6 | 0.012008
32 | 39.355 | 8.7534 | 8.7504 | 30.3199 | 30.3217 | 6.6894 | 1.8214 | 9 | 9 | 116 | 6.8179 | 0.62761 | 4.279 |  | 0.012231
51 | 30.252 | 8.9777 | 8.9781 | 29.2822 | 29.283 | 6.7007 | 1.924 | 12 | 12 | 180 | 6.8241 | 0.67657 | 4.617 |  | 0.012345
75 | 123.889 | 8.7477 | 8.7496 | 30.0528 | 30.0581 | 6.7019 | 1.6216 | 2 | 2 | 273 | 6.9047 | 0.53138 | 3.669 |  | 0.012853
1 | 40.317 | 9.0128 | 9.0132 | 30.0291 | 30.0301 | 6.6632 | 1.6809 | 4 | 4 | 4 | 6.795 | 0.58705 | 3.989 |  | 0.014456
55 | 20.103 | 8.358 | 8.3585 | 28.3568 | 28.3573 | 6.8359 | 2.2257 | 12 | 12 | 196 | 6.9454 | 0.81881 | 5.687 |  | 0.014474
1 | 10.264 | 9.1511 | 9.1492 | 29.5123 | 29.5141 | 6.6649 | 1.765 | 7 | 7 | 7 | 6.7723 | 0.5964 | 4.039 |  | 0.016458
39 | 124.967 | 9.2393 | 9.2396 | 30.4443 | 30.4449 | 6.612 | 1.5568 | 3 | 3 | 125 | 6.8185 | 0.52988 | 3.613 | 6 | 0.016653
51 | 50.677 | 9.3597 | 9.3553 | 30.0485 | 30.044 | 6.6111 | 1.492 | 10 | 10 | 178 | 6.7548 | 0.49891 | 3.37 | 6 | 0.01671
55 | 125.565 | 9.2504 | 9.2508 | 30.6865 | 30.687 | 6.6 | 1.4381 | 6 | 6 | 190 | 6.8068 | 0.43794 | 2.981 |  | 0.018498
70 | 9.702 | 8.1881 | 8.1897 | 28.756 | 28.7544 | 6.8445 | 2.2853 | 13 | 13 | 269 | 6.943 | 0.84244 | 5.849 |  | 0.019376
1 | 101.242 | 9.1587 | 9.1589 | 31.141 | 31.1422 | 6.5941 | 0.5275 | 1 | 1 | 1 | 6.7787 | 0 | 0 |  | 0.020709
1 | 4.859 | 9.2033 | 9.2344 | 28.9362 | 28.9856 | 6.6819 | 1.7831 | 8 | 8 | 8 | 6.7855 | 0.59907 | 4.065 |  | 0.022451
55 | 99.847 | 9.3762 | 9.3765 | 30.4298 | 30.4309 | 6.5925 | 1.4714 | 7 | 7 | 191 | 6.7782 | 0.49939 | 3.385 |  | 0.027022
39 | 10.11 | 8.6746 | 8.6737 | 28.2004 | 28.1981 | 6.7939 | 2.024 | 10 | 10 | 132 | 6.8977 | 0.70777 | 4.882 | 6 | 0.029002
39 | 2.652 | 7.6373 | 7.6177 | 24.5539 | 24.4395 | 7.1242 | 2.2818 | 12 | 12 | 134 | 7.2134 | 0.82956 | 5.984 |  | 0.030564
1 | 2.771 | 8.291 | 8.3067 | 26.7798 | 26.8757 | 6.9168 | 2.0702 | 9 | 9 | 9 | 7.0113 | 0.72198 | 5.062 |  | 0.036911
48 | 10.132 | 7.893 | 7.8901 | 27.6366 | 27.6349 | 6.9412 | 2.3792 | 15 | 15 | 166 | 7.0379 | 0.85295 | 6.003 |  | 0.053775
1 | 75.755 | 8.862 | 8.8634 | 30.6399 | 30.6415 | 6.6594 | 1.365 | 2 | 2 | 2 | 6.8202 | 0.35307 | 2.408 |  | 0.068367
Flagged bottles
1 | 50.391 | 8.8381 | 8.8386 | 30.1928 | 30.1937 | 6.6823 | 1.6971 | 3 | 3 | 3 | 6.8212 | 0.58963 | 4.022 | 36 | 0.0092753
7 | 10.25 | 10.0819 | 10.0811 | 29.9909 | 29.9904 | 6.5091 | 2.4451 | 12 | 12 | 35 | 6.6244 | 0.94378 | 6.252 | 3 | 0.0054916
13 | 150.104 | 7.8516 | 7.8502 | 32.9088 | 32.9143 | 6.7128 | 1.4121 | 3 | 3 | 54 | 6.929 | 0.44869 | 3.109 | 36 | 0.0047261
13 | 50.669 | 8.8077 | 8.8098 | 31.1207 | 31.1196 | 6.6468 | 2.0886 | 7 | 7 | 58 | 6.7848 | 0.77275 | 5.243 | 3 | 0.005081
19 | 75.019 | 8.5006 | 8.5005 | 31.2429 | 31.2442 | 6.6874 | 2.02 | 3 | 3 | 79 | 6.8441 | 0.73816 | 5.052 | 3 | 0.0033079
32 | 200.139 | 8.4655 | 8.4659 | 30.7647 | 30.7652 | 6.7135 | 1.9049 | 2 | 2 | 109 | 6.9811 | 0.68012 | 4.748 | 3 | 0.00031995
48 | 251.901 | 9.0832 | 9.0836 | 31.0508 | 31.0515 | 6.609 | 1.3305 | 4 | 4 | 155 | 6.9249 | 0.40795 | 2.825 | 3 | 0.0030202
48 | 40.271 | 9.103 | 9.1035 | 29.5914 | 29.5924 | 6.6687 | 1.7298 | 12 | 12 | 163 | 6.8015 | 0.60016 | 4.082 | 3 | 0.0041339
48 | 19.971 | 8.6706 | 8.6708 | 28.7878 | 28.7876 | 6.7687 | 2.0203 | 14 | 14 | 165 | 6.8807 | 0.71475 | 4.918 | 36 | 0.020274
60 | 20.273 | 8.4403 | 8.44 | 28.6068 | 28.6044 | 6.8121 | 2.2355 | 10 | 10 | 225 | 6.9223 | 0.8461 | 5.857 | 3 | 0.0080908
75 | 40.086 | 8.3193 | 8.3189 | 29.056 | 29.0557 | 6.8109 | 2.1098 | 6 | 6 | 277 | 6.9373 | 0.77293 | 5.362 | 3 | 0.0048926

## ctd + bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comments
1 | 101.242 | 9.1587 | 9.1589 | 31.141 | 31.1422 | 6.5941 | 0.5275 | 1 | 1 | 1 | 1 | 0 |  | 9 | OXY: Anoxic
1 | 75.755 | 8.862 | 8.8634 | 30.6399 | 30.6415 | 6.6594 | 1.365 | 2 | 2 | 2 | 2 | 2.408 |  | 8.9
1 | 50.391 | 8.8381 | 8.8386 | 30.1928 | 30.1937 | 6.6823 | 1.6971 | 3 | 3 | 3 | 3 | 4.022 | 36 | 8.7 | OXY: Replicate outliers
1 | 40.317 | 9.0128 | 9.0132 | 30.0291 | 30.0301 | 6.6632 | 1.6809 | 4 | 4 | 4 | 4 | 3.989 |  | 8.7
1 | 29.971 | 8.9799 | 8.9803 | 29.8673 | 29.8685 | 6.6751 | 1.7346 | 5 | 5 | 5 | 5 | 4.049 |  | 8.7
1 | 20.013 | 9.0455 | 9.0457 | 29.7891 | 29.7904 | 6.6687 | 1.7357 | 6 | 6 | 6 | 6 | 4.104 | 6 | 8.7
1 | 10.264 | 9.1511 | 9.1492 | 29.5123 | 29.5141 | 6.6649 | 1.765 | 7 | 7 | 7 | 7 | 4.039 |  | 8.7
1 | 4.859 | 9.2033 | 9.2344 | 28.9362 | 28.9856 | 6.6819 | 1.7831 | 8 | 8 | 8 | 8 | 4.065 |  | 8.9
1 | 2.771 | 8.291 | 8.3067 | 26.7798 | 26.8757 | 6.9168 | 2.0702 | 9 | 9 | 9 | 9 | 5.062 |  | 8.4
3 | 222.97 | 8.3697 | 8.37 | 31.1327 | 31.133 | 6.7119 | 1.9597 | 1 | 1 | 10 | 10 | 4.955 |  | 8.4
3 | 201.317 | 8.3675 | 8.3678 | 31.1301 | 31.1305 | 6.7124 | 1.9644 | 2 | 2 | 11 | 11 | 4.955 |  | 8.3
3 | 177.799 | 8.3662 | 8.3665 | 31.0884 | 31.0886 | 6.7144 | 1.9508 | 3 | 3 | 12 | 12 | 4.898 |  | 8.3
3 | 151.683 | 8.3715 | 8.3717 | 31.0364 | 31.0379 | 6.7158 | 1.9314 | 4 | 4 | 13 | 13 | 4.799 |  | 8.3
3 | 126.227 | 8.3812 | 8.3817 | 30.9859 | 30.9864 | 6.7166 | 1.9266 | 5 | 5 | 14 | 14 | 4.792 | 6 | 8.2
3 | 100.975 | 8.3933 | 8.3936 | 30.9245 | 30.925 | 6.7174 | 1.9173 | 6 | 6 | 15 | 15 | 4.75 |  | 8.2
3 | 75.406 | 8.3915 | 8.3919 | 30.9184 | 30.9193 | 6.718 | 1.9202 | 7 | 7 | 16 | 16 | 4.733 |  | 8.1
3 | 50.563 | 8.393 | 8.3938 | 30.9047 | 30.9047 | 6.7183 | 1.9223 | 8 | 8 | 17 | 17 | 4.735 |  | 8.1
3 | 40.238 | 8.4056 | 8.4066 | 30.8756 | 30.8747 | 6.7177 | 1.9183 | 9 | 9 | 18 | 18 | 4.72 |  | 8.1
3 | 30.475 | 8.4141 | 8.4145 | 30.8528 | 30.8531 | 6.7174 | 1.9183 | 10 | 10 | 19 | 19 | -99 | 5 | 8 | OXY: Program error. Continuous dosing after selecting a random endpoint
3 | 20.095 | 8.4293 | 8.4297 | 30.7941 | 30.7948 | 6.7177 | 1.9191 | 11 | 11 | 20 | 20 | 4.705 |  | 8.1
3 | 9.813 | 8.4729 | 8.4747 | 30.5372 | 30.5314 | 6.7223 | 1.922 | 12 | 12 | 21 | 21 | 4.691 |  | 8.1
3 | 4.804 | 8.5006 | 8.5011 | 30.3882 | 30.3882 | 6.7246 | 1.9179 | 13 | 13 | 22 | 22 | 4.696 |  | 8.1
3 | 2.741 | 8.5009 | 8.5012 | 30.3864 | 30.3871 | 6.7246 | 1.9178 | 14 | 14 | 23 | 23 | 4.692 |  | 8
7 | 239.542 | 7.4588 | 7.4593 | 33.6624 | 33.6626 | 6.7397 | 1.146 | 1 | 1 | 24 | 24 | -99 | 5 | 7.7 | OXY: Program error. Continuous dosing after selecting a random endpoint
7 | 197.278 | 7.7191 | 7.7196 | 33.5086 | 33.5087 | 6.7067 | 1.2766 | 2 | 2 | 25 | 25 | -99 | 5 | 7.9 | OXY: Program error. Continuous dosing after selecting a random endpoint
7 | 175.536 | 7.7285 | 7.7291 | 33.4685 | 33.4671 | 6.707 | 1.2771 | 3 | 3 | 26 | 26 | -99 | 5 | 7.8 | OXY: Program error. Continuous dosing after selecting a random endpoint
7 | 163.853 | 7.7733 | 7.7737 | 33.3959 | 33.3963 | 6.7034 | 1.3148 | 4 | 4 | 27 | 27 | -99 | 5 | 7.9 | OXY: Program error. Continuous dosing after selecting a random endpoint
7 | 129.571 | 8.0357 | 8.0334 | 32.3334 | 32.3422 | 6.7101 | 1.6038 | 5 | 5 | 28 | 28 | 3.74 | 6 | 8
7 | 100.739 | 8.2737 | 8.2749 | 31.6435 | 31.6419 | 6.7042 | 1.8948 | 6 | 6 | 29 | 29 | 4.708 |  | 8.2
7 | 76.691 | 9.1035 | 9.1052 | 31.2832 | 31.2832 | 6.5961 | 2.138 | 7 | 7 | 30 | 30 | 5.385 |  | 8.7
7 | 50.712 | 10.259 | 10.2603 | 30.6628 | 30.663 | 6.4561 | 2.3708 | 8 | 8 | 31 | 31 | 6.014 |  | 9.5
7 | 40.211 | 10.3284 | 10.3286 | 30.3621 | 30.362 | 6.4588 | 2.4042 | 9 | 9 | 32 | 32 | 6.093 |  | 9.5
7 | 30.226 | 10.1942 | 10.1956 | 30.0594 | 30.0586 | 6.4903 | 2.4167 | 10 | 10 | 33 | 33 | 6.18 |  | 9.4
7 | 20.166 | 10.1363 | 10.1415 | 30.0176 | 30.0204 | 6.5002 | 2.4296 | 11 | 11 | 34 | 34 | 6.21 |  | 9.4
7 | 10.25 | 10.0819 | 10.0811 | 29.9909 | 29.9904 | 6.5091 | 2.4451 | 12 | 12 | 35 | 35 | 6.252 | 3 | 9.4 | OXY: curve did not flatten, spiked with 1 mL std and corrected final value
7 | 5.008 | 10.0616 | 10.0619 | 29.9751 | 29.9762 | 6.5126 | 2.4489 | 13 | 13 | 36 | 36 | 6.255 |  | 9.3
7 | 2.601 | 10.0555 | 10.0558 | 29.9743 | 29.9754 | 6.5136 | 2.4512 | 14 | 14 | 37 | 37 | 6.261 |  | 9.3
13 | 184.21 | 7.7148 | 7.7154 | 33.3029 | 33.303 | 6.7164 | 1.2785 | 1 | 1 | 52 | 52 | -99 | 5 | 7.9 | OXY: Program crash. Nothing recorded
13 | 175.581 | 7.7141 | 7.7147 | 33.3 | 33.3004 | 6.7166 | 1.2794 | 2 | 2 | 53 | 53 | 2.629 |  | 8
13 | 150.104 | 7.8516 | 7.8502 | 32.9088 | 32.9143 | 6.7128 | 1.4121 | 3 | 3 | 54 | 54 | 3.109 | 36 | 8.1 | OXY: Leaky Niskin. Top valve did not seal.
13 | 125.092 | 8.1511 | 8.1513 | 32.1327 | 32.1334 | 6.7014 | 1.6962 | 4 | 4 | 55 | 55 | 4.023 |  | 8.3
13 | 99.481 | 8.3933 | 8.3933 | 31.6083 | 31.6094 | 6.6876 | 1.8841 | 5 | 5 | 56 | 56 | 4.632 |  | 8.5
13 | 75.96 | 8.4619 | 8.4624 | 31.4604 | 31.4608 | 6.6838 | 1.938 | 6 | 6 | 57 | 57 | 4.783 |  | 8.5
13 | 50.669 | 8.8077 | 8.8098 | 31.1207 | 31.1196 | 6.6468 | 2.0886 | 7 | 7 | 58 | 58 | 5.243 | 3 | 8.7 | OXY: Spigot open
13 | 38.866 | 8.8697 | 8.8696 | 31.0254 | 31.0264 | 6.6417 | 2.108 | 8 | 8 | 59 | 59 | 5.28 |  | 8.7
13 | 31.456 | 8.96 | 8.9607 | 30.9429 | 30.9421 | 6.6318 | 2.1364 | 9 | 9 | 60 | 60 | 5.378 | 6 | 8.8
13 | 20.394 | 9.1448 | 9.1439 | 30.7339 | 30.7356 | 6.6135 | 2.2261 | 10 | 10 | 61 | 61 | 5.647 |  | 8.9
13 | 10.059 | 9.3062 | 9.3064 | 30.3485 | 30.3496 | 6.6062 | 2.3455 | 11 | 11 | 62 | 62 | 5.996 |  | 9
13 | 5.306 | 9.2922 | 9.293 | 30.2508 | 30.2516 | 6.6124 | 2.3722 | 12 | 12 | 63 | 63 | 6.083 |  | 8.9
13 | 2.586 | 9.269 | 9.2691 | 30.2411 | 30.2414 | 6.6163 | 2.3782 | 13 | 13 | 64 | 64 | 6.104 |  | 8.9
16 | 174.404 | 7.9033 | 7.9037 | 32.7487 | 32.749 | 6.712 | 1.4625 | 1 | 1 | 65 | 65 | 3.236 |  | 8.2
16 | 150.202 | 7.9316 | 7.9321 | 32.6716 | 32.6716 | 6.7111 | 1.4923 | 2 | 2 | 66 | 66 | 3.35 |  | 8.1
16 | 125.503 | 8.2224 | 8.2223 | 32.0488 | 32.0497 | 6.6943 | 1.7322 | 3 | 3 | 67 | 67 | 4.106 |  | 8.4
16 | 101.392 | 8.5119 | 8.5125 | 31.4929 | 31.4931 | 6.6748 | 1.9432 | 4 | 4 | 68 | 68 | 4.792 |  | 8.7
16 | 76.231 | 8.4871 | 8.4875 | 31.4321 | 31.4325 | 6.6812 | 1.9476 | 5 | 5 | 69 | 69 | 4.795 |  | 8.7
16 | 49.954 | 8.7989 | 8.7982 | 31.0443 | 31.0455 | 6.6514 | 2.0894 | 6 | 6 | 70 | 70 | 5.232 |  | 8.9
16 | 40.648 | 8.9377 | 8.9376 | 30.9683 | 30.969 | 6.634 | 2.1249 | 7 | 7 | 71 | 71 | 5.327 | 6 | 9
16 | 30.343 | 9.0498 | 9.0504 | 30.9051 | 30.9054 | 6.6202 | 2.1535 | 8 | 8 | 72 | 72 | 5.406 |  | 9
16 | 19.99 | 9.1345 | 9.1428 | 30.8183 | 30.8066 | 6.6114 | 2.1797 | 9 | 9 | 73 | 73 | 5.488 |  | 9.1
16 | 10.174 | 9.9065 | 9.9086 | 30.088 | 30.0919 | 6.5301 | 2.3831 | 10 | 10 | 74 | 74 | 6.058 |  | 9.7
16 | 4.58 | 9.8633 | 9.864 | 29.8337 | 29.8346 | 6.547 | 2.4264 | 11 | 11 | 75 | 75 | 6.21 |  | 9.5
16 | 2.673 | 9.804 | 9.8052 | 29.7865 | 29.7878 | 6.5575 | 2.4309 | 12 | 12 | 76 | 76 | 6.234 |  | 9.5
19 | 116.354 | 8.4966 | 8.4968 | 31.2445 | 31.245 | 6.6879 | 2.0106 | 1 | 1 | 77 | 77 | 5.046 |  | 8.7
19 | 102.344 | 8.4965 | 8.4968 | 31.2469 | 31.2476 | 6.6878 | 2.0146 | 2 | 2 | 78 | 78 | 5.047 |  | 8.7
19 | 75.019 | 8.5006 | 8.5005 | 31.2429 | 31.2442 | 6.6874 | 2.02 | 3 | 3 | 79 | 79 | 5.052 | 3 | 8.7 | OXY: Bubble under stopper
19 | 50.783 | 8.5193 | 8.5199 | 31.2083 | 31.2091 | 6.6861 | 2.0401 | 4 | 4 | 80 | 80 | 5.115 |  | 8.7
19 | 40.648 | 8.5225 | 8.5225 | 31.172 | 31.1719 | 6.6872 | 2.0521 | 5 | 5 | 81 | 81 | 5.146 | 6 | 8.7
19 | 31.659 | 8.5093 | 8.5095 | 31.1612 | 31.1619 | 6.6896 | 2.0546 | 6 | 6 | 82 | 82 | 5.146 |  | 8.7
19 | 21.399 | 8.476 | 8.4766 | 31.1482 | 31.149 | 6.6952 | 2.0579 | 7 | 7 | 83 | 83 | 5.167 |  | 8.6
19 | 10.659 | 8.5459 | 8.5467 | 31.0852 | 31.0859 | 6.6874 | 2.0777 | 8 | 8 | 84 | 84 | 5.201 |  | 8.7
19 | 4.961 | 8.5964 | 8.5985 | 31.0652 | 31.0645 | 6.6807 | 2.0869 | 9 | 9 | 85 | 85 | 5.232 |  | 8.8
19 | 2.344 | 8.6702 | 8.6639 | 30.9835 | 30.9931 | 6.6732 | 2.1145 | 10 | 10 | 86 | 86 | 5.334 |  | 8.8
26 | 146.042 | 8.4948 | 8.4951 | 31.1083 | 31.109 | 6.6941 | 2.0293 | 1 | 1 | 97 | 97 | 5.117 |  | 8.7
26 | 125.728 | 8.5056 | 8.5062 | 31.1024 | 31.1028 | 6.6927 | 2.038 | 2 | 2 | 98 | 98 | 5.129 |  | 8.7
26 | 100.565 | 8.5159 | 8.5156 | 31.0911 | 31.0929 | 6.6917 | 2.0492 | 3 | 3 | 99 | 99 | 5.167 |  | 8.7
26 | 74.786 | 8.5581 | 8.5582 | 31.041 | 31.0419 | 6.6875 | 2.0706 | 4 | 4 | 100 | 100 | 5.22 | 6 | 8.7
26 | 50.694 | 8.6329 | 8.6313 | 30.974 | 30.9757 | 6.6792 | 2.1023 | 5 | 5 | 101 | 101 | 5.298 |  | 8.8
26 | 40.391 | 8.6067 | 8.6088 | 30.9905 | 30.9902 | 6.6824 | 2.1036 | 6 | 6 | 102 | 102 | 5.312 |  | 8.8
26 | 31.385 | 8.6103 | 8.6085 | 30.9881 | 30.9905 | 6.682 | 2.0982 | 7 | 7 | 103 | 103 | 5.272 |  | 8.7
26 | 20.514 | 8.6142 | 8.6142 | 30.9813 | 30.9822 | 6.6817 | 2.105 | 8 | 8 | 104 | 104 | 5.302 |  | 8.7
26 | 10.437 | 8.6613 | 8.6602 | 30.9577 | 30.9591 | 6.6757 | 2.1262 | 9 | 9 | 105 | 105 | 5.365 |  | 8.7
26 | 5.534 | 8.7199 | 8.7217 | 30.9299 | 30.9298 | 6.6681 | 2.1696 | 10 | 10 | 106 | 106 | 5.489 |  | 8.8
26 | 1.975 | 8.7153 | 8.7159 | 30.9297 | 30.9304 | 6.6688 | 2.1607 | 11 | 11 | 107 | 107 | 5.448 |  | 8.8
32 | 216.77 | 8.4627 | 8.4634 | 30.7827 | 30.7817 | 6.7131 | 1.9042 | 1 | 1 | 108 | 108 | 4.754 |  | 8.7
32 | 200.139 | 8.4655 | 8.4659 | 30.7647 | 30.7652 | 6.7135 | 1.9049 | 2 | 2 | 109 | 109 | 4.748 | 3 | 8.7 | OXY: Bubble from H2SO4 Dispenser
32 | 174.486 | 8.4653 | 8.4658 | 30.7471 | 30.7474 | 6.7143 | 1.9111 | 3 | 3 | 110 | 110 | 4.757 |  | 8.6
32 | 149.739 | 8.4766 | 8.477 | 30.6973 | 30.6979 | 6.7147 | 1.9135 | 4 | 4 | 111 | 111 | 4.748 | 6 | 8.6
32 | 123.134 | 8.4859 | 8.4865 | 30.6611 | 30.661 | 6.7149 | 1.9132 | 5 | 5 | 112 | 112 | 4.739 |  | 8.6
32 | 100.025 | 8.509 | 8.5092 | 30.5884 | 30.5897 | 6.7146 | 1.9129 | 6 | 6 | 113 | 113 | 4.72 |  | 8.6
32 | 76.071 | 8.5282 | 8.5286 | 30.5106 | 30.5105 | 6.7151 | 1.9112 | 7 | 7 | 114 | 114 | 4.701 |  | 8.6
32 | 48.746 | 8.8031 | 8.8037 | 30.295 | 30.2956 | 6.6831 | 1.8371 | 8 | 8 | 115 | 115 | 4.482 |  | 8.6
32 | 39.355 | 8.7534 | 8.7504 | 30.3199 | 30.3217 | 6.6894 | 1.8214 | 9 | 9 | 116 | 116 | 4.279 |  | 8.8
32 | 31.839 | 8.9183 | 8.92 | 30.2717 | 30.2719 | 6.6669 | 1.7692 | 10 | 10 | 117 | 117 | 4.174 |  | 8.8
32 | 19.823 | 8.9964 | 8.9965 | 30.1637 | 30.1645 | 6.6599 | 1.7401 | 11 | 11 | 118 | 118 | 4.152 | 6 | 9.1
32 | 10.201 | 8.5741 | 8.5767 | 28.8795 | 28.8876 | 6.7795 | 1.9695 | 12 | 12 | 119 | 119 | 4.904 |  | 8.5 | OXY: Spigot open
32 | 5.158 | 8.4054 | 8.4054 | 28.3303 | 28.3292 | 6.8297 | 2.0524 | 13 | 13 | 120 | 120 | 5.158 |  | 8.6
32 | 2.571 | 8.392 | 8.3919 | 28.2703 | 28.2691 | 6.8344 | 2.0651 | 14 | 14 | 121 | 121 | 5.222 |  | 8.5
32 | 2.535 | 8.3924 | 8.3924 | 28.2695 | 28.2688 | 6.8344 | 2.0694 | 15 | 15 | 122 | 123 | 4.506 |  | 8.9
39 | 177.901 | 8.6129 | 8.6133 | 30.5086 | 30.5091 | 6.7024 | 1.8334 | 1 | 1 | 123 | 124 | 4.527 |  | 8.7
39 | 149.733 | 8.6202 | 8.6207 | 30.4585 | 30.4589 | 6.7035 | 1.8512 | 2 | 2 | 124 | 125 | 3.613 | 6 | 9.1
39 | 124.967 | 9.2393 | 9.2396 | 30.4443 | 30.4449 | 6.612 | 1.5568 | 3 | 3 | 125 | 126 | 3.453 |  | 9.2
39 | 100.563 | 9.3462 | 9.3465 | 30.3338 | 30.3344 | 6.601 | 1.5486 | 4 | 4 | 126 | 127 | 3.462 |  | 9.2
39 | 74.288 | 9.3702 | 9.3714 | 30.1716 | 30.1715 | 6.6044 | 1.5551 | 5 | 5 | 127 | 128 | 4.083 |  | 8.9
39 | 51.306 | 9.0466 | 9.0455 | 29.8977 | 29.8978 | 6.6639 | 1.7226 | 6 | 6 | 128 | 129 | 4.437 |  | 8.6
39 | 40.69 | 8.8025 | 8.801 | 29.7103 | 29.7108 | 6.7085 | 1.8379 | 7 | 7 | 129 | 130 | 4.263 |  | 8.8
39 | 29.768 | 8.9617 | 8.9608 | 29.5863 | 29.5865 | 6.69 | 1.7908 | 8 | 8 | 130 | 131 | 4.207 |  | 8.9
39 | 20.205 | 9.0485 | 9.0453 | 29.5435 | 29.5457 | 6.6788 | 1.7673 | 9 | 9 | 131 | 132 | 4.882 | 6 | 8.8
39 | 10.11 | 8.6746 | 8.6737 | 28.2004 | 28.1981 | 6.7939 | 2.024 | 10 | 10 | 132 | 133 | 5.769 |  | 8.2
39 | 6.199 | 8.2981 | 8.2991 | 27.4514 | 27.4517 | 6.8856 | 2.1948 | 11 | 11 | 133 | 134 | 5.984 |  | 8.1
39 | 2.652 | 7.6373 | 7.6177 | 24.5539 | 24.4395 | 7.1242 | 2.2818 | 12 | 12 | 134 | 136 | 2.56 |  | 9
39 | 2.65 | 7.9358 | 7.8985 | 26.1861 | 25.953 | 7.0004 | 2.3178 | 13 | 13 | 135 | 137 | 2.596 |  | 9.1
43 | 316.402 | 9.0419 | 9.0422 | 31.1362 | 31.1367 | 6.6115 | 1.253 | 1 | 1 | 136 | 138 | 2.794 | 6 | 8.9
43 | 299.799 | 9.0466 | 9.047 | 31.1231 | 31.1235 | 6.6113 | 1.2642 | 2 | 2 | 137 | 139 | 2.991 |  | 8.9
43 | 249.978 | 9.0847 | 9.0852 | 31.0333 | 31.0338 | 6.6096 | 1.3289 | 3 | 3 | 138 | 140 | 3.11 |  | 8.9
43 | 199.575 | 9.1478 | 9.1483 | 30.8869 | 30.8873 | 6.6065 | 1.3997 | 4 | 4 | 139 | 141 | 3.302 |  | 8.8
43 | 174.572 | 9.1844 | 9.1848 | 30.7917 | 30.7924 | 6.6052 | 1.4319 | 5 | 5 | 140 | 142 | 3.578 |  | 8.8
43 | 150.057 | 9.1575 | 9.158 | 30.7099 | 30.7106 | 6.6127 | 1.4884 | 6 | 6 | 141 | 143 | 3.493 |  | 8.9
43 | 124.794 | 9.1582 | 9.1595 | 30.5001 | 30.5003 | 6.6215 | 1.575 | 7 | 7 | 142 | 144 | 3.361 |  | 8.9
43 | 100.438 | 9.2773 | 9.2777 | 30.3511 | 30.3514 | 6.6104 | 1.5583 | 8 | 8 | 143 | 145 | 3.432 |  | 8.9
43 | 75.187 | 9.386 | 9.3863 | 30.2218 | 30.222 | 6.5999 | 1.5206 | 9 | 9 | 144 | 146 | 3.933 | 6 | 8.8
43 | 50.274 | 9.3636 | 9.364 | 29.9248 | 29.9251 | 6.6159 | 1.5455 | 10 | 10 | 145 | 147 | 4.256 |  | 8.8
43 | 39.997 | 9.2045 | 9.2047 | 29.6861 | 29.6866 | 6.6495 | 1.702 | 11 | 11 | 146 | 148 | 4.575 |  | 8.6
43 | 30.495 | 9.153 | 9.1523 | 29.3623 | 29.3574 | 6.6711 | 1.7937 | 12 | 12 | 147 | 149 | 5.804 |  | 8.2
43 | 19.805 | 9.0088 | 9.0092 | 28.9232 | 28.9222 | 6.7116 | 1.8927 | 13 | 13 | 148 | 150 | 6.953 |  | 7.4
43 | 9.636 | 8.2296 | 8.2297 | 27.0374 | 27.0369 | 6.9149 | 2.2137 | 14 | 14 | 149 | 151 | 6.982 |  | 7.4
43 | 5.382 | 7.2204 | 7.2195 | 22.882 | 22.8785 | 7.273 | 2.4665 | 15 | 15 | 150 | 152 | 2.503 |  | 8.5
43 | 2.643 | 7.0746 | 7.0782 | 22.1161 | 22.1326 | 7.3348 | 2.4658 | 16 | 16 | 151 | 153 | 2.525 |  | 8.6
48 | 378.184 | 9.0143 | 9.0147 | 31.1859 | 31.1864 | 6.6134 | 1.2283 | 1 | 1 | 152 | 154 | 2.63 | 6 | 8.7
48 | 353.272 | 9.0303 | 9.0307 | 31.1644 | 31.1649 | 6.612 | 1.2361 | 2 | 2 | 153 | 155 | 2.825 | 3 | 8.6 | OXY: Bubbles under stopper
48 | 302.913 | 9.0616 | 9.0621 | 31.112 | 31.1123 | 6.6096 | 1.2767 | 3 | 3 | 154 | 156 | 3.024 |  | 8.6
48 | 251.901 | 9.0832 | 9.0836 | 31.0508 | 31.0515 | 6.609 | 1.3305 | 4 | 4 | 155 | 157 | 3.094 |  | 8.6
48 | 201.965 | 9.1059 | 9.1062 | 30.9194 | 30.92 | 6.6113 | 1.3904 | 5 | 5 | 156 | 158 | 3.55 |  | 8.5
48 | 176.432 | 9.1046 | 9.105 | 30.8697 | 30.8703 | 6.6136 | 1.42 | 6 | 6 | 157 | 159 | 3.595 |  | 8.7
48 | 151.041 | 8.9749 | 8.9764 | 30.7185 | 30.719 | 6.6392 | 1.559 | 7 | 7 | 158 | 160 | 3.414 |  | 8.9
48 | 126.49 | 9.0748 | 9.0755 | 30.5708 | 30.5712 | 6.6308 | 1.5755 | 8 | 8 | 159 | 161 | 3.174 |  | 8.9
48 | 100.638 | 9.3588 | 9.3586 | 30.3202 | 30.3213 | 6.5997 | 1.5366 | 9 | 9 | 160 | 162 | 3.831 |  | 8.7
48 | 76.2 | 9.2776 | 9.2783 | 30.1279 | 30.1285 | 6.6198 | 1.4566 | 10 | 10 | 161 | 163 | 4.082 | 3 | 8.7 | OXY: Small bubble in dose tip. Flushed before next sample
48 | 49.975 | 9.1213 | 9.1218 | 29.6865 | 29.6875 | 6.6619 | 1.6697 | 11 | 11 | 162 | 164 | 4.42 |  | 8.7
48 | 40.271 | 9.103 | 9.1035 | 29.5914 | 29.5924 | 6.6687 | 1.7298 | 12 | 12 | 163 | 165 | 4.918 | 36 | 8.7 | OXY: 1st rep: Titration ended at 89% and at a dose rate of 0.04 mL/min (Too high). Endpoint likely overshot.
48 | 29.948 | 8.9624 | 8.9634 | 29.3351 | 29.3367 | 6.7007 | 1.8455 | 13 | 13 | 164 | 166 | 6.003 |  | 8.2
48 | 19.971 | 8.6706 | 8.6708 | 28.7878 | 28.7876 | 6.7687 | 2.0203 | 14 | 14 | 165 | 167 | 6.528 |  | 7.9
48 | 10.132 | 7.893 | 7.8901 | 27.6366 | 27.6349 | 6.9412 | 2.3792 | 15 | 15 | 166 | 168 | 6.586 |  | 7.9
48 | 4.958 | 7.6891 | 7.6918 | 27.3991 | 27.4025 | 6.9846 | 2.4213 | 16 | 16 | 167 | 169 | 2.645 |  | 9
48 | 2.777 | 7.5686 | 7.5548 | 27.1257 | 27.0961 | 7.0166 | 2.4251 | 17 | 17 | 168 | 170 | 2.784 |  | 9 | OXY: Bubble under stopper
51 | 342.907 | 9.0151 | 9.0156 | 31.1658 | 31.1661 | 6.6142 | 1.2565 | 1 | 1 | 169 | 171 | 2.849 |  | 9
51 | 299.965 | 9.0707 | 9.0711 | 31.064 | 31.0645 | 6.6103 | 1.3168 | 2 | 2 | 170 | 172 | 3.349 | 6 | 8.9
51 | 249.566 | 9.0835 | 9.084 | 31.0268 | 31.0273 | 6.61 | 1.3374 | 3 | 3 | 171 | 173 | 3.369 |  | 8.9
51 | 201.255 | 9.0049 | 9.0044 | 30.8799 | 30.8798 | 6.6279 | 1.4887 | 4 | 4 | 172 | 174 | 3.382 |  | 8.9
51 | 175.851 | 9.0525 | 9.0528 | 30.8077 | 30.8083 | 6.6239 | 1.4983 | 5 | 5 | 173 | 175 | 3.359 |  | 8.9
51 | 150.283 | 9.1226 | 9.1239 | 30.7345 | 30.7349 | 6.6167 | 1.5052 | 6 | 6 | 174 | 176 | 3.165 |  | 9
51 | 125.512 | 9.2019 | 9.2103 | 30.6163 | 30.6171 | 6.6101 | 1.5067 | 7 | 7 | 175 | 177 | 3.238 |  | 9
51 | 100.961 | 9.3461 | 9.3466 | 30.5373 | 30.5378 | 6.5923 | 1.4514 | 8 | 8 | 176 | 178 | 3.37 | 6 | 8.9
51 | 75.318 | 9.4023 | 9.4028 | 30.3553 | 30.3558 | 6.5919 | 1.4761 | 9 | 9 | 177 | 179 | 3.653 |  | 8.7
51 | 50.677 | 9.3597 | 9.3553 | 30.0485 | 30.044 | 6.6111 | 1.492 | 10 | 10 | 178 | 180 | 4.617 |  | 8.6
51 | 40.545 | 9.1527 | 9.1437 | 29.7781 | 29.7703 | 6.6533 | 1.5928 | 11 | 11 | 179 | 181 | 4.95 |  | 8.5 | OXY: Bubble under stopper
51 | 30.252 | 8.9777 | 8.9781 | 29.2822 | 29.283 | 6.7007 | 1.924 | 12 | 12 | 180 | 182 | 5.933 |  | 8
51 | 20.551 | 8.8022 | 8.8005 | 28.9864 | 28.9834 | 6.7401 | 1.982 | 13 | 13 | 181 | 183 | 6.328 |  | 7.7
51 | 10.101 | 8.1691 | 8.163 | 28.1925 | 28.19 | 6.8726 | 2.2816 | 14 | 14 | 182 | 184 | 6.466 |  | 7.5
51 | 5.535 | 7.8229 | 7.8182 | 27.8336 | 27.821 | 6.9435 | 2.3654 | 15 | 15 | 183 | 185 | 2.728 |  | 9
51 | 2.691 | 7.5448 | 7.5494 | 27.1139 | 27.1221 | 7.021 | 2.3852 | 16 | 16 | 184 | 186 | 2.787 |  | 9.1
55 | 275.596 | 9.0085 | 9.0089 | 31.0813 | 31.0817 | 6.6187 | 1.2951 | 1 | 1 | 185 | 187 | 2.871 |  | 9.1
55 | 250.74 | 9.0604 | 9.0611 | 31.0039 | 31.0039 | 6.6144 | 1.3257 | 2 | 2 | 186 | 188 | 2.895 |  | 9.1
55 | 199.893 | 9.0953 | 9.0957 | 30.9319 | 30.9323 | 6.6123 | 1.3514 | 3 | 3 | 187 | 189 | 3.009 | 6 | 9.2
55 | 175.532 | 9.1268 | 9.1276 | 30.8677 | 30.8677 | 6.6104 | 1.3616 | 4 | 4 | 188 | 190 | 2.981 |  | 9.1
55 | 150.578 | 9.2067 | 9.2072 | 30.7787 | 30.779 | 6.6025 | 1.3992 | 5 | 5 | 189 | 191 | 3.385 |  | 9.1
55 | 125.565 | 9.2504 | 9.2508 | 30.6865 | 30.687 | 6.6 | 1.4381 | 6 | 6 | 190 | 192 | 3.244 |  | 9.1
55 | 99.847 | 9.3762 | 9.3765 | 30.4298 | 30.4309 | 6.5925 | 1.4714 | 7 | 7 | 191 | 193 | 5.373 | 6 | 8.7
55 | 76.002 | 9.2052 | 9.2051 | 30.1128 | 30.1132 | 6.6311 | 1.4869 | 8 | 8 | 192 | 194 | 5.341 |  | 8.6
55 | 50.679 | 8.6402 | 8.6407 | 28.9467 | 28.9474 | 6.7664 | 2.1178 | 9 | 9 | 193 | 195 | 5.458 |  | 8.6
55 | 40.2 | 8.5819 | 8.5865 | 28.7138 | 28.7203 | 6.7856 | 2.0993 | 10 | 10 | 194 | 196 | 5.687 |  | 8.4
55 | 30.038 | 8.4892 | 8.4902 | 28.6295 | 28.6305 | 6.8035 | 2.1363 | 11 | 11 | 195 | 197 | 6.045 |  | 8.2
55 | 20.103 | 8.358 | 8.3585 | 28.3568 | 28.3573 | 6.8359 | 2.2257 | 12 | 12 | 196 | 198 | 6.16 |  | 8.1
55 | 9.79 | 8.1516 | 8.1474 | 28.0734 | 28.0723 | 6.8807 | 2.3008 | 13 | 13 | 197 | 199 | 6.322 |  | 7.9
55 | 4.877 | 8.0614 | 8.0851 | 27.907 | 27.9138 | 6.9024 | 2.3159 | 14 | 14 | 198 | 216 | 2.746 |  | 9.1
55 | 2.507 | 7.7222 | 7.7279 | 27.5439 | 27.5497 | 6.9727 | 2.3627 | 15 | 15 | 199 | 217 | 2.789 |  | 9
60 | 189.595 | 8.9982 | 8.9986 | 30.9633 | 30.9638 | 6.6253 | 1.3184 | 1 | 1 | 216 | 218 | 2.824 |  | 9 | OXY: Small bubble under stopper
60 | 174.484 | 9.0335 | 9.0338 | 30.9076 | 30.9083 | 6.6225 | 1.333 | 2 | 2 | 217 | 219 | 2.947 |  | 9
60 | 150.202 | 9.0831 | 9.0834 | 30.8452 | 30.8458 | 6.6178 | 1.3483 | 3 | 3 | 218 | 220 | 2.909 | 6 | 8.9
60 | 124.916 | 9.2275 | 9.2278 | 30.6297 | 30.6306 | 6.6058 | 1.3861 | 4 | 4 | 219 | 221 | 3.765 |  | 8.7
60 | 99.941 | 9.148 | 9.1483 | 30.3127 | 30.3133 | 6.631 | 1.3781 | 5 | 5 | 220 | 222 | 5.397 |  | 8.3
60 | 75.34 | 8.9386 | 8.9394 | 29.7919 | 29.7943 | 6.6845 | 1.6438 | 6 | 6 | 221 | 223 | 5.77 |  | 8.2
60 | 49.548 | 8.5082 | 8.5091 | 28.9868 | 28.988 | 6.7848 | 2.138 | 7 | 7 | 222 | 224 | 5.834 | 6 | 8.3
60 | 39.788 | 8.4309 | 8.4306 | 28.8298 | 28.8296 | 6.8037 | 2.2196 | 8 | 8 | 223 | 225 | 5.857 | 3 | 8.1 | OXY: Odd curve. Ambiguous end point but best chosen.
60 | 29.981 | 8.4366 | 8.4372 | 28.7804 | 28.7807 | 6.805 | 2.2424 | 9 | 9 | 224 | 226 | 6.133 |  | 8
60 | 20.273 | 8.4403 | 8.44 | 28.6068 | 28.6044 | 6.8121 | 2.2355 | 10 | 10 | 225 | 227 | 6.269 |  | 7.9
60 | 9.518 | 8.1618 | 8.1483 | 28.1737 | 28.1662 | 6.8746 | 2.3253 | 11 | 11 | 226 | 228 | 6.269 |  | 8
60 | 5.131 | 8.0819 | 8.0828 | 28.1306 | 28.1314 | 6.8891 | 2.3575 | 12 | 12 | 227 | 241 | 2.73 |  | 9.1
60 | 2.383 | 8.0808 | 8.0819 | 28.1306 | 28.1312 | 6.8893 | 2.3599 | 13 | 13 | 228 | 242 | 2.772 |  | 9.1
66 | 351.312 | 9.069 | 9.0683 | 30.7952 | 30.7966 | 6.622 | 1.2997 | 1 | 1 | 241 | 243 | 2.785 | 6 | 9
66 | 299.961 | 9.0923 | 9.0922 | 30.79 | 30.7913 | 6.6188 | 1.317 | 2 | 2 | 242 | 244 | 2.844 |  | 9
66 | 250.077 | 9.0942 | 9.0941 | 30.7787 | 30.7801 | 6.619 | 1.324 | 3 | 3 | 243 | 245 | 2.846 |  | 8.9
66 | 199.666 | 9.1267 | 9.1267 | 30.7202 | 30.7216 | 6.6167 | 1.3461 | 4 | 4 | 244 | 246 | 2.864 |  | 8.9
66 | 174.996 | 9.1319 | 9.1318 | 30.7001 | 30.7015 | 6.6169 | 1.3506 | 5 | 5 | 245 | 247 | 2.889 |  | 8.8
66 | 149.819 | 9.1584 | 9.1584 | 30.6398 | 30.6412 | 6.6155 | 1.3594 | 6 | 6 | 246 | 248 | 2.923 |  | 8.8
66 | 125.028 | 9.1872 | 9.187 | 30.529 | 30.5309 | 6.616 | 1.3694 | 7 | 7 | 247 | 249 | 3.312 |  | 8.6
66 | 99.935 | 9.1651 | 9.1652 | 30.3829 | 30.3847 | 6.6255 | 1.3804 | 8 | 8 | 248 | 250 | 5.518 |  | 8.3
66 | 75.702 | 8.9979 | 8.9979 | 30.0555 | 30.0569 | 6.6643 | 1.4996 | 9 | 9 | 249 | 251 | 5.809 | 6 | 8.1
66 | 50.154 | 8.5068 | 8.5066 | 28.9105 | 28.9121 | 6.7884 | 2.1623 | 10 | 10 | 250 | 252 | 5.914 |  | 8.2
66 | 39.895 | 8.4353 | 8.4352 | 28.7499 | 28.7505 | 6.8065 | 2.2306 | 11 | 11 | 251 | 253 | 5.953 |  | 7.9
66 | 29.919 | 8.3976 | 8.3975 | 28.6928 | 28.6939 | 6.8148 | 2.2649 | 12 | 12 | 252 | 254 | 6.1 |  | 7.9
66 | 19.853 | 8.3542 | 8.3536 | 28.5597 | 28.5602 | 6.8274 | 2.2734 | 13 | 13 | 253 | 255 | 6.218 |  | 7.7
66 | 10.29 | 8.2667 | 8.2676 | 28.4458 | 28.4487 | 6.8461 | 2.3118 | 14 | 14 | 254 | 256 | 6.237 |  | 7.8
66 | 5.357 | 8.0748 | 8.0745 | 28.2962 | 28.2977 | 6.8828 | 2.3534 | 15 | 15 | 255 | 257 | 2.759 |  | 8.9
66 | 2.392 | 8.0791 | 8.0783 | 28.2963 | 28.2982 | 6.8821 | 2.3544 | 16 | 16 | 256 | 258 | 2.776 |  | 8.9
70 | 310.494 | 8.9756 | 8.976 | 30.7812 | 30.7817 | 6.6364 | 1.3028 | 1 | 1 | 257 | 259 | 2.822 |  | 8.8
70 | 249.78 | 8.9796 | 8.98 | 30.7572 | 30.7578 | 6.6369 | 1.3162 | 2 | 2 | 258 | 260 | 2.829 |  | 8.8
70 | 199.755 | 9.0312 | 9.0316 | 30.7178 | 30.7184 | 6.6309 | 1.3347 | 3 | 3 | 259 | 261 | 2.884 | 6 | 8.7
70 | 175.523 | 9.0113 | 9.0118 | 30.6574 | 30.658 | 6.6365 | 1.3389 | 4 | 4 | 260 | 262 | 2.865 |  | 8.8
70 | 149.788 | 9.0558 | 9.0554 | 30.5577 | 30.559 | 6.6342 | 1.3603 | 5 | 5 | 261 | 263 | 3.617 |  | 8.4
70 | 124.964 | 9.1032 | 9.1036 | 30.453 | 30.4536 | 6.6316 | 1.355 | 6 | 6 | 262 | 264 | 4.573 |  | 8.3
70 | 99.815 | 8.8711 | 8.8707 | 29.9825 | 29.9825 | 6.6864 | 1.579 | 7 | 7 | 263 | 265 | 5.094 | 6 | 8.3
70 | 75.045 | 8.5994 | 8.5999 | 29.4441 | 29.4447 | 6.7508 | 1.8633 | 8 | 8 | 264 | 266 | 5.285 |  | 8.2
70 | 49.788 | 8.5089 | 8.5095 | 29.1677 | 29.1674 | 6.7768 | 2.0142 | 9 | 9 | 265 | 267 | 5.415 |  | 8.1
70 | 40.441 | 8.4847 | 8.4857 | 29.1082 | 29.1086 | 6.7831 | 2.0808 | 10 | 10 | 266 | 268 | 5.599 |  | 8
70 | 29.989 | 8.4599 | 8.4603 | 29.0789 | 29.0795 | 6.7882 | 2.1174 | 11 | 11 | 267 | 269 | 5.849 |  | 8
70 | 19.562 | 8.4199 | 8.4212 | 29.048 | 29.049 | 6.7957 | 2.1696 | 12 | 12 | 268 | 270 | 6.135 |  | 7.8
70 | 9.702 | 8.1881 | 8.1897 | 28.756 | 28.7544 | 6.8445 | 2.2853 | 13 | 13 | 269 | 271 | 6.142 |  | 7.8
70 | 4.821 | 8.1232 | 8.1233 | 28.6736 | 28.6738 | 6.8583 | 2.3281 | 14 | 14 | 270 | 272 | 3.154 |  | 8.8
70 | 2.811 | 8.1311 | 8.1282 | 28.6829 | 28.6815 | 6.8567 | 2.3223 | 15 | 15 | 271 | 273 | 3.669 |  | 8.7
75 | 152.406 | 8.8889 | 8.8897 | 30.471 | 30.4729 | 6.6627 | 1.4391 | 1 | 1 | 272 | 274 | 4.835 |  | 8.5
75 | 123.889 | 8.7477 | 8.7496 | 30.0528 | 30.0581 | 6.7019 | 1.6216 | 2 | 2 | 273 | 275 | 4.998 |  | 8.5
75 | 101.363 | 8.5096 | 8.51 | 29.3416 | 29.3425 | 6.769 | 1.9376 | 3 | 3 | 274 | 276 | 5.212 | 6 | 8.4
75 | 75.393 | 8.4542 | 8.4546 | 29.2047 | 29.2051 | 6.7835 | 1.9943 | 4 | 4 | 275 | 277 | 5.362 | 3 | 8.4 | OXY: Bubble under stopper
75 | 50.292 | 8.3891 | 8.39 | 29.1252 | 29.1261 | 6.797 | 2.0512 | 5 | 5 | 276 | 278 | 5.619 |  | 8.3
75 | 40.086 | 8.3193 | 8.3189 | 29.056 | 29.0557 | 6.8109 | 2.1098 | 6 | 6 | 277 | 279 | 5.994 |  | 8.2
75 | 30.152 | 8.3239 | 8.3239 | 28.9301 | 28.9298 | 6.8157 | 2.1688 | 7 | 7 | 278 | 280 | 6.258 |  | 8
75 | 19.982 | 8.1779 | 8.1791 | 28.6439 | 28.645 | 6.8511 | 2.3049 | 8 | 8 | 279 | 281 | 6.271 |  | 8
75 | 9.464 | 8.0434 | 8.043 | 28.4896 | 28.4898 | 6.879 | 2.3554 | 9 | 9 | 280 | 282 | 6.282 |  | 7.3
75 | 4.987 | 8.0285 | 8.0291 | 28.4896 | 28.4901 | 6.8814 | 2.3584 | 10 | 10 | 281
75 | 2.611 | 8.036 | 8.0365 | 28.4912 | 28.4918 | 6.8801 | 2.3581 | 11 | 11 | 282

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers

## geographic
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
1 | 29.971 | 8.9799 | 8.9803 | 29.8673 | 29.8685 | 6.6751 | 1.7346 | 5 | 5 | 5 | 6.7978 | 0.59564 | 4.049 |  | 0.0026516 | 0.4785 | -0.2317
3 | 9.813 | 8.4729 | 8.4747 | 30.5372 | 30.5314 | 6.7223 | 1.922 | 12 | 12 | 21 | 6.8224 | 0.68759 | 4.691 |  | 0.00037942
3 | 151.683 | 8.3715 | 8.3717 | 31.0364 | 31.0379 | 6.7158 | 1.9314 | 4 | 4 | 13 | 6.9394 | 0.69156 | 4.799 |  | 0.00090734
3 | 75.406 | 8.3915 | 8.3919 | 30.9184 | 30.9193 | 6.718 | 1.9202 | 7 | 7 | 16 | 6.8745 | 0.68849 | 4.733 |  | 0.0013945
3 | 201.317 | 8.3675 | 8.3678 | 31.1301 | 31.1305 | 6.7124 | 1.9644 | 2 | 2 | 11 | 6.9799 | 0.70989 | 4.955 |  | 0.0016362
3 | 177.799 | 8.3662 | 8.3665 | 31.0884 | 31.0886 | 6.7144 | 1.9508 | 3 | 3 | 12 | 6.961 | 0.70363 | 4.898 |  | 0.0018484
3 | 222.97 | 8.3697 | 8.37 | 31.1327 | 31.133 | 6.7119 | 1.9597 | 1 | 1 | 10 | 6.9988 | 0.70797 | 4.955 |  | 0.0019576
3 | 2.741 | 8.5009 | 8.5012 | 30.3864 | 30.3871 | 6.7246 | 1.9178 | 14 | 14 | 23 | 6.8189 | 0.68809 | 4.692 |  | 0.002142
3 | 4.804 | 8.5006 | 8.5011 | 30.3882 | 30.3882 | 6.7246 | 1.9179 | 13 | 13 | 22 | 6.8206 | 0.6885 | 4.696 |  | 0.0024978
3 | 126.227 | 8.3812 | 8.3817 | 30.9859 | 30.9864 | 6.7166 | 1.9266 | 5 | 5 | 14 | 6.9177 | 0.69271 | 4.792 | 6 | 0.0025364
3 | 20.095 | 8.4293 | 8.4297 | 30.7941 | 30.7948 | 6.7177 | 1.9191 | 11 | 11 | 20 | 6.8261 | 0.68926 | 4.705 |  | 0.0026875
3 | 50.563 | 8.393 | 8.3938 | 30.9047 | 30.9047 | 6.7183 | 1.9223 | 8 | 8 | 17 | 6.853 | 0.69093 | 4.735 |  | 0.002794
3 | 100.975 | 8.3933 | 8.3936 | 30.9245 | 30.925 | 6.7174 | 1.9173 | 6 | 6 | 15 | 6.8964 | 0.68876 | 4.75 |  | 0.0030195
3 | 40.238 | 8.4056 | 8.4066 | 30.8756 | 30.8747 | 6.7177 | 1.9183 | 9 | 9 | 18 | 6.8435 | 0.68971 | 4.72 |  | 0.0034945
7 | 2.601 | 10.0555 | 10.0558 | 29.9743 | 29.9754 | 6.5136 | 2.4512 | 14 | 14 | 37 | 6.6222 | 0.94545 | 6.261 |  | 0.0042742
7 | 5.008 | 10.0616 | 10.0619 | 29.9751 | 29.9762 | 6.5126 | 2.4489 | 13 | 13 | 36 | 6.6234 | 0.94438 | 6.255 |  | 0.0042971
7 | 40.211 | 10.3284 | 10.3286 | 30.3621 | 30.362 | 6.4588 | 2.4042 | 9 | 9 | 32 | 6.6011 | 0.92303 | 6.093 |  | 0.0043344
7 | 76.691 | 9.1035 | 9.1052 | 31.2832 | 31.2832 | 6.5961 | 2.138 | 7 | 7 | 30 | 6.7589 | 0.79672 | 5.385 |  | 0.0053918
7 | 129.571 | 8.0357 | 8.0334 | 32.3334 | 32.3422 | 6.7101 | 1.6038 | 5 | 5 | 28 | 6.9101 | 0.54124 | 3.74 | 6 | 0.0054982
13 | 5.306 | 9.2922 | 9.293 | 30.2508 | 30.2516 | 6.6124 | 2.3722 | 12 | 12 | 63 | 6.7164 | 0.9057 | 6.083 |  | 0.0023108
13 | 2.586 | 9.269 | 9.2691 | 30.2411 | 30.2414 | 6.6163 | 2.3782 | 13 | 13 | 64 | 6.7176 | 0.90865 | 6.104 |  | 0.0023927
13 | 10.059 | 9.3062 | 9.3064 | 30.3485 | 30.3496 | 6.6062 | 2.3455 | 11 | 11 | 62 | 6.7143 | 0.89302 | 5.996 |  | 0.0024019
13 | 175.581 | 7.7141 | 7.7147 | 33.3 | 33.3004 | 6.7166 | 1.2794 | 2 | 2 | 53 | 6.954 | 0.37805 | 2.629 |  | 0.0024294
13 | 38.866 | 8.8697 | 8.8696 | 31.0254 | 31.0264 | 6.6417 | 2.108 | 8 | 8 | 59 | 6.7701 | 0.7799 | 5.28 |  | 0.0029229
13 | 125.092 | 8.1511 | 8.1513 | 32.1327 | 32.1334 | 6.7014 | 1.6962 | 4 | 4 | 55 | 6.8984 | 0.58318 | 4.023 |  | 0.0032483
13 | 75.96 | 8.4619 | 8.4624 | 31.4604 | 31.4608 | 6.6838 | 1.938 | 6 | 6 | 57 | 6.8407 | 0.69919 | 4.783 |  | 0.0035467
13 | 99.481 | 8.3933 | 8.3933 | 31.6083 | 31.6094 | 6.6876 | 1.8841 | 5 | 5 | 56 | 6.8645 | 0.67477 | 4.632 |  | 0.0049247
13 | 31.456 | 8.96 | 8.9607 | 30.9429 | 30.9421 | 6.6318 | 2.1364 | 9 | 9 | 60 | 6.7547 | 0.79618 | 5.378 | 6 | 0.005595
13 | 20.394 | 9.1448 | 9.1439 | 30.7339 | 30.7356 | 6.6135 | 2.2261 | 10 | 10 | 61 | 6.7287 | 0.83924 | 5.647 |  | 0.005749
16 | 76.231 | 8.4871 | 8.4875 | 31.4321 | 31.4325 | 6.6812 | 1.9476 | 5 | 5 | 69 | 6.8386 | 0.70116 | 4.795 |  | 0.00093177
16 | 101.392 | 8.5119 | 8.5125 | 31.4929 | 31.4931 | 6.6748 | 1.9432 | 4 | 4 | 68 | 6.8544 | 0.69911 | 4.792 |  | 0.00096616
16 | 150.202 | 7.9316 | 7.9321 | 32.6716 | 32.6716 | 6.7111 | 1.4923 | 2 | 2 | 66 | 6.9282 | 0.48353 | 3.35 |  | 0.001156
16 | 125.503 | 8.2224 | 8.2223 | 32.0488 | 32.0497 | 6.6943 | 1.7322 | 3 | 3 | 67 | 6.8923 | 0.59574 | 4.106 |  | 0.0014005
16 | 40.648 | 8.9377 | 8.9376 | 30.9683 | 30.969 | 6.634 | 2.1249 | 7 | 7 | 71 | 6.7647 | 0.78747 | 5.327 | 6 | 0.0024237
16 | 174.404 | 7.9033 | 7.9037 | 32.7487 | 32.749 | 6.712 | 1.4625 | 1 | 1 | 65 | 6.9503 | 0.46559 | 3.236 |  | 0.0024938
16 | 49.954 | 8.7989 | 8.7982 | 31.0443 | 31.0455 | 6.6514 | 2.0894 | 6 | 6 | 70 | 6.7888 | 0.77068 | 5.232 |  | 0.0026254
16 | 30.343 | 9.0498 | 9.0504 | 30.9051 | 30.9054 | 6.6202 | 2.1535 | 8 | 8 | 72 | 6.7429 | 0.80173 | 5.406 |  | 0.0029825
16 | 10.174 | 9.9065 | 9.9086 | 30.088 | 30.0919 | 6.5301 | 2.3831 | 10 | 10 | 74 | 6.6438 | 0.91183 | 6.058 |  | 0.0032268
16 | 4.58 | 9.8633 | 9.864 | 29.8337 | 29.8346 | 6.547 | 2.4264 | 11 | 11 | 75 | 6.6558 | 0.93303 | 6.21 |  | 0.0037085
16 | 2.673 | 9.804 | 9.8052 | 29.7865 | 29.7878 | 6.5575 | 2.4309 | 12 | 12 | 76 | 6.6642 | 0.93545 | 6.234 |  | 0.0039547
16 | 19.99 | 9.1345 | 9.1428 | 30.8183 | 30.8066 | 6.6114 | 2.1797 | 9 | 9 | 73 | 6.7261 | 0.81593 | 5.488 |  | 0.0046306
19 | 102.344 | 8.4965 | 8.4968 | 31.2469 | 31.2476 | 6.6878 | 2.0146 | 2 | 2 | 78 | 6.8684 | 0.73481 | 5.047 |  | 0.0025362
19 | 116.354 | 8.4966 | 8.4968 | 31.2445 | 31.245 | 6.6879 | 2.0106 | 1 | 1 | 77 | 6.8808 | 0.73334 | 5.046 |  | 0.0029908
19 | 10.659 | 8.5459 | 8.5467 | 31.0852 | 31.0859 | 6.6874 | 2.0777 | 8 | 8 | 84 | 6.7885 | 0.76614 | 5.201 |  | 0.0036783
19 | 31.659 | 8.5093 | 8.5095 | 31.1612 | 31.1619 | 6.6896 | 2.0546 | 6 | 6 | 82 | 6.8086 | 0.75581 | 5.146 |  | 0.0043961
19 | 40.648 | 8.5225 | 8.5225 | 31.172 | 31.1719 | 6.6872 | 2.0521 | 5 | 5 | 81 | 6.8141 | 0.7552 | 5.146 | 6 | 0.0049902
19 | 4.961 | 8.5964 | 8.5985 | 31.0652 | 31.0645 | 6.6807 | 2.0869 | 9 | 9 | 85 | 6.7774 | 0.77198 | 5.232 |  | 0.0050885
19 | 50.783 | 8.5193 | 8.5199 | 31.2083 | 31.2091 | 6.6861 | 2.0401 | 4 | 4 | 80 | 6.8218 | 0.7498 | 5.115 |  | 0.0053266
26 | 125.728 | 8.5056 | 8.5062 | 31.1024 | 31.1028 | 6.6927 | 2.038 | 2 | 2 | 98 | 6.8941 | 0.74397 | 5.129 |  | 0.00049311
26 | 146.042 | 8.4948 | 8.4951 | 31.1083 | 31.109 | 6.6941 | 2.0293 | 1 | 1 | 97 | 6.9133 | 0.74016 | 5.117 |  | 0.00084277
26 | 31.385 | 8.6103 | 8.6085 | 30.9881 | 30.9905 | 6.682 | 2.0982 | 7 | 7 | 103 | 6.8017 | 0.7751 | 5.272 |  | 0.0027924
26 | 50.694 | 8.6329 | 8.6313 | 30.974 | 30.9757 | 6.6792 | 2.1023 | 5 | 5 | 101 | 6.816 | 0.77729 | 5.298 |  | 0.0030229
26 | 100.565 | 8.5159 | 8.5156 | 31.0911 | 31.0929 | 6.6917 | 2.0492 | 3 | 3 | 99 | 6.871 | 0.752 | 5.167 |  | 0.003175
26 | 1.975 | 8.7153 | 8.7159 | 30.9297 | 30.9304 | 6.6688 | 2.1607 | 11 | 11 | 107 | 6.7641 | 0.80543 | 5.448 |  | 0.0032297
26 | 74.786 | 8.5581 | 8.5582 | 31.041 | 31.0419 | 6.6875 | 2.0706 | 4 | 4 | 100 | 6.8446 | 0.76264 | 5.22 | 6 | 0.0035691
26 | 5.534 | 8.7199 | 8.7217 | 30.9299 | 30.9298 | 6.6681 | 2.1696 | 10 | 10 | 106 | 6.7665 | 0.81121 | 5.489 |  | 0.0047473
26 | 20.514 | 8.6142 | 8.6142 | 30.9813 | 30.9822 | 6.6817 | 2.105 | 8 | 8 | 104 | 6.792 | 0.78062 | 5.302 |  | 0.0050769
26 | 40.391 | 8.6067 | 8.6088 | 30.9905 | 30.9902 | 6.6824 | 2.1036 | 6 | 6 | 102 | 6.81 | 0.78003 | 5.312 |  | 0.0051744
26 | 10.437 | 8.6613 | 8.6602 | 30.9577 | 30.9591 | 6.6757 | 2.1262 | 9 | 9 | 105 | 6.7777 | 0.79156 | 5.365 |  | 0.0058807
32 | 216.77 | 8.4627 | 8.4634 | 30.7827 | 30.7817 | 6.7131 | 1.9042 | 1 | 1 | 108 | 6.9956 | 0.67957 | 4.754 |  | 0.00010283
32 | 31.839 | 8.9183 | 8.92 | 30.2717 | 30.2719 | 6.6669 | 1.7692 | 10 | 10 | 117 | 6.7903 | 0.6147 | 4.174 |  | 0.00018519
32 | 149.739 | 8.4766 | 8.477 | 30.6973 | 30.6979 | 6.7147 | 1.9135 | 4 | 4 | 111 | 6.9377 | 0.68437 | 4.748 | 6 | 0.00045526
32 | 174.486 | 8.4653 | 8.4658 | 30.7471 | 30.7474 | 6.7143 | 1.9111 | 3 | 3 | 110 | 6.9591 | 0.68357 | 4.757 |  | 0.00079566
32 | 100.025 | 8.509 | 8.5092 | 30.5884 | 30.5897 | 6.7146 | 1.9129 | 6 | 6 | 113 | 6.894 | 0.68466 | 4.72 |  | 0.001052
32 | 76.071 | 8.5282 | 8.5286 | 30.5106 | 30.5105 | 6.7151 | 1.9112 | 7 | 7 | 114 | 6.8736 | 0.68392 | 4.701 |  | 0.0010941
32 | 123.134 | 8.4859 | 8.4865 | 30.6611 | 30.661 | 6.7149 | 1.9132 | 5 | 5 | 112 | 6.9144 | 0.68538 | 4.739 |  | 0.0016151
32 | 10.201 | 8.5741 | 8.5767 | 28.8795 | 28.8876 | 6.7795 | 1.9695 | 12 | 12 | 119 | 6.8819 | 0.71259 | 4.904 |  | 0.001901
32 | 2.571 | 8.392 | 8.3919 | 28.2703 | 28.2691 | 6.8344 | 2.0651 | 14 | 14 | 121 | 6.9288 | 0.75367 | 5.222 |  | 0.0027847
32 | 5.158 | 8.4054 | 8.4054 | 28.3303 | 28.3292 | 6.8297 | 2.0524 | 13 | 13 | 120 | 6.9265 | 0.74468 | 5.158 |  | 0.0056779
39 | 100.563 | 9.3462 | 9.3465 | 30.3338 | 30.3344 | 6.601 | 1.5486 | 4 | 4 | 126 | 6.7872 | 0.50875 | 3.453 |  | 0.00054669
39 | 29.768 | 8.9617 | 8.9608 | 29.5863 | 29.5865 | 6.69 | 1.7908 | 8 | 8 | 130 | 6.8125 | 0.62576 | 4.263 |  | 0.0005848
39 | 149.733 | 8.6202 | 8.6207 | 30.4585 | 30.4589 | 6.7035 | 1.8512 | 2 | 2 | 124 | 6.9277 | 0.65346 | 4.527 |  | 0.00061117
39 | 74.288 | 9.3702 | 9.3714 | 30.1716 | 30.1715 | 6.6044 | 1.5551 | 5 | 5 | 127 | 6.7683 | 0.5115 | 3.462 |  | 0.00092651
39 | 40.69 | 8.8025 | 8.801 | 29.7103 | 29.7108 | 6.7085 | 1.8379 | 7 | 7 | 129 | 6.8391 | 0.64877 | 4.437 |  | 0.0010207
39 | 177.901 | 8.6129 | 8.6133 | 30.5086 | 30.5091 | 6.7024 | 1.8334 | 1 | 1 | 123 | 6.9515 | 0.64821 | 4.506 |  | 0.0026503
39 | 20.205 | 9.0485 | 9.0453 | 29.5435 | 29.5457 | 6.6788 | 1.7673 | 9 | 9 | 131 | 6.7939 | 0.61923 | 4.207 |  | 0.0052549
43 | 100.438 | 9.2773 | 9.2777 | 30.3511 | 30.3514 | 6.6104 | 1.5583 | 8 | 8 | 143 | 6.796 | 0.51398 | 3.493 |  | 1.8813e-05
43 | 50.274 | 9.3636 | 9.364 | 29.9248 | 29.9251 | 6.6159 | 1.5455 | 10 | 10 | 145 | 6.7593 | 0.50775 | 3.432 |  | 5.9581e-05
43 | 9.636 | 8.2296 | 8.2297 | 27.0374 | 27.0369 | 6.9149 | 2.2137 | 14 | 14 | 149 | 7.0148 | 0.82739 | 5.804 |  | 0.00014751
43 | 174.572 | 9.1844 | 9.1848 | 30.7917 | 30.7924 | 6.6052 | 1.4319 | 5 | 5 | 140 | 6.8541 | 0.45374 | 3.11 |  | 0.00029113
43 | 30.495 | 9.153 | 9.1523 | 29.3623 | 29.3574 | 6.6711 | 1.7937 | 12 | 12 | 147 | 6.7961 | 0.62624 | 4.256 |  | 0.00032966
43 | 249.978 | 9.0847 | 9.0852 | 31.0333 | 31.0338 | 6.6096 | 1.3289 | 3 | 3 | 138 | 6.9238 | 0.40354 | 2.794 | 6 | 0.00064362
43 | 299.799 | 9.0466 | 9.047 | 31.1231 | 31.1235 | 6.6113 | 1.2642 | 2 | 2 | 137 | 6.9694 | 0.37248 | 2.596 |  | 0.00072045
43 | 75.187 | 9.386 | 9.3863 | 30.2218 | 30.222 | 6.5999 | 1.5206 | 9 | 9 | 144 | 6.7647 | 0.49684 | 3.361 |  | 0.00094715
43 | 316.402 | 9.0419 | 9.0422 | 31.1362 | 31.1367 | 6.6115 | 1.253 | 1 | 1 | 136 | 6.9843 | 0.36654 | 2.56 |  | 0.0013467
43 | 124.794 | 9.1582 | 9.1595 | 30.5001 | 30.5003 | 6.6215 | 1.575 | 7 | 7 | 142 | 6.8273 | 0.52407 | 3.578 |  | 0.0021407
43 | 150.057 | 9.1575 | 9.158 | 30.7099 | 30.7106 | 6.6127 | 1.4884 | 6 | 6 | 141 | 6.8402 | 0.48274 | 3.302 |  | 0.0022381
43 | 39.997 | 9.2045 | 9.2047 | 29.6861 | 29.6866 | 6.6495 | 1.702 | 11 | 11 | 146 | 6.7829 | 0.57984 | 3.933 | 6 | 0.002866
43 | 199.575 | 9.1478 | 9.1483 | 30.8869 | 30.8873 | 6.6065 | 1.3997 | 4 | 4 | 139 | 6.877 | 0.43493 | 2.991 |  | 0.0031425
43 | 19.805 | 9.0088 | 9.0092 | 28.9232 | 28.9222 | 6.7116 | 1.8927 | 13 | 13 | 148 | 6.8264 | 0.67019 | 4.575 |  | 0.0037637
43 | 5.382 | 7.2204 | 7.2195 | 22.882 | 22.8785 | 7.273 | 2.4665 | 15 | 15 | 150 | 7.3615 | 0.94451 | 6.953 |  | 0.0040265
48 | 353.272 | 9.0303 | 9.0307 | 31.1644 | 31.1649 | 6.612 | 1.2361 | 2 | 2 | 153 | 7.0176 | 0.35981 | 2.525 |  | 1.6807e-05
48 | 100.638 | 9.3588 | 9.3586 | 30.3202 | 30.3213 | 6.5997 | 1.5366 | 9 | 9 | 160 | 6.7861 | 0.50308 | 3.414 |  | 0.00045495
48 | 378.184 | 9.0143 | 9.0147 | 31.1859 | 31.1864 | 6.6134 | 1.2283 | 1 | 1 | 152 | 7.0413 | 0.35547 | 2.503 |  | 0.00056776
48 | 2.777 | 7.5686 | 7.5548 | 27.1257 | 27.0961 | 7.0166 | 2.4251 | 17 | 17 | 168 | 7.1038 | 0.92711 | 6.586 |  | 0.001606
48 | 302.913 | 9.0616 | 9.0621 | 31.112 | 31.1123 | 6.6096 | 1.2767 | 3 | 3 | 154 | 6.9705 | 0.3773 | 2.63 | 6 | 0.0019114
48 | 76.2 | 9.2776 | 9.2783 | 30.1279 | 30.1285 | 6.6198 | 1.4566 | 10 | 10 | 161 | 6.7847 | 0.46781 | 3.174 |  | 0.0025415
48 | 151.041 | 8.9749 | 8.9764 | 30.7185 | 30.719 | 6.6392 | 1.559 | 7 | 7 | 158 | 6.8665 | 0.51701 | 3.55 |  | 0.002725
48 | 176.432 | 9.1046 | 9.105 | 30.8697 | 30.8703 | 6.6136 | 1.42 | 6 | 6 | 157 | 6.8636 | 0.45078 | 3.094 |  | 0.0030141
48 | 126.49 | 9.0748 | 9.0755 | 30.5708 | 30.5712 | 6.6308 | 1.5755 | 8 | 8 | 159 | 6.8374 | 0.52579 | 3.595 |  | 0.0035898
48 | 29.948 | 8.9624 | 8.9634 | 29.3351 | 29.3367 | 6.7007 | 1.8455 | 13 | 13 | 164 | 6.8237 | 0.64775 | 4.42 |  | 0.00362
48 | 49.975 | 9.1213 | 9.1218 | 29.6865 | 29.6875 | 6.6619 | 1.6697 | 11 | 11 | 162 | 6.8032 | 0.56312 | 3.831 |  | 0.0041123
48 | 4.958 | 7.6891 | 7.6918 | 27.3991 | 27.4025 | 6.9846 | 2.4213 | 16 | 16 | 167 | 7.0748 | 0.92271 | 6.528 |  | 0.0041933
48 | 201.965 | 9.1059 | 9.1062 | 30.9194 | 30.92 | 6.6113 | 1.3904 | 5 | 5 | 156 | 6.8836 | 0.43931 | 3.024 |  | 0.005714
51 | 2.691 | 7.5448 | 7.5494 | 27.1139 | 27.1221 | 7.021 | 2.3852 | 16 | 16 | 184 | 7.1079 | 0.9097 | 6.466 |  | 7.8111e-05
51 | 5.535 | 7.8229 | 7.8182 | 27.8336 | 27.821 | 6.9435 | 2.3654 | 15 | 15 | 183 | 7.0352 | 0.89947 | 6.328 |  | 0.00068564
51 | 299.965 | 9.0707 | 9.0711 | 31.064 | 31.0645 | 6.6103 | 1.3168 | 2 | 2 | 170 | 6.9687 | 0.3995 | 2.784 |  | 0.0011146
51 | 249.566 | 9.0835 | 9.084 | 31.0268 | 31.0273 | 6.61 | 1.3374 | 3 | 3 | 171 | 6.9239 | 0.41147 | 2.849 |  | 0.003208
51 | 125.512 | 9.2019 | 9.2103 | 30.6163 | 30.6171 | 6.6101 | 1.5067 | 7 | 7 | 175 | 6.8166 | 0.49276 | 3.359 |  | 0.0035089
51 | 100.961 | 9.3461 | 9.3466 | 30.5373 | 30.5378 | 6.5923 | 1.4514 | 8 | 8 | 176 | 6.7787 | 0.4669 | 3.165 |  | 0.0040898
51 | 75.318 | 9.4023 | 9.4028 | 30.3553 | 30.3558 | 6.5919 | 1.4761 | 9 | 9 | 177 | 6.7567 | 0.47923 | 3.238 |  | 0.0046172
51 | 201.255 | 9.0049 | 9.0044 | 30.8799 | 30.8798 | 6.6279 | 1.4887 | 4 | 4 | 172 | 6.8991 | 0.48543 | 3.349 | 6 | 0.0047677
51 | 175.851 | 9.0525 | 9.0528 | 30.8077 | 30.8083 | 6.6239 | 1.4983 | 5 | 5 | 173 | 6.8732 | 0.49016 | 3.369 |  | 0.0049215
51 | 20.551 | 8.8022 | 8.8005 | 28.9864 | 28.9834 | 6.7401 | 1.982 | 13 | 13 | 181 | 6.8536 | 0.72225 | 4.95 |  | 0.0055613
51 | 150.283 | 9.1226 | 9.1239 | 30.7345 | 30.7349 | 6.6167 | 1.5052 | 6 | 6 | 174 | 6.8442 | 0.49414 | 3.382 |  | 0.0056192
55 | 40.2 | 8.5819 | 8.5865 | 28.7138 | 28.7203 | 6.7856 | 2.0993 | 10 | 10 | 194 | 6.9146 | 0.77242 | 5.341 |  | 0.00041547
55 | 250.74 | 9.0604 | 9.0611 | 31.0039 | 31.0039 | 6.6144 | 1.3257 | 2 | 2 | 186 | 6.9293 | 0.40221 | 2.787 |  | 0.00042716
55 | 30.038 | 8.4892 | 8.4902 | 28.6295 | 28.6305 | 6.8035 | 2.1363 | 11 | 11 | 195 | 6.9229 | 0.7884 | 5.458 |  | 0.0021369
55 | 175.532 | 9.1268 | 9.1276 | 30.8677 | 30.8677 | 6.6104 | 1.3616 | 4 | 4 | 188 | 6.8597 | 0.42203 | 2.895 |  | 0.0021921
55 | 199.893 | 9.0953 | 9.0957 | 30.9319 | 30.9323 | 6.6123 | 1.3514 | 3 | 3 | 187 | 6.8827 | 0.41713 | 2.871 |  | 0.0021957
55 | 76.002 | 9.2052 | 9.2051 | 30.1128 | 30.1132 | 6.6311 | 1.4869 | 8 | 8 | 192 | 6.7953 | 0.47739 | 3.244 |  | 0.0023807
55 | 150.578 | 9.2067 | 9.2072 | 30.7787 | 30.779 | 6.6025 | 1.3992 | 5 | 5 | 189 | 6.8306 | 0.44052 | 3.009 | 6 | 0.0027075
55 | 9.79 | 8.1516 | 8.1474 | 28.0734 | 28.0723 | 6.8807 | 2.3008 | 13 | 13 | 197 | 6.9793 | 0.86613 | 6.045 |  | 0.0031171
55 | 50.679 | 8.6402 | 8.6407 | 28.9467 | 28.9474 | 6.7664 | 2.1178 | 9 | 9 | 193 | 6.9051 | 0.77813 | 5.373 | 6 | 0.0035176
55 | 2.507 | 7.7222 | 7.7279 | 27.5439 | 27.5497 | 6.9727 | 2.3627 | 15 | 15 | 199 | 7.0609 | 0.89535 | 6.322 |  | 0.0035226
55 | 4.877 | 8.0614 | 8.0851 | 27.907 | 27.9138 | 6.9024 | 2.3159 | 14 | 14 | 198 | 6.9958 | 0.88053 | 6.16 |  | 0.0040638
55 | 275.596 | 9.0085 | 9.0089 | 31.0813 | 31.0817 | 6.6187 | 1.2951 | 1 | 1 | 185 | 6.9553 | 0.39222 | 2.728 |  | 0.0042342
60 | 174.484 | 9.0335 | 9.0338 | 30.9076 | 30.9083 | 6.6225 | 1.333 | 2 | 2 | 217 | 6.8703 | 0.40595 | 2.789 |  | 0.00017471
60 | 189.595 | 8.9982 | 8.9986 | 30.9633 | 30.9638 | 6.6253 | 1.3184 | 1 | 1 | 216 | 6.8861 | 0.39878 | 2.746 |  | 0.00036487
60 | 2.383 | 8.0808 | 8.0819 | 28.1306 | 28.1312 | 6.8893 | 2.3599 | 13 | 13 | 228 | 6.9805 | 0.89807 | 6.269 |  | 0.00057083
60 | 99.941 | 9.148 | 9.1483 | 30.3127 | 30.3133 | 6.631 | 1.3781 | 5 | 5 | 220 | 6.8153 | 0.42683 | 2.909 | 6 | 0.00087073
60 | 150.202 | 9.0831 | 9.0834 | 30.8452 | 30.8458 | 6.6178 | 1.3483 | 3 | 3 | 218 | 6.8448 | 0.41258 | 2.824 |  | 0.00089975
60 | 124.916 | 9.2275 | 9.2278 | 30.6297 | 30.6306 | 6.6058 | 1.3861 | 4 | 4 | 219 | 6.812 | 0.43262 | 2.947 |  | 0.0010491
60 | 29.981 | 8.4366 | 8.4372 | 28.7804 | 28.7807 | 6.805 | 2.2424 | 9 | 9 | 224 | 6.9237 | 0.84262 | 5.834 | 6 | 0.0013526
60 | 5.131 | 8.0819 | 8.0828 | 28.1306 | 28.1314 | 6.8891 | 2.3575 | 12 | 12 | 227 | 6.9828 | 0.89777 | 6.269 |  | 0.0013869
60 | 9.518 | 8.1618 | 8.1483 | 28.1737 | 28.1662 | 6.8746 | 2.3253 | 11 | 11 | 226 | 6.973 | 0.87953 | 6.133 |  | 0.0014461
60 | 39.788 | 8.4309 | 8.4306 | 28.8298 | 28.8296 | 6.8037 | 2.2196 | 8 | 8 | 223 | 6.931 | 0.8325 | 5.77 |  | 0.0021284
60 | 75.34 | 8.9386 | 8.9394 | 29.7919 | 29.7943 | 6.6845 | 1.6438 | 6 | 6 | 221 | 6.8464 | 0.54992 | 3.765 |  | 0.0049469
66 | 174.996 | 9.1319 | 9.1318 | 30.7001 | 30.7015 | 6.6169 | 1.3506 | 5 | 5 | 245 | 6.866 | 0.41451 | 2.846 |  | 6.5143e-05
66 | 125.028 | 9.1872 | 9.187 | 30.529 | 30.5309 | 6.616 | 1.3694 | 7 | 7 | 247 | 6.8221 | 0.42347 | 2.889 |  | 8.831e-05
66 | 250.077 | 9.0942 | 9.0941 | 30.7787 | 30.7801 | 6.619 | 1.324 | 3 | 3 | 243 | 6.9339 | 0.40165 | 2.785 | 6 | 0.00020521
66 | 149.819 | 9.1584 | 9.1584 | 30.6398 | 30.6412 | 6.6155 | 1.3594 | 6 | 6 | 246 | 6.8429 | 0.41854 | 2.864 |  | 0.00025579
66 | 99.935 | 9.1651 | 9.1652 | 30.3829 | 30.3847 | 6.6255 | 1.3804 | 8 | 8 | 248 | 6.8098 | 0.42923 | 2.923 |  | 0.0004051
66 | 2.392 | 8.0791 | 8.0783 | 28.2963 | 28.2982 | 6.8821 | 2.3544 | 16 | 16 | 256 | 6.9733 | 0.89442 | 6.237 |  | 0.0004792
66 | 199.666 | 9.1267 | 9.1267 | 30.7202 | 30.7216 | 6.6167 | 1.3461 | 4 | 4 | 244 | 6.8875 | 0.41292 | 2.844 |  | 0.00052034
66 | 75.702 | 8.9979 | 8.9979 | 30.0555 | 30.0569 | 6.6643 | 1.4996 | 9 | 9 | 249 | 6.8267 | 0.48515 | 3.312 |  | 0.00071564
66 | 29.919 | 8.3976 | 8.3975 | 28.6928 | 28.6939 | 6.8148 | 2.2649 | 12 | 12 | 252 | 6.9332 | 0.853 | 5.914 |  | 0.00095102
66 | 299.961 | 9.0923 | 9.0922 | 30.79 | 30.7913 | 6.6188 | 1.317 | 2 | 2 | 242 | 6.978 | 0.39725 | 2.772 |  | 0.0012475
66 | 351.312 | 9.069 | 9.0683 | 30.7952 | 30.7966 | 6.622 | 1.2997 | 1 | 1 | 241 | 7.027 | 0.3885 | 2.73 |  | 0.0016984
66 | 19.853 | 8.3542 | 8.3536 | 28.5597 | 28.5602 | 6.8274 | 2.2734 | 13 | 13 | 253 | 6.9366 | 0.8582 | 5.953 |  | 0.002085
66 | 39.895 | 8.4353 | 8.4352 | 28.7499 | 28.7505 | 6.8065 | 2.2306 | 11 | 11 | 251 | 6.934 | 0.83775 | 5.809 | 6 | 0.0021364
66 | 5.357 | 8.0748 | 8.0745 | 28.2962 | 28.2977 | 6.8828 | 2.3534 | 15 | 15 | 255 | 6.9766 | 0.89127 | 6.218 |  | 0.0031453
66 | 10.29 | 8.2667 | 8.2676 | 28.4458 | 28.4487 | 6.8461 | 2.3118 | 14 | 14 | 254 | 6.946 | 0.8782 | 6.1 |  | 0.0036897
70 | 4.821 | 8.1232 | 8.1233 | 28.6736 | 28.6738 | 6.8583 | 2.3281 | 14 | 14 | 270 | 6.9518 | 0.8825 | 6.135 |  | 0.00021949
70 | 40.441 | 8.4847 | 8.4857 | 29.1082 | 29.1086 | 6.7831 | 2.0808 | 10 | 10 | 266 | 6.9112 | 0.7647 | 5.285 |  | 0.00072622
70 | 149.788 | 9.0558 | 9.0554 | 30.5577 | 30.559 | 6.6342 | 1.3603 | 5 | 5 | 261 | 6.861 | 0.42035 | 2.884 | 6 | 0.0011381
70 | 249.78 | 8.9796 | 8.98 | 30.7572 | 30.7578 | 6.6369 | 1.3162 | 2 | 2 | 258 | 6.9511 | 0.39936 | 2.776 |  | 0.0012594
70 | 75.045 | 8.5994 | 8.5999 | 29.4441 | 29.4447 | 6.7508 | 1.8633 | 8 | 8 | 264 | 6.9102 | 0.66178 | 4.573 |  | 0.0018803
70 | 175.523 | 9.0113 | 9.0118 | 30.6574 | 30.658 | 6.6365 | 1.3389 | 4 | 4 | 260 | 6.8855 | 0.41087 | 2.829 |  | 0.0019256
70 | 199.755 | 9.0312 | 9.0316 | 30.7178 | 30.7184 | 6.6309 | 1.3347 | 3 | 3 | 259 | 6.9013 | 0.40891 | 2.822 |  | 0.0019524
70 | 310.494 | 8.9756 | 8.976 | 30.7812 | 30.7817 | 6.6364 | 1.3028 | 1 | 1 | 257 | 7.0047 | 0.39388 | 2.759 |  | 0.0021912
70 | 124.964 | 9.1032 | 9.1036 | 30.453 | 30.4536 | 6.6316 | 1.355 | 6 | 6 | 262 | 6.8373 | 0.41903 | 2.865 |  | 0.0023743
70 | 29.989 | 8.4599 | 8.4603 | 29.0789 | 29.0795 | 6.7882 | 2.1174 | 11 | 11 | 267 | 6.9069 | 0.784 | 5.415 |  | 0.0025007
70 | 99.815 | 8.8711 | 8.8707 | 29.9825 | 29.9825 | 6.6864 | 1.579 | 7 | 7 | 263 | 6.869 | 0.52657 | 3.617 |  | 0.0027308
70 | 2.811 | 8.1311 | 8.1282 | 28.6829 | 28.6815 | 6.8567 | 2.3223 | 15 | 15 | 271 | 6.9485 | 0.88393 | 6.142 |  | 0.0044129
70 | 19.562 | 8.4199 | 8.4212 | 29.048 | 29.049 | 6.7957 | 2.1696 | 12 | 12 | 268 | 6.9049 | 0.81088 | 5.599 |  | 0.0044297
70 | 49.788 | 8.5089 | 8.5095 | 29.1677 | 29.1674 | 6.7768 | 2.0142 | 9 | 9 | 265 | 6.9133 | 0.73684 | 5.094 | 6 | 0.004749
75 | 101.363 | 8.5096 | 8.51 | 29.3416 | 29.3425 | 6.769 | 1.9376 | 3 | 3 | 274 | 6.9511 | 0.69557 | 4.835 |  | 0.00013668
75 | 152.406 | 8.8889 | 8.8897 | 30.471 | 30.4729 | 6.6627 | 1.4391 | 1 | 1 | 272 | 6.8909 | 0.4577 | 3.154 |  | 0.00077632
75 | 9.464 | 8.0434 | 8.043 | 28.4896 | 28.4898 | 6.879 | 2.3554 | 9 | 9 | 280 | 6.9761 | 0.89707 | 6.258 |  | 0.0017043
75 | 50.292 | 8.3891 | 8.39 | 29.1252 | 29.1261 | 6.797 | 2.0512 | 5 | 5 | 276 | 6.933 | 0.75176 | 5.212 | 6 | 0.0019695
75 | 4.987 | 8.0285 | 8.0291 | 28.4896 | 28.4901 | 6.8814 | 2.3584 | 10 | 10 | 281 | 6.9743 | 0.89916 | 6.271 |  | 0.0023454
75 | 75.393 | 8.4542 | 8.4546 | 29.2047 | 29.2051 | 6.7835 | 1.9943 | 4 | 4 | 275 | 6.9423 | 0.71994 | 4.998 |  | 0.0026526
75 | 30.152 | 8.3239 | 8.3239 | 28.9301 | 28.9298 | 6.8157 | 2.1688 | 7 | 7 | 278 | 6.9334 | 0.81042 | 5.619 |  | 0.0043244
75 | 2.611 | 8.036 | 8.0365 | 28.4912 | 28.4918 | 6.8801 | 2.3581 | 11 | 11 | 282 | 6.9709 | 0.90117 | 6.282 |  | 0.0045225
1 | 20.013 | 9.0455 | 9.0457 | 29.7891 | 29.7904 | 6.6687 | 1.7357 | 6 | 6 | 6 | 6.7834 | 0.60501 | 4.104 | 6 | 0.0061704
66 | 50.154 | 8.5068 | 8.5066 | 28.9105 | 28.9121 | 6.7884 | 2.1623 | 10 | 10 | 250 | 6.9255 | 0.79677 | 5.518 |  | 0.006189
19 | 21.399 | 8.476 | 8.4766 | 31.1482 | 31.149 | 6.6952 | 2.0579 | 7 | 7 | 83 | 6.805 | 0.7593 | 5.167 |  | 0.0062683
7 | 20.166 | 10.1363 | 10.1415 | 30.0176 | 30.0204 | 6.5002 | 2.4296 | 11 | 11 | 34 | 6.6244 | 0.93745 | 6.21 |  | 0.0065818
43 | 2.643 | 7.0746 | 7.0782 | 22.1161 | 22.1326 | 7.3348 | 2.4658 | 16 | 16 | 151 | 7.4196 | 0.94102 | 6.982 |  | 0.0071483
39 | 51.306 | 9.0466 | 9.0455 | 29.8977 | 29.8978 | 6.6639 | 1.7226 | 6 | 6 | 128 | 6.8056 | 0.59995 | 4.083 |  | 0.0073753
7 | 50.712 | 10.259 | 10.2603 | 30.6628 | 30.663 | 6.4561 | 2.3708 | 8 | 8 | 31 | 6.6064 | 0.91033 | 6.014 |  | 0.0076224
39 | 6.199 | 8.2981 | 8.2991 | 27.4514 | 27.4517 | 6.8856 | 2.1948 | 11 | 11 | 133 | 6.9828 | 0.82618 | 5.769 |  | 0.0076604
51 | 342.907 | 9.0151 | 9.0156 | 31.1658 | 31.1661 | 6.6142 | 1.2565 | 1 | 1 | 169 | 7.0105 | 0.37729 | 2.645 |  | 0.0077427
51 | 40.545 | 9.1527 | 9.1437 | 29.7781 | 29.7703 | 6.6533 | 1.5928 | 11 | 11 | 179 | 6.7866 | 0.53826 | 3.653 |  | 0.007789
19 | 2.344 | 8.6702 | 8.6639 | 30.9835 | 30.9931 | 6.6732 | 2.1145 | 10 | 10 | 86 | 6.7683 | 0.78808 | 5.334 |  | 0.0079692
7 | 30.226 | 10.1942 | 10.1956 | 30.0594 | 30.0586 | 6.4903 | 2.4167 | 10 | 10 | 33 | 6.6233 | 0.93306 | 6.18 |  | 0.008377
51 | 10.101 | 8.1691 | 8.163 | 28.1925 | 28.19 | 6.8726 | 2.2816 | 14 | 14 | 182 | 6.9716 | 0.85102 | 5.933 |  | 0.0090262
7 | 100.739 | 8.2737 | 8.2749 | 31.6435 | 31.6419 | 6.7042 | 1.8948 | 6 | 6 | 29 | 6.8812 | 0.68418 | 4.708 |  | 0.0092277
32 | 48.746 | 8.8031 | 8.8037 | 30.295 | 30.2956 | 6.6831 | 1.8371 | 8 | 8 | 115 | 6.8202 | 0.65717 | 4.482 |  | 0.0098373
75 | 19.982 | 8.1779 | 8.1791 | 28.6439 | 28.645 | 6.8511 | 2.3049 | 8 | 8 | 279 | 6.9587 | 0.86137 | 5.994 |  | 0.0098466
60 | 49.548 | 8.5082 | 8.5091 | 28.9868 | 28.988 | 6.7848 | 2.138 | 7 | 7 | 222 | 6.9213 | 0.77976 | 5.397 |  | 0.011569
32 | 19.823 | 8.9964 | 8.9965 | 30.1637 | 30.1645 | 6.6599 | 1.7401 | 11 | 11 | 118 | 6.7737 | 0.61296 | 4.152 | 6 | 0.012008
32 | 39.355 | 8.7534 | 8.7504 | 30.3199 | 30.3217 | 6.6894 | 1.8214 | 9 | 9 | 116 | 6.8179 | 0.62761 | 4.279 |  | 0.012231
51 | 30.252 | 8.9777 | 8.9781 | 29.2822 | 29.283 | 6.7007 | 1.924 | 12 | 12 | 180 | 6.8241 | 0.67657 | 4.617 |  | 0.012345
75 | 123.889 | 8.7477 | 8.7496 | 30.0528 | 30.0581 | 6.7019 | 1.6216 | 2 | 2 | 273 | 6.9047 | 0.53138 | 3.669 |  | 0.012853
1 | 40.317 | 9.0128 | 9.0132 | 30.0291 | 30.0301 | 6.6632 | 1.6809 | 4 | 4 | 4 | 6.795 | 0.58705 | 3.989 |  | 0.014456
55 | 20.103 | 8.358 | 8.3585 | 28.3568 | 28.3573 | 6.8359 | 2.2257 | 12 | 12 | 196 | 6.9454 | 0.81881 | 5.687 |  | 0.014474
1 | 10.264 | 9.1511 | 9.1492 | 29.5123 | 29.5141 | 6.6649 | 1.765 | 7 | 7 | 7 | 6.7723 | 0.5964 | 4.039 |  | 0.016458
39 | 124.967 | 9.2393 | 9.2396 | 30.4443 | 30.4449 | 6.612 | 1.5568 | 3 | 3 | 125 | 6.8185 | 0.52988 | 3.613 | 6 | 0.016653
51 | 50.677 | 9.3597 | 9.3553 | 30.0485 | 30.044 | 6.6111 | 1.492 | 10 | 10 | 178 | 6.7548 | 0.49891 | 3.37 | 6 | 0.01671
1 | 101.242 | 9.1587 | 9.1589 | 31.141 | 31.1422 | 6.5941 | 0.5275 | 1 | 1 | 1 | 6.7787 | #VALUE! | #VALUE! |  | #VALUE!
55 | 125.565 | 9.2504 | 9.2508 | 30.6865 | 30.687 | 6.6 | 1.4381 | 6 | 6 | 190 | 6.8068 | 0.43794 | 2.981 |  | 0.018498
70 | 9.702 | 8.1881 | 8.1897 | 28.756 | 28.7544 | 6.8445 | 2.2853 | 13 | 13 | 269 | 6.943 | 0.84244 | 5.849 |  | 0.019376
1 | 4.859 | 9.2033 | 9.2344 | 28.9362 | 28.9856 | 6.6819 | 1.7831 | 8 | 8 | 8 | 6.7855 | 0.59907 | 4.065 |  | 0.022451
55 | 99.847 | 9.3762 | 9.3765 | 30.4298 | 30.4309 | 6.5925 | 1.4714 | 7 | 7 | 191 | 6.7782 | 0.49939 | 3.385 |  | 0.027022
39 | 10.11 | 8.6746 | 8.6737 | 28.2004 | 28.1981 | 6.7939 | 2.024 | 10 | 10 | 132 | 6.8977 | 0.70777 | 4.882 | 6 | 0.029002
39 | 2.652 | 7.6373 | 7.6177 | 24.5539 | 24.4395 | 7.1242 | 2.2818 | 12 | 12 | 134 | 7.2134 | 0.82956 | 5.984 |  | 0.030564
1 | 2.771 | 8.291 | 8.3067 | 26.7798 | 26.8757 | 6.9168 | 2.0702 | 9 | 9 | 9 | 7.0113 | 0.72198 | 5.062 |  | 0.036911
48 | 10.132 | 7.893 | 7.8901 | 27.6366 | 27.6349 | 6.9412 | 2.3792 | 15 | 15 | 166 | 7.0379 | 0.85295 | 6.003 |  | 0.053775
1 | 75.755 | 8.862 | 8.8634 | 30.6399 | 30.6415 | 6.6594 | 1.365 | 2 | 2 | 2 | 6.8202 | 0.35307 | 2.408 |  | 0.068367
Flagged bottles
1 | 50.391 | 8.8381 | 8.8386 | 30.1928 | 30.1937 | 6.6823 | 1.6971 | 3 | 3 | 3 | 6.8212 | 0.58963 | 4.022 | 36 | 0.062738
7 | 10.25 | 10.0819 | 10.0811 | 29.9909 | 29.9904 | 6.5091 | 2.4451 | 12 | 12 | 35 | 6.6244 | 0.94378 | 6.252 | 3 | 0.091643
13 | 150.104 | 7.8516 | 7.8502 | 32.9088 | 32.9143 | 6.7128 | 1.4121 | 3 | 3 | 54 | 6.929 | 0.44869 | 3.109 | 36 | 0.045733
13 | 50.669 | 8.8077 | 8.8098 | 31.1207 | 31.1196 | 6.6468 | 2.0886 | 7 | 7 | 58 | 6.7848 | 0.77275 | 5.243 | 3 | 0.075651
19 | 75.019 | 8.5006 | 8.5005 | 31.2429 | 31.2442 | 6.6874 | 2.02 | 3 | 3 | 79 | 6.8441 | 0.73816 | 5.052 | 3 | 0.07088
32 | 200.139 | 8.4655 | 8.4659 | 30.7647 | 30.7652 | 6.7135 | 1.9049 | 2 | 2 | 109 | 6.9811 | 0.68012 | 4.748 | 3 | 0.062865
48 | 251.901 | 9.0832 | 9.0836 | 31.0508 | 31.0515 | 6.609 | 1.3305 | 4 | 4 | 155 | 6.9249 | 0.40795 | 2.825 | 3 | 0.040461
48 | 40.271 | 9.103 | 9.1035 | 29.5914 | 29.5924 | 6.6687 | 1.7298 | 12 | 12 | 163 | 6.8015 | 0.60016 | 4.082 | 3 | 0.059027
48 | 19.971 | 8.6706 | 8.6708 | 28.7878 | 28.7876 | 6.7687 | 2.0203 | 14 | 14 | 165 | 6.8807 | 0.71475 | 4.918 | 36 | 0.047314
60 | 20.273 | 8.4403 | 8.44 | 28.6068 | 28.6044 | 6.8121 | 2.2355 | 10 | 10 | 225 | 6.9223 | 0.8461 | 5.857 | 3 | 0.085084
75 | 40.086 | 8.3193 | 8.3189 | 29.056 | 29.0557 | 6.8109 | 2.1098 | 6 | 6 | 277 | 6.9373 | 0.77293 | 5.362 | 3 | 0.066604