## Notes
Notes: The replicates for this cruise are for the most part all highly variable.
A reasonable explanation cannot be found to explain this at this time.
Analysis was conducted over 2 days and all usual tests were conducted
The solid standard readings before/after analysis were stable and consistent between the 2 days.
On the first day of analysis, an aliquot of chl a standard from the last calibration in May 2011 was run and gave results very close to those during calibration despite being stored in the freezer for > 6 months.
Despite these tests, values from Dec 13th appear for the most part much lower than values from Dec 16th, although in some cases they are comparable i.e. #213
Because A& B reps were randomized and often run on separate days in many cases 2 samples from the same station give data such that those run on the same day (two different sample ID's) agree better with each other than with their respective duplcaites run on a different day.
This could indicate either a big difference between the analysis between the two days (instrument problems, acidic cuvettes?), or a mixup during sampling and filtration.
It is tempting to use the data from Dec 16th which appers higher but we have no basis at this time for discarding the samples run on Dec 13th.
I am awaiting CTD fluoromter data to be available in order to compare profiles to extracted values to try to make some better judgment calls about which data should be final.
It is interesting that melissa encountered almost the same problem on the High Seas dataset (2011-45) analyzed around the same time but on a different flourometer.
Event #19 was mistakenly called #18 on the rosette log sheet; I have corrected that in this spreadsheet. G.Gatien
Mar 19, 2014:
Final chl a values adjusted for almost all samples-Nina Nemcek
HPLC processing for this cruise was finally completed and although only surface samples were available at most stations these helped shed light on the troubling chl a values.
Comparing CTD fluorescence data to Turner values and HPLC chl a values clearly showed that the resultant high variability between replicates was mostly due to unreasonably low values for some replicates (mainly those analyzed Dec 13th as noted above)
From this comparison it became clear that in most cases only one replicate was usable and the final chl and phaeo values were adjusted accordingly.
In a few cases neither replciate was good and a PAD value was used.
The most likely explanation for these results is acid contamination in either the acetone or the cuvettes used for the analysis.
See Profiles tab for comparison

## Final CHL archive
Note: flags and comments apply to chl values only. No flags or comments are assigned for phaeopigment values
HEADER: | Chlorophyll samples were filtered onto 25mm GF/F filters and stored in glass scintillation vials at -20C at sea and -80C in the lab
Samples were extracted in 90% acetone for 24 hours in the lab 10-14 days after collection
and analyzed on a Turner 10AU flourometer calibrated with commerically pure chlorophyll a standard (Sigma)
fluorescence readings taken before and after acification were used to calculate chlorophyll and phaeopigment concentrations (Holm-Hansen 1978)
For details see file QF2011-11_chl.xls
Cruise_Number | Event_Number | Sample_Number | CHL:extracted | Flag:CHL Extracted | Phaeo-Pigment:Extracted | Comments (preceded by CHL:)
2011-11 | 1 | 6 | 0.1 |  | 0.19 | CHL:dup discarded
2011-11 | 1 | 7 | 0.19 |  | 0.25 | CHL:dup discarded
2011-11 | 1 | 9 | 0.3362 | 6 | 0.39628
2011-11 | 3 | 20 | 0.12386 | 6 | 0.39934
2011-11 | 3 | 21 | 0.12 |  | 0.39 | CHL:dup discarded
2011-11 | 3 | 23 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 7 | 34 | 0.71 |  | 0.42 | CHL:dup discarded
2011-11 | 7 | 35 | 0.69 |  | 0.56 | CHL:dup discarded
2011-11 | 7 | 37 | 0.12 |  | 0.39289 | CHL:dup discarded
2011-11 | 10 | 48 | 0.18403 | 36 | 0.47122 | CHL:variability of reps >15% CV
2011-11 | 10 | 49 | 0.49546 | 6 | 0.32174
2011-11 | 10 | 51 | 1.1289 | 6 | 0.31859
2011-11 | 13 | 61 | 0.25 |  | 0.29663 | CHL:dup discarded
2011-11 | 13 | 62 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 13 | 64 | 1.3 |  | 0.53 | CHL:dup discarded
2011-11 | 16 | 73 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 16 | 74 | 0.32 | 36 | 0.40296 | CHL:variability of reps >15% CV
2011-11 | 16 | 76 | 0.39991 | 36 | 0.35002 | CHL:value low compared to HPLC
2011-11 | 19 | 83 | 0.3 |  | 0.43 | CHL:variability of reps >15% CV
2011-11 | 19 | 84 | 0.30435 | 6 | 0.41069
2011-11 | 19 | 86 | 0.38 |  | 0.31 | CHL:dup discarded
2011-11 | 23 | 93 | 0.36098 | 6 | 0.33165
2011-11 | 23 | 94 | 0.42114 | 6 | 0.34623
2011-11 | 23 | 96 | 0.5 | 36 | 0.35093 | CHL:variability of reps >15% CV
2011-11 | 26 | 104 | 0.2 |  | 0.37 | CHL:dup discarded
2011-11 | 26 | 105 | 0.25 |  | 0.35 | CHL:dup discarded
2011-11 | 26 | 107 | 0.27604 | 6 | 0.30198
2011-11 | 32 | 118 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 32 | 119 | 0.18 |  | 0.33 | CHL:dup discarded
2011-11 | 32 | 121 | 0.3 |  | 0.26 | CHL:dup discarded
2011-11 | 39 | 131 | 0.067949 | 36 | 0.238 | CHL:variability of reps >15% CV
2011-11 | 39 | 132 | 0.23 |  | 0.35 | CHL:dup discarded
2011-11 | 39 | 134 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 43 | 148 | 0.16 |  | 0.19 | CHL:dup discarded
2011-11 | 43 | 149 | 0.18049 | 6 | 0.3477
2011-11 | 43 | 151 | 0.28 |  | 0.4 | CHL:dup discarded
2011-11 | 48 | 165 | 0.10334 | 36 | 0.26515 | CHL:variability of reps >15% CV
2011-11 | 48 | 166 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 48 | 168 | 0.5 |  | 0.35 | CHL:dup discarded
2011-11 | 51 | 181 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 51 | 182 | 0.16138 | 6 | 0.27811
2011-11 | 51 | 184 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 55 | 196 | 0.12 |  | 0.24 | CHL:dup discarded
2011-11 | 55 | 197 | 0.12 |  | 0.2047 | CHL:dup discarded
2011-11 | 55 | 199 | 0.38 |  | 0.3 | CHL:dup discarded
2011-11 | 58 | 212 | 0.072549 | 36 | 0.24063 | CHL:variability of reps >15% CV
2011-11 | 58 | 213 | 0.18013 | 6 | 0.21127
2011-11 | 58 | 215 | 0.52 | 3 | 0.3 | CHL:value low compared to HPLC
2011-11 | 60 | 225 | 0.1405 | 36 | 0.26935 | CHL:variability of reps >15% CV
2011-11 | 60 | 226 | 0.27958 | 36 | 0.32335 | CHL:variability of reps >15% CV
2011-11 | 60 | 228 | 0.43 |  | 0.41 | CHL:dup discarded
2011-11 | 63 | 237 | 0.27 |  | 0.31 | CHL:dup discarded
2011-11 | 63 | 238 | 0.23711 | 6 | 0.31848
2011-11 | 63 | 240 | 0.45 |  | 0.33 | CHL:dup discarded
2011-11 | 66 | 253 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 66 | 254 | 0.23 |  | 0.33 | CHL:dup discarded
2011-11 | 66 | 256 | -99 | 5 | -99 | CHL:analytical problems
2011-11 | 70 | 268 | 0.12 |  | 0.27 | CHL:dup discarded
2011-11 | 70 | 269 | 0.17 |  | 0.29 | CHL:dup discarded
2011-11 | 70 | 271 | 0.20597 | 36 | 0.28684 | CHL:value low compared to HPLC
2011-11 | 75 | 279 | 0.2 |  | 0.28 | CHL:dup discarded
2011-11 | 75 | 280 | 0.27 |  | 0.34 | CHL:dup discarded
2011-11 | 75 | 282 | 0.38 |  | 0.36 | CHL:dup discarded

## Raw-data
Cruise: | 2011-11 | Nov 28-Dec 3, 2011
Notes: | Chlorophyll samples were filtered onto 25mm GF/F filters and stored in glass scintillation vials at -20C at sea and -80C in the lab
Samples were extracted in 90% acetone for 24 hours in the lab 10-14 days after collection
and analyzed on a Turner 10AU flourometer calibrated with commerically pure chlorophyll a standard (Sigma)
fluorescence readings taken before and after acification were used to calculate chlorophyll and phaeopigment concentrations (Holm-Hansen 1978)
Analysis: | Dec 13, 16, 2011 Nina Nemcek
Fluorometer calibration |  | May 9, 2011: chl (ug/L) = (1.704/(1.704-1)*(Fo-Fa)*(0.965)*(vol ext/vol filtered)*dilution factor
May 9, 2011: phaeo (ug/L)= (1.704/(1.704-1)*((1.704*Fa)-Fo)*0.965)*(vol ext/vol filtered)* dilution factor
Calibration linear up to a fluoresence reading of 475, no need to dilute samples
Analyzed_Date | Station | Sample_Number | rep | Vol. filt | Vol.ext | Dilution | Fo | Fa | Chl a ug/l | Phaeo ug/l | RATIO Fo/Fa | Comments
2011-12-16 | SI | 6 | A | 330 | 10 | 1 | 7.42 | 5.97 | 0.10263 | 0.19485 | 1.2429
2011-12-13 | SI | 6 | B | 330 | 10 | 1 | 4.94 | 4.47 | 0.033267 | 0.18947 | 1.1051
2011-12-13 | SI | 7 | A | 330 | 10 | 1 | 6.61 | 6.35 | 0.018403 | 0.29801 | 1.0409
2011-12-16 | SI | 7 | B | 330 | 10 | 1 | 11.5 | 8.85 | 0.18757 | 0.25342 | 1.2994
2011-12-13 | SI | 9 | A | 330 | 10 | 1 | 18.4 | 14 | 0.31143 | 0.38618 | 1.3143
2011-12-16 | SI | 9 | B | 330 | 10 | 1 | 20.5 | 15.4 | 0.36098 | 0.40639 | 1.3312
2011-12-13 | 59 | 20 | A | 330 | 10 | 1 | 11.8 | 10.2 | 0.11325 | 0.39501 | 1.1569
2011-12-16 | 59 | 20 | B | 330 | 10 | 1 | 12.7 | 10.8 | 0.13448 | 0.40367 | 1.1759
2011-12-16 | 59 | 21 | A | 330 | 10 | 1 | 12 | 10.3 | 0.12033 | 0.39291 | 1.165
2011-12-13 | 59 | 21 | B | 330 | 10 | 1 | 9.38 | 8.46 | 0.065118 | 0.35644 | 1.1087
2011-12-13 | 59 | 23 | A | 330 | 10 | 1 | 11.8 | 10.5 | 0.092014 | 0.43119 | 1.1238
2011-12-13 | 59 | 23 | B | 330 | 10 | 1 | 8.93 | 7.93 | 0.07078 | 0.32436 | 1.1261
2011-12-13 | 102 | 34 | A | 330 | 10 | 1 | 14.7 | 11.8 | 0.20526 | 0.38272 | 1.2458
2011-12-16 | 102 | 34 | B | 330 | 10 | 1 | 32.9 | 22.8 | 0.71488 | 0.42123 | 1.443
2011-12-13 | 102 | 35 | A | 330 | 10 | 1 | 15.5 | 12.3 | 0.2265 | 0.3864 | 1.2602
2011-12-16 | 102 | 35 | B | 330 | 10 | 1 | 34.8 | 25.1 | 0.68657 | 0.56414 | 1.3865
2011-12-13 | 102 | 37 | A | 330 | 10 | 1 | 19 | 14.5 | 0.31851 | 0.40401 | 1.3103
2011-12-16 | 102 | 37 | B | 330 | 10 | 1 | 21.7 | 15.9 | 0.41052 | 0.38176 | 1.3648
2011-12-13 | 75 | 48 | A | 330 | 10 | 1 | 14.9 | 12.9 | 0.14156 | 0.50124 | 1.155
2011-12-13 | 75 | 48 | B | 330 | 10 | 1 | 16.6 | 13.4 | 0.2265 | 0.44121 | 1.2388
2011-12-16 | 75 | 49 | A | 330 | 10 | 1 | 23.9 | 16.6 | 0.51669 | 0.31047 | 1.4398
2011-12-13 | 75 | 49 | B | 330 | 10 | 1 | 22.9 | 16.2 | 0.47423 | 0.33301 | 1.4136
2011-12-16 | 75 | 51 | A | 330 | 10 | 1 | 42.3 | 27.7 | 1.0334 | 0.34688 | 1.5271
2011-12-16 | 75 | 51 | B | 330 | 10 | 1 | 47.7 | 30.4 | 1.2245 | 0.29031 | 1.5691
2011-12-13 | 72 | 61 | A | 330 | 10 | 1 | 7.88 | 7.07 | 0.057332 | 0.29496 | 1.1146
2011-12-16 | 72 | 61 | B | 330 | 10 | 1 | 14.7 | 11.1 | 0.25481 | 0.2983 | 1.3243
2011-12-13 | 72 | 62 | A | 330 | 10 | 1 | 10 | 8.57 | 0.10122 | 0.32582 | 1.1669
2011-12-16 | 72 | 62 | B | 330 | 10 | 1 | 13.6 | 10.8 | 0.19818 | 0.33997 | 1.2593
2011-12-13 | 72 | 64 | A | 330 | 10 | 1 | 40.1 | 27.8 | 0.87059 | 0.51466 | 1.4424
2011-12-13 | 72 | 64 | B | 330 | 10 | 1 | 55.1 | 36.7 | 1.3024 | 0.52638 | 1.5014
2011-12-13 | 69 | 73 | A | 330 | 10 | 1 | 9.1 | 8.02 | 0.076442 | 0.32319 | 1.1347
2011-12-13 | 69 | 73 | B | 330 | 10 | 1 | 10.5 | 8.94 | 0.11042 | 0.33506 | 1.1745
2011-12-16 | 69 | 74 | A | 330 | 10 | 1 | 19.8 | 14.6 | 0.36806 | 0.35945 | 1.3562
2011-12-16 | 69 | 74 | B | 330 | 10 | 1 | 18.4 | 14.5 | 0.27604 | 0.44648 | 1.269
2011-12-13 | 69 | 76 | A | 330 | 10 | 1 | 21.4 | 15.4 | 0.42468 | 0.34269 | 1.3896
2011-12-13 | 69 | 76 | B | 330 | 10 | 1 | 20 | 14.7 | 0.37513 | 0.35735 | 1.3605
2011-12-13 | ADCP | 83 | A | 330 | 10 | 1 | 10.5 | 9.35 | 0.081397 | 0.38451 | 1.123
2011-12-16 | ADCP | 83 | B | 330 | 10 | 1 | 18.7 | 14.5 | 0.29728 | 0.42525 | 1.2897
2011-12-13 | ADCP | 84 | A | 330 | 10 | 1 | 19.4 | 14.8 | 0.32559 | 0.41188 | 1.3108
2011-12-16 | ADCP | 84 | B | 330 | 10 | 1 | 17.9 | 13.9 | 0.28312 | 0.4095 | 1.2878
2011-12-16 | ADCP | 86 | A | 330 | 10 | 1 | 19.1 | 13.8 | 0.37513 | 0.31251 | 1.3841
2011-12-13 | ADCP | 86 | B | 330 | 10 | 1 | 14.6 | 11.3 | 0.23357 | 0.32949 | 1.292
2011-12-16 | 65 | 93 | A | 330 | 10 | 1 | 19.8 | 14.3 | 0.38929 | 0.32327 | 1.3846
2011-12-13 | 65 | 93 | B | 330 | 10 | 1 | 18.2 | 13.5 | 0.33267 | 0.34003 | 1.3481
2011-12-16 | 65 | 94 | A | 330 | 10 | 1 | 21.6 | 15.5 | 0.43176 | 0.34059 | 1.3935
2011-12-16 | 65 | 94 | B | 330 | 10 | 1 | 21.1 | 15.3 | 0.41052 | 0.35186 | 1.3791
2011-12-13 | 65 | 96 | A | 330 | 10 | 1 | 15 | 11.6 | 0.24065 | 0.33737 | 1.2931
2011-12-13 | 65 | 96 | B | 330 | 10 | 1 | 24.5 | 17.4 | 0.50254 | 0.36449 | 1.408
2011-12-16 | 62 | 104 | A | 330 | 10 | 1 | 14.2 | 11.4 | 0.19818 | 0.36987 | 1.2456
2011-12-13 | 62 | 104 | B | 330 | 10 | 1 | 10.9 | 9.68 | 0.086352 | 0.39599 | 1.126
2011-12-13 | 62 | 105 | A | 330 | 10 | 1 | 11.4 | 9.64 | 0.12457 | 0.35578 | 1.1826
2011-12-13 | 62 | 105 | B | 330 | 10 | 1 | 15.4 | 11.9 | 0.24773 | 0.34524 | 1.2941
2011-12-16 | 62 | 107 | A | 330 | 10 | 1 | 15.9 | 11.9 | 0.28312 | 0.30985 | 1.3361
2011-12-16 | 62 | 107 | B | 330 | 10 | 1 | 15.1 | 11.3 | 0.26896 | 0.2941 | 1.3363
2011-12-13 | 56 | 118 | A | 330 | 10 | 1 | 7.03 | 6.47 | 0.039637 | 0.28276 | 1.0866
2011-12-13 | 56 | 118 | B | 330 | 10 | 1 | 6.21 | 5.8 | 0.02902 | 0.25999 | 1.0707
2011-12-16 | 56 | 119 | A | 330 | 10 | 1 | 12.7 | 10.2 | 0.17695 | 0.33131 | 1.2451
2011-12-13 | 56 | 119 | B | 330 | 10 | 1 | 9.3 | 7.72 | 0.11183 | 0.27285 | 1.2047
2011-12-16 | 56 | 121 | A | 330 | 10 | 1 | 15.7 | 11.4 | 0.30435 | 0.2637 | 1.3772
2011-12-13 | 56 | 121 | B | 330 | 10 | 1 | 8.66 | 7.3 | 0.096261 | 0.26749 | 1.1863
2011-12-16 | 46 | 131 | A | 330 | 10 | 1 | 8 | 6.79 | 0.085644 | 0.2527 | 1.1782
2011-12-13 | 46 | 131 | B | 330 | 10 | 1 | 6.2 | 5.49 | 0.050254 | 0.22331 | 1.1293
2011-12-13 | 46 | 132 | A | 330 | 10 | 1 | 9.67 | 7.69 | 0.14014 | 0.24304 | 1.2575
2011-12-16 | 46 | 132 | B | 330 | 10 | 1 | 14.8 | 11.6 | 0.2265 | 0.35152 | 1.2759
2011-12-13 | 46 | 134 | A | 330 | 10 | 1 | 10.4 | 9.22 | 0.08352 | 0.3759 | 1.128
2011-12-13 | 46 | 134 | B | 330 | 10 | 1 | 14.8 | 11.7 | 0.21942 | 0.36358 | 1.265
2011-12-16 | 42 | 148 | A | 330 | 10 | 1 | 9.17 | 6.97 | 0.15572 | 0.19159 | 1.3156
2011-12-13 | 42 | 148 | B | 330 | 10 | 1 | 5.28 | 4.8 | 0.033974 | 0.20521 | 1.1
2011-12-13 | 42 | 149 | A | 330 | 10 | 1 | 13 | 10.5 | 0.17695 | 0.34626 | 1.2381
2011-12-16 | 42 | 149 | B | 330 | 10 | 1 | 13.3 | 10.7 | 0.18403 | 0.34914 | 1.243
2011-12-16 | 42 | 151 | A | 330 | 10 | 1 | 17.7 | 13.7 | 0.28312 | 0.39954 | 1.292
2011-12-13 | 42 | 151 | B | 330 | 10 | 1 | 12.4 | 10.1 | 0.16279 | 0.34048 | 1.2277
2011-12-16 | 39 | 165 | A | 330 | 10 | 1 | 9.22 | 7.45 | 0.12528 | 0.24595 | 1.2376
2011-12-13 | 39 | 165 | B | 330 | 10 | 1 | 8.49 | 7.34 | 0.081397 | 0.28435 | 1.1567
2011-12-13 | 39 | 166 | A | 330 | 10 | 1 | 10.8 | 8.68 | 0.15005 | 0.28246 | 1.2442 | no replicate found
2011-12-13 | 39 | 168 | A | 330 | 10 | 1 | 13.6 | 11.2 | 0.16987 | 0.38821 | 1.2143
2011-12-16 | 39 | 168 | B | 330 | 10 | 1 | 24.3 | 17.2 | 0.50254 | 0.35452 | 1.4128
2011-12-13 | 27 | 181 | A | 330 | 10 | 1 | 6.24 | 5.32 | 0.065118 | 0.19997 | 1.1729
2011-12-16 | 27 | 181 | B | 330 | 10 | 1 | 7.06 | 5.74 | 0.09343 | 0.19259 | 1.23
2011-12-13 | 27 | 182 | A | 330 | 10 | 1 | 10.4 | 8.34 | 0.14581 | 0.26977 | 1.247
2011-12-13 | 27 | 182 | B | 330 | 10 | 1 | 11.8 | 9.3 | 0.17695 | 0.28646 | 1.2688
2011-12-16 | 27 | 184 | A | 330 | 10 | 1 | 17.5 | 12.7 | 0.33974 | 0.29309 | 1.378
2011-12-16 | 27 | 184 | B | 330 | 10 | 1 | 14.7 | 11.1 | 0.25481 | 0.2983 | 1.3243
2011-12-13 | 2 | 196 | A | 330 | 10 | 1 | 5.01 | 4.84 | 0.012033 | 0.22914 | 1.0351
2011-12-13 | 2 | 196 | B | 330 | 10 | 1 | 8.81 | 7.16 | 0.11679 | 0.23999 | 1.2304
2011-12-16 | 2 | 197 | A | 330 | 10 | 1 | 8.08 | 6.43 | 0.11679 | 0.20361 | 1.2566
2011-12-13 | 2 | 197 | B | 330 | 10 | 1 | 4.88 | 4.57 | 0.021942 | 0.20578 | 1.0678
2011-12-16 | 2 | 199 | A | 330 | 10 | 1 | 18.9 | 13.6 | 0.37513 | 0.30254 | 1.3897
2011-12-13 | 2 | 199 | B | 330 | 10 | 1 | 9.01 | 7.84 | 0.082813 | 0.30785 | 1.1492
2011-12-13 | 3 | 212 | A | 330 | 10 | 1 | 6.38 | 5.62 | 0.053793 | 0.22625 | 1.1352
2011-12-16 | 3 | 212 | B | 330 | 10 | 1 | 8.24 | 6.95 | 0.091306 | 0.25501 | 1.1856
2011-12-13 | 3 | 213 | A | 330 | 10 | 1 | 10.2 | 7.66 | 0.17978 | 0.20191 | 1.3316
2011-12-16 | 3 | 213 | B | 330 | 10 | 1 | 10.6 | 8.05 | 0.18049 | 0.22064 | 1.3168
2011-12-16 | 3 | 215 | A | 330 | 10 | 1 | 24 | 16.6 | 0.52377 | 0.30339 | 1.4458
2011-12-13 | 3 | 215 | B | 330 | 10 | 1 | 11.1 | 8.88 | 0.15713 | 0.28535 | 1.25
2011-12-13 | 6 | 225 | A | 330 | 10 | 1 | 8.52 | 7.12 | 0.099092 | 0.25569 | 1.1966
2011-12-16 | 6 | 225 | B | 330 | 10 | 1 | 11.9 | 9.33 | 0.1819 | 0.283 | 1.2755
2011-12-13 | 6 | 226 | A | 330 | 10 | 1 | 14.4 | 11.2 | 0.2265 | 0.33159 | 1.2857
2011-12-16 | 6 | 226 | B | 330 | 10 | 1 | 17.7 | 13 | 0.33267 | 0.31511 | 1.3615
2011-12-16 | 6 | 228 | A | 330 | 10 | 1 | 23 | 16.9 | 0.43176 | 0.41035 | 1.3609
2011-12-13 | 6 | 228 | B | 330 | 10 | 1 | 10.4 | 8.78 | 0.11466 | 0.32284 | 1.1845
2011-12-13 | 9 | 237 | A | 330 | 10 | 1 | 7.48 | 6.61 | 0.061579 | 0.26779 | 1.1316
2011-12-16 | 9 | 237 | B | 330 | 10 | 1 | 15.4 | 11.6 | 0.26896 | 0.30905 | 1.3276
2011-12-13 | 9 | 238 | A | 330 | 10 | 1 | 13.9 | 10.8 | 0.21942 | 0.31874 | 1.287
2011-12-13 | 9 | 238 | B | 330 | 10 | 1 | 15.1 | 11.5 | 0.25481 | 0.31823 | 1.313
2011-12-16 | 9 | 240 | A | 330 | 10 | 1 | 22.2 | 15.8 | 0.45299 | 0.33431 | 1.4051
2011-12-13 | 9 | 240 | B | 330 | 10 | 1 | 16.6 | 12.5 | 0.2902 | 0.33267 | 1.328
2011-12-13 | 12 | 253 | A | 330 | 10 | 1 | 7.11 | 6.45 | 0.046715 | 0.27468 | 1.1023
2011-12-13 | 12 | 253 | B | 330 | 10 | 1 | 5.61 | 5.29 | 0.02265 | 0.24095 | 1.0605
2011-12-13 | 12 | 254 | A | 330 | 10 | 1 | 8.21 | 7.32 | 0.062994 | 0.30175 | 1.1216
2011-12-13 | 12 | 254 | B | 330 | 10 | 1 | 14.4 | 11.2 | 0.2265 | 0.33159 | 1.2857
2011-12-13 | 12 | 256 | A | 330 | 10 | 1 | 12.5 | 10.3 | 0.15572 | 0.35752 | 1.2136
2011-12-13 | 12 | 256 | B | 330 | 10 | 1 | 14.6 | 11.2 | 0.24065 | 0.31743 | 1.3036
2011-12-16 | 14 | 268 | A | 330 | 10 | 1 | 9.52 | 7.85 | 0.1182 | 0.27296 | 1.2127
2011-12-13 | 14 | 268 | B | 330 | 10 | 1 | 6.49 | 5.85 | 0.045299 | 0.2462 | 1.1094
2011-12-13 | 14 | 269 | A | 330 | 10 | 1 | 9 | 7.76 | 0.087767 | 0.29891 | 1.1598
2011-12-13 | 14 | 269 | B | 330 | 10 | 1 | 11.6 | 9.22 | 0.16846 | 0.29097 | 1.2581
2011-12-16 | 14 | 271 | A | 330 | 10 | 1 | 13 | 10.1 | 0.20526 | 0.29801 | 1.2871
2011-12-16 | 14 | 271 | B | 330 | 10 | 1 | 12.6 | 9.68 | 0.20668 | 0.27567 | 1.3017
2011-12-13 | 16 | 279 | A | 330 | 10 | 1 | 12.3 | 9.54 | 0.19535 | 0.28002 | 1.2893
2011-12-13 | 16 | 279 | B | 330 | 10 | 1 | 8.08 | 7.15 | 0.065825 | 0.29045 | 1.1301
2011-12-13 | 16 | 280 | A | 330 | 10 | 1 | 16 | 12.2 | 0.26896 | 0.33895 | 1.3115
2011-12-13 | 16 | 280 | B | 330 | 10 | 1 | 8.16 | 7.1 | 0.075027 | 0.27876 | 1.1493
2011-12-13 | 16 | 282 | A | 330 | 10 | 1 | 10.7 | 9.12 | 0.11183 | 0.34261 | 1.1732
2011-12-16 | 16 | 282 | B | 330 | 10 | 1 | 20.2 | 14.8 | 0.38221 | 0.35526 | 1.3649

## %CV
Cruise | 2011-11
The average if two samples is typically reported
Variability is assessed as the %CV
DATE | Event | Station | Sample | rep | Depth | Chl ug/L | Phaeo ug/L | Avg Chl | CV% | FINAL CHL | Avg Phaeo | CV% | FLAG:CHL | Flag Comments for Archive
2011-12-16 |  | SI | 6 | A | 20 | 0.10263 | 0.19485 | 0.067949 | 72.1838 | 0.067949 | 0.19216 | 1.9795 | 36 | variability of reps >15% CV | 6 | 20
2011-12-13 |  | SI | 6 | B | 20 | 0.033267 | 0.18947 | [7 empty] | 6 | 20
2011-12-13 |  | SI | 7 | A | 10 | 0.018403 | 0.29801 | 0.10298 | 116.1502 | 0.10298 | 0.27572 | 11.436 | 36 | variability of reps >15% CV
2011-12-16 |  | SI | 7 | B | 10 | 0.18757 | 0.25342
2011-12-13 |  | SI | 9 | A | 0 | 0.31143 | 0.38618 | 0.3362 | 10.4205 | 0.3362 | 0.39628 | 3.607 | 6
2011-12-16 |  | SI | 9 | B | 0 | 0.36098 | 0.40639
2011-12-13 |  | 59 | 20 | A | 20 | 0.11325 | 0.39501 | 0.12386 | 12.1218 | 0.12386 | 0.39934 | 1.534 | 6
2011-12-16 |  | 59 | 20 | B | 20 | 0.13448 | 0.40367
2011-12-16 |  | 59 | 21 | A | 10 | 0.12033 | 0.39291 | 0.092722 | 42.1025 | 0.092722 | 0.37468 | 6.8842 | 36 | variability of reps >15% CV
2011-12-13 |  | 59 | 21 | B | 10 | 0.065118 | 0.35644
2011-12-13 |  | 59 | 23 | A | 0 | 0.092014 | 0.43119 | 0.081397 | 18.4463 | 0.081397 | 0.37778 | 19.9953 | 36 | variability of reps >15% CV
2011-12-13 |  | 59 | 23 | B | 0 | 0.07078 | 0.32436 | [7 empty] | 23 | 0
2011-12-13 |  | 102 | 34 | A | 20 | 0.20526 | 0.38272 | 0.46007 | 78.3257 | 0.46007 | 0.40197 | 6.7732 | 36 | variability of reps >15% CV | 35 | 20
2011-12-16 |  | 102 | 34 | B | 20 | 0.71488 | 0.42123 | [7 empty] | 35 | 20
2011-12-13 |  | 102 | 35 | A | 10 | 0.2265 | 0.3864 | 0.45653 | 71.2588 | 0.45653 | 0.47527 | 26.4444 | 36 | variability of reps >15% CV | 36 | 10
2011-12-16 |  | 102 | 35 | B | 10 | 0.68657 | 0.56414 | [7 empty] | 36 | 10
2011-12-13 |  | 102 | 37 | A | 0 | 0.31851 | 0.40401 | 0.36452 | 17.8493 | 0.36452 | 0.39289 | 4.0051 | 36 | variability of reps >15% CV | 38 | 0
2011-12-16 |  | 102 | 37 | B | 0 | 0.41052 | 0.38176 | [7 empty] | 38 | 0
2011-12-13 |  | 75 | 48 | A | 20 | 0.14156 | 0.50124 | 0.18403 | 32.6357 | 0.18403 | 0.47122 | 9.0066 | 36 | variability of reps >15% CV | 49 | 20
2011-12-13 |  | 75 | 48 | B | 20 | 0.2265 | 0.44121 | [7 empty] | 49 | 20
2011-12-16 |  | 75 | 49 | A | 10 | 0.51669 | 0.31047 | 0.49546 | 6.0609 | 0.49546 | 0.32174 | 4.953 | 6 |  | 50 | 10
2011-12-13 |  | 75 | 49 | B | 10 | 0.47423 | 0.33301 | [7 empty] | 50 | 10
2011-12-16 |  | 75 | 51 | A | 0 | 1.0334 | 0.34688 | 1.1289 | 11.9698 | 1.1289 | 0.31859 | 12.5549 | 6 |  | 52 | 0
2011-12-16 |  | 75 | 51 | B | 0 | 1.2245 | 0.29031 | [7 empty] | 52 | 0
2011-12-13 |  | 72 | 61 | A | 20 | 0.057332 | 0.29496 | 0.15607 | 89.4707 | 0.15607 | 0.29663 | 0.79504 | 36 | variability of reps >15% CV | 53 | 5
2011-12-16 |  | 72 | 61 | B | 20 | 0.25481 | 0.2983 | [7 empty] | 53 | 5
2011-12-13 |  | 72 | 62 | A | 10 | 0.10122 | 0.32582 | 0.1497 | 45.8031 | 0.1497 | 0.3329 | 3.0057 | 36 | variability of reps >15% CV | 63 | 20
2011-12-16 |  | 72 | 62 | B | 10 | 0.19818 | 0.33997 | [7 empty] | 63 | 20
2011-12-13 |  | 72 | 64 | A | 0 | 0.87059 | 0.51466 | 1.0865 | 28.1 | 1.0865 | 0.52052 | 1.5923 | 36 | variability of reps >15% CV | 64 | 10
2011-12-13 |  | 72 | 64 | B | 0 | 1.3024 | 0.52638 | [7 empty] | 64 | 10
2011-12-13 |  | 69 | 73 | A | 20 | 0.076442 | 0.32319 | 0.09343 | 25.713 | 0.09343 | 0.32912 | 2.5499 | 36 | variability of reps >15% CV | 65 | 5
2011-12-13 |  | 69 | 73 | B | 20 | 0.11042 | 0.33506 | [7 empty] | 65 | 5
2011-12-16 |  | 69 | 74 | A | 10 | 0.36806 | 0.35945 | 0.32205 | 20.2031 | 0.32205 | 0.40296 | 15.2719 | 36 | variability of reps >15% CV | 66 | 0
2011-12-16 |  | 69 | 74 | B | 10 | 0.27604 | 0.44648 | [7 empty] | 66 | 0
2011-12-13 |  | 69 | 76 | A | 0 | 0.42468 | 0.34269 | 0.39991 | 8.7606 | 0.39991 | 0.35002 | 2.9627 | 6 |  | 67 | 5
2011-12-13 |  | 69 | 76 | B | 0 | 0.37513 | 0.35735 | [7 empty] | 67 | 5
2011-12-13 |  | ADCP | 83 | A | 20 | 0.081397 | 0.38451 | 0.18934 | 80.6234 | 0.18934 | 0.40488 | 7.1153 | 36 | variability of reps >15% CV | 76 | 20
2011-12-16 |  | ADCP | 83 | B | 20 | 0.29728 | 0.42525 | [7 empty] | 76 | 20
2011-12-13 |  | ADCP | 84 | A | 10 | 0.32559 | 0.41188 | 0.30435 | 9.8666 | 0.30435 | 0.41069 | 0.40946 | 6 |  | 77 | 10
2011-12-16 |  | ADCP | 84 | B | 10 | 0.28312 | 0.4095 | [7 empty] | 77 | 10
2011-12-16 |  | ADCP | 86 | A | 0 | 0.37513 | 0.31251 | 0.30435 | 32.8887 | 0.30435 | 0.321 | 3.742 | 36 | variability of reps >15% CV | 79 | 0
2011-12-13 |  | ADCP | 86 | B | 0 | 0.23357 | 0.32949 | [7 empty] | 79 | 0
2011-12-16 |  | 65 | 93 | A | 20 | 0.38929 | 0.32327 | 0.36098 | 11.0919 | 0.36098 | 0.33165 | 3.5736 | 6 |  | 86 | 20
2011-12-13 |  | 65 | 93 | B | 20 | 0.33267 | 0.34003 | [7 empty] | 86 | 20
2011-12-16 |  | 65 | 94 | A | 10 | 0.43176 | 0.34059 | 0.42114 | 3.5652 | 0.42114 | 0.34623 | 2.3013 | 6 |  | 87 | 10
2011-12-16 |  | 65 | 94 | B | 10 | 0.41052 | 0.35186 | [7 empty] | 87 | 10
2011-12-13 |  | 65 | 96 | A | 0 | 0.24065 | 0.33737 | 0.37159 | 49.8342 | 0.37159 | 0.35093 | 5.4652 | 36 | variability of reps >15% CV | 89 | 0
2011-12-13 |  | 65 | 96 | B | 0 | 0.50254 | 0.36449 | [7 empty] | 89 | 0
2011-12-16 |  | 62 | 104 | A | 20 | 0.19818 | 0.36987 | 0.14227 | 55.5835 | 0.14227 | 0.38293 | 4.8244 | 36 | variability of reps >15% CV | 96 | 20
2011-12-13 |  | 62 | 104 | B | 20 | 0.086352 | 0.39599 | [7 empty] | 96 | 20
2011-12-13 |  | 62 | 105 | A | 10 | 0.12457 | 0.35578 | 0.18615 | 46.782 | 0.18615 | 0.35051 | 2.127 | 36 | variability of reps >15% CV | 97 | 10
2011-12-13 |  | 62 | 105 | B | 10 | 0.24773 | 0.34524 | [7 empty] | 97 | 10
2011-12-16 |  | 62 | 107 | A | 0 | 0.28312 | 0.30985 | 0.27604 | 3.6262 | 0.27604 | 0.30198 | 3.686 | 6 |  | 99 | 0
2011-12-16 |  | 62 | 107 | B | 0 | 0.26896 | 0.2941 | [7 empty] | 99 | 0
2011-12-13 |  | 56 | 118 | A | 20 | 0.039637 | 0.28276 | 0.034328 | 21.8693 | 0.034328 | 0.27137 | 5.9327 | 36 | variability of reps >15% CV | 107 | 20
2011-12-13 |  | 56 | 118 | B | 20 | 0.02902 | 0.25999 | [7 empty] | 107 | 20
2011-12-16 |  | 56 | 119 | A | 10 | 0.17695 | 0.33131 | 0.14439 | 31.8891 | 0.14439 | 0.30208 | 13.6841 | 36 | variability of reps >15% CV | 108 | 10
2011-12-13 |  | 56 | 119 | B | 10 | 0.11183 | 0.27285 | [7 empty] | 108 | 10
2011-12-16 |  | 56 | 121 | A | 0 | 0.30435 | 0.2637 | 0.20031 | 73.4591 | 0.20031 | 0.26559 | 1.01 | 36 | variability of reps >15% CV | 110 | 0
2011-12-13 |  | 56 | 121 | B | 0 | 0.096261 | 0.26749 | [7 empty] | 110 | 0
2011-12-16 |  | 46 | 131 | A | 20 | 0.085644 | 0.2527 | 0.067949 | 36.8285 | 0.067949 | 0.238 | 8.7312 | 36 | variability of reps >15% CV | 120 | 20
2011-12-13 |  | 46 | 131 | B | 20 | 0.050254 | 0.22331 | [7 empty] | 120 | 20
2011-12-13 |  | 46 | 132 | A | 10 | 0.14014 | 0.24304 | 0.18332 | 33.3077 | 0.18332 | 0.29728 | 25.8028 | 36 | variability of reps >15% CV | 121 | 10
2011-12-16 |  | 46 | 132 | B | 10 | 0.2265 | 0.35152 | [7 empty] | 121 | 10
2011-12-13 |  | 46 | 134 | A | 0 | 0.08352 | 0.3759 | 0.15147 | 63.4414 | 0.15147 | 0.36974 | 2.3564 | 36 | variability of reps >15% CV | 123 | 0
2011-12-13 |  | 46 | 134 | B | 0 | 0.21942 | 0.36358 | [7 empty] | 123 | 0
2011-12-16 |  | 42 | 148 | A | 20 | 0.15572 | 0.19159 | 0.094845 | 90.763 | 0.094845 | 0.1984 | 4.8515 | 36 | variability of reps >15% CV | 124 | 5
2011-12-13 |  | 42 | 148 | B | 20 | 0.033974 | 0.20521 | [7 empty] | 124 | 5
2011-12-13 |  | 42 | 149 | A | 10 | 0.17695 | 0.34626 | 0.18049 | 2.773 | 0.18049 | 0.3477 | 0.58729 | 6 |  | 125 | 5
2011-12-16 |  | 42 | 149 | B | 10 | 0.18403 | 0.34914 | [7 empty] | 125 | 5
2011-12-16 |  | 42 | 151 | A | 0 | 0.28312 | 0.39954 | 0.22296 | 38.1613 | 0.22296 | 0.37001 | 11.2864 | 36 | variability of reps >15% CV | 126 | 5
2011-12-13 |  | 42 | 151 | B | 0 | 0.16279 | 0.34048 | [7 empty] | 126 | 5
2011-12-16 |  | 39 | 165 | A | 20 | 0.12528 | 0.24595 | 0.10334 | 30.0278 | 0.10334 | 0.26515 | 10.2413 | 36 | variability of reps >15% CV | 135 | 20
2011-12-13 |  | 39 | 165 | B | 20 | 0.081397 | 0.28435 | [7 empty] | 135 | 20
2011-12-13 |  | 39 | 166 | A | 10 | 0.15005 | 0.28246 |  |  | 0.15005 |  |  |  |  | 136 | 10
2011-12-13 |  | 39 | 168 | A | 0 | 0.16987 | 0.38821 | 0.3362 | 69.9664 | 0.3362 | 0.37137 | 6.415 | 36 | variability of reps >15% CV | 136 | 10
2011-12-16 |  | 39 | 168 | B | 0 | 0.50254 | 0.35452 | [7 empty] | 137 | 5
2011-12-13 |  | 27 | 181 | A | 20 | 0.065118 | 0.19997 | 0.079274 | 25.2538 | 0.079274 | 0.19628 | 2.66 | 36 | variability of reps >15% CV | 137 | 5
2011-12-16 |  | 27 | 181 | B | 20 | 0.09343 | 0.19259 | [7 empty] | 138 | 0
2011-12-13 |  | 27 | 182 | A | 10 | 0.14581 | 0.26977 | 0.16138 | 13.6459 | 0.16138 | 0.27811 | 4.2441 | 6 |  | 138 | 0
2011-12-13 |  | 27 | 182 | B | 10 | 0.17695 | 0.28646 | [7 empty] | 139 | 5
2011-12-16 |  | 27 | 184 | A | 0 | 0.33974 | 0.29309 | 0.29728 | 20.2031 | 0.29728 | 0.29569 | 1.2458 | 36 | variability of reps >15% CV | 139 | 5
2011-12-16 |  | 27 | 184 | B | 0 | 0.25481 | 0.2983
2011-12-13 |  | 2 | 196 | A | 20 | 0.012033 | 0.22914 | 0.06441 | 115.002 | 0.06441 | 0.23456 | 3.2705 | 36 | variability of reps >15% CV
2011-12-13 |  | 2 | 196 | B | 20 | 0.11679 | 0.23999
2011-12-16 |  | 2 | 197 | A | 10 | 0.11679 | 0.20361 | 0.069364 | 96.686 | 0.069364 | 0.2047 | 0.74721 | 36 | variability of reps >15% CV
2011-12-13 |  | 2 | 197 | B | 10 | 0.021942 | 0.20578
2011-12-16 |  | 2 | 199 | A | 0 | 0.37513 | 0.30254 | 0.22897 | 90.2736 | 0.22897 | 0.30519 | 1.2293 | 36 | variability of reps >15% CV
2011-12-13 |  | 2 | 199 | B | 0 | 0.082813 | 0.30785
2011-12-13 |  | 3 | 212 | A | 20 | 0.053793 | 0.22625 | 0.072549 | 36.5626 | 0.072549 | 0.24063 | 8.4512 | 36 | variability of reps >15% CV
2011-12-16 |  | 3 | 212 | B | 20 | 0.091306 | 0.25501
2011-12-13 |  | 3 | 213 | A | 10 | 0.17978 | 0.20191 | 0.18013 | 0.27784 | 0.18013 | 0.21127 | 6.2672 | 6
2011-12-16 |  | 3 | 213 | B | 10 | 0.18049 | 0.22064
2011-12-16 |  | 3 | 215 | A | 0 | 0.52377 | 0.30339 | 0.34045 | 76.15 | 0.34045 | 0.29437 | 4.3335 | 36 | variability of reps >15% CV
2011-12-13 |  | 3 | 215 | B | 0 | 0.15713 | 0.28535
2011-12-13 |  | 6 | 225 | A | 20 | 0.099092 | 0.25569 | 0.1405 | 41.6783 | 0.1405 | 0.26935 | 7.1696 | 36 | variability of reps >15% CV
2011-12-16 |  | 6 | 225 | B | 20 | 0.1819 | 0.283
2011-12-13 |  | 6 | 226 | A | 10 | 0.2265 | 0.33159 | 0.27958 | 26.8522 | 0.27958 | 0.32335 | 3.6033 | 36 | variability of reps >15% CV
2011-12-16 |  | 6 | 226 | B | 10 | 0.33267 | 0.31511
2011-12-16 |  | 6 | 228 | A | 0 | 0.43176 | 0.41035 | 0.27321 | 82.0684 | 0.27321 | 0.36659 | 16.8809 | 36 | variability of reps >15% CV
2011-12-13 |  | 6 | 228 | B | 0 | 0.11466 | 0.32284
2011-12-13 |  | 9 | 237 | A | 20 | 0.061579 | 0.26779 | 0.16527 | 88.729 | 0.16527 | 0.28842 | 10.1159 | 36 | variability of reps >15% CV
2011-12-16 |  | 9 | 237 | B | 20 | 0.26896 | 0.30905
2011-12-13 |  | 9 | 238 | A | 10 | 0.21942 | 0.31874 | 0.23711 | 10.5538 | 0.23711 | 0.31848 | 0.11315 | 6
2011-12-13 |  | 9 | 238 | B | 10 | 0.25481 | 0.31823
2011-12-16 |  | 9 | 240 | A | 0 | 0.45299 | 0.33431 | 0.37159 | 30.978 | 0.37159 | 0.33349 | 0.34818 | 36 | variability of reps >15% CV
2011-12-13 |  | 9 | 240 | B | 0 | 0.2902 | 0.33267
2011-12-13 |  | 12 | 253 | A | 20 | 0.046715 | 0.27468 | 0.034682 | 49.0646 | 0.034682 | 0.25781 | 9.2529 | 36 | variability of reps >15% CV
2011-12-13 |  | 12 | 253 | B | 20 | 0.02265 | 0.24095
2011-12-13 |  | 12 | 254 | A | 10 | 0.062994 | 0.30175 | 0.14475 | 79.8737 | 0.14475 | 0.31667 | 6.662 | 36 | variability of reps >15% CV
2011-12-13 |  | 12 | 254 | B | 10 | 0.2265 | 0.33159
2011-12-13 |  | 12 | 256 | A | 0 | 0.15572 | 0.35752 | 0.19818 | 30.3046 | 0.19818 | 0.33748 | 8.3999 | 36 | variability of reps >15% CV
2011-12-13 |  | 12 | 256 | B | 0 | 0.24065 | 0.31743
2011-12-16 |  | 14 | 268 | A | 20 | 0.1182 | 0.27296 | 0.081751 | 63.058 | 0.081751 | 0.25958 | 7.2882 | 36 | variability of reps >15% CV
2011-12-13 |  | 14 | 268 | B | 20 | 0.045299 | 0.2462
2011-12-13 |  | 14 | 269 | A | 10 | 0.087767 | 0.29891 | 0.12811 | 44.536 | 0.12811 | 0.29494 | 1.9033 | 36 | variability of reps >15% CV
2011-12-13 |  | 14 | 269 | B | 10 | 0.16846 | 0.29097
2011-12-16 |  | 14 | 271 | A | 0 | 0.20526 | 0.29801 | 0.20597 | 0.48598 | 0.20597 | 0.28684 | 5.5081 | 6
2011-12-16 |  | 14 | 271 | B | 0 | 0.20668 | 0.27567
2011-12-13 |  | 16 | 279 | A | 20 | 0.19535 | 0.28002 | 0.13059 | 70.1358 | 0.13059 | 0.28523 | 2.5871 | 36 | variability of reps >15% CV
2011-12-13 |  | 16 | 279 | B | 20 | 0.065825 | 0.29045
2011-12-13 |  | 16 | 280 | A | 10 | 0.26896 | 0.33895 | 0.172 | 79.7314 | 0.172 | 0.30886 | 13.7804 | 36 | variability of reps >15% CV
2011-12-13 |  | 16 | 280 | B | 10 | 0.075027 | 0.27876
2011-12-13 |  | 16 | 282 | A | 0 | 0.11183 | 0.34261 | 0.24702 | 77.3968 | 0.24702 | 0.34893 | 2.5635 | 36 | variability of reps >15% CV
2011-12-16 |  | 16 | 282 | B | 0 | 0.38221 | 0.35526
[8 empty] | mean | 44.2488

## profiles
|  |  | extracted values |  | suspect values
Analysis date | Station | Depth | Chl ug/L | Sample # | HPLC | final chl | flag
2011-12-16 | SI | 20 | 0.10263 | 6 | 0.11125 | 0.10263
2011-12-13 | SI | 20 | 0.033267 | 6 | 0.089971
2011-12-13 | SI | 10 | 0.018403 | 7 | 0.22338 | 0.18757
2011-12-16 | SI | 10 | 0.18757 | 7 | 0.2173
2011-12-13 | SI | 2.5 | 0.31143 | 9 | 0.53691 | 0.3362 | 6
2011-12-16 | SI | 2.5 | 0.36098 | 9 | 0.39163
2011-12-13 | 59 | 20 | 0.11325 | 20 |  | 0.12386 | 6
2011-12-16 | 59 | 20 | 0.13448 | 20
2011-12-16 | 59 | 10 | 0.12033 | 21 |  | 0.12033
2011-12-13 | 59 | 10 | 0.065118 | 21
2011-12-13 | 59 | 2.5 | 0.092014 | 23 | 0.26603 | -99 | 5
2011-12-13 | 59 | 2.5 | 0.07078 | 23 | 0.25988
2011-12-13 | 102 | 20 | 0.20526 | 34 |  | 0.71488
2011-12-16 | 102 | 20 | 0.71488 | 34
2011-12-13 | 102 | 10 | 0.2265 | 35 |  | 0.68657
2011-12-16 | 102 | 10 | 0.68657 | 35
2011-12-13 | 102 | 2.5 | 0.31851 | 37 | 0.5724 | -99 | 5
2011-12-16 | 102 | 2.5 | 0.41052 | 37 | 0.76533
2011-12-13 | 75 | 20 | 0.14156 | 48 |  | 0.18403 | 6
2011-12-13 | 75 | 20 | 0.2265 | 48
2011-12-16 | 75 | 10 | 0.51669 | 49 |  | 0.49546 | 6
2011-12-13 | 75 | 10 | 0.47423 | 49
2011-12-16 | 75 | 2.5 | 1.0334 | 51 | 1.0836 | 1.1289 | 6
2011-12-16 | 75 | 2.5 | 1.2245 | 51 | 0.97525
2011-12-13 | 72 | 20 | 0.057332 | 61 |  | 0.25481
2011-12-16 | 72 | 20 | 0.25481 | 61
2011-12-13 | 72 | 10 | 0.10122 | 62 |  | -99 | 5
2011-12-16 | 72 | 10 | 0.19818 | 62
2011-12-13 | 72 | 2.5 | 0.87059 | 64 | 1.5109 | 1.3024
2011-12-13 | 72 | 2.5 | 1.3024 | 64 | 1.3088
2011-12-13 | 69 | 20 | 0.076442 | 73 |  | -99 | 5
2011-12-13 | 69 | 20 | 0.11042 | 73
2011-12-16 | 69 | 10 | 0.36806 | 74 |  | 0.32205 | 6
2011-12-16 | 69 | 10 | 0.27604 | 74
2011-12-13 | 69 | 2.5 | 0.42468 | 76 | 0.7239 | 0.39991 | 46 | low compared to HPLC
2011-12-13 | 69 | 2.5 | 0.37513 | 76 | 0.71061
2011-12-13 | ADCP | 20 | 0.081397 | 83 |  | 0.29728
2011-12-16 | ADCP | 20 | 0.29728 | 83
2011-12-13 | ADCP | 10 | 0.32559 | 84 |  | 0.30435 | 6
2011-12-16 | ADCP | 10 | 0.28312 | 84
2011-12-16 | ADCP | 2.5 | 0.37513 | 86 | 0.39773 | 0.37513
2011-12-13 | ADCP | 2.5 | 0.23357 | 86 | 0.42005
2011-12-16 | 65 | 20 | 0.38929 | 93 |  | 0.36098 | 6
2011-12-13 | 65 | 20 | 0.33267 | 93
2011-12-16 | 65 | 10 | 0.43176 | 94 |  | 0.42114 | 6
2011-12-16 | 65 | 10 | 0.41052 | 94
2011-12-13 | 65 | 2.5 | 0.24065 | 96 | 0.55412 | 0.50254
2011-12-13 | 65 | 2.5 | 0.50254 | 96 | 0.61472
2011-12-16 | 62 | 20 | 0.19818 | 104 |  | 0.19818
2011-12-13 | 62 | 20 | 0.086352 | 104
2011-12-13 | 62 | 10 | 0.12457 | 105 |  | 0.24773
2011-12-13 | 62 | 10 | 0.24773 | 105
2011-12-16 | 62 | 2.5 | 0.28312 | 107 | 0.35317 | 0.27604 | 6
2011-12-16 | 62 | 2.5 | 0.26896 | 107 | 0.38704
2011-12-13 | 56 | 20 | 0.039637 | 118 |  | -99 | 5
2011-12-13 | 56 | 20 | 0.02902 | 118
2011-12-16 | 56 | 10 | 0.17695 | 119 |  | 0.17695
2011-12-13 | 56 | 10 | 0.11183 | 119
2011-12-16 | 56 | 2.5 | 0.30435 | 121 |  | 0.30435
2011-12-13 | 56 | 2.5 | 0.096261 | 121
2011-12-16 | 46 | 20 | 0.085644 | 131 |  | 0.067949 | 6
2011-12-13 | 46 | 20 | 0.050254 | 131
2011-12-13 | 46 | 10 | 0.14014 | 132 |  | 0.2265
2011-12-16 | 46 | 10 | 0.2265 | 132
2011-12-13 | 46 | 2.5 | 0.08352 | 134 |  | -99 | 5
2011-12-13 | 46 | 2.5 | 0.21942 | 134
2011-12-16 | 42 | 20 | 0.15572 | 148 |  | 0.15572
2011-12-13 | 42 | 20 | 0.033974 | 148
2011-12-13 | 42 | 10 | 0.17695 | 149 |  | 0.18049 | 6
2011-12-16 | 42 | 10 | 0.18403 | 149
2011-12-16 | 42 | 2.5 | 0.28312 | 151 | 0.4656 | 0.28312
2011-12-13 | 42 | 2.5 | 0.16279 | 151 | 0.446
2011-12-16 | 39 | 20 | 0.12528 | 165 |  | 0.10334 | 6
2011-12-13 | 39 | 20 | 0.081397 | 165
2011-12-13 | 39 | 10 | 0.15005 | 166 |  | -99 | 5
2011-12-13 | 39 | 2.5 | 0.16987 | 168 | 0.48093
2011-12-16 | 39 | 2.5 | 0.50254 | 168 | 0.52631 | 0.50254
2011-12-13 | 27 | 20 | 0.065118 | 181 |  | -99 | 5
2011-12-16 | 27 | 20 | 0.09343 | 181
2011-12-13 | 27 | 10 | 0.14581 | 182 |  | 0.16138 | 6
2011-12-13 | 27 | 10 | 0.17695 | 182
2011-12-16 | 27 | 2.5 | 0.33974 | 184 | 0.55379 | -99 | 5
2011-12-16 | 27 | 2.5 | 0.25481 | 184 | 0.54906
2011-12-13 | 2 | 20 | 0.012033 | 196 |  | 0.11679
2011-12-13 | 2 | 20 | 0.11679 | 196
2011-12-16 | 2 | 10 | 0.11679 | 197 |  | 0.11679
2011-12-13 | 2 | 10 | 0.021942 | 197
2011-12-16 | 2 | 2.5 | 0.37513 | 199 | 0.4243 | 0.37513
2011-12-13 | 2 | 2.5 | 0.082813 | 199 | 0.44589
2011-12-13 | 3 | 20 | 0.053793 | 212 |  | 0.072549 | 6
2011-12-16 | 3 | 20 | 0.091306 | 212
2011-12-13 | 3 | 10 | 0.17978 | 213 |  | 0.18013 | 6
2011-12-16 | 3 | 10 | 0.18049 | 213
2011-12-16 | 3 | 2.5 | 0.52377 | 215 | 0.80534 | 0.52377 | 4 | value low compared to HPLC
2011-12-13 | 3 | 2.5 | 0.15713 | 215 | 0.75828
2011-12-13 | 6 | 20 | 0.099092 | 225 |  | 0.1405 | 6
2011-12-16 | 6 | 20 | 0.1819 | 225
2011-12-13 | 6 | 10 | 0.2265 | 226 |  | 0.27958 | 6
2011-12-16 | 6 | 10 | 0.33267 | 226
2011-12-16 | 6 | 2.5 | 0.43176 | 228 | 0.60996 | 0.43176
2011-12-13 | 6 | 2.5 | 0.11466 | 228 | 0.54631
2011-12-13 | 9 | 20 | 0.061579 | 237 |  | 0.26896
2011-12-16 | 9 | 20 | 0.26896 | 237
2011-12-13 | 9 | 10 | 0.21942 | 238 |  | 0.23711 | 6
2011-12-13 | 9 | 10 | 0.25481 | 238
2011-12-16 | 9 | 2.5 | 0.45299 | 240 | 0.4649 | 0.45299
2011-12-13 | 9 | 2.5 | 0.2902 | 240 | 0.45223
2011-12-13 | 12 | 20 | 0.046715 | 253 |  | -99 | 5
2011-12-13 | 12 | 20 | 0.02265 | 253
2011-12-13 | 12 | 10 | 0.062994 | 254 |  | 0.2265
2011-12-13 | 12 | 10 | 0.2265 | 254
2011-12-13 | 12 | 2.5 | 0.15572 | 256 | 0.48344 | -99 | 5
2011-12-13 | 12 | 2.5 | 0.24065 | 256 | 0.51617
2011-12-16 | 14 | 20 | 0.1182 | 268 |  | 0.1182
2011-12-13 | 14 | 20 | 0.045299 | 268
2011-12-13 | 14 | 10 | 0.087767 | 269 |  | 0.16846
2011-12-13 | 14 | 10 | 0.16846 | 269
2011-12-16 | 14 | 2.5 | 0.20526 | 271 | 0.36149 | 0.20597 | 36 | value low compared to HPLC
2011-12-16 | 14 | 2.5 | 0.20668 | 271 | 0.36525
2011-12-13 | 16 | 20 | 0.19535 | 279 |  | 0.19535
2011-12-13 | 16 | 20 | 0.065825 | 279
2011-12-13 | 16 | 10 | 0.26896 | 280 |  | 0.26896
2011-12-13 | 16 | 10 | 0.075027 | 280
2011-12-13 | 16 | 2.5 | 0.11183 | 282 | 0.42462 | 0.38221
2011-12-16 | 16 | 2.5 | 0.38221 | 282 | 0.37811