## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2011-12

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Project: Discovery Islands Circulation

Chief Scientist: Tuele D.

Platform: Haida Joye

Date: 22 July 2011 – 23 July 2011

Processed by: Germaine Gatien

Date of Processing: 27 June 2012 – 27 June 2012

Number of original CTD casts:  7   	Number of casts processed: 7

# INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#1294) was used with no external sensors mountted.

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard and needed to be corrected.

The log book included details from 2 cruises. A copy is in the cruise folder, but for the originals see the folder for 2012-36.

The soak period was long enough for the first 4 casts, but rather short for the last 3.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The CTD Daily Log was obtained.

The configuration file was saved as 1294 Feb2011.xmlcon; the parameters were checked and no errors were found.

The Cruise summary sheet was completed.

The times in the log book are in local time, PST. The computer times appear to be in PDT.

- Conversion of Raw Data

- The raw data were converted using conversion file 1294 Feb2011.con.

The file names were non-standard so were renamed after conversion using the log event numbers.

The CTD was soaked at the surface for about 1 minute for the first 4 casts, but only about 30s for the last 3 casts. There is too much variability to judge if 30s was sufficient for the conductivity to reach equilibrium.

- The descent rate of the CTD was low but very steady.

- Temperature and conductivity look ok.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

The conductivity was not aligned at this point – that will be done later using SHIFT.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

CLEAN was run to add event numbers.

The station names, bottom depths and positions are missing from the headers. Those were entered in file 2011-12-header.csv with the format for latitude and longitude being  50 01.572 N !(deg min) and 127 19.521 W !(deg min) being careful to put a space before the latitude but not longitude. The “Merge:csv to header” routine was used to add those header lines to the CLN files.

8. Checking Headers

Track plots and header checks turned up no header errors. The plots were added to the end of this report.

A cross reference list was produced and details were checked against the log records; no errors were found except for the known discrepancy in times.

The surface check was run. The average surface pressure is -0.23db. At -0.3db the CTD is clearly out of water. During 2011-59 in July 2011 the average surface pressure was -0.02db. The pressure values from the first few scans are probably not reliable and pressure from this sensor is considered accurate to only 1db..So no adjustment will be made to pressure.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. There are differences, but both look reasonable.

ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

10. SHIFT

Tests were run to determine the best parameter for alignment of the conductivity channel. In the past a shift of -1.5 or -2 records has worked well for this sensor, though for some of the data sets little difference was noticed. For these data there is a significant difference between the choices with -2.5records looking best, though some unstable features remain.

SHIFT was run using -2.5 records for all casts.

11. DELETE

The following DELETE parameters were used for both runs:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.1m/s over 5 records between 10db and 10db above maximum pressure ***

There were no warnings in the DELETE log.

*** The minimum drop rate feature was relaxed for these casts due to the slow steady descent rate.

12. CTDEDIT

The DEL files were opened in CTDEDIT was used to clean salinity, remove bad surface records for all casts and bottom records and/or records corrupted by shed wakes for some.

Notes about editing were made in the headers.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined. There were some small unstable features, but nothing that justified further editing.

14. Calibration information

COMPARE – There was no calibration sampling.

Previous Use of CTD – This CTD was used in April 2010, February 2011 and July 2011. During the earliest cruise there were only 3 salinity samples were available, from about 80m. The CTD was found to be higher than the samples by an average of either 0.0001 or 0.002 depending on whether the Niskin was 1.5m or 1m above the CTD. The February cruise included many samples from about 10m. There was a lot of scatter in the results; when large outliers were excluded the CTD salinity was found to be low by 0.008 with a standard deviation of 0.02. Due to the scatter and the method used for comparison, this result was not considered reliable enough to justify recalibration. During 2011-59 there was only 1 salinity sample and there was a long wait before it was analyzed so it was not considered reliable.

Historic ranges – No local climatology was available for this area.

15. CALIBRATION

There is no indication that the salinity or pressure data need recalibration. None was applied.

16. REMOVE and HEADEDIT

REMOVE was run to remove the following channels from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The standards check routine was run and no problems were found.

Header Edit was used to make header names and formats standard and to add a comment.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Particulars including relevant notes from log book

7. Touched bottom.

## Institute of Ocean Sciences

## CRUISE SUMMARY

Cruise ID#: 2011-12

Dates: Start: 22 February 2011 End: 23 February 2011

Location: Kyuquot Sound

Vessel: Haida Joye

Party Chief: Tuele D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/19 SEACAT / 1294       Cruise ID#:	2011-12

| Calibration Information SBE 19 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 1 Apr 2008 | Factory |  |  |
| Conductivity | 1294 | 1 Apr. 2008 | Factory |  |  |
| Pressure Sensor | 163223 | 2 Apr. 2008 | Factory |  |  |