## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 23 August 2020 | ### Added HPLC Data. S.H. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2011-16

Agency: OSD

Location: WCVI

Project: La Perouse/WCVI

Party Chief: Yelland D.

Platform: John P. Tully

Date: May 24, 2011 – June 3, 2011

Processed by: Germaine Gatien

Date of Processing: 6 October 2011 – 27 October 2011

Number of original HEX files: 97

Number of CTD files:   96 (Cast #88 not archived on downcast; Cast #89 done as rerun.)

Number of bottle casts:	46

Number of original TSG files:  unknown (not yet processed)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1119), a Wet Labs Eco-AFL/FL Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4601) and an altimeter (no serial # available). For some casts a second fluorometer was mounted, Seapoint (#2356) with a 10X cable.

The deck unit is unknown.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS is believed to have been a Guildline model 8400B Autosal, serial # 68572, but the analysis sheet is missing.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #2416 and a flow meter. (not confirmed – data not yet available.)

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but it was missing transmissometer, altimeter and TSG information. The altimeter serial number was also missing from the configuration file. The TSG serial number was not recorded, though the sampling notes make it clear that one was used. The rosette logs were in good order.

Errors in format in the entries of station name and water depth prevented that information from appearing in the IOS headers, so the hex files had to be edited. The colon is critical in entries like STATION: B1 and Depth (m): 125. Entries that apply to the whole cruise are easy to fix, but those that vary from cast to cast are not.

Firing Niskin bottles out of order works well only if there is a clear record of which samples come from which Niskin bottle. The rosette log for cast #30 was very confusing. There were notes indicating that the normal method for drawing samples had not been followed but the notes were hard to interpret. Several analysts noted doubts about the labelling of the samples. It appears that someone decided to assign sample numbers in the order that bottles were fired, rather than in the order of Niskin bottle number. So the sampling pressures indicated on the log sheet do not match where the samples actually came from. The information on the chlorophyll sample labels had been altered, but incorrectly. This must have been very confusing at sea, and was extremely difficult to sort out back at IOS. For full details see section 5.

Both CTD salinity channels were found to be lower than bottles by about 0.003 with no obvious time or pressure dependence. However, there is a lot of noise in the fits and some suggestion of Autosal problems since most of the most extreme outliers were analyzed sequentially. This is the second cruise providing similar results and both had salinity analyzed in late August/early September. The results of the other cruise were consistent with a non-linearity problem with the Autosal, but the evidence was weak. From this cruise there is evidence that non-linearity was not significant. With recently improved sampling protocols and newly serviced T and C sensors, we should get better results. No internal standards were run for this data set, so the addition of that step to track the Autosal performance may provide the last piece of the plan to provide better salinity calibration information.

The WetLabs ECO fluorometer was used for the first time at IOS and for 11 casts was mounted with the SeaPoint fluorometer so some comparisons could be made. The ECO sensor has a poorer time response so the traces are much smoother, no filtering is needed, and the alignment needs a larger correction.

From 3 casts with both fluorometers and extracted chlorophyll sampling, it looks like the ECO fluorometer compares to bottles as well, or better than, the Seapoint. The comparison of bottles with downcast fluorescence shows the usual pattern of fluorescence being too high in low-chlorophyll conditions (<2ug/L), and quite close at higher values. The ECO values are likely to be too low above the peak due to slow response. If the peak is broad the ECO can “catch up” so that it is likely to be too high below the peak. If the peak is narrow then the ECO is likely to be too low below the peak.  For identification of the peak and to see how variable fluorescence is, the Seapoint looks better. If the general shape is all that is required, then the ECO is a good choice since it does not need to be pumped. However, when using fluorescence traces at sea to determine the level of maximum chlorophyll for upcast HPLC sampling, assume that the peak is really shallower than it appears by at least 2m, but the offset could be greater if the descent rate of the CTD is >1m/s. The smoothing of the peak can make estimates of the peak depth even more uncertain.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as sampling notes summarizing problems and points of interest with reference to processing. An error in file name format for 4 casts was corrected in the raw files. It is noted that the draw temperature was recorded for DO sampling so that concentration can be calculated in mass units as well as mL/L.

Extracted chlorophyll, nutrients, dissolved oxygen, salinity and NH4 data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated shortly before this cruise.

The calibration constants were checked for all instruments. There were a number of errors in the con files:

- The PAR sensor parameters are not correct for the serial number given, #4601. They would fit sensor #4565, so either the serial number is wrong or the parameters and it is impossible to be sure which. There was a sticky note in the folder that the PAR sensor was #4601, so obviously someone was aware of the inconsistency and believed the serial number to be correct. The same equipment was used on 2011-26, a Line P cruise when the equipment list is usually accurate, and that also gave the serial # as 4601, so it will be assumed that it was #4601. The parameters were changed to match #4601.

- The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

- The calibration for the Wetlabs Eco fluorometer could not be confirmed. The parameters entered are close to those on the characterization sheet from the factory, but the slight difference in dark value suggests that a field test was done, but the test date is not entered in the configuration file. This value will be assumed correct, since no one is available to confirm this.

Many versions of the configuration files exist, but there are actually only 2 different set-ups. Configuration files 2011-16-ctd1 and 2011-16-ctd2 were prepared, the first with the Wetlabs Eco Fluorometer and PAR sensor (casts 1-96, 104-111, the second with two fluorometers and no PAR. The PAR sensor was not always mounted for the casts for which 2011-16-ctd1 will be used. At the end of processing that channel will be removed from casts for which there was no PAR mounted.

##### Initial Rosette File Conversion and DO Calibration Study

In the first attempt at rosette conversion it was discovered that some files could not be converted. There was a note in the log book about an error message “header line length exceeds buffer length”. This provided a clue that there was a problem in the header files. While they looked fine, there was a space after the Depth entry. A blank line is not usually a problem, but in fact there were many spaces after the Depth entry. When those were removed, the files could be converted. If this problem occurs in future, try checking the header file and see if moving forward 1 space after a line takes you to the next line. If not, then get rid of the spaces between. The last 23 hex files needed to be edited to remove the spaces. Then conversion worked.

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. (Since bottles were fired out of order, the file was 1st order on bottle position, sample #s added and it was then reordered on bottle number.) Sample numbers were added to the ADDSAMP file based on rosette log records.

There were a few problems encountered in preparing the ADDSAMP file:

- Sample #126 was never used

- For cast #53 there is a note that there was no water for Niskin #7; according to the rosette file, that bottle was never fired.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-16-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. A problem was found in the analysis spreadsheet where the duplicates for sample #149 were entered separately. There was duplicate sampling but one value was flagged due to a bubble, so only one should have been used and there should not be a flag, just a comment. That was fixed in the 2011-16-oxy.csv file. The data for most of casts #11 and #21 were also removed because there were bad oxygen solubility values. A calculations was made of  Ф  using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.4625 and B = -0.2352 with an R2 value of 0.9863. Only 2 outliers were severe and in both cases the titrated values had been flagged “3” by the analyst who suspected a Niskin mis-trip. For moderate outliers an efficient method was needed for identification. The initial estimates of M and B were used to calculate a difference between each point and the fit:
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. When removing a little more data has little effect on the fit, it may be judged that a reasonable value has been found unless, in so doing, a whole class of points has been removed such as all high values, or all values from late in the cruise.

Among the points with larger differences were another group that had been flagged because of problems with labeling; those were later found to have wrong sample number/ depth correspondence.

Another 12 points with differences >0.04 were noted, and will be examined carefully later in the COMPARE stage so that it can be determined if the CTD data are noisy, or if there are other likely mislabeled samples; then a judgment can be made about whether flags should be assigned to the bottle data.

Using a difference of 0.02 as a cut off, there are 331 points in the fit (out of a total of 371 data points available) and the result is m=0.4661, b= -0.2471 and an R2 value of 0.9997. There was only a small difference between using 0.03 and 0.02 as the cut-off difference. Examining the data removed in going from 0.03 to 0.02 as cut-off shows that no particular range of DO or pressure values were preferentially removed. There were more points from cast #1 than might be expected from a single cast, but that was a Saanich Inlet cast where the CTD DO sensor often has difficulty keeping up with the high DO gradient. Going the extra step to 0.01 as cut-off does improve the fit, but at the cost of removing the higher values of DO Voltage, so that looks like going too far. The following table gives the values found for Soc and Voffset which are the parameters to be entered in the configuration file:

|  | M | b | Soc | Voffset | R^2 |
| --- | --- | --- | --- | --- | --- |
| Original factory cal | 0.4453 | -0.2321 | 0.4453 | -0.5212 |  |
| All data (371 pts) | 0.4625 | -0.2352 | 0.4625 | -0.5085 | 0.9863 |
| Diff<0.05 (358 pts) | 0.4670 | -0.2479 | 0.4670 | -0.5308 | 0.9992 |
| Diff<0.03 (348 pts) | 0.4668 | -0.2483 | 0.4668 | -0.5319 | 0.9996 |
| Diff <0.02 (331 pts) | 0.4661 | -0.2471 | 0.4661 | -0.5301 | 0.9997 |
| Diff <0.01 (240 pts) | 0.4650 | -0.2437 | 0.4650 | -0.5241 | 0.9998 |
| Diff<0.007 (161 pts) | 0.4638 | -0.2401 | 0.4638 | -0.5177 | 0.9999 |

A plot of differences from the fit versus pressure shows that most all the outliers are near the surface; that is expected since there are large gradients at the surface. There is no suggestion of hysteresis in the deep differences. Plots against event # and sample # show no significant patterns. A plot of differences against Niskin bottle # showed no consistent problem with any one bottle.

The configuration files were updated with the new parameters Soc and Voffset and saved with names 2011-16-ctd1-new.xmlcon and 2011-15-ctd2-new.xmlcon.

The ROS files will need to be reconverted.

##### Hysteresis Study

This sensor has been recalibrated since the last hysteresis checks were done, but there was little deep sampling during this cruise. There is no evidence from the checks in the last section that hysteresis is a problem and there is not enough deep data to enable reliable tests to be done to fine-tune H1, H3 and E.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters.

An error in the station name format in the Sea-Bird headers prevented that information from appearing in the IOS format headers, so the ROS files were edited to add “:” between the word Station and the entry. Since the same error will affect the full files as well, it is more efficient to fix this in the hex files and reconvert.

The ROS files were converted to IOS header format. They were put through CLEAN to create BOT files.

A problem arose with file #5 – there was no data in the ROS file. There is no obvious problem in either the hex, bl or hdr file. When the same steps were tried an hour later conversion was successful!

Temperature and salinity were plotted for all BOT files and only 1 significant outlier was found in the primary salinity for cast #49. There was some very odd data in cast #154, but that is because bottles were closed without stopping the CTD just to get them all closed at the end of the cruise. Primary salinity was cleaned lightly using CTDEDIT on cast #49 and the output files were then copied to BOT.

The addsamp.csv file prepared in the DO calibration step was ordered on sample #s and converted to CST files to be used as a framework for the bottle files. It was also used to create SAM files with bottle #s and bottle positions.

The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-16-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2011-16chl.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2011-16-chl.csv which was then converted to individual CHL files. There was no data for cast #1 though sampling is indicated on the rosette sheet. The raw data was checked and it was not there either.

NH4

NH4 data were obtained in file 2011-16_NH4.xls. Event numbers were added to the file, header names were changed to standard format and “NH4:” was entered before the comments to make it clear when the different analyses are merged. The file was saved as 2011-16-NH4.csv and converted to individual NH4 files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-16nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers, header names were changed to standard format and the file was saved as 2011-16-nuts.csv. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. In a few cases the comment appeared to affect all sampling, so “ALL” was placed in front instead of “Nuts”. A few comments were removed as they are not relevant to the merged bottle files. Lines were removed where no nutrient samples were taken. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-16oxy.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified by removing a few unnecessary columns, “DO:” was entered in front of comments and the file was then saved as 2011-16-oxy.csv. That file was converted into individual *.ADD files.

SALINITY

Salinity analysis was done at IOS and received in file 2011-16.xls. This was renamed as 2011-16-sal*.xls (* indicates the date of file creation).  Comments and were included but no quality flags. The file was then saved as 2011-16-sal.csv; and for that version a flag channel was added and flag “4” was added to the sample with comment “leaky liner”. Some columns were removed and a column was added for the event number – the latter information was found in the log book. The station name and sample numbers were combined in a single column; those had to be separated. The data were ordered on event number and sample number. There was one loop sample which was removed. The file was converted into individual SAL files. There were no duplicates.

It was later discovered that 3 of the samples were from another cruise, so the spreadsheet was edited to remove them and the steps above were rerun.

The SAL, CHL, ADD, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

There are problems with cast #30. Some of the bottle values seem wrong above 50m. There are notes on the rosette sheet about samples being drawn in firing order rather than Niskin order, but that seemed very difficult to do, and without a record of pressures to go with the sample numbers it was all very confusing. One note was more explicit, that sample #100 may have come from Niskin #20. The labels on the chlorophyll samples were altered, but not in a way that fit the firing-order comment. The dissolved oxygen analyst also entered comments expressing doubts about the labels.

The MRG file was opened in EXCEL to experiment with reassigning the sample #s. The result that worked does fit the firing-order comment. It appears that someone decided to sample bottles in the order that they were fired, rather than by the Niskin bottle number. So the sampling pressures indicated (and desired) are not the ones that were selected. For example, the 16th bottle fired was Niskin #20 at 40m. The intention was to gather only HPLC samples and assign them sample #98. What actually happened is that sample #98 (HPLC) was gathered from Niskin #16 which was the 18th bottle fired and closed at 30m. And Niskin #20 was sampled for dissolved oxygen, nutrients and extracted chlorophyll and given sample #100. This must have been very confusing at sea, and was extremely difficult to sort out back at IOS.

The following table outlines what samples came from what bottles:

| Sample # | Firing Order # | Niskin # | Pressure | Sampling |
| --- | --- | --- | --- | --- |
| 100 | 18 | 16 | 40 | Chl,Nuts,DO,NH4 |
| 101 | 19 | 17 | 30.5 | Sal,Chl,Nuts,DO,NH4 |
| 102 | 20 | 18 | 30 | HPLC only |
| 103 | 21 | 19 | 21 | HPLC only |
| 98 | 16 | 20 | 10 | Nuts, DO |
| 99 | 17 | 21 | 5 | Chl, Nuts, DO |

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

All the file names 2011-16-0047.* were changed to 2011-16-0046.* to match the downcast file and the rosette log entry.

##### Compare

Salinity

Compare was run with pressure as reference channel. The scatter in the fit is unusually high. While a fairly flat fit can be achieved it involves salinity being low by ~0.003 for both sensors. This is remarkably like the results of 2011-08 with different sensors. In both cases the T & C sensors had been recently calibrated so that it was surprising to see large errors and the same errors in both. Only one sample had any comments about problems.

For 2011-08 there was some suspicion that linearity problems might be involved, but for this cruise when samples are removed that have high noise level in the CTD salinity there is no suggestion of differences varying with salinity values. There are 5 cases with differences >0.006 and 3 of those have remarkably similar differences (~0.011) and 4 of the 5 bottles involved were analyzed sequentially. This does suggest some sort of Autosal problem. When points with high standard deviation in the CTD salinity and differences >0.006 are excluded the fits are quite flat, but the variability is still quite high and the two sensor pairs are remarkably close with the primary low by an average  of 0.0032 and the secondary low by 0.0030. Since the bottles were analyzed in order of sample number, a plot of differences against cast # cannot distinguish between changes through the cruise or changes during analysis. This result is disappointing since other cruises have raised the hope that when sampling protocols improved we could get better calibrations. The use of internal standards may help keep track of how the Autosal is doing.

Most of the major outliers were at the surface in very high gradients. Three deeper outliers were investigated and two were associated with high standard deviation in the CTD data and one was in a complex area where the 1.5m vertical offset between bottle and CTD was likely to lead to a significant difference in salinity. No flags were assigned.

For details see 2011-16-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected the fit of differences against DO concentration is quite flat and most differences are within ±0.1mL/L. No further recalibration is warranted.

The fit of differences against pressure does not suggest any hysteresis but there is no data below 2000db. The plot against file pair number does not indicate any time-dependence in the calibration.

The outliers, differences >0.2mL/L were investigated. Many had been identified as outliers in the calibration study described in section 3.

- Cast #1 – There were two outliers. This was a cast in Saanich Inlet which greatly challenges the SBE DO sensor, so those data are not trusted. No flags were assigned.

- Casts 2, 3 and 5 had a few outliers, but they were in high gradient regions where the distance between CTD and bottle was likely significant.

- Cast #27 – Outlier in original step seems ok in COMPARE – slightly off, but not notably so.

- Cast #30, Sample #97 – outlier in COMPARE and while not flagged, there is a note on the rosette sheet that there was a bubble in the flask. Flag “4” was added with a comment. Other samples from this cast that looked bad in the calibration stage are much better because the sample numbers were rearranged in the MERGE stage. None are outliers now.

- Cast #44, Sample 147 looks out of line. It was fired at 150db but looks like it closed around 125db so it resembles sample #148. The nutrient analyst remarked these could be possible mis-trips or mis-sampling for 4 bottles and for #147 that looks to be the case. The nutrient analyst also questioned whether samples #145 and 146 were mis-trips, but judging by the CTD temperature and salinity there was little difference between conditions at 190 and 171db. The titrated DO sample suggests there was no mis-trip here.

- Cast #47, sample #158 – out of line with SBE DO, but a plot of SBE DO shows a complex structure between 30 and 50db. The nutrients and draw temperature do not suggest a mis-sample. A comment will be entered about it being an outlier in COMPARE but it does not appear to be a bottle problem so no flag will be added.

- Cast #52 and 53, samples #183 and 187 – in fairly high gradient region, likely SBE DO problem

- Cast #61, sample 219 – all samples look like from top 20m, should be 50m. CTD data ok for 50m and values don’t look like any other bottle, so assumed to be mis-trip of Niskin #4. Label DO “3”; nutrients already flagged “3”.

- Cast 67, Sample 253 – all samples look like from above 100m, should be 300m. CTD data ok for 300m and values don’t look like any other bottle, so assumed to be mis-trip of Niskin #8. Label DO “3”; nutrients already flagged “3

- Cast 78, sample 305 – DO out of line, but DO profile complex. Other bottle samples very slightly out of line, but not suggestive of misfire or mistake in sampling. Add comment but no flag.

- Cast #102, sample 381 – Mis-trip. Bottle closed at depth, but all samples look like from above 75db but not from any other bottle. Flag “3” already assigned.

- Cast #88, sample 347 – outlier in compare, already flagged due to bubble.

- Cast #92, sample 363 – outlier replicates, one value selected as better than the other in COMPARE and “6” flag removed.

- Cast #102, sample 392 – mild outlier, but high gradient. Bottle likely ok. No flag attached.

- Cast 109 sample 422 – mild outlier, but lot of SBE DO variability at bottom. No flag justified.

The flags assigned by the analyst were also examined to see if any look inappropriate in the light of the COMPARE results and no further changes were required.

(See 2011-16-dox-comp1*.xls.).

Fluorescence

COMPARE was run using the Wet Labs ECO CTD Fluorescence and the Extracted Chlorophyll from bottles. The ratio of CTD Fluorescence to extracted CHL was high for CHL<1.5 but close to 1 for CHL>1.5. (See 2011-16-chl-fluor-comp1-ECO.xls.)

A second run of COMPARE used the Seapoint CTD fluorometry which was available for just 3 casts. The fluorometer was a little low for the higher CHL values, but read 2 to 3 times too high for CHL<1. (See 2011-16-chl-fluor-comp1-Seapoint.xls.)

The ECO Fluorometry from the 3 casts with both fluorometers mounted were added to the 2nd comparison and a plot was made of the two fluorometer readings versus the bottle CHL. They look quite similar, sometimes one closer, sometimes the other. The range of CHL was low for this comparison with only 3 of the 9 bottles having CHL >1ug/L. A brief check of the downcast fluorescence traces shows that the dark values differ by about 0.03ug/L which is not enough to make a significant difference in the results.

The merge process was repeated to capture the new flags and comments.

Once again data were exported to spreadsheet 2011-16-bottles.xls and compared to the rosette sheets and all expected data are present. A few errors were found and corrected. One bottle file had not been created and while there are no IOS data to be entered there may be a need for this file for other researchers so it was processed. A few samples had been labelled as cast #22 that should have been #23 so were missed in the merges. The merging process was repeated as needed.

The MRG files were edited to remove some bottles that were fired but not needed in the final files: Casts 1 (Niskins 13-24), 52 (18), 53 (9), 93 (none needed so file deleted), 144(1,3,5), 144 (4-24) and 154(1-22).

The MRG files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file and to adjust the record numbers and maxima and minima in the headers.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-16-ctd1-new.con or 2011-16-ctd2-new.con, as appropriate.

A few casts were examined and all expected channels are present. The two temperature and conductivity channels are fairly close during the downcasts but the upcast traces differ much more, and the primary channels are noisier especially during the upcasts.

The two fluorescence traces are similar in a broad sense but the ECO sensor has much less detail. The response time is obviously much slower as the vertical offset between downcast and upcast is twice as large as for the Seapoint sensor.

The DO voltage looks as expected with a vertical offset, but some reasonable detail that should help with alignment.

Transmissivity looks normal.

PAR looks fine. The altimetry looks ok though there are some bottom spikes that may cause trouble.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 3 casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge in these data because of noisy upcast temperatures and frequent stops for bottles, but 4s looks reasonable and is what was found appropriate for 2010-22 when this sensor was used.

ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using 4 casts. The goal is to make upcasts look closer to downcasts on a T-S surface. The temperature traces are extremely noisy on the upcasts making it difficult to judge which settings are best. The choice (α = 0.02, β=7) for the primary and (α = 0.03, β=9) looked best overall. CELLTM was run on all casts using those settings.

In the course of doing these tests it was noted that cast #64 has unusually large differences between the primary and secondary temperature traces. That will be investigated further at the editing phase.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-16-0034 | 800 1000 1950 | -0.0003 -0.0004 -0.0007 | +0.00004 +0.00007 +0.00005 | +0.0008 +0.0011 +0.0015 | Extremely noisy |
| 2011-16-0092 | 800 1000 1950 | -0.0003 -0.0003 -0.0007 | +0.00005 +0.00006 +0.00007 | +0.0007 +0.0010 +0.0014 | Very noisy |
| 2011-16-0104 | 800 1000 1950 | -0.0002 -0.0006 -0.0008 | +0.00007 +0.00007 +0.00007 | +0.0011 +0.0014 +0.0017 | Noisy |
| 2011-16-0111 | 800 1000 1950 | -0.0001 -0.0003 -0.0007 | +0.00010 +0.00009 +0.00010 | +0.0014 +0.0015 +0.0019 | Noisy |

The differences are small but the last one checked had higher conductivity and salinity differences. This may be one odd cast, or this could mark the beginning of drift in one of the sensors. The bottle comparison does not suggest a change, but the comparison was very noisy and there was only 1 salinity bottle taken after cast #104 and it was from 175db. These sensors are both good enough for archiving, but the drift should be checked in future cruises to see if one of these sensors drifts further. Note was made of this in the sensor history.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.

There was a problem with large negative pressures and a few unbelievable positive values nearby in many casts. These problems are at the surface in records that will likely be removed in the DELETE process, but those values might cause some problems in how DELETE works. A few casts were edited by replacing the bad values with pad values in the IOS files and CLEAN was rerun. But it soon became apparent that it was a problem with many casts, so CLEAN was run again on the IOS files in 2 runs. The first run was used to add event numbers and replace pressures <-0.1db with pad values. A second run was used to replace pad values in the pressure channel with interpolated values based on record number. Care will be needed in the editing phase to ensure that this process worked well.

The header check was then rerun.

Checks for off-scale fluorescence were made and the Seapoint fluorescence is found to be near the maximum in cast #119; however, the peaks vary in size so don’t look like more than a few points could have reached or exceeded the maximum. The ECO fluorometer maximum is about 38ug/L except for cast #138 where it reaches almost 130ug/L but only at the surface where many variables are spiking; that data will likely not survive the DELETE step. The peak at cast #10 can’t be compared with bottles since none were taken there. The nearest cast with extracted chlorophyll data is #23 where the peak was 15ug/L. It is interesting to note that the peak for cast #10 is at a level where CHL is not usually sampled, 15ug/L. HPLC was sampled at 15db, presumably based on the downcast fluorescence peak.

There are values of dissolved oxygen that look too high but they are likely in surface spikes that will be removed in processing, but this should be checked later.

The descent rate has some unbelievable values, but again these appear to be surface spikes, probably related to the pressure problems noted above.

Surface check was run and shows an average surface pressure for the cruise was 2.0db. This is a little low for the Tully, but the corresponding salinity values are fairly low, so this may be ok. A few casts were examined carefully to see if negative pressures look believable. In many cases when pressures were around 0 conductivity was very low, the two salinity channels differed significantly as though one sensor was in water and the other not, and transmissivity varied from 75%/m to 0 in just a few scans. All this suggests the CTD being very close to the surface, so the pressure readings are ok.

The cross-reference check was compared with the log book and the only problems were 3 casts with the wrong station names; those were corrected in both bottle files and full profile files.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. The values look reasonable; a selection of casts was checked and no problems were noted.

The Water Depth header was also examined. There were deviations from the log book entries for 6 cases; upon investigation for 2 cases the header looks more reasonable than the log book. In 3 cases the header entry looks wrong and in one cast the log entry had been updated but that information had not been changed in the header. So the header was corrected for 4 casts. (31, 90, 96 & 132). None of these casts were rosette casts so the bottle files do not need correction.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. A selection of casts was checked and no problems were found.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. For the Seapoint fluorometer, the shift applied is almost always +24 records, and for this cruise it does appear to be the best choice again.

For the ECO sensor the estimates give highly variable results, from about 1s to 5s, with 2.5s looking about average. This is partly because of the “steppy” nature of the traces and the fact that peaks are not resolved well. There were some casts with both sensors mounted, so it may be easier to choose a shift for the ECO sensor by seeing how it compares to the Seapoint. By comparing particular features it does look like the difference in response time is >1s, maybe 1.5s. This is still hard to judge because if the ECO sensor does not reach a peak value, it will look like it is well-aligned with the Seapoint just below the peak. But the overall offset is consistent with the judgment based on temperature traces, with differences ranging from 1.8 to 4.6s.

Tests were done on a few casts using settings from +2s to +3s and +2.5s seems best well below the peak, but often too much at the peak. The temperature data are very noisy on the upcasts, so this really is a best guess and +2s looks reasonable and best near the peak.

SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records. (Output: SHFFL1)

Those files were copied to files *.SHFFL2 and then a second run of SHIFT was used on only the casts that had 2 fluorometers, to advance the Seapoint fluorescence channel by +24 records. This results in a complete set of SHFFL2 files that have both fluorometers aligned.

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of +0.5s to the secondary conductivity. SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concern upcast sections of two casts.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in late March 2011 and this was the first known use since that time.

2. Dissolved Oxygen

The DO sensor was repaired and recalibrated in April 2011. This was the first known use after that date.

3. Pressure

The sensor was recalibrated in April 2011 and this was the first known use since that time.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Temperatures were a little above the historic maxima at some depths for casts near the 1000m contour in the southern part of the cruise area and seaward of the 2000m contour in the north. Salinity was low for one cast at the bottom. These excursions look real, not indicative of instrument malfunction.

Repeat Casts –

The only repeat cast was run because of a failure to acquire the downcast data on the first cast.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

At this stage a decision has to be made about whether to edit primary or secondary channels. A few casts were opened in CTDEDIT to compare the salinity channels. In general the secondary looks better, with fewer spikes, but for cast #64 the secondary temperature looks odd, as though the response was poor. The secondary salinity looks much noisier than the primary, and the upcast is no better. It is likely that there was some sort of obstruction affecting the secondary channels. The choice of primary for cast #64 and secondary for all other casts will provide the best data and since the two channels were very close, no special recalibration will be required to make them match.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as they appear to be corrupted by either ship effects or too short a wait at the top.

Cast #89 at LBP5 had very odd T-S features around 245 and 315db. This was a cast that was repeated because no data were available for the first downcast, so there are 2 upcasts and 1 downcast to compare. The features vary through that time, but in all cases there are signs of active mixing. The features are stable and no editing was applied.

All EDU files were copied to EDT.

In the course of editing, a few casts were investigated to see if DELETE selected the best data near the surface. Where it had not, the SHFC1 file was edited to remove data acquired during the soak period, then DELETE was rerun. Casts edited were: 25, 40, 42, 86, 117, 119 and 121. These casts were ones that did not follow the standard deployment method described in the sampling notes; acquisition was started before the soak period so that DELETE selected some of the initial lowering and patched it to the main lowering. Even after editing to remove the initial lowering the near-surface data are very noisy, probably due to both ship effects and the mixing caused by the CTD bumping up and down during the wait before the main cast was run.

For cast #9 more than 1db of bottom data were removed, so the altimetry header was adjusted to reflect the fact that the deepest data available are further from the bottom than before and an explanatory note was added.

##### Initial Recalibration

The pressure looks ok. COMPARE shows no need to recalibrate dissolved oxygen. The salinity comparison would suggest the values are low, but the results are not trusted. When 2011-26 has been processed there will be more information. For now no correction will be applied to the salinity.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure the deepest values looked a little high, but shallower values look ok, though there is a lot of scatter. When the differences are plotted against DO concentration and outliers excluded, a reasonable fit is found with SBE DO values high by about 0.06mL/L at low DO and ~0 at high values:

CTD DO corrected = 1.0082 * CTD DO - 0.0624mL/L

File 2011-16-recal.ccf was prepared to apply the above correction to the Oxygen:Dissolved:SBE channel. It was first applied to the thinned files and COMPARE was rerun; the results are satisfactory.

CALIBRATE was then run on the EDT files.

This correction does not apply to the bottle files.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering the Seapoint fluorescence before bin-averaging; this applies to only 11 casts. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

Only the Seapoint fluorescence needs filtering, so all COR1 files were copied to *.FIL, and then the casts with a Seapoint fluorometer were treated as follows: a median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness, with output *.FIL. A few casts were examined before and after this step and showed that the filter was effective.

At this point the two fluorometers were compared. First, unfiltered data were examined, with special focus on the 3 casts which included both fluorometers and extracted CHL sampling. Plots of downcast data show that the peaks are closer after alignment than they were before, though in a few cases the ECO peak is still lower than the Seapoint. For cast #119 the fluorescence peak was at the surface and in that case the chlorophyll values are close to those of both fluorometers, being especially close to the Seapoint at 5m and the Eco at 10m and 20m. We should not expect great matches since there are several sharp peaks in the Seapoint that would not likely match the upcast, and the ECO cannot resolve fine details. For casts #102 and 126 the peak is around 20-30m and lower than for cast #119. The peak values are quite close for both sensors, being about 2X that of the extracted CHL. The peak for cast #102 is broader in the ECO fluorometer than the Seapoint.

During a recent cruise it was noted that the upcast and downcast fluorescence traces were offset by more than expected, and in some cases the extracted chlorophyll values show no peak whereas the fluorescence trace does. Those data were raw, so the fluorescence had not been aligned, which would cause a vertical offset of about 4m if the descent rate is 1m/s and more if that rate were higher. Moreover, the ECO will likely fail to resolve a narrow peak which might confuse the issue. It is also clear that both types of fluorometers do a poor job in low chlorophyll conditions, which can complicate that shape.

HPLC sampling is done during the upcast at levels 5m below and 5m above the depth at which the downcast fluorescence reached its maximum value. When making this judgment at sea, it would be best to estimate the depth of the peak by subtracting at least 2m from that seen in the downcast trace.

Further tests were run using both types of fluorometer on at least one other cruise, so conclusions about which fluorometer is better are premature. But from this cruise it looks like the ECO fluorometer compares to bottles as well, or better than, the Seapoint. The comparison of bottles with downcast fluorescence shows the usual pattern of fluorescence being too high for chlorophyll <2ug/L, and quite close at higher CHL values. The ECO is probably too low above the peak due to slow response. If the peak is broad the ECO can “catch up” so that it is likely to be too high below the peak. If the peak is narrow then the ECO is likely to be too low below the peak.  For identification of the peak and to see how variable fluorescence is, the Seapoint is probably better. If the general shape is all that is required, then the ECO is a good choice since it does not need to be pumped. However, it is important to consider the alignment offset in deciding the depth of the peak CHL.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

The following casts had a PAR sensor mounted: 1-29, 45-63, 73-82, 96, 106-109, 128-154. Plots were made to ensure no PAR casts had been missed.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor (except for #64) to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

For cast #64 only the following casts were removed:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4mL/L from 0 to 200db

±0.2mL/L from 200db to 500db

±0.1mL/L below 500db

For details on the processing see processing report: 2011-16-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

Profile plots were made. Some of the casts from Rivers Inlet had very low surface salinity, but this is likely real given the locations; no further editing was applied.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from ~85% to ~175%, the maximum being in Saanich Inlet. A check of the bottle files shows that the SBE values were close to titrated samples at 5m, so this is presumed to reflect real conditions.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run on the final files and a problem was found with the order of channels, the fluorescence channel (Wet labs) had not been moved. This is likely due to an error in the way the new channel was converted to IOS HEADERS. This is not important for the CTD files, but for the CHE it means the bottle data come between the fluorescence and all the other CTD channels. To fix this, the channel name was changed manually in the MRGOX2 files to Fluorescence:Wetlabs:ECO-AFL. Then Reorder was rerun successfully. HEADEDIT was rerun.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. An error was found for one chlorophyll/phaeo sample. It had been labelled as sample #191 (Niskin #7), but in fact there were no samples with that number because there was no water in the bottle. In the merging process, because there was no line for sample #191, the CHL got added to #192. Using a text editor, the data were re-assigned to sample #194 (Niskin #10), since the 5m samples came from that bottle according to a note on the rosette log sheet.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and a format error was found; that was fixed and no further errors were found.

##### 25. Thermosalinograph Data

NOTE: As of October 2011 the raw Thermosalinograph data were unavailable. It was noted at sea that there were errors in the time channel. The chief scientist believes this can be corrected using ship position records. The data will be processed if and when they become available.

NOTE There was just one loop sample, salinity = 31.6469 taken at 1353 on May 25, 2011.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (Notes from log book and sampling notes)

1-96. Used PAR and Wet labs Eco fluorometer

97-102. Used 2 fluorometers, no PAR

104-111. Used PAR and Wetlabs Eco fluorometer

112-127. Used 2 fluorometers, no PAR

128-154. Used PAR and Wetlabs Eco fluorometer

1. 24 Niskins closed but only 12 sampled. Rest closed only to check rosette.

52. Niskin 18 closed to test lanyard. Not needed in CHE file.

53. Niskin 7 closed but was empty. Niskin 10 was also closed at 5db so all samples were taken from that bottle. Labels are ok.

67. Likely mis-trip Niskin #8.

68. Spike in salinity at 1350db on upcast. Perhaps a blockage since salinity low.

79. Should be called 80.

84. Possible mis-trip Niskin #9.

86. Bottom depth reading jumped around a lot from 1075 to 1120. Altimeter never kicked in.

88. No archiving of downcast data. Upcast ok.

89. Downcast can substitute for missing downcast data of 88.

88-92. Error in file name format – corrected in raw files.

93. All Niskins closed just so that they are closed before long sail; no sampling.

102. Niskin 4 (1000db) tripped without the 30s wait. Niskin #3 probable mis-trip.

104. Pressure spikes at 161db on way down.

111. Winch slowed down due to bad wire wraps

144. Not sure if top valves closed for Niskin 1,2,3; so closed 1 and 4 at 30db, 2 and 5 at 5db and 6 at surface. Samples came from 4,5,6 so Niskins 1,2,3 can be deleted from CHE file.

146. Only Niskins 1,2,3 needed in CHE file. Rest closed before long sail.

154. Only Niskins 23 and 24 at 10 and 5db were sampled. Rest closed but not needed in CHE file.

TSG – problem with time.

Files with PAR: 1-29, 45-63, 73-82, 96, 106-109, 128-154.

Files with 2 fluorometers: 96-102, 113-127.

There were many out-of-order bottle firings.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3396 | 1Apr2011 | Factory |  |  |
| Conductivity | 2374 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2754 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2668 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug2010 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2216 | ? |  |  |  |
| Seapoint Fluorometer | 2356 | ? |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

# TSG (data not yet available, so not confirmed)

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2011-16

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24Apr09 | Factory |  |  |
| Conductivity | 2488 | 24Apr09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 2416 | 23Dec06 |  |  |  |