## 2011-16-DO-cal
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Summary of Soc Voffset including the original values in the factory calibration
88 | 10.467 | 10.9118 | 10.9115 | 30.8301 | 30.831 | 6.3586 | 2.8471 | 17 | 19 | 351 | 6.2359 | 1.0816 | 6.745 |  | 7.3624e-05 | 0.4625 | -0.2352
49 | 10.495 | 10.9087 | 10.9124 | 30.8958 | 30.8948 | 6.3564 | 3.1319 | 8 | 10 | 173 | 6.2337 | 1.2134 | 7.564 |  | 8.6037e-05 | 0.4625 | -0.2352 |  | m | b | Soc | Voffset | R^2
63 | 201.379 | 6.7197 | 6.7197 | 33.9598 | 33.9609 | 6.8414 | 1.1422 | 6 | 6 | 232 | 6.9061 | 0.29336 | 2.026 |  | 0.00027779 | 0.4625 | -0.2352 | Original factory cal | 0.4453 | -0.23209 | 0.4453 | -0.5212
88 | 200.944 | 6.3513 | 6.3482 | 33.9265 | 33.9274 | 6.9015 | 1.2656 | 8 | 8 | 342 | 6.9706 | 0.35047 | 2.443 |  | 0.00032006 | 0.4625 | -0.2352 | All data (371 pts) | 0.4625 | -0.2352 | 0.4625 | -0.50854 | 0.9863
[18 empty] | Diff<0.05 (358 pts) | 0.467 | -0.2479 | 0.467 | -0.53084 | 0.9992
106 | 150.599 | 7.1681 | 7.1689 | 33.8086 | 33.8091 | 6.778 | 1.5295 | 9 | 9 | 410 | 6.793 | 0.47254 | 3.21 | 6 | 0.00035418 | 0.4625 | -0.2352 | Diff<0.03 (348 pts) | 0.4668 | -0.2483 | 0.4668 | -0.53192 | 0.9996
47 | 20.615 | 9.4958 | 9.4833 | 31.9108 | 31.9167 | 6.5127 | 2.737 | 5 | 7 | 160 | 6.4021 | 1.0311 | 6.601 |  | 0.00037508 | 0.4625 | -0.2352 | Diff <0.02 (331 pts) | 0.4661 | -0.2471 | 0.4661 | -0.53014 | 0.9997
67 | 10.537 | 11.4523 | 11.457 | 31.8701 | 31.8695 | 6.2441 | 2.839 | 20 | 22 | 265 | 6.1216 | 1.0783 | 6.601 |  | 0.00048551 | 0.4625 | -0.2352 | Diff <0.01 (240 pts) | 0.465 | -0.2437 | 0.465 | -0.52409 | 0.9998
63 | 125.275 | 7.5736 | 7.5729 | 33.8577 | 33.857 | 6.7135 | 1.3368 | 9 | 9 | 235 | 6.7027 | 0.38357 | 2.571 | 6 | 0.00049563 | 0.4625 | -0.2352 | Diff<0.007 (161 pts) | 0.4638 | -0.2401 | 0.4638 | -0.51768 | 0.9999
102 | 10.268 | 10.9431 | 10.9433 | 31.8169 | 31.8171 | 6.3146 | 2.8486 | 19 | 21 | 397 | 6.1924 | 1.0828 | 6.705 |  | 0.00051613 | 0.4625 | -0.2352
5 | 20.345 | 8.349 | 8.3528 | 31.7344 | 31.7311 | 6.6888 | 2.1302 | 5 | 5 | 25 | 6.5826 | 0.75061 | 4.941 |  | 0.00057468 | 0.4625 | -0.2352
28 | 49.937 | 8.3414 | 8.3413 | 32.5289 | 32.5293 | 6.6555 | 2.6955 | 11 | 11 | 76 | 6.5747 | 1.0121 | 6.654 |  | 0.00057716 | 0.4625 | -0.2352
73 | 99.519 | 7.9747 | 7.9681 | 33.5377 | 33.543 | 6.6668 | 1.672 | 11 | 11 | 283 | 6.6307 | 0.5375 | 3.564 |  | 0.00059254 | 0.4625 | -0.2352
92 | 100.265 | 7.6017 | 7.6034 | 33.283 | 33.2819 | 6.7345 | 2.018 | 14 | 14 | 369 | 6.7019 | 0.69876 | 4.683 |  | 0.00063119 | 0.4625 | -0.2352
53 | 5.029 | 10.5237 | 10.5238 | 30.7143 | 30.7109 | 6.4169 | 3.3897 | 10 | 10 | 194 | 6.2904 | 1.3319 | 8.378 |  | 0.00063842 | 0.4625 | -0.2352
28 | 299.556 | 6.0299 | 6.0304 | 34.0087 | 34.0087 | 6.9495 | 0.95811 | 3 | 3 | 68 | 7.1128 | 0.20723 | 1.474 |  | 0.00069465 | 0.4625 | -0.2352
109 | 4.795 | 10.9407 | 10.9444 | 31.828 | 31.8287 | 6.3145 | 2.9417 | 7 | 9 | 428 | 6.188 | 1.1244 | 6.958 |  | 0.00089164 | 0.4625 | -0.2352
73 | 50.851 | 8.3211 | 8.3222 | 32.5642 | 32.5647 | 6.657 | 2.6273 | 13 | 13 | 285 | 6.5772 | 0.97899 | 6.439 |  | 0.00090966 | 0.4625 | -0.2352
79 | 9.69 | 11.5484 | 11.5496 | 30.9055 | 30.9095 | 6.2695 | 3.0601 | 4 | 4 | 317 | 6.1456 | 1.181 | 7.258 |  | 0.00091917 | 0.4625 | -0.2352
78 | 91.042 | 7.8609 | 7.8605 | 33.0947 | 33.0969 | 6.7033 | 1.7348 | 1 | 1 | 304 | 6.6607 | 0.56811 | 3.784 |  | 0.00094489 | 0.4625 | -0.2352
24 | 20.551 | 9.5061 | 9.5129 | 32.0363 | 32.033 | 6.506 | 2.878 | 6 | 6 | 52 | 6.3954 | 1.0969 | 7.015 |  | 0.00099645 | 0.4625 | -0.2352
9 | 19.621 | 8.5454 | 8.5456 | 31.5648 | 31.5636 | 6.6667 | 2.25 | 5 | 5 | 32 | 6.5588 | 0.80441 | 5.276 |  | 0.0010099 | 0.4625 | -0.2352
119 | 20.239 | 8.4089 | 8.4192 | 31.2793 | 31.2784 | 6.6996 | 2.6121 | 4 | 4 | 435 | 6.5927 | 0.97395 | 6.421 | 6 | 0.0010618 | 0.4625 | -0.2352
92 | 20.071 | 10.3441 | 10.3426 | 32.5221 | 32.5224 | 6.3681 | 2.8461 | 18 | 20 | 373 | 6.2552 | 1.0801 | 6.756 |  | 0.0010738 | 0.4625 | -0.2352
67 | 124.483 | 7.4871 | 7.4873 | 33.6844 | 33.6835 | 6.7344 | 1.6512 | 14 | 14 | 259 | 6.7236 | 0.52739 | 3.546 |  | 0.0010782 | 0.4625 | -0.2352
73 | 74.647 | 8.016 | 8.0182 | 32.8353 | 32.8386 | 6.6911 | 2.3068 | 12 | 12 | 284 | 6.6334 | 0.83049 | 5.509 |  | 0.0011931 | 0.4625 | -0.2352
40 | 100.069 | 7.6442 | 7.6445 | 33.3926 | 33.3916 | 6.7232 | 1.829 | 10 | 10 | 136 | 6.6901 | 0.61195 | 4.094 |  | 0.0012181 | 0.4625 | -0.2352
106 | 20.242 | 10.5926 | 10.5935 | 31.9515 | 31.952 | 6.357 | 2.8538 | 15 | 17 | 416 | 6.2434 | 1.086 | 6.78 |  | 0.0012847 | 0.4625 | -0.2352
26 | 30.743 | 8.8743 | 8.8777 | 32.2945 | 32.2935 | 6.5867 | 2.5708 | 8 | 8 | 62 | 6.4871 | 0.95251 | 6.179 |  | 0.0012883 | 0.4625 | -0.2352
30 | 150.075 | 7.5282 | 7.5272 | 33.8295 | 33.8285 | 6.7217 | 1.4583 | 11 | 11 | 93 | 6.7327 | 0.44054 | 2.966 |  | 0.0012937 | 0.4625 | -0.2352
63 | 75.015 | 7.7627 | 7.7605 | 33.0712 | 33.0704 | 6.7192 | 2.1429 | 11 | 11 | 237 | 6.6637 | 0.75454 | 5.028 |  | 0.0013642 | 0.4625 | -0.2352
61 | 19.715 | 8.8416 | 8.8486 | 32.1223 | 32.1252 | 6.5988 | 2.3586 | 6 | 8 | 221 | 6.4901 | 0.85422 | 5.544 |  | 0.0014379 | 0.4625 | -0.2352
106 | 10.093 | 10.8644 | 10.8661 | 31.9583 | 31.9591 | 6.3196 | 2.86 | 16 | 18 | 417 | 6.1975 | 1.0861 | 6.731 |  | 0.0014667 | 0.4625 | -0.2352
126 | 124.76 | 6.6407 | 6.6409 | 33.6127 | 33.6134 | 6.8696 | 1.2875 | 4 | 4 | 444 | 6.8674 | 0.3588 | 2.464 |  | 0.0014723 | 0.4625 | -0.2352
49 | 4.98 | 10.9872 | 10.9869 | 30.827 | 30.8268 | 6.3484 | 3.1872 | 9 | 11 | 174 | 6.2213 | 1.2374 | 7.698 |  | 0.0015248 | 0.4625 | -0.2352
126 | 100.477 | 7.1695 | 7.1691 | 32.9502 | 32.9518 | 6.8161 | 1.7319 | 5 | 5 | 445 | 6.7874 | 0.56738 | 3.851 |  | 0.0015561 | 0.4625 | -0.2352
119 | 10.112 | 9.1038 | 9.1174 | 30.868 | 30.8641 | 6.6138 | 2.847 | 5 | 6 | 436 | 6.4952 | 1.0831 | 7.035 |  | 0.0015614 | 0.4625 | -0.2352
61 | 75.341 | 7.8121 | 7.8122 | 33.0013 | 33.0009 | 6.7148 | 2.165 | 3 | 3 | 218 | 6.6592 | 0.76782 | 5.113 | 6 | 0.0016939 | 0.4625 | -0.2352
61 | 49.739 | 8.3327 | 8.3344 | 32.5914 | 32.5912 | 6.6541 | 2.5997 | 4 | 4 | 219 | 6.5733 | 0.96543 | 6.346 |  | 0.0017217 | 0.4625 | -0.2352
67 | 50.102 | 8.5094 | 8.5097 | 32.4979 | 32.4984 | 6.6318 | 2.5915 | 17 | 17 | 262 | 6.5502 | 0.96164 | 6.299 | 6 | 0.0017431 | 0.4625 | -0.2352
9 | 29.842 | 8.2865 | 8.2845 | 31.9804 | 31.9822 | 6.6876 | 2.0817 | 4 | 4 | 31 | 6.5899 | 0.72581 | 4.783 | 6 | 0.0017599 | 0.4625 | -0.2352
30 | 302.144 | 6.2307 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 7.0803 | 0.23982 | 1.698 |  | 0.0017805 | 0.4625 | -0.2352
30 | 250.628 | 6.6667 | 6.6669 | 33.9585 | 33.9571 | 6.8498 | 1.169 | 7 | 7 | 89 | 6.9592 | 0.30363 | 2.113 |  | 0.0018153 | 0.4625 | -0.2352
73 | 29.994 | 8.9575 | 8.9633 | 32.4574 | 32.4571 | 6.5676 | 2.8164 | 14 | 16 | 286 | 6.4671 | 1.0656 | 6.891 |  | 0.0018157 | 0.4625 | -0.2352
49 | 29.575 | 8.6523 | 8.6486 | 31.7829 | 31.7867 | 6.6414 | 2.175 | 6 | 6 | 171 | 6.5415 | 0.76894 | 5.03 | 6 | 0.0018162 | 0.4625 | -0.2352
102 | 75.198 | 7.676 | 7.6748 | 32.7672 | 32.7672 | 6.7459 | 2.3864 | 15 | 15 | 393 | 6.691 | 0.86669 | 5.799 | 6 | 0.0018432 | 0.4625 | -0.2352
102 | 199.538 | 6.7257 | 6.7264 | 33.8724 | 33.8724 | 6.8444 | 1.4558 | 10 | 10 | 388 | 6.9074 | 0.43996 | 3.039 |  | 0.0018629 | 0.4625 | -0.2352
88 | 252.15 | 6.0085 | 6.0136 | 33.976 | 33.9779 | 6.9545 | 1.0399 | 7 | 7 | 341 | 7.0748 | 0.24764 | 1.752 |  | 0.0018664 | 0.4625 | -0.2352
92 | 5.11 | 10.5406 | 10.5436 | 31.6444 | 31.6564 | 6.3766 | 2.8434 | 20 | 22 | 375 | 6.2508 | 1.0779 | 6.738 |  | 0.001929 | 0.4625 | -0.2352
106 | 4.991 | 11.0243 | 11.0219 | 31.953 | 31.954 | 6.2981 | 2.8607 | 17 | 19 | 418 | 6.1718 | 1.0859 | 6.702 |  | 0.0019644 | 0.4625 | -0.2352
84 | 20.967 | 10.0283 | 10.0287 | 31.2698 | 31.2701 | 6.4636 | 2.6146 | 8 | 9 | 329 | 6.3513 | 0.97209 | 6.174 | 6 | 0.0019665 | 0.4625 | -0.2352
23 | 20.296 | 8.7379 | 8.7355 | 31.4964 | 31.4973 | 6.641 | 2.4244 | 7 | 8 | 42 | 6.5328 | 0.88401 | 5.775 |  | 0.0020547 | 0.4625 | -0.2352
92 | 124.636 | 7.2073 | 7.2074 | 33.6363 | 33.6374 | 6.7796 | 1.8691 | 13 | 13 | 368 | 6.7716 | 0.63131 | 4.275 |  | 0.002067 | 0.4625 | -0.2352
88 | 149.618 | 6.7559 | 6.756 | 33.8524 | 33.8529 | 6.8405 | 1.6069 | 10 | 10 | 344 | 6.859 | 0.51013 | 3.499 |  | 0.0021211 | 0.4625 | -0.2352
102 | 150.085 | 7.0192 | 7.0195 | 33.7499 | 33.7504 | 6.8038 | 1.7797 | 12 | 12 | 390 | 6.8199 | 0.58579 | 3.995 |  | 0.0021215 | 0.4625 | -0.2352
26 | 76.477 | 7.8819 | 7.8825 | 32.9178 | 32.9195 | 6.7078 | 2.1599 | 6 | 6 | 60 | 6.6526 | 0.76586 | 5.095 |  | 0.0021225 | 0.4625 | -0.2352
47 | 5.045 | 11.5066 | 11.5064 | 30.6824 | 30.6823 | 6.2839 | 3.0581 | 7 | 9 | 162 | 6.1562 | 1.177 | 7.246 |  | 0.0021736 | 0.4625 | -0.2352
102 | 5.419 | 10.9396 | 10.94 | 31.8171 | 31.8176 | 6.3151 | 2.856 | 20 | 22 | 398 | 6.1891 | 1.0835 | 6.706 |  | 0.0021976 | 0.4625 | -0.2352
106 | 50.671 | 7.7815 | 7.7819 | 32.5418 | 32.5426 | 6.7396 | 2.6362 | 13 | 13 | 414 | 6.663 | 0.98185 | 6.542 |  | 0.0022009 | 0.4625 | -0.2352
55 | 4.809 | 11.388 | 11.3906 | 30.8316 | 30.8325 | 6.2939 | 3.1254 | 5 | 5 | 206 | 6.1662 | 1.2126 | 7.477 | 6 | 0.0022911 | 0.4625 | -0.2352
37 | 100.077 | 8.0226 | 8.0299 | 33.4058 | 33.3984 | 6.6653 | 1.7555 | 13 | 13 | 119 | 6.6293 | 0.5791 | 3.839 |  | 0.0023871 | 0.4625 | -0.2352
37 | 5.259 | 11.5081 | 11.5083 | 31.7488 | 31.7491 | 6.2415 | 2.8048 | 19 | 19 | 125 | 6.1148 | 1.0596 | 6.479 |  | 0.0024409 | 0.4625 | -0.2352
84 | 5.433 | 10.8034 | 10.8042 | 30.8934 | 30.8938 | 6.3709 | 2.8552 | 10 | 12 | 331 | 6.2444 | 1.0829 | 6.762 |  | 0.0024527 | 0.4625 | -0.2352
102 | 50.333 | 8.132 | 8.1322 | 32.519 | 32.5195 | 6.6874 | 2.6329 | 16 | 16 | 394 | 6.6082 | 0.98499 | 6.509 |  | 0.0024693 | 0.4625 | -0.2352
28 | 75.143 | 8.1022 | 8.1024 | 33.2553 | 33.2559 | 6.6599 | 1.9242 | 10 | 10 | 75 | 6.6022 | 0.65736 | 4.34 |  | 0.002632 | 0.4625 | -0.2352
61 | 4.369 | 11.9998 | 12.0055 | 31.1361 | 31.1366 | 6.2007 | 2.8564 | 8 | 10 | 223 | 6.0727 | 1.0832 | 6.578 |  | 0.002684 | 0.4625 | -0.2352
63 | 20.501 | 9.193 | 9.1941 | 32.4637 | 32.4638 | 6.5331 | 2.8445 | 14 | 16 | 240 | 6.4239 | 1.0777 | 6.923 |  | 0.0026897 | 0.4625 | -0.2352
84 | 30.228 | 8.9961 | 8.9959 | 31.8356 | 31.8364 | 6.5884 | 2.3931 | 7 | 7 | 328 | 6.4875 | 0.8689 | 5.637 |  | 0.0026936 | 0.4625 | -0.2352
84 | 75.617 | 7.8723 | 7.8713 | 32.8466 | 32.8484 | 6.7124 | 2.1117 | 5 | 5 | 326 | 6.6565 | 0.73868 | 4.917 |  | 0.0027859 | 0.4625 | -0.2352
40 | 175.588 | 6.9961 | 6.9961 | 33.8761 | 33.8764 | 6.8018 | 1.362 | 7 | 7 | 133 | 6.8405 | 0.39193 | 2.681 |  | 0.0027927 | 0.4625 | -0.2352
102 | 126.078 | 7.3288 | 7.33 | 33.6634 | 33.6641 | 6.7596 | 1.6523 | 13 | 13 | 391 | 6.7517 | 0.52609 | 3.552 | 4 | 0.0029202 | 0.4625 | -0.2352
84 | 50.12 | 8.2994 | 8.3005 | 32.4457 | 32.4454 | 6.6654 | 2.284 | 6 | 6 | 327 | 6.585 | 0.81822 | 5.388 |  | 0.0029395 | 0.4625 | -0.2352
26 | 124.446 | 7.2863 | 7.2864 | 33.8078 | 33.8077 | 6.7598 | 1.3528 | 4 | 4 | 58 | 6.7509 | 0.38751 | 2.616 | 6 | 0.0029733 | 0.4625 | -0.2352
67 | 29.494 | 8.9716 | 8.9685 | 32.4051 | 32.406 | 6.5677 | 2.7796 | 18 | 20 | 263 | 6.4667 | 1.0474 | 6.773 |  | 0.0030106 | 0.4625 | -0.2352
67 | 200.635 | 7.0949 | 7.093 | 33.9465 | 33.9477 | 6.7832 | 1.1693 | 10 | 10 | 255 | 6.8428 | 0.30251 | 2.07 |  | 0.0030851 | 0.4625 | -0.2352
126 | 49.797 | 8.3539 | 8.3573 | 31.7946 | 31.7951 | 6.6855 | 2.6342 | 7 | 7 | 447 | 6.6041 | 0.9862 | 6.513 |  | 0.0030907 | 0.4625 | -0.2352
73 | 10.073 | 11.3663 | 11.3645 | 32.0468 | 32.0476 | 6.2485 | 2.8361 | 16 | 18 | 288 | 6.1259 | 1.0733 | 6.575 |  | 0.0032064 | 0.4625 | -0.2352
73 | 5.321 | 11.3857 | 11.3857 | 32.047 | 32.0475 | 6.246 | 2.8474 | 17 | 19 | 289 | 6.1197 | 1.0785 | 6.6 |  | 0.0032541 | 0.4625 | -0.2352
37 | 10.963 | 11.5104 | 11.5105 | 31.7505 | 31.7506 | 6.2411 | 2.8054 | 18 | 18 | 124 | 6.1189 | 1.059 | 6.48 |  | 0.0032797 | 0.4625 | -0.2352
26 | 100.42 | 7.525 | 7.5236 | 33.587 | 33.5891 | 6.7329 | 1.6773 | 5 | 5 | 59 | 6.7011 | 0.53723 | 3.6 |  | 0.0033149 | 0.4625 | -0.2352
40 | 49.627 | 8.2142 | 8.2132 | 32.5642 | 32.5666 | 6.673 | 2.3043 | 12 | 12 | 138 | 6.5928 | 0.83394 | 5.498 |  | 0.0034138 | 0.4625 | -0.2352
109 | 29.679 | 8.4654 | 8.4642 | 32.0873 | 32.0863 | 6.6561 | 2.6181 | 4 | 4 | 425 | 6.5574 | 0.97219 | 6.375 | 6 | 0.0034612 | 0.4625 | -0.2352
126 | 9.769 | 11.2679 | 11.268 | 30.7778 | 30.7793 | 6.3123 | 3.1134 | 10 | 12 | 450 | 6.1885 | 1.2013 | 7.434 |  | 0.0034921 | 0.4625 | -0.2352
30 | 175.609 | 7.2473 | 7.2473 | 33.9007 | 33.8994 | 6.7617 | 1.3549 | 9 | 9 | 91 | 6.7977 | 0.38793 | 2.637 | 6 | 0.0035038 | 0.4625 | -0.2352
102 | 20.432 | 10.2306 | 10.2298 | 31.9709 | 31.9717 | 6.4063 | 2.875 | 18 | 20 | 396 | 6.2936 | 1.091 | 6.866 |  | 0.0035203 | 0.4625 | -0.2352
52 | 19.576 | 8.5936 | 8.5933 | 31.5746 | 31.5757 | 6.6591 | 2.2236 | 5 | 5 | 182 | 6.551 | 0.78965 | 5.173 |  | 0.0035743 | 0.4625 | -0.2352
67 | 100.805 | 7.6958 | 7.6961 | 33.4659 | 33.4651 | 6.7121 | 1.7509 | 15 | 15 | 260 | 6.6792 | 0.57102 | 3.814 |  | 0.0035822 | 0.4625 | -0.2352
44 | 50.514 | 8.4414 | 8.4415 | 32.5825 | 32.583 | 6.6383 | 2.5605 | 7 | 7 | 151 | 6.5575 | 0.95265 | 6.247 |  | 0.0036132 | 0.4625 | -0.2352
106 | 75.57 | 7.6725 | 7.6725 | 32.965 | 32.9657 | 6.7377 | 2.2005 | 12 | 12 | 413 | 6.6832 | 0.78614 | 5.254 |  | 0.0036369 | 0.4625 | -0.2352
30 | 400.611 | 5.6252 | 5.6258 | 34.0594 | 34.0581 | 7.0134 | 0.82051 | 5 | 5 | 87 | 7.2779 | 0.14056 | 1.023 |  | 0.0037241 | 0.4625 | -0.2352
28 | 4.332 | 11.5269 | 11.5264 | 31.5486 | 31.5484 | 6.2469 | 2.8306 | 15 | 17 | 80 | 6.1193 | 1.0702 | 6.549 |  | 0.0037556 | 0.4625 | -0.2352
88 | 20.145 | 9.9696 | 9.9715 | 31.4447 | 31.4441 | 6.4646 | 2.669 | 16 | 18 | 350 | 6.3519 | 1.003 | 6.371 |  | 0.0037888 | 0.4625 | -0.2352
40 | 150.256 | 7.2658 | 7.2655 | 33.8237 | 33.8242 | 6.7622 | 1.4871 | 8 | 8 | 134 | 6.7759 | 0.44879 | 3.041 | 6 | 0.0037993 | 0.4625 | -0.2352
28 | 30.103 | 8.9932 | 8.9931 | 32.3556 | 32.3555 | 6.5667 | 2.7754 | 12 | 14 | 77 | 6.4661 | 1.0523 | 6.804 |  | 0.0038302 | 0.4625 | -0.2352
92 | 300.977 | 5.6492 | 5.6495 | 33.9318 | 33.9321 | 7.0154 | 1.1518 | 8 | 8 | 363 | 7.1866 | 0.2936 | 2.11 | 36 | 0.0038878 | 0.4625 | -0.2352
126 | 5.405 | 11.5753 | 11.5731 | 30.4181 | 30.4217 | 6.2852 | 3.0342 | 11 | 13 | 451 | 6.1576 | 1.1721 | 7.217 |  | 0.0039169 | 0.4625 | -0.2352
40 | 10.789 | 11.6493 | 11.6505 | 31.3907 | 31.3905 | 6.2369 | 2.8185 | 15 | 17 | 141 | 6.1141 | 1.0644 | 6.508 |  | 0.0039359 | 0.4625 | -0.2352
88 | 31.405 | 8.7947 | 8.7924 | 32.0745 | 32.0772 | 6.6078 | 2.3777 | 15 | 16 | 349 | 6.5089 | 0.86051 | 5.601 |  | 0.0039618 | 0.4625 | -0.2352
37 | 75.229 | 8.3835 | 8.3833 | 32.9214 | 32.925 | 6.6323 | 2.1523 | 14 | 14 | 120 | 6.5727 | 0.76422 | 5.023 |  | 0.0039631 | 0.4625 | -0.2352
102 | 30.133 | 9.3127 | 9.3135 | 32.2026 | 32.2031 | 6.5268 | 2.8486 | 17 | 19 | 395 | 6.4248 | 1.0783 | 6.928 |  | 0.0039636 | 0.4625 | -0.2352
40 | 124.746 | 7.603 | 7.6054 | 33.7149 | 33.7158 | 6.7153 | 1.5744 | 9 | 9 | 135 | 6.7038 | 0.48898 | 3.278 |  | 0.0039829 | 0.4625 | -0.2352
88 | 4.746 | 10.9157 | 10.9157 | 30.8207 | 30.8212 | 6.3585 | 2.858 | 18 | 20 | 352 | 6.2312 | 1.0826 | 6.746 |  | 0.0040112 | 0.4625 | -0.2352
84 | 9.999 | 10.7636 | 10.7632 | 30.9101 | 30.9109 | 6.3757 | 2.85 | 9 | 11 | 330 | 6.2529 | 1.0789 | 6.746 |  | 0.0040688 | 0.4625 | -0.2352
28 | 175.453 | 7.144 | 7.1447 | 33.8836 | 33.8834 | 6.7784 | 1.4284 | 6 | 6 | 71 | 6.8154 | 0.42125 | 2.871 | 6 | 0.0041732 | 0.4625 | -0.2352
88 | 750.5 | 4.2579 | 4.2577 | 34.2465 | 34.2473 | 7.2364 | 0.61555 | 3 | 3 | 337 | 7.8813 | 0.045297 | 0.357 |  | 0.0041949 | 0.4625 | -0.2352
37 | 199.08 | 7.0578 | 7.0569 | 33.922 | 33.9223 | 6.7901 | 1.253 | 9 | 9 | 115 | 6.8488 | 0.34006 | 2.329 |  | 0.0042412 | 0.4625 | -0.2352
30 | 125.487 | 7.9765 | 7.9754 | 33.6614 | 33.66 | 6.6611 | 1.569 | 12 | 12 | 94 | 6.6471 | 0.48623 | 3.232 |  | 0.0042451 | 0.4625 | -0.2352
102 | 175.186 | 6.7367 | 6.7372 | 33.8164 | 33.817 | 6.8452 | 1.7366 | 11 | 11 | 389 | 6.8865 | 0.56371 | 3.882 |  | 0.0042475 | 0.4625 | -0.2352
44 | 4.313 | 11.5368 | 11.5363 | 31.5174 | 31.5181 | 6.2468 | 2.8432 | 11 | 11 | 155 | 6.1192 | 1.0755 | 6.581 |  | 0.0043179 | 0.4625 | -0.2352
30 | 748.912 | 4.4111 | 4.4111 | 34.2842 | 34.2832 | 7.208 | 0.60088 | 3 | 3 | 85 | 7.8457 | 0.038365 | 0.301 |  | 0.0043421 | 0.4625 | -0.2352
92 | 74.615 | 7.3185 | 7.3186 | 32.6472 | 32.6476 | 6.8064 | 2.6808 | 15 | 15 | 370 | 6.7538 | 1.0003 | 6.756 |  | 0.0043561 | 0.4625 | -0.2352
92 | 31.241 | 9.2571 | 9.2574 | 32.5495 | 32.5499 | 6.5202 | 2.8786 | 17 | 19 | 372 | 6.4196 | 1.0918 | 7.009 |  | 0.0043647 | 0.4625 | -0.2352
47 | 30.465 | 8.5848 | 8.5804 | 32.4018 | 32.4059 | 6.6248 | 2.4357 | 4 | 4 | 159 | 6.5263 | 0.88672 | 5.787 | 6 | 0.0045907 | 0.4625 | -0.2352
40 | 29.229 | 9.3081 | 9.3043 | 32.3846 | 32.3855 | 6.5198 | 2.8484 | 13 | 15 | 139 | 6.4173 | 1.0776 | 6.915 |  | 0.0046043 | 0.4625 | -0.2352
23 | 75.047 | 7.8692 | 7.8692 | 32.7091 | 32.7094 | 6.7189 | 1.764 | 4 | 4 | 39 | 6.6625 | 0.57606 | 3.838 | 36 | 0.0046043 | 0.4625 | -0.2352
67 | 74.763 | 7.9951 | 7.9947 | 32.7929 | 32.7932 | 6.6961 | 2.284 | 16 | 16 | 261 | 6.6386 | 0.81643 | 5.42 |  | 0.0047035 | 0.4625 | -0.2352
28 | 248.785 | 6.4099 | 6.405 | 33.9694 | 33.9701 | 6.8901 | 1.1132 | 4 | 4 | 69 | 7.0014 | 0.27494 | 1.925 |  | 0.0047111 | 0.4625 | -0.2352
40 | 499.532 | 5.3069 | 5.3068 | 34.1064 | 34.1072 | 7.064 | 0.73761 | 2 | 2 | 128 | 7.4296 | 0.10122 | 0.752 |  | 0.0047276 | 0.4625 | -0.2352
92 | 175.546 | 6.7091 | 6.7105 | 33.8553 | 33.8555 | 6.8478 | 1.6192 | 11 | 11 | 366 | 6.8897 | 0.50887 | 3.506 |  | 0.0048031 | 0.4625 | -0.2352
106 | 99.79 | 7.5339 | 7.5342 | 33.3709 | 33.3713 | 6.741 | 1.8798 | 11 | 11 | 412 | 6.7086 | 0.62934 | 4.222 |  | 0.0048564 | 0.4625 | -0.2352
52 | 5.65 | 11.202 | 11.2021 | 30.7666 | 30.7669 | 6.3216 | 3.2472 | 7 | 7 | 184 | 6.1947 | 1.2617 | 7.816 |  | 0.0049177 | 0.4625 | -0.2352
58 | 9.468 | 11.3802 | 11.3871 | 30.939 | 30.938 | 6.2907 | 3.1642 | 5 | 7 | 211 | 6.1667 | 1.2233 | 7.544 |  | 0.0049191 | 0.4625 | -0.2352
54 | 5.065 | 11.3928 | 11.4027 | 30.7533 | 30.7505 | 6.2964 | 3.0965 | 4 | 6 | 198 | 6.1689 | 1.192 | 7.353 |  | 0.0049757 | 0.4625 | -0.2352
75 | 101.103 | 7.657 | 7.6556 | 33.2684 | 33.2724 | 6.7267 | 1.9014 | 2 | 2 | 294 | 6.6944 | 0.6492 | 4.346 | 6 | 0.0049866 | 0.4625 | -0.2352
58 | 5.407 | 11.7781 | 11.7744 | 30.8823 | 30.8836 | 6.2398 | 3.0724 | 6 | 8 | 212 | 6.1125 | 1.1807 | 7.217 |  | 0.0050829 | 0.4625 | -0.2352
5 | 50.724 | 7.8742 | 7.876 | 32.4692 | 32.4637 | 6.7287 | 1.7935 | 3 | 3 | 23 | 6.6514 | 0.5892 | 3.919 |  | 0.0050885 | 0.4625 | -0.2352
73 | 19.885 | 11.015 | 11.024 | 32.0962 | 32.0963 | 6.2937 | 2.8627 | 15 | 17 | 287 | 6.1792 | 1.0835 | 6.695 |  | 0.0053121 | 0.4625 | -0.2352
49 | 20.593 | 8.7112 | 8.7113 | 31.7274 | 31.7283 | 6.635 | 2.2062 | 7 | 8 | 172 | 6.5273 | 0.7798 | 5.09 |  | 0.0053599 | 0.4625 | -0.2352
75 | 4.923 | 11.2668 | 11.256 | 31.2494 | 31.2524 | 6.2936 | 2.8717 | 8 | 10 | 300 | 6.1664 | 1.0984 | 6.773 |  | 0.0054175 | 0.4625 | -0.2352
106 | 301.028 | 5.8643 | 5.8624 | 33.9801 | 33.9812 | 6.9778 | 1.0029 | 5 | 5 | 406 | 7.1453 | 0.22322 | 1.595 |  | 0.0054268 | 0.4625 | -0.2352
28 | 100.593 | 7.837 | 7.8378 | 33.6106 | 33.6094 | 6.6844 | 1.6658 | 9 | 9 | 74 | 6.6502 | 0.52961 | 3.522 |  | 0.0056225 | 0.4625 | -0.2352
92 | 150.263 | 7.0517 | 7.0595 | 33.7916 | 33.79 | 6.7969 | 1.733 | 12 | 12 | 367 | 6.8128 | 0.56056 | 3.819 |  | 0.0057343 | 0.4625 | -0.2352
24 | 51.207 | 8.1158 | 8.1164 | 32.6638 | 32.6626 | 6.6835 | 2.2577 | 4 | 4 | 50 | 6.6052 | 0.80315 | 5.305 |  | 0.0058394 | 0.4625 | -0.2352
40 | 400.762 | 5.4764 | 5.4767 | 34.0834 | 34.0839 | 7.0369 | 0.77523 | 3 | 3 | 129 | 7.3046 | 0.11746 | 0.858 |  | 0.0058839 | 0.4625 | -0.2352
67 | 5.329 | 11.4879 | 11.4907 | 31.867 | 31.8665 | 6.2395 | 2.8445 | 21 | 23 | 266 | 6.113 | 1.0744 | 6.568 |  | 0.0059248 | 0.4625 | -0.2352
92 | 200.896 | 6.3983 | 6.4002 | 33.8713 | 33.8716 | 6.8965 | 1.5342 | 10 | 10 | 365 | 6.9649 | 0.46835 | 3.262 |  | 0.0060106 | 0.4625 | -0.2352
119 | 50.605 | 7.8088 | 7.8092 | 31.9507 | 31.9513 | 6.7615 | 2.3126 | 2 | 2 | 433 | 6.6843 | 0.82836 | 5.537 |  | 0.0060315 | 0.4625 | -0.2352
9 | 50.713 | 7.8866 | 7.8885 | 32.6546 | 32.6535 | 6.7187 | 1.7848 | 3 | 3 | 30 | 6.6414 | 0.58422 | 3.88 |  | 0.0060623 | 0.4625 | -0.2352
73 | 174.891 | 7.2555 | 7.2561 | 33.9038 | 33.9041 | 6.7603 | 1.2921 | 8 | 8 | 280 | 6.7955 | 0.35626 | 2.421 | 6 | 0.0061191 | 0.4625 | -0.2352
23 | 29.533 | 8.56 | 8.5615 | 31.6686 | 31.6687 | 6.66 | 2.2419 | 6 | 6 | 41 | 6.5605 | 0.79552 | 5.219 |  | 0.0061484 | 0.4625 | -0.2352
26 | 20.743 | 9.4315 | 9.4322 | 32.0912 | 32.0899 | 6.5144 | 2.7244 | 9 | 9 | 63 | 6.4043 | 1.031 | 6.603 |  | 0.0062127 | 0.4625 | -0.2352
84 | 98.947 | 7.5471 | 7.5453 | 33.2538 | 33.2567 | 6.7442 | 1.818 | 4 | 4 | 325 | 6.7109 | 0.61184 | 4.106 |  | 0.0062148 | 0.4625 | -0.2352
37 | 50.433 | 8.5466 | 8.5447 | 32.4949 | 32.4961 | 6.6264 | 2.5768 | 15 | 15 | 121 | 6.5449 | 0.95035 | 6.22 |  | 0.0062339 | 0.4625 | -0.2352
37 | 300.722 | 6.1178 | 6.1181 | 34.0003 | 34.0006 | 6.9357 | 0.96266 | 7 | 7 | 113 | 7.0986 | 0.2037 | 1.446 |  | 0.0063282 | 0.4625 | -0.2352
44 | 125.589 | 7.0689 | 7.0691 | 33.8024 | 33.803 | 6.7937 | 1.3485 | 4 | 4 | 148 | 6.7879 | 0.382 | 2.593 |  | 0.0064865 | 0.4625 | -0.2352
92 | 250.134 | 5.9489 | 5.949 | 33.8915 | 33.8918 | 6.9681 | 1.3655 | 9 | 9 | 364 | 7.0876 | 0.38984 | 2.763 |  | 0.0064928 | 0.4625 | -0.2352
47 | 85.382 | 7.5876 | 7.5947 | 33.3371 | 33.3303 | 6.7343 | 1.6317 | 1 | 1 | 156 | 6.689 | 0.51293 | 3.431 |  | 0.0065085 | 0.4625 | -0.2352
61 | 127.991 | 6.8968 | 6.897 | 33.8709 | 33.8721 | 6.8175 | 1.0594 | 1 | 1 | 216 | 6.8155 | 0.24826 | 1.692 |  | 0.0065248 | 0.4625 | -0.2352
9 | 75.837 | 7.741 | 7.7415 | 33.0155 | 33.0155 | 6.725 | 1.6779 | 1 | 1 | 28 | 6.6703 | 0.53431 | 3.564 |  | 0.0065281 | 0.4625 | -0.2352
28 | 200.547 | 6.8613 | 6.8611 | 33.9332 | 33.9332 | 6.8203 | 1.2944 | 5 | 5 | 70 | 6.8826 | 0.35684 | 2.456 |  | 0.0066217 | 0.4625 | -0.2352
78 | 50.488 | 8.3942 | 8.3944 | 32.4704 | 32.4717 | 6.6502 | 2.009 | 3 | 3 | 306 | 6.5695 | 0.70066 | 4.603 |  | 0.0066785 | 0.4625 | -0.2352
67 | 250.724 | 6.6385 | 6.6384 | 33.9833 | 33.9836 | 6.8532 | 1.0472 | 9 | 9 | 254 | 6.9629 | 0.24243 | 1.688 |  | 0.0066854 | 0.4625 | -0.2352
102 | 300.175 | 6.0238 | 6.0241 | 33.9585 | 33.959 | 6.9528 | 1.0839 | 8 | 8 | 386 | 7.1169 | 0.25938 | 1.846 |  | 0.0067013 | 0.4625 | -0.2352
63 | 9.797 | 11.2291 | 11.2442 | 31.8271 | 31.8213 | 6.2756 | 2.8485 | 15 | 17 | 241 | 6.1527 | 1.0889 | 6.7 |  | 0.0067135 | 0.4625 | -0.2352
88 | 49.044 | 8.2255 | 8.2248 | 32.4718 | 32.4729 | 6.6754 | 2.4362 | 14 | 14 | 348 | 6.5945 | 0.88483 | 5.835 | 6 | 0.0067264 | 0.4625 | -0.2352
55 | 30.406 | 8.3799 | 8.3803 | 32.2088 | 32.2099 | 6.6636 | 1.9753 | 2 | 2 | 203 | 6.566 | 0.67164 | 4.41 |  | 0.0067334 | 0.4625 | -0.2352
1 | 100.614 | 8.5346 | 8.5358 | 30.6067 | 30.6092 | 6.7099 | 0.92443 | 6 | 6 | 6 | 6.6701 | 0.1856 | 1.238 | 6 | 0.0067446 | 0.4625 | -0.2352
73 | 150.352 | 7.4485 | 7.4484 | 33.8543 | 33.8548 | 6.7328 | 1.3823 | 9 | 9 | 281 | 6.7448 | 0.39734 | 2.68 |  | 0.0067735 | 0.4625 | -0.2352
37 | 175.628 | 7.2637 | 7.2644 | 33.8846 | 33.8848 | 6.7599 | 1.3412 | 10 | 10 | 116 | 6.7957 | 0.37833 | 2.571 |  | 0.0067808 | 0.4625 | -0.2352
102 | 250.957 | 6.2669 | 6.267 | 33.9219 | 33.9223 | 6.9152 | 1.2429 | 9 | 9 | 387 | 7.0306 | 0.33283 | 2.34 |  | 0.0068109 | 0.4625 | -0.2352
106 | 250.274 | 6.2298 | 6.2301 | 33.926 | 33.9268 | 6.921 | 1.2202 | 6 | 6 | 407 | 7.0363 | 0.32233 | 2.268 |  | 0.0068143 | 0.4625 | -0.2352
73 | 125.667 | 7.7501 | 7.7506 | 33.7785 | 33.7789 | 6.6902 | 1.4652 | 10 | 10 | 282 | 6.6782 | 0.4356 | 2.909 |  | 0.0068806 | 0.4625 | -0.2352
119 | 5.264 | 10.0913 | 10.0877 | 30.6071 | 30.6083 | 6.4821 | 3.2847 | 6 | 8 | 437 | 6.3565 | 1.291 | 8.206 |  | 0.0069986 | 0.4625 | -0.2352
88 | 397.129 | 5.4963 | 5.4964 | 34.053 | 34.0535 | 7.035 | 0.81813 | 5 | 5 | 339 | 7.2989 | 0.13618 | 0.994 | 6 | 0.0070011 | 0.4625 | -0.2352
52 | 30.812 | 8.321 | 8.3215 | 31.818 | 31.8185 | 6.6894 | 2.0543 | 4 | 4 | 181 | 6.5922 | 0.7078 | 4.666 |  | 0.0071127 | 0.4625 | -0.2352
49 | 156.959 | 7.7248 | 7.7244 | 33.0494 | 33.0519 | 6.726 | 1.5035 | 1 | 1 | 166 | 6.7411 | 0.45304 | 3.054 |  | 0.0071346 | 0.4625 | -0.2352
67 | 149.21 | 7.3059 | 7.3052 | 33.8152 | 33.8157 | 6.7564 | 1.4931 | 13 | 13 | 258 | 6.7688 | 0.44808 | 3.033 |  | 0.0072611 | 0.4625 | -0.2352
75 | 9.575 | 10.9299 | 10.9495 | 31.3587 | 31.35 | 6.3348 | 2.8747 | 7 | 9 | 299 | 6.2118 | 1.1016 | 6.843 |  | 0.0072674 | 0.4625 | -0.2352
63 | 251.208 | 6.3442 | 6.345 | 33.9916 | 33.992 | 6.8996 | 0.99976 | 5 | 5 | 231 | 7.014 | 0.21984 | 1.542 |  | 0.0073442 | 0.4625 | -0.2352
61 | 100.578 | 7.2842 | 7.2876 | 33.7058 | 33.7043 | 6.7646 | 1.3507 | 2 | 2 | 217 | 6.7351 | 0.38203 | 2.573 |  | 0.0074917 | 0.4625 | -0.2352
9 | 75.983 | 7.7415 | 7.742 | 33.0167 | 33.0155 | 6.7249 | 1.68 | 2 | 2 | 29 | 6.6703 | 0.53431 | 3.564 |  | 0.0075002 | 0.4625 | -0.2352
84 | 125.632 | 7.135 | 7.1345 | 33.6517 | 33.6524 | 6.7902 | 1.5222 | 3 | 3 | 324 | 6.7837 | 0.46125 | 3.129 |  | 0.0075618 | 0.4625 | -0.2352
53 | 29.928 | 8.2629 | 8.2629 | 31.9575 | 31.9582 | 6.6921 | 1.9645 | 4 | 4 | 188 | 6.5946 | 0.6657 | 4.39 |  | 0.0076934 | 0.4625 | -0.2352
24 | 75.865 | 7.7505 | 7.7555 | 33.1345 | 33.128 | 6.7183 | 1.6927 | 3 | 3 | 49 | 6.6636 | 0.5554 | 3.701 |  | 0.0077483 | 0.4625 | -0.2352
37 | 123.931 | 7.7419 | 7.7414 | 33.698 | 33.6979 | 6.6949 | 1.5628 | 12 | 12 | 118 | 6.6815 | 0.47983 | 3.206 |  | 0.0077826 | 0.4625 | -0.2352
73 | 251.678 | 6.4981 | 6.4984 | 33.982 | 33.9826 | 6.8755 | 1.0396 | 6 | 6 | 278 | 6.9881 | 0.23783 | 1.662 |  | 0.0077872 | 0.4625 | -0.2352
28 | 506.458 | 5.1322 | 5.1316 | 34.1337 | 34.1345 | 7.092 | 0.71003 | 1 | 1 | 66 | 7.4685 | 0.085292 | 0.637 |  | 0.0078969 | 0.4625 | -0.2352
63 | 29.486 | 8.9537 | 8.955 | 32.5106 | 32.5102 | 6.5659 | 2.8102 | 13 | 15 | 239 | 6.465 | 1.0566 | 6.831 |  | 0.0079019 | 0.4625 | -0.2352
106 | 125.632 | 7.3811 | 7.3817 | 33.6549 | 33.6551 | 6.7519 | 1.6762 | 10 | 10 | 411 | 6.7432 | 0.53209 | 3.588 |  | 0.0079536 | 0.4625 | -0.2352
67 | 1559.725 | 2.4657 | 2.4644 | 34.5312 | 34.5329 | 7.5468 | 0.71981 | 1 | 1 | 246 | 9.1812 | 0.089749 | 0.824 |  | 0.0079634 | 0.4625 | -0.2352
58 | 30.025 | 8.3022 | 8.3058 | 32.0682 | 32.0677 | 6.6814 | 2.0253 | 3 | 3 | 209 | 6.5838 | 0.69352 | 4.566 |  | 0.0079662 | 0.4625 | -0.2352
63 | 100.149 | 7.7651 | 7.7677 | 33.6945 | 33.6933 | 6.6916 | 1.5548 | 10 | 10 | 236 | 6.6577 | 0.49192 | 3.275 |  | 0.008034 | 0.4625 | -0.2352
58 | 50.316 | 8.1542 | 8.1543 | 32.5052 | 32.5063 | 6.6847 | 1.8736 | 2 | 2 | 208 | 6.6053 | 0.62329 | 4.117 |  | 0.0080464 | 0.4625 | -0.2352
106 | 30.589 | 8.726 | 8.7331 | 32.2125 | 32.2099 | 6.612 | 2.7197 | 14 | 15 | 415 | 6.5129 | 1.0146 | 6.608 |  | 0.0080473 | 0.4625 | -0.2352
119 | 57.425 | 7.8041 | 7.8042 | 31.9983 | 31.9998 | 6.7601 | 2.3255 | 1 | 1 | 432 | 6.6888 | 0.83229 | 5.567 |  | 0.0080582 | 0.4625 | -0.2352
106 | 175.648 | 6.9469 | 6.9471 | 33.8789 | 33.8797 | 6.8093 | 1.3766 | 8 | 8 | 409 | 6.8486 | 0.39336 | 2.694 |  | 0.0081189 | 0.4625 | -0.2352
92 | 50.408 | 8.1829 | 8.1954 | 32.6073 | 32.607 | 6.6759 | 2.8684 | 16 | 17 | 371 | 6.5965 | 1.0833 | 7.146 | 6 | 0.0081225 | 0.4625 | -0.2352
88 | 299.545 | 5.8022 | 5.8023 | 33.9985 | 33.999 | 6.9871 | 0.9592 | 6 | 6 | 340 | 7.1543 | 0.2003 | 1.433 |  | 0.0081305 | 0.4625 | -0.2352
37 | 250.343 | 6.6056 | 6.6059 | 33.9535 | 33.9539 | 6.8597 | 1.1451 | 8 | 8 | 114 | 6.9696 | 0.28624 | 1.995 |  | 0.0081489 | 0.4625 | -0.2352
63 | 50.395 | 8.3713 | 8.3794 | 32.5772 | 32.5748 | 6.649 | 2.675 | 12 | 12 | 238 | 6.5684 | 1.0101 | 6.635 |  | 0.0081549 | 0.4625 | -0.2352
63 | 151.42 | 7.3333 | 7.3337 | 33.9001 | 33.9004 | 6.7485 | 1.2744 | 8 | 8 | 234 | 6.7625 | 0.34603 | 2.34 |  | 0.0081695 | 0.4625 | -0.2352
63 | 300.898 | 6.0001 | 6.0005 | 34.017 | 34.0175 | 6.954 | 0.93725 | 4 | 4 | 230 | 7.119 | 0.19005 | 1.353 |  | 0.0082235 | 0.4625 | -0.2352
102 | 400.783 | 5.1308 | 5.131 | 33.9986 | 33.9992 | 7.0987 | 0.89609 | 7 | 7 | 385 | 7.374 | 0.17101 | 1.261 |  | 0.0082352 | 0.4625 | -0.2352
67 | 175.056 | 7.2389 | 7.2373 | 33.9153 | 33.9156 | 6.7623 | 1.2516 | 12 | 12 | 257 | 6.7979 | 0.3354 | 2.28 |  | 0.0082814 | 0.4625 | -0.2352
63 | 174.939 | 7.055 | 7.0543 | 33.9336 | 33.9342 | 6.79 | 1.2203 | 7 | 7 | 233 | 6.8275 | 0.32091 | 2.191 |  | 0.0082899 | 0.4625 | -0.2352
44 | 100.498 | 7.6799 | 7.6795 | 33.4928 | 33.4936 | 6.7133 | 1.773 | 5 | 5 | 149 | 6.6804 | 0.57647 | 3.851 |  | 0.0083227 | 0.4625 | -0.2352
73 | 301.065 | 6.179 | 6.1794 | 34.0099 | 34.0107 | 6.9253 | 0.94637 | 5 | 5 | 277 | 7.0876 | 0.19414 | 1.376 |  | 0.0083535 | 0.4625 | -0.2352
106 | 200.647 | 6.6106 | 6.6102 | 33.9006 | 33.9014 | 6.8613 | 1.346 | 7 | 7 | 408 | 6.9268 | 0.37896 | 2.625 |  | 0.0083757 | 0.4625 | -0.2352
30 | 199.628 | 7.0976 | 7.0978 | 33.9246 | 33.9233 | 6.7838 | 1.2906 | 8 | 8 | 90 | 6.8425 | 0.35323 | 2.417 |  | 0.0084768 | 0.4625 | -0.2352
49 | 49.272 | 8.3431 | 8.3427 | 32.2438 | 32.2451 | 6.6676 | 1.9626 | 5 | 5 | 170 | 6.5861 | 0.66397 | 4.373 |  | 0.0085212 | 0.4625 | -0.2352
44 | 30.254 | 9.2166 | 9.2034 | 32.2788 | 32.2842 | 6.5375 | 2.7423 | 8 | 8 | 152 | 6.436 | 1.0416 | 6.704 |  | 0.0085312 | 0.4625 | -0.2352
79 | 30.803 | 8.5833 | 8.5836 | 32.1229 | 32.1235 | 6.637 | 2.0223 | 2 | 2 | 315 | 6.5387 | 0.69158 | 4.522 |  | 0.0085533 | 0.4625 | -0.2352
102 | 498.962 | 4.7973 | 4.7976 | 34.0653 | 34.0659 | 7.1521 | 0.76451 | 6 | 6 | 384 | 7.5299 | 0.10983 | 0.827 | 6 | 0.0085572 | 0.4625 | -0.2352
84 | 150.306 | 6.7911 | 6.7912 | 33.8047 | 33.8056 | 6.8371 | 1.3972 | 2 | 2 | 323 | 6.8558 | 0.40243 | 2.759 |  | 0.008561 | 0.4625 | -0.2352
26 | 174.598 | 6.8759 | 6.8737 | 33.9117 | 33.9123 | 6.819 | 1.1964 | 2 | 2 | 56 | 6.8582 | 0.30956 | 2.123 |  | 0.0085723 | 0.4625 | -0.2352
109 | 10.167 | 10.532 | 10.533 | 31.842 | 31.8424 | 6.3698 | 2.9419 | 6 | 8 | 427 | 6.2482 | 1.1168 | 6.978 |  | 0.0086246 | 0.4625 | -0.2352
88 | 174.998 | 6.6192 | 6.618 | 33.8929 | 33.8935 | 6.8603 | 1.4368 | 9 | 9 | 343 | 6.9028 | 0.43794 | 3.023 |  | 0.0086251 | 0.4625 | -0.2352
24 | 29.859 | 8.6781 | 8.6787 | 32.3282 | 32.3279 | 6.6141 | 2.5773 | 5 | 5 | 51 | 6.5147 | 0.94801 | 6.176 |  | 0.0087763 | 0.4625 | -0.2352
30 | 1000.117 | 3.7562 | 3.7557 | 34.3712 | 34.3706 | 7.3187 | 0.60115 | 2 | 2 | 84 | 8.2445 | 0.033962 | 0.28 |  | 0.0088699 | 0.4625 | -0.2352
102 | 999.625 | 3.5277 | 3.5276 | 34.3754 | 34.3763 | 7.3594 | 0.60372 | 4 | 4 | 382 | 8.2951 | 0.035081 | 0.291 |  | 0.0089397 | 0.4625 | -0.2352
84 | 164.165 | 6.7782 | 6.7781 | 33.8015 | 33.8024 | 6.8393 | 1.3962 | 1 | 1 | 322 | 6.8704 | 0.40158 | 2.759 |  | 0.0089789 | 0.4625 | -0.2352
109 | 20.395 | 9.0108 | 9.0094 | 31.9278 | 31.9285 | 6.5823 | 2.7632 | 5 | 6 | 426 | 6.4733 | 1.0338 | 6.692 |  | 0.0090005 | 0.4625 | -0.2352
49 | 74.684 | 8.1084 | 8.1086 | 32.7087 | 32.7093 | 6.6827 | 1.8163 | 4 | 4 | 169 | 6.6243 | 0.59583 | 3.947 |  | 0.0090053 | 0.4625 | -0.2352
73 | 200.018 | 7.0296 | 7.033 | 33.9365 | 33.9366 | 6.7939 | 1.213 | 7 | 7 | 279 | 6.8536 | 0.31677 | 2.171 |  | 0.0090654 | 0.4625 | -0.2352
40 | 299.738 | 5.9892 | 5.9894 | 34.0131 | 34.0136 | 6.9559 | 0.94904 | 4 | 4 | 130 | 7.1201 | 0.19466 | 1.386 |  | 0.0090702 | 0.4625 | -0.2352
23 | 99.594 | 7.3506 | 7.353 | 33.4081 | 33.4065 | 6.7676 | 1.4186 | 3 | 3 | 38 | 6.7365 | 0.41178 | 2.774 |  | 0.0091093 | 0.4625 | -0.2352
30 | 100.358 | 8.4116 | 8.4122 | 33.4005 | 33.399 | 6.6075 | 1.7935 | 13 | 13 | 95 | 6.569 | 0.58517 | 3.844 |  | 0.0091409 | 0.4625 | -0.2352
37 | 400.978 | 5.6442 | 5.6458 | 34.071 | 34.0712 | 7.0097 | 0.80431 | 6 | 6 | 112 | 7.2742 | 0.12757 | 0.928 | 6 | 0.0092193 | 0.4625 | -0.2352
30 | 73.988 | 8.2025 | 8.2067 | 32.6943 | 32.6932 | 6.6692 | 2.4358 | 14 | 14 | 96 | 6.6096 | 0.88205 | 5.83 |  | 0.0093305 | 0.4625 | -0.2352
40 | 199.812 | 6.7605 | 6.761 | 33.9427 | 33.943 | 6.8357 | 1.1763 | 6 | 6 | 132 | 6.8986 | 0.29948 | 2.066 |  | 0.0093603 | 0.4625 | -0.2352
92 | 399.739 | 5.1255 | 5.1257 | 34.003 | 34.0035 | 7.0994 | 0.87688 | 7 | 7 | 362 | 7.3738 | 0.16098 | 1.187 |  | 0.0093813 | 0.4625 | -0.2352
23 | 50.358 | 8.1543 | 8.154 | 32.0573 | 32.0586 | 6.7042 | 1.9424 | 5 | 5 | 40 | 6.6246 | 0.65378 | 4.331 |  | 0.0094028 | 0.4625 | -0.2352
102 | 2005.31 | 1.9608 | 1.9602 | 34.584 | 34.5858 | 7.64 | 0.82471 | 1 | 1 | 379 | 9.8702 | 0.13678 | 1.35 |  | 0.009453 | 0.4625 | -0.2352
44 | 10.569 | 11.497 | 11.491 | 31.5761 | 31.5773 | 6.2498 | 2.8647 | 10 | 10 | 154 | 6.1271 | 1.0803 | 6.619 |  | 0.0094642 | 0.4625 | -0.2352
67 | 400.278 | 5.61 | 5.6099 | 34.0643 | 34.0648 | 7.0157 | 0.79662 | 7 | 7 | 252 | 7.2802 | 0.12376 | 0.901 |  | 0.0094766 | 0.4625 | -0.2352
26 | 5.84 | 11.5041 | 11.504 | 31.2934 | 31.2931 | 6.26 | 2.8618 | 11 | 11 | 65 | 6.1334 | 1.0788 | 6.617 |  | 0.0095604 | 0.4625 | -0.2352
5 | 10.097 | 8.6343 | 8.6367 | 31.4972 | 31.4957 | 6.6564 | 2.3245 | 6 | 6 | 26 | 6.5401 | 0.83027 | 5.43 |  | 0.009638 | 0.4625 | -0.2352
79 | 40.57 | 8.5675 | 8.5683 | 32.1738 | 32.1749 | 6.6371 | 2.0363 | 1 | 1 | 314 | 6.5471 | 0.69695 | 4.563 |  | 0.0096468 | 0.4625 | -0.2352
28 | 125.787 | 7.6165 | 7.6148 | 33.7487 | 33.7489 | 6.7118 | 1.5606 | 8 | 8 | 73 | 6.701 | 0.47694 | 3.196 |  | 0.0096544 | 0.4625 | -0.2352
58 | 19.755 | 8.7064 | 8.703 | 31.5241 | 31.5271 | 6.6445 | 2.2663 | 4 | 5 | 210 | 6.5359 | 0.82268 | 5.377 | 6 | 0.0097229 | 0.4625 | -0.2352
67 | 499.109 | 5.113 | 5.1133 | 34.127 | 34.1274 | 7.0956 | 0.70193 | 6 | 6 | 251 | 7.4654 | 0.079701 | 0.595 |  | 0.0097417 | 0.4625 | -0.2352
88 | 501.665 | 5.04 | 5.0396 | 34.092 | 34.0926 | 7.1096 | 0.74745 | 4 | 4 | 338 | 7.4838 | 0.10075 | 0.754 |  | 0.0097443 | 0.4625 | -0.2352
37 | 30.181 | 9.192 | 9.1889 | 32.392 | 32.3928 | 6.5362 | 2.8626 | 16 | 16 | 122 | 6.4349 | 1.0985 | 7.069 |  | 0.0097708 | 0.4625 | -0.2352
26 | 149.943 | 7.0828 | 7.0831 | 33.8688 | 33.8689 | 6.7886 | 1.2651 | 3 | 3 | 57 | 6.8039 | 0.3401 | 2.314 |  | 0.0097998 | 0.4625 | -0.2352
106 | 400.967 | 5.198 | 5.1982 | 34.031 | 34.0319 | 7.0858 | 0.83751 | 4 | 4 | 405 | 7.3598 | 0.14226 | 1.047 |  | 0.0098889 | 0.4625 | -0.2352
28 | 10.303 | 11.5098 | 11.5095 | 31.5477 | 31.5474 | 6.2492 | 2.8441 | 14 | 16 | 79 | 6.1263 | 1.0703 | 6.557 |  | 0.0099019 | 0.4625 | -0.2352
119 | 30.307 | 8.1736 | 8.1738 | 31.4655 | 31.4663 | 6.7271 | 2.5173 | 3 | 3 | 434 | 6.6301 | 0.91914 | 6.094 |  | 0.0099068 | 0.4625 | -0.2352
30 | 1193.511 | 3.3438 | 3.3431 | 34.4305 | 34.4301 | 7.3899 | 0.62484 | 1 | 1 | 83 | 8.5469 | 0.043876 | 0.375 |  | 0.0099129 | 0.4625 | -0.2352
55 | 34.226 | 8.3615 | 8.3588 | 32.2628 | 32.2665 | 6.664 | 1.9538 | 1 | 1 | 202 | 6.5698 | 0.65847 | 4.326 |  | 0.0099422 | 0.4625 | -0.2352
67 | 751.184 | 4.3734 | 4.3735 | 34.2664 | 34.267 | 7.2154 | 0.60105 | 5 | 5 | 250 | 7.8568 | 0.032838 | 0.258 |  | 0.0099479 | 0.4625 | -0.2352
88 | 99.959 | 7.305 | 7.3048 | 33.6071 | 33.6075 | 6.7658 | 1.8254 | 12 | 12 | 346 | 6.7355 | 0.59906 | 4.035 |  | 0.0099646 | 0.4625 | -0.2352
106 | 500.368 | 4.8697 | 4.8698 | 34.1098 | 34.1105 | 7.1376 | 0.71286 | 3 | 3 | 404 | 7.5148 | 0.0845 | 0.635 |  | 0.0099982 | 0.4625 | -0.2352
24 | 100.3 | 7.2536 | 7.2543 | 33.6152 | 33.6149 | 6.7734 | 1.2593 | 2 | 2 | 48 | 6.7438 | 0.3372 | 2.274 | 6 | 0.010029 | 0.4625 | -0.2352
37 | 499.92 | 5.0674 | 5.0679 | 34.1307 | 34.131 | 7.1031 | 0.69954 | 5 | 5 | 111 | 7.4748 | 0.078263 | 0.585 |  | 0.010075 | 0.4625 | -0.2352
28 | 398.637 | 5.6244 | 5.624 | 34.0677 | 34.0681 | 7.0131 | 0.8105 | 2 | 2 | 67 | 7.2758 | 0.12933 | 0.941 |  | 0.010324 | 0.4625 | -0.2352
55 | 10.243 | 11.1252 | 11.1261 | 30.8331 | 30.8334 | 6.3294 | 3.1487 | 4 | 4 | 205 | 6.2062 | 1.2107 | 7.514 |  | 0.010355 | 0.4625 | -0.2352
63 | 5.409 | 11.6173 | 11.6148 | 31.7626 | 31.7629 | 6.2265 | 2.863 | 16 | 18 | 242 | 6.0999 | 1.0785 | 6.579 |  | 0.010393 | 0.4625 | -0.2352
67 | 1500.318 | 2.6659 | 2.6658 | 34.5085 | 34.5098 | 7.5104 | 0.68508 | 2 | 2 | 247 | 9.0603 | 0.07119 | 0.645 |  | 0.010459 | 0.4625 | -0.2352
126 | 149.421 | 6.5278 | 6.528 | 33.7207 | 33.7213 | 6.8826 | 1.1934 | 3 | 3 | 443 | 6.9036 | 0.30622 | 2.114 | 6 | 0.010525 | 0.4625 | -0.2352
37 | 150.905 | 7.5178 | 7.5167 | 33.8233 | 33.824 | 6.7235 | 1.4328 | 11 | 11 | 117 | 6.7353 | 0.4169 | 2.808 |  | 0.010551 | 0.4625 | -0.2352
61 | 9.795 | 11.7215 | 11.7242 | 31.1883 | 31.1914 | 6.2353 | 2.8698 | 7 | 9 | 222 | 6.1116 | 1.0814 | 6.609 |  | 0.010673 | 0.4625 | -0.2352
75 | 135.406 | 6.9675 | 6.9678 | 33.7849 | 33.7859 | 6.8103 | 1.3462 | 1 | 1 | 293 | 6.8141 | 0.37672 | 2.567 |  | 0.010713 | 0.4625 | -0.2352
63 | 400.594 | 5.5181 | 5.5187 | 34.0787 | 34.0792 | 7.0302 | 0.7827 | 3 | 3 | 229 | 7.2969 | 0.11608 | 0.847 |  | 0.010722 | 0.4625 | -0.2352
54 | 20.673 | 11.1461 | 11.1464 | 30.9358 | 30.9367 | 6.3224 | 3.1266 | 2 | 2 | 196 | 6.2075 | 1.2 | 7.449 |  | 0.010869 | 0.4625 | -0.2352
52 | 50.167 | 8.0816 | 8.084 | 32.2584 | 32.2589 | 6.7064 | 1.8661 | 3 | 3 | 180 | 6.6272 | 0.617 | 4.089 |  | 0.010888 | 0.4625 | -0.2352
1 | 125.171 | 9.242 | 9.2416 | 31.0897 | 31.0903 | 6.5841 | 0.5542 | 5 | 5 | 5 | 6.5607 | 0.010212 | 0.067 |  | 0.010905 | 0.4625 | -0.2352
63 | 500.432 | 5.1785 | 5.1787 | 34.1223 | 34.1231 | 7.0848 | 0.71267 | 2 | 2 | 228 | 7.4543 | 0.083442 | 0.622 |  | 0.010968 | 0.4625 | -0.2352
40 | 5.309 | 11.6449 | 11.6452 | 31.3842 | 31.3845 | 6.2377 | 2.8303 | 16 | 18 | 142 | 6.1107 | 1.0626 | 6.493 |  | 0.011264 | 0.4625 | -0.2352
37 | 1253.515 | 3.125 | 3.1245 | 34.4606 | 34.4618 | 7.4281 | 0.63643 | 2 | 2 | 108 | 8.6643 | 0.047782 | 0.414 |  | 0.011367 | 0.4625 | -0.2352
58 | 52.363 | 8.1574 | 8.1579 | 32.5045 | 32.5061 | 6.6842 | 1.8797 | 1 | 1 | 207 | 6.6065 | 0.62272 | 4.114 |  | 0.011445 | 0.4625 | -0.2352
67 | 1001.104 | 3.7686 | 3.7684 | 34.3736 | 34.3745 | 7.3164 | 0.59731 | 4 | 4 | 249 | 8.2427 | 0.029602 | 0.244 | 6 | 0.011454 | 0.4625 | -0.2352
92 | 2003.823 | 1.9616 | 1.9609 | 34.5897 | 34.5913 | 7.6395 | 0.82934 | 1 | 1 | 356 | 9.8677 | 0.13681 | 1.35 |  | 0.011559 | 0.4625 | -0.2352
1 | 5.011 | 9.6559 | 9.6465 | 28.989 | 28.9939 | 6.6128 | 3.2103 | 12 | 12 | 12 | 6.4867 | 1.2612 | 8.181 |  | 0.011646 | 0.4625 | -0.2352
2 | 25.907 | 8.1223 | 8.1226 | 32.0473 | 32.0478 | 6.7095 | 1.9801 | 1 | 1 | 13 | 6.6094 | 0.6689 | 4.421 |  | 0.011701 | 0.4625 | -0.2352
73 | 499.927 | 5.1671 | 5.1671 | 34.1113 | 34.1119 | 7.0872 | 0.71886 | 3 | 3 | 275 | 7.4566 | 0.085562 | 0.638 |  | 0.01171 | 0.4625 | -0.2352
40 | 594.753 | 4.8427 | 4.8431 | 34.1759 | 34.1768 | 7.139 | 0.66168 | 1 | 1 | 127 | 7.6092 | 0.059007 | 0.449 |  | 0.01182 | 0.4625 | -0.2352
52 | 76.171 | 8.0735 | 8.074 | 32.7801 | 32.7808 | 6.6848 | 1.7594 | 2 | 2 | 179 | 6.628 | 0.56669 | 3.756 | 6 | 0.011823 | 0.4625 | -0.2352
5 | 30.552 | 8.1214 | 8.1212 | 32.0164 | 32.0173 | 6.7109 | 1.9698 | 4 | 4 | 24 | 6.6147 | 0.66397 | 4.392 | 6 | 0.011843 | 0.4625 | -0.2352
126 | 188.258 | 6.3379 | 6.3386 | 33.889 | 33.8898 | 6.9053 | 1.1137 | 1 | 1 | 441 | 6.9632 | 0.26798 | 1.866 |  | 0.011884 | 0.4625 | -0.2352
73 | 400.358 | 5.6085 | 5.6087 | 34.0628 | 34.0634 | 7.016 | 0.80182 | 4 | 4 | 276 | 7.2806 | 0.12375 | 0.901 |  | 0.011889 | 0.4625 | -0.2352
3 | 45.459 | 7.8868 | 7.8873 | 32.3914 | 32.3915 | 6.7302 | 1.8008 | 1 | 1 | 16 | 6.6483 | 0.58572 | 3.894 | 6 | 0.011968 | 0.4625 | -0.2352
9 | 9.724 | 10.0692 | 10.0604 | 30.7834 | 30.7852 | 6.4779 | 3.8458 | 6 | 7 | 33 | 6.3561 | 1.5555 | 9.887 |  | 0.012042 | 0.4625 | -0.2352
106 | 750.328 | 4.1848 | 4.1856 | 34.268 | 34.2689 | 7.2481 | 0.60197 | 2 | 2 | 403 | 7.8954 | 0.031158 | 0.246 |  | 0.012054 | 0.4625 | -0.2352
37 | 1441.923 | 2.7827 | 2.7823 | 34.4993 | 34.5007 | 7.4892 | 0.67505 | 1 | 1 | 107 | 8.9626 | 0.064937 | 0.582 |  | 0.012074 | 0.4625 | -0.2352
54 | 9.916 | 11.2999 | 11.3006 | 30.8049 | 30.8047 | 6.3068 | 3.1089 | 3 | 4 | 197 | 6.1832 | 1.1905 | 7.361 | 6 | 0.012185 | 0.4625 | -0.2352
47 | 10.261 | 11.4606 | 11.4624 | 30.8179 | 30.8165 | 6.2847 | 3.055 | 6 | 8 | 161 | 6.1612 | 1.1654 | 7.18 |  | 0.012393 | 0.4625 | -0.2352
126 | 175.04 | 6.3498 | 6.35 | 33.887 | 33.8877 | 6.9035 | 1.11 | 2 | 2 | 442 | 6.9494 | 0.26578 | 1.847 |  | 0.012406 | 0.4625 | -0.2352
49 | 125.593 | 7.9453 | 7.9444 | 32.9043 | 32.9054 | 6.6988 | 1.6426 | 2 | 2 | 167 | 6.6851 | 0.51204 | 3.423 |  | 0.012462 | 0.4625 | -0.2352
78 | 19.962 | 10.5594 | 10.5572 | 31.9156 | 31.9177 | 6.3631 | 2.8678 | 5 | 7 | 308 | 6.2492 | 1.1037 | 6.897 |  | 0.012489 | 0.4625 | -0.2352
24 | 5.669 | 11.8018 | 11.8014 | 31.2308 | 31.2306 | 6.223 | 2.901 | 8 | 8 | 54 | 6.0961 | 1.094 | 6.669 |  | 0.012531 | 0.4625 | -0.2352
37 | 750.234 | 4.3161 | 4.3165 | 34.2784 | 34.2789 | 7.2247 | 0.59724 | 4 | 4 | 110 | 7.8672 | 0.028473 | 0.224 |  | 0.012551 | 0.4625 | -0.2352
53 | 10.115 | 10.4348 | 10.4331 | 30.9093 | 30.9132 | 6.4213 | 3.4527 | 6 | 8 | 190 | 6.2991 | 1.3491 | 8.498 | 6 | 0.012613 | 0.4625 | -0.2352
49 | 99.334 | 8.0001 | 8.0003 | 32.8526 | 32.8531 | 6.6928 | 1.6988 | 3 | 3 | 168 | 6.6562 | 0.53785 | 3.58 |  | 0.012635 | 0.4625 | -0.2352
44 | 20.533 | 10.6453 | 10.6168 | 31.9066 | 31.912 | 6.3516 | 2.9095 | 9 | 9 | 153 | 6.2381 | 1.0978 | 6.848 |  | 0.012656 | 0.4625 | -0.2352
102 | 1499.5 | 2.5741 | 2.5737 | 34.4933 | 34.4946 | 7.5284 | 0.67421 | 2 | 2 | 380 | 9.0837 | 0.063961 | 0.581 |  | 0.012662 | 0.4625 | -0.2352
47 | 75.568 | 7.6985 | 7.6987 | 33.201 | 33.2017 | 6.7233 | 1.674 | 2 | 2 | 157 | 6.6688 | 0.52633 | 3.51 |  | 0.012673 | 0.4625 | -0.2352
102 | 749.989 | 4.0506 | 4.0507 | 34.2708 | 34.2715 | 7.2716 | 0.60798 | 5 | 5 | 383 | 7.9232 | 0.033194 | 0.263 |  | 0.012797 | 0.4625 | -0.2352
92 | 499.646 | 4.793 | 4.7929 | 34.0893 | 34.0899 | 7.1517 | 0.71748 | 6 | 6 | 361 | 7.5302 | 0.083663 | 0.63 |  | 0.012971 | 0.4625 | -0.2352
88 | 1270.702 | 3.0449 | 3.0445 | 34.4429 | 34.4442 | 7.4437 | 0.64818 | 1 | 1 | 335 | 8.7041 | 0.051585 | 0.449 |  | 0.012999 | 0.4625 | -0.2352
3 | 30.087 | 8.1737 | 8.1741 | 32.0277 | 32.0275 | 6.7025 | 1.9824 | 2 | 2 | 17 | 6.6057 | 0.66867 | 4.417 |  | 0.013005 | 0.4625 | -0.2352
26 | 9.872 | 11.5012 | 11.5015 | 31.2998 | 31.2985 | 6.2602 | 2.8709 | 10 | 10 | 64 | 6.1367 | 1.0796 | 6.625 |  | 0.013032 | 0.4625 | -0.2352
88 | 1000.311 | 3.6439 | 3.6444 | 34.3504 | 34.3512 | 7.3398 | 0.61341 | 2 | 2 | 336 | 8.2711 | 0.035425 | 0.293 |  | 0.013077 | 0.4625 | -0.2352
53 | 74.917 | 7.9844 | 7.984 | 32.5944 | 32.5947 | 6.7064 | 1.7413 | 2 | 2 | 186 | 6.6491 | 0.55707 | 3.704 |  | 0.013081 | 0.4625 | -0.2352
5 | 75.224 | 7.6483 | 7.648 | 33.0106 | 33.01 | 6.7394 | 1.6294 | 2 | 2 | 22 | 6.6848 | 0.50532 | 3.378 |  | 0.013089 | 0.4625 | -0.2352
92 | 1000.129 | 3.6627 | 3.6634 | 34.3706 | 34.3712 | 7.3354 | 0.60187 | 4 | 4 | 359 | 8.2655 | 0.030004 | 0.248 |  | 0.013161 | 0.4625 | -0.2352
37 | 1001.384 | 3.697 | 3.697 | 34.3819 | 34.3826 | 7.3288 | 0.60171 | 3 | 3 | 109 | 8.2585 | 0.029908 | 0.247 |  | 0.013182 | 0.4625 | -0.2352
92 | 1251.247 | 3.0872 | 3.0871 | 34.449 | 34.45 | 7.4356 | 0.62753 | 3 | 3 | 358 | 8.6715 | 0.041746 | 0.362 |  | 0.013287 | 0.4625 | -0.2352
63 | 750.768 | 4.3787 | 4.3785 | 34.2433 | 34.2445 | 7.2156 | 0.62366 | 1 | 1 | 227 | 7.8565 | 0.03984 | 0.313 |  | 0.013403 | 0.4625 | -0.2352
5 | 103.752 | 7.3441 | 7.3445 | 33.5217 | 33.5223 | 6.7636 | 1.4469 | 1 | 1 | 21 | 6.7362 | 0.42057 | 2.833 |  | 0.013435 | 0.4625 | -0.2352
92 | 748.901 | 4.0943 | 4.0942 | 34.2512 | 34.2518 | 7.2649 | 0.60615 | 5 | 5 | 360 | 7.9139 | 0.03159 | 0.25 |  | 0.013554 | 0.4625 | -0.2352
73 | 750.028 | 4.3358 | 4.3361 | 34.2546 | 34.2555 | 7.2225 | 0.61 | 2 | 2 | 274 | 7.8641 | 0.033316 | 0.262 |  | 0.013609 | 0.4625 | -0.2352
26 | 194.545 | 6.7678 | 6.7677 | 33.9373 | 33.9377 | 6.8348 | 1.0984 | 1 | 1 | 55 | 6.8929 | 0.25911 | 1.786 |  | 0.013721 | 0.4625 | -0.2352
24 | 9.999 | 11.6477 | 11.6908 | 31.2792 | 31.2624 | 6.2415 | 2.9002 | 7 | 7 | 53 | 6.118 | 1.0922 | 6.682 |  | 0.013944 | 0.4625 | -0.2352
44 | 175.934 | 6.4578 | 6.4583 | 33.9706 | 33.9712 | 6.8824 | 1.042 | 2 | 2 | 146 | 6.9278 | 0.23254 | 1.611 |  | 0.014178 | 0.4625 | -0.2352
73 | 974.345 | 3.8924 | 3.8919 | 34.3388 | 34.34 | 7.2961 | 0.60075 | 1 | 1 | 273 | 8.1885 | 0.028088 | 0.23 |  | 0.014559 | 0.4625 | -0.2352
53 | 85.413 | 7.9383 | 7.9382 | 32.7199 | 32.7209 | 6.7079 | 1.675 | 1 | 1 | 185 | 6.6599 | 0.52478 | 3.495 |  | 0.014722 | 0.4625 | -0.2352
92 | 9.982 | 10.6667 | 10.6648 | 32.2068 | 32.1912 | 6.3365 | 2.8451 | 19 | 21 | 374 | 6.2148 | 1.0954 | 6.808 |  | 0.014767 | 0.4625 | -0.2352
106 | 999.864 | 3.5253 | 3.5252 | 34.3758 | 34.3773 | 7.3598 | 0.61003 | 1 | 1 | 402 | 8.2959 | 0.032064 | 0.266 |  | 0.014875 | 0.4625 | -0.2352
126 | 20.201 | 8.9594 | 8.9626 | 31.4329 | 31.4324 | 6.6109 | 2.8014 | 9 | 11 | 449 | 6.5017 | 1.0756 | 6.993 |  | 0.01513 | 0.4625 | -0.2352
88 | 74.772 | 7.7071 | 7.7069 | 32.99 | 32.9907 | 6.7313 | 2.0842 | 13 | 13 | 347 | 6.6759 | 0.74417 | 4.968 | 4 | 0.015424 | 0.4625 | -0.2352
1 | 74.346 | 8.1055 | 8.1025 | 30.235 | 30.2344 | 6.7917 | 1.7746 | 7 | 7 | 7 | 6.7321 | 0.5701 | 3.838 |  | 0.015448 | 0.4625 | -0.2352
92 | 1499.184 | 2.6221 | 2.6221 | 34.5066 | 34.5077 | 7.5187 | 0.68753 | 2 | 2 | 357 | 9.0702 | 0.066812 | 0.606 |  | 0.015971 | 0.4625 | -0.2352
75 | 20.7 | 8.654 | 8.6479 | 32.3879 | 32.3921 | 6.6151 | 2.728 | 6 | 8 | 298 | 6.5082 | 1.0428 | 6.787 |  | 0.016348 | 0.4625 | -0.2352
3 | 19.208 | 8.324 | 8.3274 | 31.7731 | 31.7705 | 6.6909 | 2.1238 | 3 | 3 | 18 | 6.5839 | 0.73057 | 4.81 |  | 0.016495 | 0.4625 | -0.2352
23 | 4.636 | 10.423 | 10.4231 | 31.1951 | 31.1951 | 6.4113 | 3.4853 | 9 | 11 | 44 | 6.2849 | 1.3602 | 8.549 |  | 0.016498 | 0.4625 | -0.2352
23 | 156.087 | 6.8174 | 6.8173 | 33.8756 | 33.876 | 6.8298 | 1.0856 | 1 | 1 | 36 | 6.8533 | 0.25039 | 1.716 |  | 0.0165 | 0.4625 | -0.2352
78 | 9.801 | 11.7865 | 11.7868 | 31.1716 | 31.1712 | 6.2274 | 2.8716 | 6 | 8 | 309 | 6.1036 | 1.1095 | 6.772 |  | 0.016573 | 0.4625 | -0.2352
75 | 29.896 | 8.158 | 8.1567 | 32.532 | 32.5339 | 6.6829 | 2.5221 | 5 | 6 | 297 | 6.5863 | 0.94803 | 6.244 |  | 0.016748 | 0.4625 | -0.2352
67 | 19.905 | 9.6656 | 9.6756 | 32.2696 | 32.2687 | 6.4735 | 2.9417 | 19 | 21 | 264 | 6.362 | 1.1084 | 7.052 |  | 0.016879 | 0.4625 | -0.2352
24 | 140.535 | 6.6956 | 6.6962 | 33.9434 | 33.9436 | 6.8459 | 1.1144 | 1 | 1 | 47 | 6.8571 | 0.26309 | 1.804 |  | 0.017129 | 0.4625 | -0.2352
44 | 191.688 | 6.4552 | 6.4556 | 33.9724 | 33.9732 | 6.8828 | 1.0449 | 1 | 1 | 145 | 6.9422 | 0.23076 | 1.602 |  | 0.017289 | 0.4625 | -0.2352
26 | 51.163 | 8.2598 | 8.2587 | 32.5156 | 32.5156 | 6.6683 | 2.4958 | 7 | 7 | 61 | 6.589 | 0.9018 | 5.942 |  | 0.017312 | 0.4625 | -0.2352
55 | 20.3 | 9.044 | 9.0288 | 31.6149 | 31.6205 | 6.5907 | 2.3757 | 3 | 3 | 204 | 6.4813 | 0.88099 | 5.71 |  | 0.017416 | 0.4625 | -0.2352
79 | 4.426 | 11.5805 | 11.5798 | 30.541 | 30.5457 | 6.2797 | 3.1925 | 5 | 5 | 318 | 6.1513 | 1.2588 | 7.743 |  | 0.017428 | 0.4625 | -0.2352
109 | 75.143 | 7.6514 | 7.6504 | 32.8192 | 32.8205 | 6.7473 | 2.0082 | 2 | 2 | 423 | 6.6927 | 0.71123 | 4.76 | 4 | 0.017653 | 0.4625 | -0.2352
61 | 30.285 | 8.4453 | 8.4448 | 32.3152 | 32.3059 | 6.6492 | 2.1092 | 5 | 5 | 220 | 6.5513 | 0.72261 | 4.734 |  | 0.017685 | 0.4625 | -0.2352
9 | 5.429 | 10.6208 | 10.6393 | 30.6106 | 30.6079 | 6.4077 | 4.041 | 7 | 8 | 34 | 6.2812 | 1.6514 | 10.373 |  | 0.01769 | 0.4625 | -0.2352
1 | 50.619 | 7.6865 | 7.6863 | 29.5225 | 29.5236 | 6.8888 | 2.0037 | 8 | 8 | 8 | 6.8112 | 0.67359 | 4.588 |  | 0.017926 | 0.4625 | -0.2352
109 | 49.295 | 7.9879 | 7.9914 | 32.3746 | 32.3756 | 6.7155 | 2.5178 | 3 | 3 | 424 | 6.6362 | 0.91121 | 6.047 |  | 0.018074 | 0.4625 | -0.2352
1 | 20.019 | 8.8231 | 8.8237 | 29.2913 | 29.2922 | 6.7236 | 2.5299 | 10 | 10 | 10 | 6.6132 | 0.91559 | 6.055 |  | 0.019271 | 0.4625 | -0.2352
40 | 19.108 | 10.7647 | 10.7113 | 32.0959 | 32.123 | 6.3277 | 2.9388 | 14 | 16 | 140 | 6.2129 | 1.1043 | 6.861 |  | 0.019679 | 0.4625 | -0.2352
126 | 30.038 | 8.452 | 8.4461 | 31.5128 | 31.5127 | 6.683 | 2.619 | 8 | 8 | 448 | 6.5843 | 0.99601 | 6.558 |  | 0.019902 | 0.4625 | -0.2352
78 | 4.396 | 11.8162 | 11.8151 | 31.1661 | 31.1676 | 6.2237 | 2.8642 | 7 | 9 | 310 | 6.0957 | 1.1101 | 6.767 |  | 0.02062 | 0.4625 | -0.2352
44 | 74.502 | 7.8623 | 7.8624 | 33.0136 | 33.0133 | 6.7066 | 2.1243 | 6 | 6 | 150 | 6.6499 | 0.76798 | 5.107 |  | 0.020688 | 0.4625 | -0.2352
78 | 30.613 | 9.2817 | 9.2813 | 32.3514 | 32.3527 | 6.525 | 2.6149 | 4 | 5 | 307 | 6.4237 | 0.99523 | 6.393 |  | 0.021031 | 0.4625 | -0.2352
1 | 150.119 | 9.3571 | 9.3566 | 31.2352 | 31.2363 | 6.5612 | 0.55522 | 4 | 4 | 4 | 6.5579 | 0 | 0 |  | 0.021589 | 0.4625 | -0.2352
40 | 75.682 | 7.895 | 7.897 | 32.9423 | 32.941 | 6.7048 | 2.1903 | 11 | 11 | 137 | 6.6488 | 0.75592 | 5.026 |  | 0.021894 | 0.4625 | -0.2352
1 | 175.8 | 9.3491 | 9.3492 | 31.2643 | 31.2654 | 6.5611 | 0.55748 | 3 | 3 | 3 | 6.5794 | 0 | 0 |  | 0.022634 | 0.4625 | -0.2352
54 | 27.652 | 9.3586 | 9.3748 | 31.357 | 31.3555 | 6.5558 | 2.8371 | 1 | 1 | 195 | 6.4511 | 1.1 | 7.096 |  | 0.023023 | 0.4625 | -0.2352
126 | 75.954 | 7.7701 | 7.7717 | 32.1256 | 32.1248 | 6.7597 | 2.4162 | 6 | 6 | 446 | 6.7046 | 0.90624 | 6.076 |  | 0.023931 | 0.4625 | -0.2352
88 | 125.337 | 7.0635 | 7.0638 | 33.797 | 33.7973 | 6.7948 | 1.6994 | 11 | 11 | 345 | 6.7888 | 0.52675 | 3.576 |  | 0.024023 | 0.4625 | -0.2352
1 | 199.987 | 9.354 | 9.3539 | 31.2769 | 31.2782 | 6.5599 | 0.56088 | 2 | 2 | 2 | 6.5984 | 0 | 0 |  | 0.024207 | 0.4625 | -0.2352
75 | 75.068 | 7.8617 | 7.8617 | 32.7758 | 32.7768 | 6.7171 | 2.2536 | 3 | 3 | 295 | 6.6608 | 0.83143 | 5.538 |  | 0.024336 | 0.4625 | -0.2352
37 | 19.912 | 10.7096 | 10.636 | 31.9466 | 31.9635 | 6.3413 | 2.8983 | 17 | 17 | 123 | 6.2271 | 1.0803 | 6.727 |  | 0.024993 | 0.4625 | -0.2352
109 | 94.58 | 7.3519 | 7.3484 | 33.2949 | 33.3001 | 6.7724 | 1.7308 | 1 | 1 | 422 | 6.737 | 0.59062 | 3.979 |  | 0.025315 | 0.4625 | -0.2352
1 | 216.971 | 9.3563 | 9.3564 | 31.2827 | 31.2841 | 6.5593 | 0.5642 | 1 | 1 | 1 | 6.6121 | 0 | 0 |  | 0.025743 | 0.4625 | -0.2352
75 | 49.843 | 8.0684 | 8.0676 | 32.5727 | 32.5746 | 6.6947 | 2.425 | 4 | 4 | 296 | 6.6154 | 0.91271 | 6.038 |  | 0.02634 | 0.4625 | -0.2352
67 | 1248.531 | 3.1842 | 3.1843 | 34.4563 | 34.4573 | 7.4176 | 0.63502 | 3 | 3 | 248 | 8.6448 | 0.029266 | 0.253 |  | 0.029231 | 0.4625 | -0.2352
102 | 100.506 | 7.5549 | 7.5552 | 33.2997 | 33.3001 | 6.7409 | 1.9114 | 14 | 14 | 392 | 6.7089 | 0.67835 | 4.551 |  | 0.029508 | 0.4625 | -0.2352
28 | 19.807 | 10.5361 | 10.5485 | 31.7522 | 31.749 | 6.3729 | 2.9208 | 13 | 15 | 78 | 6.2589 | 1.0847 | 6.789 |  | 0.030959 | 0.4625 | -0.2352
79 | 20.313 | 9.9509 | 9.9219 | 31.536 | 31.5489 | 6.4635 | 2.6148 | 3 | 3 | 316 | 6.3511 | 1.0057 | 6.387 | 6 | 0.031514 | 0.4625 | -0.2352
1 | 30.148 | 8.6623 | 8.6628 | 29.3606 | 29.3614 | 6.7449 | 2.4606 | 9 | 9 | 9 | 6.6439 | 0.86847 | 5.77 |  | 0.034357 | 0.4625 | -0.2352
53 | 20.653 | 8.6903 | 8.685 | 31.5945 | 31.5982 | 6.6438 | 2.2167 | 5 | 6 | 189 | 6.5362 | 0.82479 | 5.391 |  | 0.034775 | 0.4625 | -0.2352
23 | 9.7 | 9.7443 | 9.8014 | 31.2852 | 31.2765 | 6.5033 | 3.3544 | 8 | 10 | 43 | 6.3827 | 1.3563 | 8.657 |  | 0.040146 | 0.4625 | -0.2352
3 | 4.865 | 10.3217 | 10.3203 | 31.1303 | 31.1323 | 6.4281 | 3.9176 | 5 | 5 | 20 | 6.3021 | 1.5314 | 9.651 |  | 0.045298 | 0.4625 | -0.2352
78 | 76.852 | 7.9783 | 7.9778 | 32.8823 | 32.8919 | 6.6948 | 1.9794 | 2 | 2 | 305 | 6.6392 | 0.7278 | 4.832 | 6 | 0.047525 | 0.4625 | -0.2352
30 | 9.859 | 11.486 | 11.4867 | 31.8013 | 31.7997 | 6.2423 | 2.8009 | 18 | 20 | 100 | 6.1193 | 1.1081 | 6.781 | 3 | 0.047922 | 0.4625 | -0.2352
52 | 11.239 | 10.0917 | 10.1407 | 31.0157 | 30.9956 | 6.4652 | 3.1385 | 6 | 6 | 183 | 6.3447 | 1.2647 | 8.024 |  | 0.048333 | 0.4625 | -0.2352
53 | 50.102 | 8.0047 | 8.005 | 32.3066 | 32.3071 | 6.716 | 1.823 | 3 | 3 | 187 | 6.6372 | 0.65645 | 4.357 |  | 0.048517 | 0.4625 | -0.2352
2 | 10.688 | 8.9486 | 8.9603 | 31.4924 | 31.4893 | 6.61 | 3.1592 | 2 | 2 | 14 | 6.4929 | 1.1619 | 7.544 | 6 | 0.06405 | 0.4625 | -0.2352
2 | 6.043 | 9.9761 | 9.9759 | 31.3503 | 31.3498 | 6.4676 | 3.7854 | 3 | 3 | 15 | 6.3434 | 1.5807 | 10.027 |  | 0.065153 | 0.4625 | -0.2352
30 | 21.163 | 10.2057 | 10.2052 | 32.122 | 32.1209 | 6.4036 | 2.9159 | 17 | 19 | 99 | 6.2916 | 1.0285 | 6.471 | 3 | 0.0849 | 0.4625 | -0.2352
44 | 151.466 | 6.687 | 6.6873 | 33.9221 | 33.9226 | 6.8483 | 1.139 | 3 | 3 | 147 | 6.8691 | 0.37719 | 2.591 |  | 0.085606 | 0.4625 | -0.2352
30 | 5.261 | 11.4936 | 11.493 | 31.7995 | 31.7982 | 6.2414 | 2.8137 | 19 | 21 | 101 | 6.1148 | 1.1577 | 7.079 | 3 | 0.091549 | 0.4625 | -0.2352
30 | 30.37 | 9.6709 | 9.6683 | 32.4319 | 32.4315 | 6.466 | 2.9164 | 16 | 18 | 98 | 6.3631 | 1.0206 | 6.494 | 3 | 0.093071 | 0.4625 | -0.2352
47 | 50.456 | 8.169 | 8.1691 | 32.7264 | 32.7259 | 6.6728 | 2.0963 | 3 | 3 | 158 | 6.5936 | 0.84476 | 5.57 |  | 0.11041 | 0.4625 | -0.2352
1 | 10.532 | 9.2311 | 9.2432 | 29.1443 | 29.1411 | 6.6688 | 2.7536 | 11 | 11 | 11 | 6.5488 | 1.1627 | 7.614 |  | 0.12431 | 0.4625 | -0.2352
5 | 4.672 | 9.9217 | 10.07 | 30.7664 | 30.689 | 6.4997 | 3.2139 | 7 | 7 | 27 | 6.3741 | 1.3883 | 8.849 |  | 0.13706 | 0.4625 | -0.2352
30 | 50.136 | 8.9105 | 8.9106 | 32.5004 | 32.4993 | 6.5726 | 2.6962 | 15 | 15 | 97 | 6.489 | 1.158 | 7.514 |  | 0.14618 | 0.4625 | -0.2352
3 | 10.022 | 9.2365 | 9.3434 | 31.4517 | 31.4152 | 6.5696 | 3.0509 | 4 | 4 | 19 | 6.4509 | 1.0296 | 6.642 |  | 0.14622 | 0.4625 | -0.2352
67 | 299.665 | 6.2603 | 6.2599 | 34.0102 | 34.0109 | 6.9123 | 0.95258 | 8 | 8 | 253 | 7.0719 | 0.60818 | 4.301 | 3 | 0.40281 | 0.4625 | -0.2352
102 | 1251.369 | 3.0226 | 3.0218 | 34.4421 | 34.4431 | 7.4478 | 0.62231 | 3 | 3 | 381 | 8.6877 | 0.71653 | 6.225 | 3 | 0.66392 | 0.4625 | -0.2352