## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | Updated channel names & formats in TOB files. GG |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 23 August 2020 | ### Added HPLC Data. S.H. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2011-17

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: September 1, 2011 – September 13, 2011

Processed by: Germaine Gatien

Date of Processing: 15 February 2012 – 29 February 2012

Number of original HEX files: 111

Number of CTD files: 111

Number of bottle casts:	39

Number of original TSG files: 2

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1176), a Wet Labs Eco-AFL/FL Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4601) and an altimeter (no serial # available).

The deck unit was a Seabird model 11, serial #0471. The logging computer was #3.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The Oxygen kit was Metrohm Model 765 Dosimat Kit Number 2.

There were 24 10L bottles mounted on an IOS Rosette.

A thermosalinograph (SeaBird 21 S/N 3363) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but the altimeter serial number was missing from the log and from the configuration file. Details were given for the set-ups for the CTD, rosette, oxygen kit and Thermosalinograph.

Sampling notes were provided with a list of casts that had the PAR sensor mounted and note of a few problems in file headers. The cruise report was available and indicated that the CTD and computer worked well throughout the cruise.

The rosette logs were generally in good order. For most casts some rosette bottles were fired out of order. There is a list of the order at the bottom of the rosette log sheets, but in some cases there were last-minute changes or misfires and some errors were made in sampling for cast #47.

The WetLabs ECO fluorometer was used for this cruise. It does not need to be pumped but has a poorer time response (though adjusting the range may improve this). The traces are thus much smoother than from the SeaPoint and the alignment needs a larger correction.

The Dissolved Oxygen sensor calibration showed no significant drift through the cruise.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1      mL/L	 above 10db

±0.2   mL/L	 from 10 to 100db

±0.1   mL/L	 from 100db to 200db

±0.04 mL/L 	below 200db

The TSG intake temperature compared very well with the CTD temperature. TSG salinity was found to be lower than the CTD by about 0.024 when outliers were excluded. Those outliers were all 1-sided with the TSG salinity being lower than that of the CTD by more than 0.05. This is suggestive of problems noted during 2011-26 where it looks like the water in the loop came from a shallower level than expected. This is likely to be more significant when the ship is moving. There was only 1 underway loop sample, so recalibration of the TSG salinity was based only on the comparison with the CTD salinity and the history of the sensor. The fluorometer was higher than that on the CTD by a median factor of 1.6

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as sampling notes. Errors in station names for 2 casts were corrected in the raw files. No CTD equipment problems were noted and the cruise report mentions that the new CTD computer worked very well.

Extracted chlorophyll, nutrients, dissolved oxygen, salinity and NH4 data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated in March or April 2011.

The PAR sensor was not always mounted. Cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

The calibration constants were checked for all instruments. The only error found was a small mistake in the PAR calibration constant. That was corrected and the file was saved as 2011-17-ctd.xmlcon. A few files were converted; all channels were present and the data look reasonable, though (as observed during 2011-27) the difference between T and C channels is much higher when in motion during the upcasts than for the downcast or while stopped for bottles. This suggests an alignment problem that might be caused by the rosette package set-up. The primary sensor looks spikier than the secondary as found during the previous cruise.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. (Since bottles were fired out of order, the file was 1st ordered on bottle position, sample #s added and it was then reordered on bottle number.) Sample numbers were added to the ADDSAMP file based on rosette log records.

There were some repeats of sample numbers and some non-standard sample numbers, but all concern casts with no IOS sampling, so this should not be a problem.

Note that while in general the sample numbers were entered in order of Niskin bottle #, there are a few cases were that was not the case, and a few that are uncertain. Checks of casts 29, 47 and 66 should be made after the bottle files are merged to see if the bottle values fit the pressures. This problem generally arises because a bottle is accidentally tripped at the wrong depth. An example of a problem cast is #47 which will be discussed later in this section.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-17-DO-cal.csv. A template file was named 2011-17-DO_study.xls. The exported CTD data was entered on sheet “CTD data” and the bottle DO data were added to sheet “bottle data”. The bottle data were temporarily added to the “CTD data” sheet as well, so a comparison could be done and any CTD data removed for which there are no DO samples. Then the CTD data were moved to the main page.

The titrated DO values were added to that file and lines removed for which there was no DO sampling.   A calculation was made of Ф using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

When all values are included the R2 value was 0.9866 but there are a few obvious outliers. Removing values flagged “3” or “4” produced only slightly better results. Next, fits were done by gradually removing outliers. It is difficult and tedious to pick our outliers on the plots and find and eliminate them from the fit. A simpler approach was to use the M and B values from the factory calibration to determine the difference from the fit for each sample, as follows: 
  	Difference = M*Voltage – B – DO/Phi

Watching how the plots change, the process stops when obvious outliers have been removed as judged by visual inspection and the R2 value, being careful to stop before the DO range is significantly reduced.

When the data were sorted on that difference, fits were done excluding points with differences > a chosen amount. Here is a summary of the results:

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4453 | -0.2321 | 0.4453 | -0.5212 |  |
| 343 | all | 0.4726 | -0.2435 | 0.4726 | -0.5152 | 0.9866 |
| 333 | all data except 3 and 4 flags | 0.4716 | -0.2428 | 0.4716 | -0.5148 | 0.9894 |
| 323 | exclude flagged and diff>0.1 | 0.4738 | -0.2462 | 0.4738 | -0.5196 | 0.9993 |
| 240 | excl. outliers diff>0.05 | 0.4676 | -0.2401 | 0.4676 | -0.5135 | 0.9991 |
| 176 | excl. outliers diff>0.03 | 0.4622 | -0.2362 | 0.4622 | -0.5110 | 0.9991 |
| 117 | excl. outliers diff>0.02 | 0.4585 | -0.2350 | 0.4585 | -0.5125 | 0.9992 |
| 62 | excl. outliers diff>0.01 | 0.4545 | -0.2344 | 0.4545 | -0.5157 | 0.9994 |
| casts 1-59 | exc. Flagged "3" & diff>0.1 | 0.4728 | -0.2449 | 0.4728 | -0.5180 | 0.9992 |
| casts 61-135 | exc. Flagged "3" & diff>0.1 | 0.4747 | -0.2474 | 0.4747 | -0.5212 | 0.9995 |

Excluding the differences >0.1 looks like a good choice as it includes a lot of bottles, has one of the highest R2 values and maintains the full range of DO  values. The fits excluding differences >0.02 and 0.01 remove data from the high end of the DO range. Here is a summary of recent results for this sensor:

| History of Calibration Parameters |  |  |  |
| --- | --- | --- | --- |
|  |  | Soc | Voffset |
| Factory | April | 0.4453 | -0.5212 |
| 2011-09 | June | 0.4622 | -0.4935 |
| 2011-27 | Aug. | 0.4694 | -0.5075 |
| 2011-17 | Sept. | 0.4738 | -0.5196 |

This shows a fairly steady rate of change in Soc, but that may well reflect different conditions rather than be totally due to calibration drift. During 2011-27 there was some apparent time-dependence, though the change was in the opposite direction one would expect for calibration drift, so may be due to different DO range or pressure effects since the deepest casts were later in the cruise. For this cruise there is also a slight hint of change but it is not significant enough to justify a time-dependent recalibration, and again may reflect different DO ranges and pressures. A time-dependent recalibration requires very time-consuming study and the errors due to using fits based on much smaller data sets would likely be larger. There will be chances later to refine the calibration, if necessary.

(For full details see 2011-17-do-cal-study.xls.)

The only significant outlier that had not been flagged by the analyst is sample #164. This is from a confusing cast, #47. It was originally intended to fire Niskin #7 at 15db, but Niskin #12 was inadvertently fired instead at that level. Niskin #7 was later fired at 5db. Most of the samples from Niskin #12 were labeled as sample #164, but it appears the DO sample came from Niskin #12 and should have been given sample #169. Sample #164 should be excluded from the comparison.

The configuration files were updated with the new parameters Soc and Voffset and saved with names 2011-17-ctd-new.xmlcon.

##### Hysteresis Study

Hysteresis tests were run on this sensor during 2011-27 and no changes were found appropriate.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. The only data that looked odd were primary salinity near the bottom of cast #59. The noisy points were removed using CTDEDIT. The output file was copied to *.BOT.

A preliminary header check turned up no problems and the maximum fluorescence value is ~49 so there are no off-scale fluorescence data, however, the minimum values are slightly negative, as was found for other recent cruises. A decision on what offset to apply to remove negative data will be made later, but the minimum in the BOT files was -0.092ug/L.

The addsamp.csv file prepared in the DO calibration step was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-17-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2011-17chl.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2011-17chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-17oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2011-17oxy.csv. That file was converted into individual *.OXY files.

NH4

NH4 data were obtained in file QF2011-17NH4.xls which included a report on precision. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was saved as 2011-17NH4.csv. The file was then converted to individual NH4 files. (It was later found that the sample that should have been named #164 was actually labelled as 169. The values make sense for #164.)

SALINITY

Salinity analysis was obtained in 2011-17SAL.xls. The analysis was done within a month of the beginning of the cruise. The files were simplified and saved as 2011-17sal.csv. A loop sample was saved in a separate file, 2011-17-loop-sal.csv and then removed from the main file. Event numbers were added. There was one pair of duplicates that had not been averaged. They differed by 0.0004. The 2 values were replaced by the average and flag “6” was added. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-17nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers, header names were changed to standard format, reordered on sample number and the file was saved as 2011-17-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, ADD, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

Cast #47 was examined to see if there was a labelling error in the DO sample noted in section 3. Niskin #7 was meant to be fired at 15db but due to an error Niskin #12 was fired there instead. So Niskin #7 was closed at 5db. The DO sample labelled as #164 is clearly not from 5db – it matches the CTD DO at 15db where sample #169 is appropriate. The draw temperature falls between those at 10db and 25db. There was no other sampling from that Niskin bottle and it is likely that the DO sample was meant to come from the 5db bottle. Other samples labelled #164 are clearly from near the surface. The DO sample was renamed as #169, flagged “3” and an explanatory comment entered. The merge process was repeated.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was 1 major outlier and there are a lot of near-surface samples with large differences. When data from above 20db are excluded the CTD primary salinity was found to be high by an average of 0.0014 and the secondary high by 0.0009, with standard deviations of 0.006 for both. When data above 100db are excluded the primary is high by an average of 0.0001 and the secondary is low by 0.0004 and standard deviations of 0.001.

The major outlier was sample #92 from cast #29. That sample should have been from 5db according to the rosette log, but the value looks close to the CTD salinity for sample #80 at 400db. Given how the bottles are mounted this looks like a mis-sample, choosing from a nearby bottle. The other samples labelled #92 look fine, with a draw temperature that is consistent with surface samples. There is already a salinity sample #80 which is not noted on the rosette log, so sample #92 will not be relabelled, but will be replaced with a pad value and flagged “5”.

A full profile was run at cast #93 enabling an assessment of pressure dependence and that was very slight when bottles above 200db were excluded. Plots of differences versus file pair numbers indicate that there is no significant time dependence.

No recalibration is required for these sensors.

For full details for the COMPARE run see file 2011-17-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected the fit of differences against DO concentration is quite flat and most differences are within ±0.2mL/L. The only outliers from below 12db were from event #40 near the bottom and event #93 at 24db. For the former it is possible that the CTD was in a bottom layer so that the distance between bottle and CTD made a significant difference. For the second case the CTD data were very noisy. There is no evidence of a problem with the bottle samples. No further flags will be added.

A plot of differences against file pair # showed no time dependence.

For full details of the comparison see 2011-17-dox-comp1.xls.

Fluorescence

COMPARE was run using the Wet Labs ECO CTD Fluorescence and the Extracted Chlorophyll from bottles. There was a wide range of values with no Extracted CHL values <0.15ug/L and the ratio of CTD Fluorescence to Extracted CHL is generally ~1 except for where CHL <2.2ug/L where the ratio gets noisier and sometimes higher. This is a better fit than for some previous cruises, likely because the chlorophyll levels are higher. No notable outliers were found.

See 2011-17-chl-fluor-comp1.xls for the full COMPARE results.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. A few problems were discovered:

- There was no sample # for Niskin #21 for cast #29 – this line was removed from the SAMAVG file.

- There were no sample #s for any bottles for cast #30, so that cast will not be processed further.

- There was no sample # for Niskin #18 for cast #93 – this line was removed from the SAMAVG file.

- There was no sampling and no sample #s for any bottles for cast #103, so that cast will not be processed further.

- There was no sample # for Niskin #24 for cast #113 – this line was removed from the SAMAVG file.

- Many of the later casts had no sample numbers, or non-standard sample numbers, but sampling was done for other institutions, so files will be prepared.

- There was no sampling for Niskins #4-24 for cast #158 – those lines were removed from the SAMAVG file.

- NH4 Sample #169 looks like it is mislabelled. The rosette sheet indicates the sample was #164 and the value looks out of place as #169. The sample # was changed to #164 and flag “3” added.

MERGE was rerun. CLEAN was rerun on the MRG files. Data were exported to a spreadsheet again and the results look ok.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-17-ctd-new.con.

A few casts were examined and all expected channels are present. The descent rate varies greatly with some casts steady and others very noisy with obvious shed wake corruption.

The two temperature channels are fairly close during the downcasts though the primary has some small spikes and odd excursions. During the upcasts traces differ much more and again the primary looks noisiest. The conductivity channels are similar to temperature. Fortunately, during stops the noise mostly disappears, though some primary conductivity spikes were seen during stops. For some casts there was a lot of vertical motion during stops producing a lot of noise in the signals.

Altimetry looks useful at the bottom and fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge because the temperature is so noisy on the upcast, but 4s appears to align features best and that setting was found appropriate for other recent cruises for which this sensor was used.

ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### CELLTM

The upcast data are extremely noisy so the usual tests for CELLTM settings are not helpful. The same equipment was used during 5 other recent cruises. The two most recent ones had the same problem, but the earlier 3, while a little noisy, did provide some results. The tests for 2011-26 in June 2011 were reasonably clear and the best choice was found to be (α = 0.02, β=7) for the primary and (α = 0.03, β=9) looked best overall. Those results also appeared useful for 2011-44, 2011-16 and 2011-27 when the same equipment was used. So they were applied to these casts as well.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen. Data from one cast from each of the 2 previous cruises that used the same sensors are also shown for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-26-0082 | 800 1000 1950 3200 | ~0 ~0 -0.0006 -0.0008 | +0.00006 +0.00007 +0.00007 +0.00008 | +0.0007 +0.0008 +0.0013 +0.0018 | VNoisy, VHigh |
| 2011-27-0070 | 800 1000 1950 3200 | -0.0002 -0.0005 -0.0006 -0.0012 | ~0 +0.00001 -0.00001 ~0 | +0.0002 +0.0003 +0.0005 +0.0012 | High, X Noisy |
| 2011-17-0032 | 800 1000 1950 | -0.0001 -0.0004 VN -0.0010 N | -0.00003 -0.00003 -0.00003 | -0.0002 VN ~0 VN +0.0005 N | High, V Noisy |
| 2011-17-0068 | 800 1000 1950 | -0.0005 N -0.0005 N -0.0009 XN | -0.00005 XN -0.00005 XN -0.00005 XN | -0.0005 XN -0.0003 XN +0.0003 VN | High, Noisy |
| 2011-17-0121 | 800 1000 1950 | -0.0004 -0.0004 XN -0.0008 VN | -0.00008 -0.0001 VN -0.00006 N | -0.0005 -0.0006 XN +0.0001 XN | High, V Noisy |

The differences are small. Temperature shows little change with time, but conductivity may be drifting though that does not look significant in light of the noisiness of the data. There is some pressure dependence but not enough to raise concern about the sensors.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some negative values in pressure early in the file for cast #96. The minimum fluorescence value is -0.091ug/L.

Surface check was run and shows an average surface pressure for the cruise was 2.5db which looks about right for the Tully. A few casts with negative pressure readings were examined and there were only single slightly negative values or a few such values together, but there was no evidence that the CTD had really come out of the water. The negative pressures look like small spikes, not evidence of trouble with the pressure sensor.

The cross-reference check was compared with the log book and the only problem found was 1 cast with the wrong station name (#165); that was corrected in the full profile files. There was no bottle file for that cast.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet. For 2 CLN files and 1 bottle file (MRGCLN2 and SAMAVG) a spike was misinterpreted, so those readings were removed from the headers.

The Water Depth header was also examined. There were deviations from the log book entries or entries that were less than the maximum depth sampled; most look like errors in entry but some may have been due to the ship drifting between the CTD reaching the bottom and the bottom depth reading. For 8 cases the entries were adjusted (casts #27, 30, 71, 87, 106, 108, 121 & 142) in the IOS and CLN files and, where they existed, the MRGCLN2, SAM and SAMAVG files.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence trace look like that of the temperature trace. The noisiness of the upcast temperature traces makes this a difficult judgment, but the value used on other recent uses of the ECO sensor +48 records produces reasonable results. SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -1s was applied to the primary and a shift of +0.5s to the secondary conductivity. The shift is larger for the primary and the same for the secondary as found when the same equipment was used during 2011-16, 2011-26 and 2011-27.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings are for cast #96 and were caused by pressure spikes from early in the cast. Removal of the first 2224 records from the SHFC1 file and DELETE was then rerun successfully for that file.

Header Check was repeated on the DEL files and the minimum fluorescence value is -0.091mL/L, so that setting should be used for recalibration. No negative pressures were found.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in late March 2011 and were used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27. For the first cruise, they were few calibration samples and for the others the bottle calibration was not trusted. No corrections were applied.

2. Dissolved Oxygen

The DO sensor was repaired and recalibrated in April 2011. It was used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27. There appeared to be some time dependence in the 2011-26 data, but not in any of the others. The variation in the slope and offset does not follow a simple relation with time. This may be a sign that the DO range, pressure range and number of samples available are significant in the fits. And the 2nd calibration to 2011-26 complicates the issue. A 2nd calibration was also applied to 2011-27 to remove some remaining hysteresis.

3. Pressure

The sensor was recalibrated in April 2011 and was used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27. No further offset was applied to any of those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions in the temperature data were some values a little above the historic maxima at the surface of cast #61. Salinity was slightly low in the top 10 to 20m at many inshore casts along the LC and LD lines and at a few in Queen Charlotte Sound and 1 in the Strait of Georgia. Salinity was a little high near the surface at cast #82. These excursions look real, not indicative of instrument malfunction. The 3-standard deviation limit is severe for near-shore casts.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### DETAILED EDITING

The bottle comparison shows little difference between the primary and secondary salinity channels, but the primary is noisier than the secondary. The secondary sensors were chosen for archiving for all 2011 cruises using this equipment. So the secondary T and S channels were chosen for archiving, and so, editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. Many of the casts had extremely noisy descent rates and required a lot of editing. Where more than 1db of data were removed from the bottom of the cast, the altimetry header was adjusted by increasing it by the amount removed.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure, salinity and dissolved oxygen channels do not need recalibration.

The fluorescence channel needs adjusting by adding 0.091ug/L.

CALIBRATE was run using file 2011-17-recal1.ccf to apply Fluorescence correction. It was applied to the EDT, SAM and MRGCLN2 files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against DO concentration and a few outliers were excluded, there was very little pressure dependence and the CTD data were higher than the bottles by an average of 0.003mL/L. A plot against Dissolved Oxygen (bottles) show a similar flat distribution. There is no need to calibrate DO further. (See 2011-17-dox-comp2.xls for details.)

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD and saved. Since there was only an ECO fluorometer which does not need filtering, the usual second set was not prepared. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

Since the ECO fluorometer was used no filtering is required.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1     mL/L above 10db

±0.2  mL/L from 10 to 100db

±0.1  mL/L from 100db to 200db

±0.04 mL/L below 200db

For details on the processing see processing report: 2011-17-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was more variability than usual, with values from 70 to 90% in the near-shore area of the LB line and one value even lower in Queen Charlotte Sound. The highest values (>130%) were also near-shore, but on the LC, LD, LG and LJ lines. Most of the values were between 100 and 120%. A few cases of very high and very low saturations were checked by comparing the CTD and bottle DO. The lowest values are associated with high surface DO gradients and the CTD values do not match the bottles well, so the low saturations are likely not reliable. For the casts with high saturation the CTD values are often lower than the bottles, but occasionally higher. There appears to be a lot of near-surface variability, but the high saturations appear to be real.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and a few errors were found and fixed.

A header check was run on the final files and no problems were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

##### 25. Thermosalinograph Data

Data were provided in 2 hex files. There was 1 loop sample (Salinity and Chlorophyll); those were combined in file 2011-17-loop.csv. Times were added to the spreadsheet based on the log record.

a.) Checking calibrations

The calibrations were checked and the only problems were in the fluorometer entry which had the wrong scale factor. After that correction the CON file was saved as 2011-17-tsg.xmlcon. The remote temperature calibration parameters could not be checked. They are not listed in the configuration file or in the header records. This is because it is connected serially with the new TSG system and details can only be checked through SEATERM when it is connected to a computer directly. This is awkward and we have no way of being sure it was done correctly. This should be checked at the beginning of every cruise to be sure no-one has changed it.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. Isolated spikes in salinity were noted, but no major problems.

c.)  Checking Time Channel

The CTD data, after editing and recalibration, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.3db of 4db and exported to a spreadsheet which was saved as 2011-17-ctd-tsg-comp.xls.

The files ATC files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 108 matches. TSG values were found in the same way for times of the underway loop sampling and added to file and it was saved as 2011-17-loop-tsg-comp.xls.

The flow rate was mostly ~1.1 except for the first 1000 scans of the first file when it was ~1.0.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.0002° and in longitude <0.0005°. The median differences were 0.00001° for both. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. The median difference between the two temperatures was between 0.17 and 0.18Cº, but the traces are very noisy.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. When all data were included the standard deviations were very high, so the casts were sorted in order of the standard deviation in the temperature differences, and the 44 casts with lowest standard deviation were selected. This eliminated the largest outliers. For each of the temperature and salinity further outliers were identified and excluded from the comparison. (6 for T with differences >0.17 and 9 for S with differences >0.05)

When all data were included the TSG intake temperature was higher than the CTD by an average of 0.143Cº, median of 0.020Cº with a standard deviation of 0.49ºC. From the reduced data set, the temperature was high by a median value of 0.0005Cº, but the standard deviation was 0.025Cº, so this is a rough estimate.

Using the same reduced data set the median heating in the loop is 0.17Cº with a standard deviation of 0.06Cº.

For salinity the reduced set showed the TSG salinity to be low by a median of 0.024 with a standard deviation of 0.147. The most severe outliers are from inlet casts.

The ratio of TSG fluorescence to CTD fluorescence ranges from 1.1 to 12.6 with a median of 1.6 and standard deviations of 1.7. (See 2011-17-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons  Spreadsheet 2011-17-loop-TSG-comp.xls was prepared by combining the loop samples and adding dates and times and then adding the TSG data from that time (median values over 2 minutes). The TSG salinity was lower than the loop by 0.12 and the TSG Fluorescence was 2.25 times the extracted CHL value.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and were used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27, but there are not yet any data available for 2 of those.

2011-44: The intake temperature looked unbelievable being higher than the CTD by 0.41Cº and higher than the lab temperature by 0.2 Cº. The lab temperature was higher than the CTD by 0.25Cº which is about the amount of heating we would expect to find in the loop at that time of year. The TSG Salinity was lower than the CTD salinity by ~0.007 during offshore casts though it was lower by much more in inlets with high near-surface gradients. The TSG fluorescence was higher than the CTD fluorescence by a median factor of 2.

2011-26: The intake temperature was very close to the CTD temperature. The TSG salinity was poor, especially for offshore casts; an estimate was made that it was low by 0.02 based on 1 loop sample and comparisons with CTD salinity in some sections where salinity looked better. TSG fluorescence was higher than the CTD fluorescence by a median factor of 2.

2011-09: The intake temperature was higher than the CTD temperature by 0.02C°. The TSG salinity was low by a median value of 0.21, but near-surface CTD salinity was very noisy and there were no loop samples, so the results of 2011-26 were used. TSG fluorescence was higher than the CTD fluorescence by a median factor of 1.6.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was fairly steady.

3. The temperature in the loop increases by a median of 0.17Cº between intake and lab which is about what would we expect at this time of year.

4. The TSG intake temperature is as close to the CTD temperatures as we can expect.

5. The TSG Salinity is lower than the CTD salinity by a median value of 0.024 which is about what is expected from the history of this sensor. Calibrate should be run to add 0.024 to salinity.

6. During cruise 2011-26 there were problems with TSG salinity being anomalously low during the casts that were well offshore. This was mostly observed while the ship was moving, but not exclusively. No clear explanation was found for why the TSG would behave differently in deeper water than close to shore, though there could be a relationship to higher ship speeds on the long runs of Line P, which might affect the depth from which intake waters were drawn.  For this cruise there are some minor outliers (with low salinity) at the casts furthest offshore of the LB and LC lines, but deeper casts to the north do not appear out of line. The 1 loop salinity sample available indicates that the TSG salinity was low by 0.12 and that was from the offshore while the ship was moving, so is consistent with the Line P observations, but very weak evidence.

7. The TSG fluorometer had values higher than the CTD fluorometer by a median factor of 1.6 which is typical of this instrument. The fluorometer should go to the factory for calibration, but the data will not be adjusted due to the large uncertainties in the comparison.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. The initial 15 records were removed from the first file as the flow was off or recently turned on so the data were noisy. In the 1st file there were some off-scale fluorescence data (maximum values are 81.97ug/L), so those points were removed since the values were likely higher than they appear. In both files single-point spikes in salinity were cleaned by interpolation where there was no significant temperature variation to explain them.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

File 2011-17-tsg-recal1.ccf was prepared to adjust salinity by adding 0.024. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag. Position:New was removed since there were very few 0 values, so it is not very useful.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Intake and Temperature:Lab. The reorder is to ensure that programs pick the intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

29. Bottle 16 (instead of 19) was fired at 16m – sample number was changed

30. Sampled only for nutrient standard – no bottle file required

44. Bottle #9 not fired – sample taken from #6.

50. Very goopy, reddish

65. Wrong station name – has been corrected

91. Wrong station name – has been corrected

93. Salinity test

103. Bottles fired to close and soak, no bottle file required

116. LARS fixed – had been a leak

137. Pressure on deck 0.7-0.8-0.6-0.5

146. Stop at 40m for crew change

158. Niskins 4-24 closed at surface, no sampling - only 1-3 required for bottle file.

165. Wrong station name – has been corrected.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3396 | 1Apr2011 | Factory |  |  |
| Conductivity | 2374 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2754 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2668 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1396DR | 15Aug2011 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2216 | ? |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2011-17

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 23Mar11 | Factory |  |  |
| Conductivity | 3363 | 23Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 0603 | 03Mar11 | “ |  |  |