## Notes and Updates
Entered MS Nov2011
| Cruise: | SWL 2011-18 Arctic Cruise
Analysing Equipment
| > | Guildline Autosal Model 8400B, serial # 68572
|  | > | bath temperature of 24OC
|  | > | accuracy of < +/- 0.0001 conductivity ratio
|  | > | resolution <0.00001 conductivity ratio
| > | The calibration checks were done using IAPSO Standard Seawater
|  | > | K15 Value = 0.99979 Batch# P153 Expiration Date-March 8, 2014
|  | > | The Salinometer remained stable and within the Guildline spec's of 0.002 PSU
Data Interpretation Equipment
| > | Ocean Scientific International Salinometer Data Logger software was used to log the conductivity readings from the Guildline Autosal calculating salinity value.
Lab and Sample Conditions
| > | Analysed in a temperature controlled lab with the ambient temperature changing between 21.0oc-23.4oc overall with a 1.3oc maximum change during analysing session.
| > | Sample temperatures ranged from 21.0oc to 22.4oc overall with a 1.4oc maximum change during analysing session
| > | The sample bottles were inverted and mixed prior to analysis.
| > | all samples had minimum 3 flushes and minimum 2 reading
|  | > | each reading is the average of 5 readings taken in a 5 second period
Analyst Observation
| 1 | The front panel of the Autosal was covered with salt residue on every knob, window, and most importantly the bung.
Autosal washed with clean water before any analysing was started.
| 2 | The rubber bung was deformed and would not form a vacuum so sample water could be drawn up into the cells
rubber bung was replaced and calibration was done.
The Autosal was stable resulting in a smooth analysing process
|  | > | Autosal accuracy for electronics is < 0.002 PSU equivalents
|  | > | the electronics of the Autosal 8400B drifted within the manufactures spec's during analysing
| 3 | A calibration was done at the start of each day before analysing started.
|  | > | correction values were within manufactures spec's for both days.
|  | > | the Rs value 5.68 and the Standby value read 24+5996 for both days
| 4 | A Deepwater reference sample was analysed every 24 samples to check for drift in the electronics -all looks good -see DeepWaterCheck Records sheet
| 5 | Most of the samples were stored in proper salinity bottles with separate inserts and caps which helps in minimizing evaporation and leakage
|  | > | there were however incorrect inserts used on some samples resulting in leakage and probably evaporation this may have compromised the integrity of the salinity value making it suspect
|  | > | there were also some samples with no inserts used and this too may result in sample being suspect.
|  | > | ***these suspect samples are identified in the comments on the Autosal-Data by Time worksheet
Analyst Suggestions
| 1 | department responsible for maintaining the Autosal should do regular checks on machine performance and parts replacement
| 2 | all users should wash down the front of the Autosalinometer and specially the rubber bung after analysing sessions
this practice will minimize contamination of your samples and the rubber bung from drying out, deforming, and breaking down
it is also considerate to other users
(Note: Mary had found the autosal and bung had been covered in salt before she started)
| 3 | continue to use proper salinity bottles with separate insert and cap mechanisms
| 4 | make sure a sufficient number of inserts for the samples are packed
| 5 | train samplers on the importance of using proper inserts and have them readily available at time of sampling
| 6 | arctic waters are VERY cold -after sampling when the sample warms to ambient temp check that inserts have not popped up
| 7 | assure that transport of samples are protected from freezing
| 22 Nov 2011 SZ |  | Clarified "Batch Date" is "Batch Expiration Date" in the above notes
Suggest an At-Sea MANUAL for next season with pictures.

## flags
IOS QA/QC code from Joe - to be paired with final data
Flag = definition
a = Sample for this measurement was drawn from water bottle but not analyzed (not normally used)
b = Acceptable measurement (not normally used) | although data look odd they are real
c = Questionable measurement (no problem observed in sampling or analysis, but value is not trusted, nonetheless; includes outlyers) | keep - but questionable
d = Bad measurement (known problem with sampling or analysis, but not serious enough to completely discard the value)
e = Not reported (lost sample; unredeemably bad measurement) | don't use data point (left in working sheet but not transferred to final data)
f = Mean of replicate measurements (summary of replicates and table of replicate data should be included in the metadata file)
g = Manual chromatographic peak measurement
h = Irregular digital chromatographic peak integration
i = Sample not drawn for this measurement from this bottle (not normally used)
Other QA/QC
q = data point changed but requires another round of plotting (check again!)
qc = questionable measurment, checked by analyst & looks good from analysts view
j = missing sample

## Autosal-Data by Time
Entered MS Nov2011
Salinometer ID | Station | Sample ID (bottle label) | Duplicates | DateTime | Bath Temp | Uncorrected Ratio | Uncorrected Ratio StandDev | Original Correction | Original Adjusted Ratio | Original Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 |  | K15 |  | 2011-11-07 10:11 | 24C | 0.99984 | 4.1893e-06 | -4.5955e-05 | 0.99979 |  |  | fill rate too slow -stopped to check autosal tubing and change bung
Calibration#2 |  | K15 |  | 2011-11-07 10:17 | 24C | 0.9998 | 4.1682e-06 | -8.9407e-06 | 0.99979 |  |  | new bung replacement -new/proper flow rate=adjust standardize button to 5.68
SWL2011-18#1 | Slip-1 | 1 |  | 2011-11-07 11:03 | 24C | 0.93162 | 2.4918e-06 | -9.0003e-06 | 0.93161 | 32.3246 | 9.5561e-05
SWL2011-18#2 | Slip-1 | 2 |  | 2011-11-07 11:07 | 24C | 0.93164 | 3.6111e-06 | -9.0599e-06 | 0.93163 | 32.3254 | 0.00014188
SWL2011-18#3 | Slip-1 | 3 |  | 2011-11-07 11:09 | 24C | 0.91689 | 4.7141e-06 | -8.9407e-06 | 0.91688 | 31.7528 | 0.00018394
SWL2011-18#4 | Slip-1 | 4 |  | 2011-11-07 11:11 | 24C | 0.91362 | 4.3836e-06 | -9.0003e-06 | 0.91361 | 31.626 | 0.00016776
SWL2011-18#5 | Slip-1 | 5 |  | 2011-11-07 11:14 | 24C | 0.90938 | 4.4251e-06 | -9.0599e-06 | 0.90937 | 31.4621 | 0.00017322
SWL2011-18#6 | Slip-1 |  | 5 dup | 2011-11-07 11:15 | 24C | 0.90957 | 1.4046e-06 | -8.9407e-06 | 0.90956 | 31.4692 | 5.4857e-05
SWL2011-18#7 | Slip-1 | 6 |  | 2011-11-07 11:19 | 24C | 0.89134 | 3.6104e-06 | -8.9407e-06 | 0.89133 | 30.7645 | 0.00013994
SWL2011-18#8 | Slip-1 | 7 |  | 2011-11-07 11:21 | 24C | 0.88852 | 2.7425e-06 | -9.0003e-06 | 0.88851 | 30.6558 | 0.00010656
SWL2011-18#9 | Slip-2 | 8 |  | 2011-11-07 11:23 | 24C | 0.93067 | 4.4651e-06 | -9.0599e-06 | 0.93066 | 32.2877 | 0.00017522
SWL2011-18#10 | Slip-2 | 9 |  | 2011-11-07 11:26 | 24C | 0.9303 | 2.2392e-06 | -9.0599e-06 | 0.93029 | 32.2735 | 8.7009e-05
SWL2011-18#11 | Slip-2 | 10 |  | 2011-11-07 11:29 | 24C | 0.91717 | 4.1136e-06 | -9.0003e-06 | 0.91716 | 31.7639 | 0.00016142
SWL2011-18#12 | Slip-2 | 11 |  | 2011-11-07 11:31 | 24C | 0.91453 | 3.8497e-06 | -8.9407e-06 | 0.91452 | 31.6615 | 0.00014871
SWL2011-18#13 | Slip-2 | 12 |  | 2011-11-07 11:33 | 24C | 0.90995 | 3.9655e-06 | -9.0599e-06 | 0.90994 | 31.484 | 0.00015242
SWL2011-18#14 | Slip-2 |  | 12 dup | 2011-11-07 11:35 | 24C | 0.91006 | 3.999e-06 | -8.9407e-06 | 0.91005 | 31.4881 | 0.00015427
SWL2011-18#15 | Slip-2 | 13 |  | 2011-11-07 11:37 | 24C | 0.89392 | 5.1292e-06 | -8.9407e-06 | 0.89391 | 30.8643 | 0.00019704
SWL2011-18#16 | Slip-2 | 14 |  | 2011-11-07 11:39 | 24C | 0.88599 | 4.6142e-06 | -9.0599e-06 | 0.88598 | 30.5583 | 0.00017872
SWL2011-18#17 | Slip-3 | 15 |  | 2011-11-07 11:42 | 24C | 0.92375 | 3.5941e-06 | -8.9407e-06 | 0.92374 | 32.0188 | 0.00013703
SWL2011-18#18 | Slip-3 | 16 |  | 2011-11-07 11:44 | 24C | 0.92386 | 5.5967e-06 | -9.0003e-06 | 0.92385 | 32.0233 | 0.00021611
SWL2011-18#19 | Slip-3 | 17 |  | 2011-11-07 11:46 | 24C | 0.92401 | 3.1312e-06 | -9.0003e-06 | 0.924 | 32.029 | 0.00012419
SWL2011-18#20 | Slip-3 | 18 |  | 2011-11-07 11:48 | 24C | 0.91955 | 5.7797e-06 | -9.0003e-06 | 0.91955 | 31.8562 | 0.0002229
SWL2011-18#21 | Slip-3 | 19 |  | 2011-11-07 11:50 | 24C | 0.89793 | 7.6227e-06 | -9.0599e-06 | 0.89792 | 31.0192 | 0.00029181
SWL2011-18#22 | Slip-3 | 20 |  | 2011-11-07 11:52 | 24C | 0.89616 | 3.466e-06 | -9.0599e-06 | 0.89615 | 30.9507 | 0.00013157
SWL2011-18#23 | Slip-5 | 21 |  | 2011-11-07 11:55 | 24C | 0.92584 | 4.6189e-06 | -8.9407e-06 | 0.92583 | 32.1002 | 0.00017971
SWL2011-18#24 | Slip-5 | 22 |  | 2011-11-07 11:57 | 24C | 0.925 | 4.1137e-06 | -8.9407e-06 | 0.92499 | 32.0675 | 0.00016089
SWL2011-18#25 | CB18 | 1 R |  | 2011-11-07 11:59 | 24C | 0.99889 | 4.4633e-06 | -8.9407e-06 | 0.99888 | 34.9559 | 0.00017402 | DeepWaterReference
SWL2011-18#26 | Slip-5 | 23 |  | 2011-11-07 12:03 | 24C | 0.92203 | 2.2271e-06 | -9.0003e-06 | 0.92202 | 31.9521 | 8.7027e-05
SWL2011-18#27 | Slip-5 | 24 |  | 2011-11-07 12:06 | 24C | 0.92097 | 4.3232e-06 | -9.0003e-06 | 0.92096 | 31.9113 | 0.00016662
SWL2011-18#28 | Slip-5 | 25 |  | 2011-11-07 12:08 | 24C | 0.89364 | 6.2835e-06 | -9.0003e-06 | 0.89363 | 30.8536 | 0.00024157
SWL2011-18#29 | Slip-5 | 26 |  | 2011-11-07 12:10 | 24C | 0.89011 | 3.974e-06 | -8.9407e-06 | 0.8901 | 30.717 | 0.00015466
SWL2011-18#30 | Slip-4 | 27 |  | 2011-11-07 12:13 | 24C | 0.93948 | 6.1184e-06 | -9.0003e-06 | 0.93947 | 32.6306 | 0.00023691
SWL2011-18#31 | Slip-4 | 28 |  | 2011-11-07 12:15 | 24C | 0.93497 | 5.4781e-06 | -8.9407e-06 | 0.93496 | 32.4552 | 0.00021292
SWL2011-18#32 | Slip-4 | 30 |  | 2011-11-07 12:17 | 24C | 0.92553 | 2.4908e-06 | -9.0003e-06 | 0.92552 | 32.088 | 9.6599e-05
SWL2011-18#33 | Slip-4 | 32 |  | 2011-11-07 12:20 | 24C | 0.90923 | 4.0096e-06 | -9.0599e-06 | 0.90922 | 31.456 | 0.00015593
SWL2011-18#34 | Slip-4 | 33 |  | 2011-11-07 12:22 | 24C | 0.90659 | 5.8793e-06 | -8.9407e-06 | 0.90658 | 31.3539 | 0.00022756
SWL2011-18#35 | Slip-4 | 39 |  | 2011-11-07 12:24 | 24C | 0.90528 | 5.4535e-06 | -9.0599e-06 | 0.90527 | 31.3033 | 0.00021065
SWL2011-18#36 | BCL-6a | 40 |  | 2011-11-07 12:26 | 24C | 0.92599 | 3.6105e-06 | -9.0599e-06 | 0.92598 | 32.1061 | 0.00013925
SWL2011-18#37 | BCL-6a | 41 |  | 2011-11-07 12:28 | 24C | 0.92611 | 6.9588e-06 | -9.0599e-06 | 0.9261 | 32.1106 | 0.0002715
SWL2011-18#38 | BCL-6a | 42 |  | 2011-11-07 12:30 | 24C | 0.92583 | 4.4373e-06 | -9.0003e-06 | 0.92582 | 32.0997 | 0.0001706
SWL2011-18#39 | BCL-6a | 43 |  | 2011-11-07 12:32 | 24C | 0.91254 | 5.7678e-06 | -9.0599e-06 | 0.91253 | 31.5842 | 0.00022045
SWL2011-18#40 | BCL-6a | 44 |  | 2011-11-07 12:34 | 24C | 0.91303 | 4.3194e-06 | -9.0003e-06 | 0.91302 | 31.6032 | 0.00016796
SWL2011-18#41 | UTBS-5 | 50 |  | 2011-11-07 12:37 | 24C | 0.92351 | 1.8035e-06 | -8.9407e-06 | 0.9235 | 32.0098 | 6.9481e-05
SWL2011-18#42 | UTBS-5 | 51 |  | 2011-11-07 12:58 | 24C | 0.92337 | 4.6354e-06 | -9.0599e-06 | 0.92336 | 32.0041 | 0.00018037
SWL2011-18#43 | UTBS-5 | 52 |  | 2011-11-07 13:00 | 24C | 0.9231 | 3.1417e-06 | -8.9407e-06 | 0.92309 | 31.9937 | 0.00012138
SWL2011-18#44 | UTBS-5 | 53 |  | 2011-11-07 13:02 | 24C | 0.92331 | 7.7046e-06 | -9.0003e-06 | 0.9233 | 32.0019 | 0.00029785
SWL2011-18#45 | UTBS-5 | 54 |  | 2011-11-07 13:04 | 24C | 0.92147 | 1.4046e-06 | -9.0599e-06 | 0.92146 | 31.9307 | 5.2235e-05
SWL2011-18#46 | UTSB-1 | 55 |  | 2011-11-07 13:07 | 24C | 0.92541 | 3.1905e-06 | -9.0003e-06 | 0.9254 | 32.0835 | 0.00012094
SWL2011-18#47 | UTSB-1 | 56 |  | 2011-11-07 13:09 | 24C | 0.91829 | 5.3587e-06 | -9.0003e-06 | 0.91828 | 31.807 | 0.00020765
SWL2011-18#48 | UTSB-1 | 57 |  | 2011-11-07 13:11 | 24C | 0.9162 | 6.7813e-06 | -9.0599e-06 | 0.91619 | 31.7262 | 0.00026323
SWL2011-18#49 | UTSB-1 | 58 |  | 2011-11-07 13:13 | 24C | 0.91268 | 5.5701e-06 | -8.9407e-06 | 0.91267 | 31.5897 | 0.00021193
SWL2011-18#50 | CB18 | 2 R |  | 2011-11-07 13:16 | 24C | 0.99889 | 5.0935e-06 | -8.9407e-06 | 0.99888 | 34.9561 | 0.00019805
SWL2011-18#51 | UTSB-1 | 59 |  | 2011-11-07 13:19 | 24C | 0.91246 | 2.4751e-06 | -9.0003e-06 | 0.91246 | 31.5814 | 9.6591e-05
SWL2011-18#52 | UTSB-2 | 60 |  | 2011-11-07 13:21 | 24C | 0.91931 | 3.344e-06 | -9.0599e-06 | 0.9193 | 31.8469 | 0.00012993
SWL2011-18#53 | UTSB-2 | 61 |  | 2011-11-07 13:23 | 24C | 0.91918 | 3.996e-06 | -8.9407e-06 | 0.91918 | 31.8419 | 0.00015614
SWL2011-18#54 | UTSB-2 | 62 |  | 2011-11-07 13:26 | 24C | 0.91928 | 3.862e-06 | -8.9407e-06 | 0.91928 | 31.8458 | 0.00015214
SWL2011-18#55 | UTSB-2 | 63 |  | 2011-11-07 13:27 | 24C | 0.91839 | 3.9511e-06 | -9.0599e-06 | 0.91838 | 31.8111 | 0.00015294
SWL2011-18#56 | UTSB-2 | 64 |  | 2011-11-07 13:30 | 24C | 0.91626 | 2.2401e-06 | -9.0599e-06 | 0.91625 | 31.7285 | 8.5326e-05
SWL2011-18#57 | UTSB-4 | 65 |  | 2011-11-07 13:32 | 24C | 0.92321 | 4.0811e-06 | -9.0003e-06 | 0.92321 | 31.9982 | 0.00015933
SWL2011-18#58 | UTSB-4 | 66 |  | 2011-11-07 13:34 | 24C | 0.9228 | 4.9116e-06 | -9.0599e-06 | 0.92279 | 31.9821 | 0.00019099
SWL2011-18#59 | UTSB-4 | 67 |  | 2011-11-07 13:36 | 24C | 0.92191 | 5.5671e-06 | -8.9407e-06 | 0.9219 | 31.9475 | 0.00021728
SWL2011-18#60 | UTSB-4 | 68 |  | 2011-11-07 13:38 | 24C | 0.92165 | 4.4374e-06 | -8.9407e-06 | 0.92164 | 31.9374 | 0.00017322
SWL2011-18#61 | UTSB-4 | 69 |  | 2011-11-07 13:40 | 24C | 0.92048 | 3.3587e-06 | -9.0003e-06 | 0.92047 | 31.892 | 0.00012996
SWL2011-18#62 | BRS-3 | 70 |  | 2011-11-07 13:43 | 24C | 0.91834 | 5.7528e-06 | -9.0003e-06 | 0.91833 | 31.8091 | 0.00022295
SWL2011-18#63 | BRS-3 | 71 |  | 2011-11-07 13:45 | 24C | 0.91796 | 7.0607e-06 | -9.0599e-06 | 0.91795 | 31.7945 | 0.0002766
SWL2011-18#64 | BRS-3 | 72 |  | 2011-11-07 13:47 | 24C | 0.91725 | 5.348e-06 | -9.0599e-06 | 0.91724 | 31.767 | 0.00020791
SWL2011-18#65 | BRS-3 | 73 |  | 2011-11-07 13:52 | 24C | 0.91426 | 2.4915e-06 | -9.0003e-06 | 0.91425 | 31.6509 | 9.9629e-05
SWL2011-18#66 | BRS-3 | 74 |  | 2011-11-07 13:55 | 24C | 0.91388 | 3.3594e-06 | -9.0003e-06 | 0.91387 | 31.6363 | 0.00013249
SWL2011-18#67 | UTN-1 | 75 |  | 2011-11-07 13:57 | 24C | 0.89602 | 5.246e-06 | -9.0003e-06 | 0.89602 | 30.9455 | 0.00020372
SWL2011-18#68 | UTN-1 | 76 |  | 2011-11-07 13:59 | 24C | 0.89782 | 5.0935e-06 | -9.0599e-06 | 0.89781 | 31.0149 | 0.0001947
SWL2011-18#69 | UTN-1 | 77 |  | 2011-11-07 14:01 | 24C | 0.8928 | 3.4516e-06 | -8.9407e-06 | 0.89279 | 30.821 | 0.00013203
SWL2011-18#70 | UTN-1 | 78 |  | 2011-11-07 14:03 | 24C | 0.88769 | 3.8622e-06 | -8.9407e-06 | 0.88768 | 30.6239 | 0.00014674
SWL2011-18#71 | UTN-2 | 79 |  | 2011-11-07 14:06 | 24C | 0.91614 | 5.8911e-06 | -9.0599e-06 | 0.91613 | 31.7237 | 0.00022879
SWL2011-18#72 | UTN-2 | 80 |  | 2011-11-07 14:08 | 24C | 0.91629 | 1.3713e-06 | -9.0599e-06 | 0.91628 | 31.7298 | 5.1717e-05
SWL2011-18#73 | UTN-2 | 81 |  | 2011-11-07 14:10 | 24C | 0.91562 | 4.3579e-06 | -9.0003e-06 | 0.91562 | 31.7039 | 0.00016875
SWL2011-18#74 | UTN-2 | 82 |  | 2011-11-07 14:12 | 24C | 0.91352 | 4.3737e-06 | -8.9407e-06 | 0.91351 | 31.6222 | 0.00016996
SWL2011-18#75 | CB18 | 3 R |  | 2011-11-07 14:14 | 24C | 0.99888 | 4.7275e-06 | -9.0003e-06 | 0.99887 | 34.9554 | 0.00018504
SWL2011-18#76 | UTN-2 | 83 |  | 2011-11-07 14:16 | 24C | 0.91236 | 4.3336e-06 | -8.9407e-06 | 0.91236 | 31.5775 | 0.00016701
SWL2011-18#77 | UTN-3 | 84 |  | 2011-11-07 14:19 | 24C | 0.92505 | 4.6353e-06 | -9.0003e-06 | 0.92504 | 32.0694 | 0.00017835
SWL2011-18#78 | UTN-3 | 85 |  | 2011-11-07 14:21 | 24C | 0.92469 | 5.1607e-06 | -9.0003e-06 | 0.92469 | 32.0557 | 0.00020091
SWL2011-18#79 | UTN-3 | 86 |  | 2011-11-07 14:23 | 24C | 0.9252 | 4.189e-06 | -9.0003e-06 | 0.92519 | 32.0754 | 0.00016214
SWL2011-18#80 | UTN-3 | 87 |  | 2011-11-07 14:25 | 24C | 0.92396 | 4.1602e-06 | -9.0003e-06 | 0.92395 | 32.0272 | 0.00016102
SWL2011-18#81 | UTN-3 | 88 |  | 2011-11-07 14:27 | 24C | 0.92183 | 4.6601e-06 | -9.0599e-06 | 0.92182 | 31.9444 | 0.00017889
SWL2011-18#82 | UTN-4 | 89 |  | 2011-11-07 14:29 | 24C | 0.93634 | 5.3343e-06 | -8.9407e-06 | 0.93633 | 32.5082 | 0.00020727
SWL2011-18#83 | UTN-4 | 90 |  | 2011-11-07 14:32 | 24C | 0.93526 | 6.7174e-06 | -9.0599e-06 | 0.93525 | 32.4662 | 0.00026049
SWL2011-18#84 | UTN-4 | 91 |  | 2011-11-07 14:34 | 24C | 0.93368 | 5.4848e-06 | -8.9407e-06 | 0.93368 | 32.4051 | 0.00021313
SWL2011-18#85 | UTN-4 | 93 |  | 2011-11-07 14:36 | 24C | 0.93208 | 3.6104e-06 | -9.0599e-06 | 0.93207 | 32.3425 | 0.00013781
SWL2011-18#86 | UTN-4 | 97 |  | 2011-11-07 14:38 | 24C | 0.92284 | 3.2142e-06 | -8.9407e-06 | 0.92283 | 31.9839 | 0.00012413
SWL2011-18#87 | UTN-5 | 101 |  | 2011-11-07 14:41 | 24C | 0.93445 | 4.7277e-06 | -9.0003e-06 | 0.93444 | 32.4349 | 0.00018429
SWL2011-18#88 | UTN-5 | 102 |  | 2011-11-07 14:43 | 24C | 0.93393 | 2.1136e-06 | -9.0599e-06 | 0.93392 | 32.4147 | 8.1174e-05
SWL2011-18#89 | UTN-5 | 103 |  | 2011-11-07 14:45 | 24C | 0.93311 | 6.1184e-06 | -9.0003e-06 | 0.9331 | 32.3828 | 0.00023683
SWL2011-18#90 | UTN-5 | 104 |  | 2011-11-07 14:47 | 24C | 0.93246 | 3.2141e-06 | -9.0003e-06 | 0.93245 | 32.3575 | 0.00012382
SWL2011-18#91 | UTN-5 | 105 |  | 2011-11-07 14:49 | 24C | 0.9254 | 3.8546e-06 | -9.0599e-06 | 0.92539 | 32.083 | 0.00015067
SWL2011-18#92 | UTN-6 | 106 |  | 2011-11-07 14:51 | 24C | 0.92663 | 6.664e-08 | -9.0003e-06 | 0.92662 | 32.1309 | 0
SWL2011-18#93 | UTN-6 | 107 |  | 2011-11-07 14:53 | 24C | 0.92693 | 5.7086e-06 | -9.0599e-06 | 0.92692 | 32.1424 | 0.00022065
SWL2011-18#94 | UTN-6 | 108 |  | 2011-11-07 14:55 | 24C | 0.92463 | 4.3337e-06 | -9.0003e-06 | 0.92462 | 32.0533 | 0.00016797
SWL2011-18#95 | UTN-6 | 109 |  | 2011-11-07 14:57 | 24C | 0.92191 | 4.6353e-06 | -9.0003e-06 | 0.9219 | 31.9476 | 0.00017854
SWL2011-18#96 | UTN-6 | 110 |  | 2011-11-07 14:59 | 24C | 0.92099 | 4.4783e-06 | -9.0003e-06 | 0.92098 | 31.912 | 0.00017381
SWL2011-18#97 | SEC-2 | 111 |  | 2011-11-07 15:01 | 24C | 0.92853 | 4.1856e-06 | -8.9407e-06 | 0.92852 | 32.2048 | 0.00016593
SWL2011-18#98 | SEC-2 | 112 |  | 2011-11-07 15:03 | 24C | 0.92702 | 5.5967e-06 | -8.9407e-06 | 0.92701 | 32.146 | 0.00021675
SWL2011-18#99 | SEC-2 | 113 |  | 2011-11-07 15:05 | 24C | 0.92606 | 5.9928e-06 | -9.0599e-06 | 0.92605 | 32.1088 | 0.00023178
SWL2011-18#100 | CB18 | 4 R |  | 2011-11-07 15:07 | 24C | 0.99889 | 4.9004e-06 | -8.9407e-06 | 0.99888 | 34.9558 | 0.00019102
SWL2011-18#101 | SEC-2 | 114 |  | 2011-11-07 15:43 | 24C | 0.9237 | 4.5688e-06 | -9.0003e-06 | 0.92369 | 32.0171 | 0.00017845
SWL2011-18#102 | SEC-2 | 115 |  | 2011-11-07 15:45 | 24C | 0.92015 | 6.1982e-06 | -8.9407e-06 | 0.92014 | 31.8794 | 0.00024151
SWL2011-18#103 | UTN-7 | 116 |  | 2011-11-07 15:47 | 24C | 0.94166 | 4.3337e-06 | -9.0003e-06 | 0.94165 | 32.7154 | 0.00016344
SWL2011-18#104 | UTN-7 | 117 |  | 2011-11-07 15:49 | 24C | 0.94175 | 4.8368e-06 | -8.9407e-06 | 0.94174 | 32.7188 | 0.00018923
SWL2011-18#105 | UTN-7 | 118 |  | 2011-11-07 15:51 | 24C | 0.9411 | 4.2266e-06 | -9.0599e-06 | 0.94109 | 32.6936 | 0.00016697
SWL2011-18#106 | UTN-7 | 119 |  | 2011-11-07 15:53 | 24C | 0.9412 | 4.4254e-06 | -9.0003e-06 | 0.94119 | 32.6976 | 0.00016959
SWL2011-18#107 | UTN-7 | 120 |  | 2011-11-07 15:55 | 24C | 0.94036 | 3.0995e-06 | -8.9407e-06 | 0.94036 | 32.665 | 0.00011836
SWL2011-18#108 | UTN-7 | 121 |  | 2011-11-07 15:57 | 24C | 0.93983 | 2.8544e-06 | -8.9407e-06 | 0.93982 | 32.6441 | 0.0001126
SWL2011-18#109 | SEC-3 | 122 |  | 2011-11-07 16:00 | 24C | 0.92947 | 4.189e-06 | -9.0003e-06 | 0.92946 | 32.2414 | 0.00016375
SWL2011-18#110 | SEC-3 | 123 |  | 2011-11-07 16:02 | 24C | 0.92791 | 5.8469e-06 | -8.9407e-06 | 0.9279 | 32.1807 | 0.00022723
SWL2011-18#111 | SEC-3 | 124 |  | 2011-11-07 16:04 | 24C | 0.92772 | 3.2141e-06 | -9.0003e-06 | 0.92771 | 32.1731 | 0.00012429
SWL2011-18#112 | SEC-3 | 125 |  | 2011-11-07 16:06 | 24C | 0.92541 | 4.081e-06 | -9.0599e-06 | 0.9254 | 32.0835 | 0.00015463
SWL2011-18#113 | SEC-3 | 126 |  | 2011-11-07 16:08 | 24C | 0.92401 | 6.6079e-06 | -9.0003e-06 | 0.924 | 32.029 | 0.00026009
SWL2011-18#114 | SEC-4 | 127 |  | 2011-11-07 16:10 | 24C | 0.92585 | 6.2368e-06 | -9.0003e-06 | 0.92584 | 32.1005 | 0.00024341
SWL2011-18#115 | SEC-4 | 128 |  | 2011-11-07 16:13 | 24C | 0.92462 | 3.1319e-06 | -9.0003e-06 | 0.92461 | 32.0527 | 0.00012108
SWL2011-18#116 | SEC-4 | 129 |  | 2011-11-07 16:15 | 24C | 0.92222 | 5.1292e-06 | -9.0599e-06 | 0.92221 | 31.9596 | 0.00019704
SWL2011-18#117 | SEC-4 | 130 |  | 2011-11-07 16:17 | 24C | 0.9207 | 5.8623e-06 | -9.0003e-06 | 0.92069 | 31.9006 | 0.00022379
SWL2011-18#118 | SEC-4 | 131 |  | 2011-11-07 16:19 | 24C | 0.91966 | 4.452e-06 | -9.0599e-06 | 0.91965 | 31.8601 | 0.00017061
SWL2011-18#119 | SEC-5 | 132 |  | 2011-11-07 16:21 | 24C | 0.92029 | 4.1137e-06 | -9.0003e-06 | 0.92028 | 31.8848 | 0.0001622
SWL2011-18#120 | SEC-5 | 133 |  | 2011-11-07 16:23 | 24C | 0.92012 | 4.8368e-06 | -8.9407e-06 | 0.92011 | 31.8782 | 0.0001886
SWL2011-18#121 | SEC-5 | 134 |  | 2011-11-07 16:26 | 24C | 0.91995 | 7.1553e-06 | -8.9407e-06 | 0.91994 | 31.8714 | 0.00027592
SWL2011-18#122 | SEC-5 | 135 |  | 2011-11-07 16:29 | 24C | 0.9201 | 5.6049e-06 | -8.9407e-06 | 0.92009 | 31.8773 | 0.00021796
SWL2011-18#123 | SEC-5 | 136 |  | 2011-11-07 16:30 | 24C | 0.91872 | 3.8625e-06 | -9.0003e-06 | 0.91871 | 31.8239 | 0.00015183
SWL2011-18#124 | SEC-6 | 137 |  | 2011-11-07 16:33 | 24C | 0.91786 | 3.6104e-06 | -8.9407e-06 | 0.91786 | 31.7907 | 0.00014229
SWL2011-18#125 | CB18 | 5 R |  | 2011-11-07 16:35 | 24C | 0.99889 | 7.0228e-06 | -8.9407e-06 | 0.99888 | 34.956 | 0.00027324
Calibration#3 |  | K15 |  | 2011-11-08 09:11 | 24C | 0.9998 | 2.7425e-06 | -7.987e-06 | 0.99979 |  |  | perfect!
SWL2011-18#126 | SEC-6 | 138 |  | 2011-11-08 09:13 | 24C | 0.91805 | 2.2391e-06 | -7.987e-06 | 0.91805 | 31.7981 | 8.7441e-05
SWL2011-18#127 | SEC-6 | 139 |  | 2011-11-08 09:16 | 24C | 0.91708 | 2.2394e-06 | -8.0466e-06 | 0.91707 | 31.7601 | 8.617e-05
SWL2011-18#128 | SEC-6 | 140 |  | 2011-11-08 09:18 | 24C | 0.91631 | 2.2394e-06 | -7.8678e-06 | 0.9163 | 31.7305 | 8.7009e-05
SWL2011-18#129 | SEC-6 | 141 |  | 2011-11-08 09:20 | 24C | 0.91653 | 2.89e-06 | -7.987e-06 | 0.91652 | 31.7388 | 0.00011229
SWL2011-18#130 | SEC-7 | 142 |  | 2011-11-08 09:22 | 24C | 0.90774 | 1.3712e-06 | -8.0466e-06 | 0.90774 | 31.3986 | 5.328e-05
SWL2011-18#131 | SEC-7 | 143 |  | 2011-11-08 09:24 | 24C | 0.90864 | 2.7102e-06 | -7.9274e-06 | 0.90864 | 31.4335 | 0.00010656
SWL2011-18#132 | SEC-7 | 144 |  | 2011-11-08 09:26 | 24C | 0.90723 | 1.1063e-06 | -7.987e-06 | 0.90723 | 31.3789 | 4.3507e-05
SWL2011-18#133 | SEC-7 | 145 |  | 2011-11-08 09:29 | 24C | 0.90593 | 1.3713e-06 | -7.987e-06 | 0.90593 | 31.3286 | 5.2761e-05
SWL2011-18#134 | SEC-7 | 146 |  | 2011-11-08 09:31 | 24C | 0.88296 | 3.1418e-06 | -7.987e-06 | 0.88295 | 30.4415 | 0.00012259
SWL2011-18#135 | SEC-8 | 147 |  | 2011-11-08 09:33 | 24C | 0.90358 | 6.664e-08 | -8.0466e-06 | 0.90357 | 31.2374 | 2.1325e-06 | wrong stopper - poor seal - loose
SWL2011-18#136 | SEC-8 | 148 |  | 2011-11-08 09:35 | 24C | 0.90644 | 1.3712e-06 | -8.0466e-06 | 0.90643 | 31.3481 | 5.4346e-05 | wrong stopper - poor seal - loose
SWL2011-18#137 | SEC-8 | 149 |  | 2011-11-08 09:37 | 24C | 0.88725 | 1.1196e-06 | -8.1062e-06 | 0.88724 | 30.6071 | 4.4997e-05 | wrong stopper - poor seal - loose
SWL2011-18#138 | SEC-8 | 150 |  | 2011-11-08 09:39 | 24C | 0.76378 | 2.4908e-06 | -7.987e-06 | 0.76377 | 25.9011 | 9.4267e-05 | wrong stopper - poor seal - loose
SWL2011-18#139 | CS-2011 | 151 |  | 2011-11-08 09:41 | 24C | 0.90606 | 1.3713e-06 | -7.987e-06 | 0.90606 | 31.3336 | 5.2759e-05 | wrong stopper - poor seal - loose
SWL2011-18#140 | CS-2011 | 152 |  | 2011-11-08 09:45 | 24C | 0.90681 | 2.8321e-06 | -7.9274e-06 | 0.9068 | 31.3624 | 0.00010953 | wrong stopper - poor seal - loose
SWL2011-18#141 | CS-2011 | 153 |  | 2011-11-08 09:47 | 24C | 0.90545 | 1.1199e-06 | -7.987e-06 | 0.90544 | 31.31 | 4.4356e-05 | wrong stopper - poor seal - loose
SWL2011-18#142 | CS-2011 | 154 |  | 2011-11-08 09:50 | 24C | 0.89884 | 1.1529e-06 | -8.0466e-06 | 0.89883 | 31.0543 | 4.3095e-05 | wrong stopper - poor seal - loose //salt on rim
SWL2011-18#143 | BarC-10 | 155 |  | 2011-11-08 09:52 | 24C | 0.93479 | 4.4866e-06 | -8.0466e-06 | 0.93479 | 32.4482 | 0.00017291 | wrong stopper - poor seal - loose
SWL2011-18#144 | BarC-10 | 156 |  | 2011-11-08 09:55 | 24C | 0.93169 | 1.3712e-06 | -7.987e-06 | 0.93169 | 32.3277 | 5.2235e-05 | wrong stopper - poor seal - loose
SWL2011-18#145 | BarC-10 | 157 |  | 2011-11-08 09:57 | 24C | 0.91912 | 2.7425e-06 | -7.987e-06 | 0.91911 | 31.8394 | 0.0001071 | wrong stopper - poor seal - loose //salt on rim
SWL2011-18#146 | BarC-10 | 158 |  | 2011-11-08 09:59 | 24C | 0.89778 | 2.2391e-06 | -8.0466e-06 | 0.89777 | 31.0132 | 8.5734e-05 | wrong stopper - poor seal - loose
SWL2011-18#147 | BarC-10 | 159 |  | 2011-11-08 10:01 | 24C | 0.89464 | 3.1147e-06 | -7.987e-06 | 0.89463 | 30.8921 | 0.00012002 | wrong stopper - poor seal - loose
SWL2011-18#148 | BarC-10 | 160 |  | 2011-11-08 10:03 | 24C | 0.89316 | 2.4916e-06 | -7.987e-06 | 0.89315 | 30.835 | 9.638e-05 | wrong stopper - poor seal - loose
SWL2011-18#149 | BarC-9 | 161 |  | 2011-11-08 10:06 | 24C | 0.93243 | 2.8662e-06 | -7.9274e-06 | 0.93243 | 32.3564 | 0.00011084 | wrong stopper - poor seal - loose
SWL2011-18#150 | CB18 | 6 R |  | 2011-11-08 10:09 | 24C | 0.99889 | 4.8259e-06 | -8.0466e-06 | 0.99888 | 34.9558 | 0.00019085
SWL2011-18#151 | BarC-9 | 162 |  | 2011-11-08 10:27 | 24C | 0.92972 | 3.0413e-06 | -7.9274e-06 | 0.92971 | 32.2508 | 0.00011959 | wrong stopper - poor seal - loose
SWL2011-18#152 | BarC-9 | 163 |  | 2011-11-08 10:32 | 24C | 0.91472 | 2.4909e-06 | -7.9274e-06 | 0.91471 | 31.6689 | 9.669e-05 | wrong stopper - poor seal - loose
SWL2011-18#153 | BarC-9 | 164 |  | 2011-11-08 10:34 | 24C | 0.89897 | 3.2142e-06 | -7.987e-06 | 0.89897 | 31.0595 | 0.00012392 | wrong stopper - poor seal - loose
SWL2011-18#154 | BarC-9 | 165 |  | 2011-11-08 10:36 | 24C | 0.89539 | 2.2259e-06 | -7.987e-06 | 0.89538 | 30.9209 | 8.5747e-05 | wrong stopper - poor seal - loose
SWL2011-18#155 | BarC-9 | 166 |  | 2011-11-08 10:39 | 24C | 0.88188 | 2.7315e-06 | -7.9274e-06 | 0.88187 | 30.3999 | 0.00010552 | wrong stopper - poor seal - loose
SWL2011-18#156 | BarC-8 | 167 |  | 2011-11-08 10:42 | 24C | 0.93418 | 3.4657e-06 | -7.9274e-06 | 0.93417 | 32.4243 | 0.00013469
SWL2011-18#157 | BarC-8 | 168 |  | 2011-11-08 10:45 | 24C | 0.92913 | 2.4777e-06 | -7.987e-06 | 0.92913 | 32.2282 | 9.8074e-05 | wrong stopper - poor seal - loose
SWL2011-18#158 | BarC-8 | 169 |  | 2011-11-08 10:47 | 24C | 0.91641 | 1.3713e-06 | -7.987e-06 | 0.9164 | 31.7342 | 5.328e-05 | wrong stopper - poor seal - loose
SWL2011-18#159 | BarC-8 | 170 |  | 2011-11-08 10:49 | 24C | 0.90711 | 2.4744e-06 | -7.9274e-06 | 0.9071 | 31.374 | 9.5844e-05 | wrong stopper - poor seal - loose
SWL2011-18#160 | BarC-8 | 171 |  | 2011-11-08 10:51 | 24C | 0.89216 | 1.3713e-06 | -7.987e-06 | 0.89215 | 30.7962 | 5.3295e-05 | wrong stopper - poor seal - loose
SWL2011-18#161 | BarC-8 | 172 |  | 2011-11-08 10:54 | 24C | 0.8683 | 3.7068e-06 | -7.987e-06 | 0.86829 | 29.8774 | 0.00014217 | wrong stopper - poor seal - loose
SWL2011-18#162 | BarC-7 | 173 |  | 2011-11-08 10:57 | 24C | 0.94611 | 4.0096e-06 | -7.987e-06 | 0.9461 | 32.8888 | 0.00015676 | wrong stopper - poor seal - loose
SWL2011-18#163 | BarC-7 | 174 |  | 2011-11-08 10:59 | 24C | 0.93227 | 2.8901e-06 | -7.8678e-06 | 0.93226 | 32.3502 | 0.00011384
SWL2011-18#164 | BarC-7 | 175 |  | 2011-11-08 11:01 | 24C | 0.91987 | 1.1199e-06 | -8.0466e-06 | 0.91986 | 31.8685 | 4.3507e-05 | wrong stopper - poor seal - loose
SWL2011-18#165 | BarC-7 | 176 |  | 2011-11-08 11:03 | 24C | 0.92431 | 1.1196e-06 | -8.0466e-06 | 0.9243 | 32.0408 | 4.3536e-05 | NO INSERT
SWL2011-18#166 | BarC-7 | 177 |  | 2011-11-08 11:06 | 24C | 0.92282 | 2.7425e-06 | -7.987e-06 | 0.92282 | 31.9832 | 0.00010709 | NO INSERT
SWL2011-18#167 | BarC-7 | 178 |  | 2011-11-08 11:08 | 24C | 0.89022 | 1.3712e-06 | -7.987e-06 | 0.89021 | 30.7215 | 5.4346e-05 | NO INSERT
SWL2011-18#168 | BarC-6 | 179 |  | 2011-11-08 11:10 | 24C | 0.94713 | 2.4908e-06 | -7.987e-06 | 0.94712 | 32.9285 | 9.6822e-05
SWL2011-18#169 | BarC-6 | 180 |  | 2011-11-08 11:13 | 24C | 0.9461 | 1.4046e-06 | -7.987e-06 | 0.94609 | 32.8885 | 5.2235e-05
SWL2011-18#170 | BarC-6 | 181 |  | 2011-11-08 11:16 | 24C | 0.93567 | 1.1196e-06 | -7.9274e-06 | 0.93566 | 32.4823 | 4.3505e-05
SWL2011-18#171 | BarC-6 | 182 |  | 2011-11-08 11:18 | 24C | 0.93354 | 3.2144e-06 | -7.987e-06 | 0.93353 | 32.3994 | 0.00012416
SWL2011-18#172 | BarC-6 | 183 |  | 2011-11-08 11:21 | 24C | 0.93088 | 3.4659e-06 | -7.987e-06 | 0.93087 | 32.2959 | 0.00013597
SWL2011-18#173 | BarC-6 | 184 |  | 2011-11-08 11:23 | 24C | 0.93163 | 3.4658e-06 | -7.987e-06 | 0.93163 | 32.3253 | 0.00013376 | wrong stopper - poor seal - loose
SWL2011-18#174 | BarC-6 | 185 |  | 2011-11-08 11:26 | 24C | 0.91128 | 5.6799e-06 | -8.0466e-06 | 0.91127 | 31.5354 | 0.00021932 | wrong stopper - poor seal - loose -leaking
SWL2011-18#175 | CB18 | R7 |  | 2011-11-08 11:29 | 24C | 0.99886 | 2.8897e-06 | -7.9274e-06 | 0.99885 | 34.9548 | 0.00011297
SWL2011-18#176 | BarC-5 | 186 |  | 2011-11-08 11:32 | 24C | 0.94781 | 3.1414e-06 | -7.987e-06 | 0.9478 | 32.9551 | 0.00012143 | wrong stopper - poor seal - loose -leaking
SWL2011-18#177 | BarC-5 | 187 |  | 2011-11-08 11:34 | 24C | 0.94823 | 2.2259e-06 | -7.987e-06 | 0.94823 | 32.9717 | 8.7009e-05 | wrong stopper - poor seal - loose -leaking
SWL2011-18#178 | BarC-5 | 188 |  | 2011-11-08 11:36 | 24C | 0.94677 | 4.0093e-06 | -7.987e-06 | 0.94677 | 32.9147 | 0.00015552 | wrong stopper - poor seal - loose -leaking
SWL2011-18#179 | BarC-5 | 189 |  | 2011-11-08 11:39 | 24C | 0.94426 | 2.7434e-06 | -7.9274e-06 | 0.94425 | 32.8168 | 0.00010658 | wrong stopper - poor seal - loose -leaking
SWL2011-18#180 | BarC-5 | 190 |  | 2011-11-08 11:42 | 24C | 0.93841 | 4.4617e-06 | -8.0466e-06 | 0.9384 | 32.5888 | 0.00017225 | wrong stopper - poor seal - loose -leaking
SWL2011-18#181 | BarC-5 | 191 |  | 2011-11-08 11:45 | 24C | 0.94081 | 1.1196e-06 | -8.1062e-06 | 0.9408 | 32.6823 | 4.265e-05 | NO INSERT
SWL2011-18#182 | BarC-5 | 192 |  | 2011-11-08 11:47 | 24C | 0.94018 | 5.4721e-06 | -7.9274e-06 | 0.94018 | 32.658 | 0.00021138 | NO INSERT
SWL2011-18#183 | BarC-5 | 193 |  | 2011-11-08 11:50 | 24C | 0.93398 | 3.4554e-06 | -7.987e-06 | 0.93397 | 32.4165 | 0.00013705 | NO INSERT
SWL2011-18#184 | BarC-4 | 194 |  | 2011-11-08 11:52 | 24C | 0.9423 | 2.4908e-06 | -8.0466e-06 | 0.9423 | 32.7405 | 0.00010062 | wrong stopper - poor seal - loose
SWL2011-18#185 | BarC-4 | 195 |  | 2011-11-08 11:55 | 24C | 0.93977 | 2.7425e-06 | -7.987e-06 | 0.93976 | 32.6419 | 0.00010656
SWL2011-18#186 | BarC-4 | 196 |  | 2011-11-08 11:57 | 24C | 0.93856 | 3.4661e-06 | -7.987e-06 | 0.93855 | 32.5948 | 0.0001355 | wrong stopper - poor seal - loose
SWL2011-18#187 | BarC-4 | 197 |  | 2011-11-08 12:00 | 24C | 0.93714 | 2.2391e-06 | -7.9274e-06 | 0.93714 | 32.5397 | 8.5406e-05
SWL2011-18#188 | BarC-4 | 198 |  | 2011-11-08 12:02 | 24C | 0.93627 | 5.5978e-06 | -7.987e-06 | 0.93626 | 32.5058 | 0.00021933
SWL2011-18#189 | BarC-4 | 199 |  | 2011-11-08 12:05 | 24C | 0.93596 | 2.7425e-06 | -7.987e-06 | 0.93595 | 32.4936 | 0.00010658
SWL2011-18#190 | BarC-4 | 200 |  | 2011-11-08 12:08 | 24C | 0.93627 | 2.4915e-06 | -7.987e-06 | 0.93627 | 32.5058 | 9.87e-05
SWL2011-18#191 | BarC-4 | 201 |  | 2011-11-08 12:10 | 24C | 0.93389 | 3.2142e-06 | -7.987e-06 | 0.93388 | 32.413 | 0.00012404
SWL2011-18#192 | BarC-3 | 202 |  | 2011-11-08 12:12 | 24C | 0.93441 | 2.2398e-06 | -7.987e-06 | 0.93441 | 32.4335 | 8.7868e-05
SWL2011-18#193 | BarC-3 | 203 |  | 2011-11-08 12:15 | 24C | 0.93432 | 2.2268e-06 | -7.9274e-06 | 0.93431 | 32.4299 | 8.5352e-05
SWL2011-18#194 | BarC-3 | 204 |  | 2011-11-08 12:17 | 24C | 0.93423 | 1.3549e-06 | -7.987e-06 | 0.93422 | 32.4262 | 5.2235e-05
SWL2011-18#195 | BarC-3 | 205 |  | 2011-11-08 12:19 | 24C | 0.93419 | 1.8038e-06 | -7.8678e-06 | 0.93419 | 32.4249 | 6.8797e-05
SWL2011-18#196 | BarC-3 | 206 |  | 2011-11-08 12:21 | 24C | 0.93498 | 1.3713e-06 | -7.987e-06 | 0.93497 | 32.4554 | 5.4326e-05 | wrong stopper - poor seal - loose
SWL2011-18#197 | BarC-3 | 207 |  | 2011-11-08 12:23 | 24C | 0.93033 | 2.4776e-06 | -7.987e-06 | 0.93032 | 32.2745 | 9.5936e-05
SWL2011-18#198 | BarC-3 | 208 |  | 2011-11-08 12:25 | 24C | 0.90658 | 2.4776e-06 | -7.9274e-06 | 0.90658 | 31.3537 | 9.5846e-05
SWL2011-18#199 | BarC-2 | 209 |  | 2011-11-08 12:27 | 24C | 0.92889 | 1.1203e-06 | -8.0466e-06 | 0.92888 | 32.2185 | 4.2703e-05
SWL2011-18#200 | CB18 | R8 |  | 2011-11-08 12:29 | 24C | 0.99888 | 3.5298e-06 | -7.987e-06 | 0.99887 | 34.9556 | 0.00013893
SWL2011-18#201 | BarC-2 | 210 |  | 2011-11-08 12:32 | 24C | 0.92922 | 3.4657e-06 | -7.987e-06 | 0.92921 | 32.2314 | 0.00013559
SWL2011-18#202 | BarC-2 | 211 |  | 2011-11-08 12:35 | 24C | 0.92378 | 3.115e-06 | -7.987e-06 | 0.92377 | 32.0203 | 0.00012037
SWL2011-18#203 | BarC-2 | 212 |  | 2011-11-08 12:38 | 24C | 0.90266 | 2.2394e-06 | -7.987e-06 | 0.90265 | 31.2018 | 8.6581e-05
SWL2011-18#204 | BarC-1 | 213 |  | 2011-11-08 12:40 | 24C | 0.91958 | 2.4751e-06 | -7.987e-06 | 0.91957 | 31.8572 | 9.4459e-05 | wrong stopper - poor seal - loose
SWL2011-18#205 | BarC-1 | 214 |  | 2011-11-08 12:42 | 24C | 0.92002 | 0 | -7.987e-06 | 0.92001 | 31.8744 | 1.0662e-06
SWL2011-18#206 | BarC-1 | 215 |  | 2011-11-08 12:45 | 24C | 0.91643 | 2.4915e-06 | -8.0466e-06 | 0.91642 | 31.7351 | 9.6692e-05 | wrong stopper - poor seal - loose
SWL2011-18#207 | BarC-1 | 216 |  | 2011-11-08 12:48 | 24C | 0.90149 | 1.3882e-06 | -7.987e-06 | 0.90148 | 31.1567 | 5.2781e-05
| Samples of Mary O'Briens | [10 empty] | Written on Label
SWL2011-18#208 | VAG 1 | 1 |  | 2011-11-08 12:51 | 24C | 0.82693 | 2.7426e-06 | -8.0466e-06 | 0.82692 | 28.2943 | 0.00010345 | (Sfc) Littleton Island 2011-64-0001
SWL2011-18#209 | SSN01 | 2 |  | 2011-11-08 12:54 | 24C | 0.80609 | 2.4911e-06 | -7.9274e-06 | 0.80608 | 27.5011 | 9.4377e-05 | Surface 2011-64-0002
SWL2011-18#210 | SSN05 | 3 |  | 2011-11-08 12:55 | 24C | 0.8291 | 2.7425e-06 | -7.987e-06 | 0.82909 | 28.3768 | 0.00010552 | Surface 2011-64-0003
SWL2011-18#211 | BG04 | 4 |  | 2011-11-08 12:57 | 24C | 0.84378 | 1.1199e-06 | -7.9274e-06 | 0.84377 | 28.9374 | 4.4571e-05 | CTD's 1500-1000 Surface 2011-64-0004
SWL2011-18#212 | BG10 | 5 |  | 2011-11-08 12:59 | 24C | 0.85912 | 4.2256e-06 | -7.987e-06 | 0.85911 | 29.5252 | 0.00016269 | Belcher Glacier Surface CTD 1000 2011-64-0005
SWL2011-18#213 | JS | 6 |  | 2011-11-08 13:01 | 24C | 0.84616 | 2.4911e-06 | -7.987e-06 | 0.84616 | 29.0288 | 9.4909e-05 | Calib Surface 2011-64-0006

## Timeline Summary
Entered MS Nov2011
Summary of Analysis - One row per Cage
SWL 2011-18 Salts |  |  |  | Starting Standby # |  | Temperature
CASE | Date (d-mon-yr) | Station | Samples Analysed | Starting | Ending | Room | Sample | Comments
1 | 2011-11-07 | Slip-1, Slip-2, Slip-3, Slip-5 | 1-22 | 24+5996 | 24+5596 | 22.3 | 22
2 |  | Slip-5, Slip-4, BCL-6a, UTBS-5, UTBS-1 | 23-58 | 24+5996 | 24+5596 | 23 | 20
3 |  | UTBS-1, UTBS-2, UTBS-4, BRS-3, UTN-1, UTN-2 | 59-82 | 24+5996 | 24+5596 | 23 | 22
4 |  | UTN-2, UTN-3, UTN-4, UTN-5, UTN-6, SEC-2 | 83-113 | 24+5996 | 24+5596 | 23 | 21.2
5 |  | SEC-2, UNT-7, SEC-3, SEC-4, SEC-5, SEC-6 | 114-137 | 24+5996 | 24+5596 | 21 | 22.4
6 | 2011-11-08 | SEC-6, SEC-7, SEC-8, CS-2011, BarC-10, BarC-9 | 138-161 | 24+5995 | 24+5595 | 22.6 | 22.4
7 |  | BarC-9, BarC-8, BarC-7, BarC-6 | 162-185 | 24+5995 | 24+5595 | 23.1 | 22.4
8 |  | BarC-5, BarC-4, BarC-3, BarC-2 | 186-209 | 24+5995 | 24+5595 | 23.4 | 21
9 |  | BarC-2, BarC-1, | 210-216 | 24+5995 | 24+5595 | 23.2 | 21.9
10 |  | VAG1, SSN01, SSN05, BG04, BG10, J.S. | 001-006 | 24+5995 | 24+5595 | 23.4 | 21 | Mary O'Briens samples

## Cal Records
Entered MS Nov2011
SWL 2011-18 Salinity
|  | Standby # |  |  |  |  | Uncorrected Ratio | Correction | Adjusted Ratio | IAPSO Standard Seawater Bottle Info.
Date/Time | Sample ID | Read Value @ Beginning of Daily Run | Read Value @ End of Daily Run | Read Value @ Start Calibration | Read Value @ End Calibration | Zero |  |  |  | K15 | Batch # | Batch Date | Salinity | Comments
2011-11-07 10:11 | Calibration#1 | 24+5999 | ---- |  |  | 1e-05 | 0.99984 | -4.5955e-05 | 0.99979 | 0.99979 | P153 | 2014-03-08 | 34.999 | fill rate too slow -stopped to check autosal tubing and change bung
2011-11-07 10:17 | Calibration#2 | 24+5996 | 24+5996 | 24+5999 | 24+5996 | 1e-05 | 0.9998 | -8.9407e-06 | 0.99979 | 0.99979 | P153 | 2014-03-08 | 34.999 | new bung replacement -new/proper flow rate=adjust standardize knob to 5.68
2011-11-08 09:11 | Calibration#3 | 24+5995 | 24+5995 |  |  | 1e-05 | 0.9998 | -7.987e-06 | 0.99979 | 0.99979 | P153 | 2014-03-08 | 34.999 | perfect!

## DeepWaterCheck Records
entered: MS Nov2011
A Deepwater sample was analysed after ever 24 samples analysed for SWL 2011-18 to check for drift in the Autosalinometer
Deep Water Check (water from LSSL cruise)
Cast No. | Station | DWS ID | Sample mix | DateTime | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments | Average | STDEV
n/a | CB18 | 2011-20 | n/a | 2011-11-07 11:59 | 0.99889 | 4.4633e-06 | -8.9407e-06 | 0.99888 | 34.9559 | 0.00017402 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-07 13:16 | 0.99889 | 5.0935e-06 | -8.9407e-06 | 0.99888 | 34.9561 | 0.00019805 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-07 14:14 | 0.99888 | 4.7275e-06 | -9.0003e-06 | 0.99887 | 34.9554 | 0.00018504 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-07 15:07 | 0.99889 | 4.9004e-06 | -8.9407e-06 | 0.99888 | 34.9558 | 0.00019102 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-07 16:35 | 0.99889 | 7.0228e-06 | -8.9407e-06 | 0.99888 | 34.956 | 0.00027324 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-08 10:09 | 0.99889 | 4.8259e-06 | -8.0466e-06 | 0.99888 | 34.9558 | 0.00019085 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-08 11:29 | 0.99886 | 2.8897e-06 | -7.9274e-06 | 0.99885 | 34.9548 | 0.00011297 | DeepWaterReference
" | CB18 | 2011-20 | '' | 2011-11-08 12:29 | 0.99888 | 3.5298e-06 | -7.987e-06 | 0.99887 | 34.9556 | 0.00013893 | DeepWaterReference | 34.9557 | 0.00039874 | 8
Note (SZ): A summary of this year's check water from the deep Canada Basin
Cruise | Number of Obs | Average | STD | Station Deep Water | Analysis
2011-20 JOIS | 33 | 34.9558 | 0.0009803 | CB5-CW | Arctic Autosal during 2011-20
2011-21 UNCLOS | 9 | 34.9569 | 0.00078441 | CB5-CW | Arctic Autosal during 2011-21
2011-18 C3O | 8 | 34.9557 | 0.00039874 | CB18 | Water Prop Autosal at IOS

## Sp
Example data in pale yellow
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded
Cast No. | Station | Bottle | Pressure | Sample ID | Salt-1 | Salt-2 | diff | diff^2 | Chauvenet | Accept or Reject |  | QA/QC - Analyst | QA/QC - Data | mean diff | sd diff
|  |  |  | 6 | 34.7256 | 34.7255 | 8.0109e-05 | 6.4174e-09 | 1.0999 | ACCEPT |  |  |  | 0.00050735 | 0.00038844
|  |  |  | 11 | 34.3057 | 34.3064 | 0.00060272 | 3.6327e-07 | 0.24551 | ACCEPT
|  |  |  | 16 | 32.6789 | 32.6781 | 0.00083923 | 7.0431e-07 | 0.85438 | ACCEPT
Did not use this page
Mary Steel
0
SAL ALL
[11 empty] | N tot |  | 3
[11 empty] | N used |  | 3
[11 empty] | sum of sqrs |  | 0
0
[11 empty] | Sp |  | 0
[11 empty] | N discard |  | 0
Critical values for Chauvenet statistic
[11 empty] | n | Chauvenet
[11 empty] | 3 | 1.37
[11 empty] | 4 | 1.53
[11 empty] | 5 | 1.64
[11 empty] | 6 | 1.73
[11 empty] | 7 | 1.8
[11 empty] | 8 | 1.87
[11 empty] | 9 | 1.91
[11 empty] | 10 | 1.97
[11 empty] | 20 | 2.24
[11 empty] | 30 | 2.39
[11 empty] | 50 | 2.57
[11 empty] | 80 | 2.74
[11 empty] | 100 | 2.81
[11 empty] | 150 | 2.93
[11 empty] | 200 | 3.02
Pooled standard deviation of pairs of samples (Sp) was calculated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs

## Manual Salt Calc
entered: MS Nov2011
Worksheet in case any sample needs to be recalcualted
Sample # | Bath temp | Pressure | M | Autosal Display of Cond. Ratio (Flush 1) | Autosal Display of Cond. Ratio (Flush 2) | Autosal Display of Cond. Ratio (Flush 3) | Autosal Display of Cond. Ratio (Flush 4) | Average Conductivity Ratio | Correction to Ratio | Salinity
223 | 24 | 0 | 2 | 1.9596 | 1.9597 | 1.9598 |  | 1.9597 | -9.5367e-07 | #VALUE!
1480 | 24 | 0 | ?? | 1.9957 | 1.9958 | 1.9959 |  | 1.9958 |  | 34.9168
Did not use this to calulate salinity
Mary Steel