## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2011-18

Agency: OSD

Location: North Pacific	/ Bering Sea / Chukchi Sea

Project: C3O/IPY

Party Chief: Vagle S.

Platform: Sir Wilfrid Laurier

Date: June 24, 2010 - July 21, 2010

Processed by: Germaine Gatien

Date of Processing: 17 October 2014 – 7 March 2016

Number of original HEX files: 38 (including 1 test file)

Number of CTD files:  37	Number of bottle casts:	36

# INSTRUMENT SUMMARY

A Seabird SBE-9 profiling CTD was used with a custom built compact 24-bottle rosette water sampler which was equipped with a Seabird Carousel pylon to remotely trigger the 10-litre sample bottles. An SBE-11 deck unit was used with Seasave software to acquire real-time data from the CTD and to close the bottles at depths selected before and/or during the cast. The deck unit included a NMEA board to automatically add GPS position into the header of the data file. The CTD was mounted in the rosette and attached to it were a SeaTech transmissometer (#1050-DR), an SBE 43 DO sensor (#1117), a SeaPoint Fluorometer (#2745), a Biospherical PAR sensor (#4601) and an altimeter (#40853).  A Reference PAR sensor (#20281) was entered in the configuration file, but there was no signal recorded.

A thermosalinograph (SeaBird 21 S/N 3297) was mounted with a WetLabs WetStar fluorometer (S/N WS3S-376P) and a WetLabs ECO CDOM fluorometer (S/N WSCD-1281) and a remote temperature sensor # 0319.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572. Salinity samples were analyzed in November 2011.

# SUMMARY OF QUALITY AND CONCERNS

Photocopies of the Daily Science Log Book sheets were obtained but these did not include the equipment page or personnel list. A cruise report was available with some of that information. The event numbers are wrong in the Daily Science Log but the file names and rosette log sheets are correct.

Salinity samples were analyzed about 3½ months after the end of the cruise. The comparison of bottles with CTD salinity shows a lot of scatter. Incomplete flushing of Niskin bottles likely explains much of the scatter. Evaporation and/or adsorption of samples are other possible explanations, but those would have relatively small effects for most of the cruise. However, late in the cruise many of the bottles were flagged because of poor seals so that evaporation may be a bigger problem; those casts were not included in the main comparison. The comparison suggests that the secondary CTD salinity channel was reading low by 0.005. But the post-cruise calibration shows the secondary salinity reading low by about 0.002. The CTD salinity was recalibrated based on the post-cruise calibration.

There are two sets of nutrient data. The first set with names in the usual format were analyzed at the Institute of Ocean Sciences while a second set with the number 2 at the end, example Phosphate2, were analyzed at the University of Maryland Center for Environmental Science.

A SeaPoint fluorometer was mounted on the CTD for this cruise. As is usually noted for these sensors the CTD fluorescence is close to extracted CHL for CHL<1ug/L though there is a lot of scatter in the fit. At higher CHL values the fluorescence reads lower.

There was no dissolved oxygen sampling, so the error due to slow response time cannot be assessed. No attempt was made to estimate the error due to calibration drift, since the post-cruise check found several problems with the sensor, and we have no way of knowing when those occurred. There was no open ocean sampling, so the DO surface saturation is not very helpful in judging the quality of the DO data. No recalibration was applied and the data should be considered nominal and results are reported with fewer significant figures than usual.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Files had non-standard names, so those were fixed before conversion.

##### Preliminary Steps

Photocopies of the Daily Science Log Book sheets and the rosette log sheets were obtained but this did not include the equipment page or personnel list. A cruise report was available with some of that information. The event numbers are wrong in the Daily Science Log but the file names and rosette log are correct.

Extracted chlorophyll, nutrients, salinity, Ammonium , DIC and Alkalinity data were obtained in the chemistry spreadsheet. (Some DIC data were missing; these can be added later if found.)

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. The temperature and conductivity sensors had been recalibrated in late January 2011 and there were post-cruise calibrations from December 2011. The dissolved oxygen sensor was recalibrated in December 2010 and November 2012. There were errors in the PAR and transmissivity parameters; those were corrected. The pressure sensor offset was changed from the factory setting but this was a deliberate decision at sea, so that value will be used for conversion and tests will be done later to see if that was appropriate. The pre-cruise calibrations were chosen for all sensors. The corrected configuration file was saved as 2011-18-ctd.xmlcon.

In processing cruise 2010-05 an attempt was made to estimate dissolved oxygen calibration drift based on converting and then comparing some files using both pre-cruise and post-cruise calibrations. That study did not provide clear enough evidence to justify a recalibration due to a high noise level that was likely due to the temperature-dependence of the calibration drift. Tests were not attempted for this cruise as the casts are all shallow with very complex temperature profiles.

All files were converted using file 2011-18-ctd.xmlcon. The hysteresis factor was not applied.

Voltage channels were converted for Transmissivity, SeaPoint Fluorescence, altimeter and PAR channels. Plots were made to ensure that the voltage channels and the channels derived from them have the same shapes and they do.

All channels were plotted for a few casts. The temperature and conductivity channels track well on downcasts but are much noisier during upcasts. Altimetry looks useful. The transmissivity, PAR, fluorescence and dissolved oxygen have the right shapes. Descent rates are mostly high and steady. There was no SPAR signal, so conversion was rerun without it.

##### Hysteresis Study

The nominal hysteresis factor was applied but there was no deep sampling, so this is of no importance.

##### BOTTLE FILE PREPARATION

The HEX files were converted in the same way as the full files except that bottle number and bottle position channels were included and SBE oxygen concentration and salinity were derived.

The files were then converted to IOS Header format. The resulting files either lacked station names (file #1 had a name but it was incorrect and there was no sampling from that cast.)

To add the missing headers a CSV file was prepared with headers: FileName (without extension), LOC:Water Depth and LOC:Station. The information was found from a chemistry spreadsheet and entered in the CSV file. Then program “MERGE: CSV files to Headers” was run.

The output files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files. There were a spiky points in both salinity channels, more in the primary than secondary; these were mostly in shallow water, so may not be very significant, but CTDEDIT was used to remove such spikes from casts #9, 13, 20, 22, 28 and 34. The edited files were copied to BOT.

A preliminary header check turned up no problems.

The BOT files were averaged on bottle number to enable preparation of file ADDSAMP.csv. That file was edited to add sample numbers. Event #1 was a test cast; all bottles were fired but none were sampled, so it will not be processed further.

The ADDSAMP file was then used to add sample numbers to the BOT files, creating SAM files. Those were bin-averaged on bottle number to create SAMAVGSTD files and standard deviations were included. Depth, theta, sigma-T and sigma-theta were derived for those files, the Salinity:T1:C1 channel was recalibrated and the files were saved as *.BOTSTD. Those files will be stored for reference, but not used for archiving.

The binning step was repeated without standard deviations included and extension SAMAVG.

All bottle sample data were presented in a single spreadsheet, 2011-18-SWL_Chem.xls. This was simplified to include only the columns needed for addition to the BOT files and saved as 2011-18-chem.csv.

File 2011-18-chem.csv was then converted to individual MRG1 files. Those were put through CLEAN and SORT (on bottle #) and then merged with the SAMAVG files. These files were put through CLEAN to remove SeaBird headers and comments and to remove channels with no data values.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was a lot of scatter in the plot of (CTD SAL – BOTTLE SAL) versus pressure for both sensors. This is not surprising given the shallow sampling and likelihood of either poor flushing and/or evaporation of samples. Analysis was completed on 7 November 2011, 3.5 months after the end of the cruise. Some evaporation might be expected, but not enough to explain the differences seen above 50m. The major source of trouble is more likely poor flushing in an area of high gradients. When a plot is made against file number, the greatest scatter is for files 24 to 37. Those were near-shore casts where gradients tend to be very high near the surface, and for casts 27 to 37 there were many flags on bottles due to the wrong stoppers being used or missing inserts.

If we exclude casts #24-37 we still have 8 outliers with differences between bottles and CTD Sal0 >0.05 (and 7 outliers for Sal1.) Those large outliers which were investigated:

Cast #2 – 15m – High temperature and salinity gradient at this level;

Cast #3 – 15m – High temperature and salinity gradient at this level

Cast #5 – 15m – High temperature and salinity gradient at this level

Cast #7 – 15m – High temperature and salinity gradient at this level

Cast # 9 – 43m(bottom) – Quite well mixed near surface (outlier in opposite direction to most outliers)

Cast # 13 – 30m (bottom) – Fairly well mixed near surface (outlier in opposite direction to most outliers)

Cast #17 – 5m – Sal0 outlier, but not Sal1. CTD Sal0 standard deviation high

Cast #25 – 5m – Sal0 outlier, but not Sal1. Surface gradient fairly high.

The two bottom sample outliers have differences of opposite sign to the shallow outliers. Poor flushing would have that effect on bottom bottles.

Excluding those outliers and using only files 2-25 the primary salinity is found to be low by 0.0037 and the secondary by 0.0050. The difference between these two estimates is 0.0013, but the standard deviations in both data sets was >0.01. For cases where the standard deviation in the CTD salinity is <0.001, the primary is low by 0.0035 and the secondary low by 0.0058 for a difference of 0.0023. When the sensors were recalibrated after the cruise the primary was found to be high by 0.0005 and the secondary low by ~0.002, so the difference between the post-cruise calibration differences is 0.0025. The post-cruise factory calibration shows little change in the primary. If the secondary salinity is selected for archiving it will be necessary to recalibrate.

None of the bottles need flagging since the problems are not with analysis or sampling, but a warning should go into the headers to indicate that there is clear evidence that Niskin bottles did not flush completely so that water likely comes from deeper than the bottle firing depth except at the bottom where water likely comes from shallower levels.

A study was made of casts #26 to 37 because there were many bottles flagged 3 or 5 due problems noted with how they were sealed. File 2018-11-sal-comp-flag-study.xls shows that most of the unflagged samples have small differences from the CTD. Exceptions are one bottom bottle with a high standard deviation in the CTD data and some bottles above 20m. It is impossible to separate the effects of poor flushing from evaporation of samples. It is notable that the cases flagged 5 look better than many of those flagged 3, but the “5” flags are all from near the surface from casts with fairly low vertical salinity gradients so that flushing effects would be small. The situation is too complex to analyze. The flags should be kept, but the 5 flags were changed to 4 so that the values will not have to be replaced with pad values.

COMPARE was rerun with recalibrated SAM files to ensure the correction was applied correctly and it was. The results show that the secondary salinity now compares well with the primary salinity, though both are lower than the bottles.

For full details for the COMPARE run see file 2011-18-sal-comp1.xls and 2011-18-sal-compo2.xls.

Fluorescence

COMPARE was run using the SeaPoint Fluorescence and the Extracted Chlorophyll from bottles. As usual the fit is poor for very low chlorophyll where small differences may lead to very different ratios. For high CHL values the ratio varies from 0.5 to 1. The problems with Niskin bottle flushing discussed in the salinity comparison means that the CHL values come from deeper than the FL, and since the depth of the CHL maximum varies greatly, there is a lot of variability in the results. We can conclude only that the general shape of the fit is reasonable and the fluorometer appears to have performed properly.

For more detail see file 2011-18-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-18-ctd.xmlcon.

A few casts were examined and all expected channels are present.

The descent rate is mostly steady and high on downcasts.

The two temperature and conductivity channels are fairly close during the downcasts though both channels have occasional spikes. During the upcasts traces differ more and there are odd excursions that are often seen in upcasts. This is likely something to do with how the CTD is mounted. The conductivity channels are similar to temperature.

PAR, Transmissivity and Fluorescence look normal. Altimetry looks useful.

The dissolved oxygen voltage is hard to judge with so many bottle stops.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage channel and the primary temperature. It is always hard to judge this when there are many stops for bottles, but these data have a very unusual character. Estimates were made based on the offset of the DO and temperature traces taking into account the descent rate. The best results appeared to be with 1 to 2s which was surprising based on previous results with this sensor. When those values were applied and Dissolved Oxygen was derived in mL/L, the results were very puzzling with values almost constant except for spikes near the maximum gradient. The calibration history of the sensor was checked and it was found that major repairs were made:

Slow dry tau response found. Repair:  3. Installed New 4-pin wetplug bulkhead connector.  4. Replaced the lid and membrane assembly (0.5 mil). Replaced the electrolyte with new fluid.

To see if the problem with the DO data is because the TAU correction was used in deriving dissolved oxygen, DERIVE was run without that correction. The results are slightly different, but not much. If DO is quite well mixed from top to bottom, then the results are ok, but we have no DO sampling to confirm this. Comparisons were made with some 2010 casts at the same sites and values were different, but not in a consistent way. For now the DO will be derived with TAU correction. Later, the surface saturation will be examined and a decision made about whether to archive these data. At the very least, data should be reported with fewer significant figures given the lack of calibration sampling and the problem noted at the factory with sensor response.

ALIGNCTD was run with setting +2s. Plots made after DO volume was derived suggest that setting is about right with some cases where it looks too low and others too high.

##### CELLTM

The upcast data are extremely noisy, the casts are shallow, there were many stops for bottles and the cell thermal mass correction is likely to be small due to the small temperature range in the profiles. The usual tests are too noisy to interpret. The default choice of (α = 0.0245, β=9.5) usually produces good results and was used for 2010-05 when the same sensors were used.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The Tau correction was applied.

The usual calculation on a few casts to find differences between primary and secondary channels for temperature, conductivity and salinity were not run because all casts were very shallow and the results would be meaningless.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

The problem of missing header information that was described in section 4 also affects these files, so routine “MERGE: CSV to Headers” was used to add missing information to the IOS Headers.

##### Checking Headers

The header check was run.  There are some negative pressures and negative values for many channels as well as unbelievably high salinity values. These are likely surface spikes. The fluorescence maximum was 33.6ug/L, so it never went off-scale.

Speeds look reasonable.

Surface check was run and shows an average surface pressure for the cruise was 0.06db. The associated salinity values were mostly very low. A few casts were examined in detail and the pumps were not on at the beginning of the record in all cases examined. After some time most salinity values came up to “in-water” values even before the pumps were turned on. At the end of casts data were recorded with pumps off near the surface. Again the values look like the CTD was in the water but the pumps were off. When pressure is about 0.1db for one cast, it looks like the sensors were very close to the surface with one T/C pair occasionally going to very low salinity. So the pressure looks reasonably accurate. The pressure offset had been adjusted at sea and that value was used in conversion of the raw data. Had the factory setting been used the pressures would have been too high.

The cross-reference check was compared with the log book. The station names were all correct. The log times are usually a little later than the file header times. The largest difference was ~11 minutes, so that cast was examined. There was a 10 minute soak periods, so the log time was likely recorded when the downcast actually began.

The altimeter readings and bottom depths from the headers of the CLN and MRG files were exported to spreadsheets. A sample of casts was checked and no errors were found in the altimetry or bottom depths.

The cruise track was plotted and added to the end of this report. No problems were found.

##### 13. Shift

Fluorescence

There are many stops for all upcasts, so the usual test of comparing vertical offsets between upcast and downcast temperature and fluorescence do not work very well. The usual shift applied is +24 records so that was chosen for these data. A check was run afterwards and there was an improvement in the alignment as measured against temperature.

SHIFT was run on all casts to advance the SeaPoint fluorescence channel by +24 records.

Conductivity

Tests were run on the primary and secondary conductivity channels to see what shifts had the best effect on the T-S plots. The results were very difficult to interpret varying from feature to feature with none having a particularly good effect. This is likely to be at least partly due to real instabilities in active mixing areas. The best results appeared to be close to those applied to the 2010-05 data which used the same sensors, so those were applied.

SHIFT was run on all casts using an advance of -0.4 records for the primary and -0.6 records for the secondary.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. Plots were examined to see if a further shift is needed and as far as can be judged from these noisy traces, no further alignment is needed.

##### 14. DELETE

Before running DELETE it is necessary to remove the initial soak period for all casts except #3 and #6, because data were recorded as the CTD was lowered to about 5m and then brought up to 2m before running the full cast. The data removal was done by recording a suitable scan number based on plots and then using CLIP. CLIP was run on casts #3 and 6 as well, but no records were removed. Records were removed from cast #1 by using a text editor.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings, but there was a problem with cast #3. There was a long stop at about 7.5db and data above that was lost due to the SURFACE RECORD REMOVAL step. When that was turned off, the results were better, though it is likely that the stop will lead to a mismatch between data either side of the stop.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in January 2011 and they were used for 3 inshore cruise between then and 2011-18. In all 3 cases the sensors appeared to be reading low by from 0.005 to 0.01. This seemed unlikely given they had been recently serviced and because both T/C pairs showed the same drift. Problems with sampling or analysis or poor flushing of bottles may have been factors.

2. Dissolved Oxygen

The DO sensor was recalibrated in Dec. 2010 and there is no record of any other use between the pre-cruise and post-cruise calibrations.

3. Pressure

The sensor was recalibrated in December 2009 and was used for 4 cruises since then using the factory offset.

Post-Cruise Calibrations

The temperature and conductivity sensors were calibrated in early January 2011 and again in early December 2011. The drift between those two calibrations was approximately -0.0005C° for the primary temperature, -0.0003 for the secondary temperature, 0 salinity units for the primary conductivity and -0.002 salinity units for the secondary conductivity. The combined effect would mean that by December 2011 the primary salinity was high by ~0.0005 and secondary salinity was low by 0.002.

Historic ranges – There was no local climatology available.

Repeat Casts –

There were no repeat casts.

##### DETAILED EDITING

The bottle comparison shows that the primary salinity is closest to bottles, but it is much noisier than the secondary. A revisit was made to the SHIFT stage to see if there was a better choice for the primary conductivity, but while improvements could be made at one level, noise was worse at other levels. So the secondary channels were selected for editing.

CTDEDIT was used to remove large spikes, smaller excursions that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. All casts required some editing especially in areas of high vertical temperature gradient. There were many casts that had unstable sections in the top 10m that were not clearly instrumental, so were left unedited. Very heavy editing would be required to remove all unstable features and it is not clear that they are not real.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

We have insufficient information on which to base a recalibration of dissolved oxygen.

The secondary salinity was found to be lower than the primary by 0.0023 in COMPARE. During the post-cruise calibration the combined effects of conductivity and temperature drift indicate that the primary salinity was high by ~0.0005psu and the secondary was low by ~0.002psu.

CALIBRATE was run using file 2011-18-recal1.ccf to add 0.002 to channel Salinity:T1:C1.

##### Fluorometer Processing

There were no off-scale fluorescence data.

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 0.5	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Many unstable features remain as described in section 16. Very heavy editing would be needed to remove them and often the unstable feature is due to a very small reversal in salinity. No further editing was applied.

Bin Average was run a second time to include standard deviations for all channels. Those files have extensions AVGSTD.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on the AVG files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate,Voltage:0, Voltage:1, Voltage:3, Voltage:4 and Flag

REMOVE was run on the AVGSTD files to remove (Output: *.REMSTD):

Scan_Number, Scan_Number:STD, Status:Pump:STD, Flag and Flag:STD

Derived Quantities was run to calculate Theta, Depth, Sigma-T and Sigma-Theta on the REMSTD files which were saved as *.CTDSTD. They will be stored, but not go in the OSD Data Archive.

A second SBE DO channel (with umol/kg units) was added to the REM files.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:SeaPoint and PAR are nominal and unedited except that some

records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, so values are nominal and

are reported with 1 less significant figure than usual.

Salinity was recalibrated based on a post-cruise calibration.

For further processing details see the processing report 2011-18_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok, though the SBE:Oxygen:Dissolved profile looked odd for one cast with a narrow range because with only 1 digit included, the results were constant.

The track plot looks ok.

The sensor history files were updated.

At the end of processing it was discovered that one file was missing. Event #1 had no bottle sampling but there is a profile, so it was put through all the earlier steps except that a text editor was used rather than by running programs CLIP and MERGE Headers.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 95% to 160% with the higher values at stations UTN-5 and UTN-7. Most values were between 100% and 110%. There are no casts from offshore areas where the surface saturation is fairly consistent, so no conclusion can be made from this except to say the values are not unreasonable.

The highest values were associated with high near-surface chlorophyll values.

##### 24. Final Bottle Files

CALIBRATE was run on files MRGCLN2 using file 2011-18-recal1.ccf.

The MRGCOR1 files were run through SORT to order on increasing pressure.

REMOVE was run on the MRGSORT files to remove the following channels (Output *.REM):

Scan Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Descent_Rate, Status:Pump, Voltage:0, Voltage:1, Voltage:3, Voltage:4 and Flag .

Dissolved oxygen was derived in mass units for the CTD DO channel.

As no draw temperature was available titrated dissolved oxygen could not be derived in mass units.

REORDER was run to get the 2 CTD DO channels together.

A preliminary header check turned up a few problems.

- Two values of -0.01 for Nitrate_plus_Nitrate entries (casts #28 and 36) were replaced with a 0 value.

- Three casts (Events 9, 18, 33) have some bottles with no sample numbers and no sampling, so the lines for those bottles were removed.

- Two casts (Events 6, 16) have some bottles with sample numbers but no sampling; the rosette log confirms there was no sampling for those bottles, so the lines for those bottles were removed.

- Two bottles from event #6 have no sampling indicated in the log but values were reported for nutrients. These were cases where 2 bottles were fired at roughly the same depth and it appears that the samples for nutrients analyzed in Maryland came from different bottles from those analyzed at IOS or else the sample numbers are wrong. It was assumed that the reported data are correct and no changes were made to those entries.

The header comments for samples from casts #26 and #27 were edited for clarity, separating remarks for salinity and nutrients.

CLEAN was run on all files to correct the # of records and variable ranges in the headers.

File 2011-18-bot-hdr.txt was prepared with explanations of quality flags and comments from analysts.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add comments from file 2011-18-bot-hdr.txt.

A header check was run on the final files and no errors were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no errors were found.

Standards check was run on all files. Some problems were noted and fixed by rerunning Header Edit. Some of the warnings of non-standard formats were due to deliberate choices, so they were not changed.

A second set of files was saved with standard deviations and original voltages included, as well all T, C and S channels. These have extension SAMAVGSTD and include depth, sigma-T and theta. They have been recalibrated.

# Particulars

# 6. Pumps were not turned on until CTD was at 7db.

# Institute of Ocean Sciences Cruise Summary

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 6Jan 2011 | Factory | 2 Dec 2011 |  |
| Conductivity | 3579 | 8Jan2011 | Factory | 1 Dec 2011 | Factory |
| Secondary Temp. | 5073 | 6Jan 2011 | Factory | 2 Dec 2011 |  |
| Secondary Cond. | 3581 | 8Jan2011 | Factory | 1 Dec 2011 | Factory |
| Transmissometer | CST662DR | 26May2008 |  |  |  |
| SBE 43 DO sensor | 1117 | 29Dec2010 | Factory | 2 Nov 2012 | Factory |
| SeaPoint Fluorometer | 2745 |  |  |  |  |
| PAR | 4601 | 16Mar2011 |  |  |  |
| Pressure Sensor | 0941 | 07Dec2009 | Factory |  |  |
| Altimeter | 40853 |  |  |  |  |