## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2011-24

Agency: MEAD

Location: Strait of Georgia

Project: Baynes Sound

Party Chief: Sutherland T.

Platform: Vector

Date: April 7, 2011 – April 13, 2011

Processed by: Germaine Gatien

Date of Processing: 7 November 2011 – 15 November 2011

Number of original HEX files: 45

Number of CTD files:   43	Number of bottle casts:  24

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1483), a Seapoint Fluorometer (#2745) with a 3X cable and an altimeter (#1252).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list and notes about problems with computer crashes which caused 2 split casts (2 files per cast.)

There pressure offset (-0.9db) had not been entered in the configuration file used at sea.

A draw temperature taken from Niskin bottles was available, allowing for the calculation of dissolved oxygen concentration in mass units for the bottle samples.

There were nutrient samples collected from a bucket. These are not included in the CHE files because there was no corresponding bottle firing and CTD data is not available that close to the surface. The bucket data are included in the nutrient spreadsheet, QF2011-24nuts*.xls.

CTD salinity data were not recalibrated. The temperature and conductivity sensors had been serviced recently and the agreement between the sensor pairs was excellent. The comparison with bottles was very noisy, with 11 out of 23 bottles differing from the CTD by >0.01. Only 1 of those outlier bottles was lower than the CTD salinity, and 8 bottles were higher than the CTD salinity by from 0.02 to 0.09. This may due to the fact that these samples were analyzed 5 months after the cruise. The 2 deepest samples did agree well with the CTD salinity. All salinity bottle data were flagged “3”.

The Oxygen:Dissolved:SBE values were generally higher than the titrated dissolved oxygen samples in the top 10m. The temperature and dissolved oxygen did not reach equilibrium during many of the bottle stops due to local variability. The bottle comparison did not prove useful for recalibration purposes. The data were recalibrated based on observations during 2011-08 when the same equipment was used.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1.5mL/L from 0 to 15db

±0.3mL/L from 15db to1500db

±0.03mL/L below 150db

Starting with 2011 cruises a new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel was compared with bottles to find the slope and offset to enter in the configuration files. While a little less convenient than calibrating oxygen concentration against bottles, it is the standard approach and is recommended by SeaBird, and should produce better results. For this cruise the comparison was extremely noisy and the results of 2011-08 were used instead.

Starting with 2011, the transmissivity conversion has been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder:

OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were good notes about problems with computer crashes. In 2 cases the crashes led to split casts with 2 files each.

- 10/11 – Downcast ok; upcast requires merge. Save as 10. There was some confusion in the 2nd file as it was forgotten that Niskin #1 had already been closed in the first file. So it was fired again at 150m but couldn’t close, so there was no sample #44 taken. Niskin #2 contains water from 125m and is called sample #45.

- 84/85 – Downcast ok; upcast requires merge. Save as 84.

Nutrients, dissolved oxygen and salinity were obtained in spreadsheet format from the analysts. The file creation date was added to the file names to avoid confusion in case some changes need to be made later. The nutrient file contains bucket samples as well as rosette samples. Only the rosette samples will be included in the CHE files. Even if we wanted to match this to surface CTD data it would generally be impossible since CTD data above 2m are usually unavailable or unreliable.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

Only one configuration was used for the cruise. The calibration constants were checked and 2 changes were made. The pressure offset had not been entered so that was added. The transmissometer parameters were entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

There is an entry for a Rinko III Temperature sensor and Surface PAR but it is believed these were not connected to the CTD.

The file was saved as 2011-24-ctd-orignal.con. This will be updated after the DO calibration study.

The event # between 13 and 15 is entered in the log and file name as #914. This is assumed to be an error (confirmed by Hugh Maclean), so will be renamed as 14.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT).

The BOT files for files 10 & 11 and 84 & 85 were changed to 10.botx & 10.boty, 84.botx & 84.boty; JOIN was used to produce new files named 2011-24-0010.BOT and 2011-24-0084.BOT. The files were then edited to remove the 150m stop from event #10 and the 3 bottles fired at 30m from event #84.

The BOT files were then averaged to enable an ADDSAMP file to be. Sample numbers were added to the ADDSAMP file based on rosette log records, skipping the bucket sample numbers. The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-24-DO-cal.csv. The titrated DO values and flags were added to that file; there were entries for every line so none needed to be removed. A calculation was made of Ф using equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P, T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

When all data were included M= 0.4073, B= -0.1625 and R2 =0.9716. The scatter is much higher than when the same instrument was used during 2011-08 right after this cruise.

To determine outliers the differences from the fit were calculated as:

(DO Volt * M – B – DObot/Phi)

where M and B are the results of the fit using all data. For 2011-08 which used the same instrument, there were only 6 out of 260 samples with differences >0.02. For these data about one third of the cases had differences >0.02. To reduce the noise a little, the 11 most severe outliers were removed and the resulting values of M and B were used in the above calculation. Gradually more severe restrictions were applied to see when the R value settled down. The trick is to remove enough outliers to get a reasonable fit without reducing the range of values significantly.

The following table summarizes the results and includes the values from the factory calibration and the results of 2011-08.

|  | # of bottles in fit | m | b | Soc | Voffset | R2 |
| --- | --- | --- | --- | --- | --- | --- |
| Original | Factory | 0.4329 | -0.2218 | 0.4329 | -0.5123 |  |
| all data | 181 | 0.4073 | -0.1625 | 0.4073 | -0.3990 | 0.9716 |
| diff<0.1 | 170 | 0.4244 | -0.1945 | 0.4244 | -0.4583 | 0.9935 |
| diff<0.06 | 160 | 0.4260 | -0.1969 | 0.4260 | -0.4622 | 0.9968 |
| diff<0.04 | 150 | 0.4285 | -0.2021 | 0.4285 | -0.4716 | 0.9981 |
| diff<0.03 | 142 | 0.4281 | -0.2001 | 0.4281 | -0.4674 | 0.9986 |
| diff<0.02 | 127 | 0.4298 | -0.2044 | 0.4298 | -0.4756 | 0.9992 |
| diff<0.01 | 79 | 0.4259 | -0.1948 | 0.4259 | -0.4574 | 0.9997 |
| 2011-08 | 255 used out of 260 | 0.4346 | -0.2135 | 0.4346 | -0.4913 | 0.9993 |

These are small corrections as is expected since this sensor had been recalibrated in late 2010. However, it seems odd that the slopes found for this cruise are lower than both those found at the factory and those found for 2011-08. A lot of the samples are from near-shore sites where temporal variability kept even temperature from equilibrating during some bottle stops, so dissolved oxygen certainly could not do so. For now, it looks best to use the 2011-08 values.

New values of Soc and Voffset based on 2011-08 were entered and the corrected file was saved as 2011-24-ctd-cor.con. (See file 2011-24-do-cal-test.xls for details.)

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. Those files were converted to IOS header format. They were put through CLEAN to create BOT files. Again, the BOT files for files 11 & 11 and 84 & 85 were combined; first the names were changed to 10.botx and 10.boty etc., then they were edited to remove the 150m stop from event #10 and the 3 bottles fired at 30m from event #84 and finally JOIN was run to add the boty files to the botx files.

Temperature and salinity were plotted for all BOT files and no significant outliers were found.

A header check was run and shows that there are no off-scale fluorescence values in the BOT files.

The addsamp.csv file prepared in the DO calibration step was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files.

The SAM files were then bin-averaged. The files for 10 and 84 were checked to make sure the JOIN was done right and all looks well.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-24-bot-hdr.txt.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-24nuts.xls which included a report on precisions. The file was saved as 2011-24-nuts.csv. Headers were changed to standard format and bucket samples were removed. Comments had “Nuts:” placed in front of them so they will be clear when merged with other comments. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files. Event #3 was changed to Event #4 to match the log book entry; the rosette log gives #3 but that was really a GRAB not a ROS event.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-24oxy.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified by removing a few unnecessary columns, removing bucket samples and removing one of 2 lines entered for a pair of duplicates. The duplicates were Chauvenet outliers and one had been identified as having a bad titration, so only the other value was used; the “3” flag was left. As explained for the nutrients event #3 was changed to event #4.  The file was then saved as 2011-24-oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. This file was saved as 2011-24-sal.csv. There were no comments or quality flags, and no problems were noted on the log sheet. Some columns were removed, header names were changed to standard format and a column was added for the event number. There were no duplicates. As explained for the nutrients event #3 was changed to event #4. The file was converted to individual SAL files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

##### Compare and other data checks

Salinity

Compare was run with pressure as reference channel.

In an initial run many problems were found:

- Sample #45, event #10, 125m. The bottle value is very low compared to the CTD and looks much closer to Niskin #4 than Niskin #2. This was the event where there was confusion over Niskin firing after a computer crash, so other variables were examined to see if they were also out of line, and no problems were found. The DO values match the SBE DO quite well and the nutrients look quite reasonable. The extracted DO has an odd profile, but so does the SBE DO profile. So this is likely a case of mis-sampling of salinity only. The CTD value is lower than the bottle by 0.008 when the sample is reassigned to sample #47. The sample was reassigned with flag “4”.

- Sample #130, event 42 was entered in the salinity spreadsheet as 114. On the rosette sheet there was no sample of that # for that cast. It makes sense as #130 though it is not really close to the CTD. Sample #114 belongs to another cast and there was no salinity sample taken there. The value was reassigned to #130.

- Event #4 was identified as #3 on the rosette log, so the samples did not get linked to the correct CTD data. Event #3 was a GRAB, not a Rosette cast.

- There were a few errors in the ADDSAMP file that led to a few other samples being missed in the MERGE with CTD data.

COMPARE was rerun after fixing those errors. There was a large scatter in the results. When cases with differences >0.01 were excluded, the primary salinity was low by 0.0016 with a standard deviation of 0.004 and the secondary salinity was low by 0.0014 with a standard deviation of 0.003. The near-surface data do not stand out in the fits. There were only 3 samples below 100m and one of those was a major outlier. The 2 bottles below 175m showed the primary salinity to be high by 0.00005 and low by 0.0018 while they indicate that the secondary salinity was high by 0.0004 and low by 0.0014. These suggest there is no significant error, but the evidence is weak.

It is difficult to make a judgment about the quality of the bottles from this cruise because the casts were close to shore and shallow so it is possible the local gradients might be significant. But a careful examination of the outliers did not turn up any cases where the difference could be explained by vertical gradients. Another possibility is that the bottles did not flush as well as usual since the pressure generally varied little through many of the stops. It is possible that might lead to the Niskin bottles containing water with higher salinity as the shed wake caused by the stop might fill the bottles and with no motion after that, they might not flush well.

Another possibility is that the bottles sat too long before being analyzed. After 5.2 months some evaporation may have occurred. This would explain why some comparisons look excellent and others not – variations in seals and tightness of caps would lead to some randomness. It has been reported by Sarah Zimmermann that there is a lot of variability in fits if analysis is delayed, and recent observations from early 2011 cruises shows the results of comparisons were suspicious when bottles sat for 3 months or longer and the other cruises were not ones for which poor flushing seemed a possible explanation. The fact that the T and C sensors had all been recently calibrated, and that the differences between them are very small, does lend weight to the conclusion that the bottle values are unreliable. There is no indication of time-dependence in the fits and a plot against salinity value shows no sign of a linearity problem in the Autosal.

Flag “3” was assigned to all bottles. Some look ok, but judging that the small differences are ok and the large differences are bad, is not justified. We cannot assume the CTD values are correct; we can only say that the bottle data appear to be unreliable. The same decision was made for 2011-28, for which the evidence was even stronger.

For details see 2011-24-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The plot of differences against pressure shows that the CTD readings are mostly much higher than bottles in the top 5m. Examination of a few casts shows that the CTD temperature and dissolved oxygen often did not equilibrate during the stops even when the pressure was very steady. With this much variability the dissolved oxygen sensor cannot produce a good match to bottle data because of the vertical offset, temporal variations and slow response of the DO sensor.

When the data from above 6m are excluded most of the noise disappears. When another 8 bottles with differences >0.3mL/L are excluded, the average difference is <0.0001mL/L. Outliers with differences > 0.3mL/L from below 15m were investigated, as were duplicates that had been flagged to see if one value was better than the others.

- Event #2, sample #7, 141db. The SBE DO is much lower than the bottle value. There is no evidence of a large shed wake. The vertical gradient is high but not enough to explain the difference as due to the vertical offset between bottle and CTD. The CTD DO had lots of time to settle down and appears to have done so. Flag “3” was attached to this bottle.

- Event #15, sample #71, 124db. CTD DO data very noisy.

- Event #27, sample #96, 131db. CTD DO data very noisy.

- Event #42, sample #128, 44db. CTD DO data very noisy.

- Event #49, sample #136, 48db. CTD DO data very noisy.

- Event #56, sample #144, 49db. CTD DO data very noisy.

- Event #77, sample #169, 38db. CTD DO data very noisy.

Two replicate outliers were also investigated:

- Event #2, sample3 – the average looks better than either value when compared to SBE DO. So the flag should be left.  Note about COMPARE result added to comment.

- Event #35, sample 123 – one value looks fine, the other had a problem with titration, so the average was not used and the flag was removed, but a comment about the duplicates was added.

Most of the outliers from below the surface were from the bottom of casts where there may be bottom currents leading to significant gradients.

For details see 2011-24-dox-comp1*.xls.

Other checks on bottle files

The comments “NO3 rerun” was removed from the headers of the nut files for casts #56 – 77.

The merge process was repeated to capture the new flags and comments.

Data were exported to spreadsheet 2011-24-bottles.xls and compared to the rosette sheets and all expected data are present.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### DO sensor hysteresis study

This sensor has not been fine-tuned for hysteresis parameters, but there is no deep sampling from this cruise so the tests cannot be done. This does not matter for this cruise since sampling was not deep.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-24-ctd-cor.xmlcon.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are in good agreement for downcasts, but very noisy in the upcasts in a manner that suggests the sensors are slightly out of alignment. There are some small spikes in the primary conductivity.

The altimetry looks fine near the bottom, and the transmissivity and fluorescence look ok.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

For the cruises just before and after this one, tests were done to determine the best choice for the offset between the DO voltage and the primary temperature, with 4.4s looking best overall. That setting was applied to these data because tests on these casts were inconclusive due to high variability.

ALIGNCTD was used to advance the DO Voltage by 4.4s relative to the pressure.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM during 2011-08 and 2011-28 in the same region and same month. The goal is to make upcasts look closer to downcasts on a T-S surface. As for the earlier cruises, no setting was found that improved this data, possibly because the temperature range was small and there is so much small-scale variation that it is hard to detect an improvement if there is one. CELLTM was not run.

##### DERIVE

Program DERIVE was run

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. One cast from each of cruises 2011-28 and 2001-08 are also listed for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-28-0025 | 200 | +0.0005 | <+0.0001 | +0.0006 | X Noisy, High |
| 2011-24-0008 | 220 | +0.0001 | <+0.0001 | +0.0002 | Steady, High |
| 2011-24-0014 | 220 | +0.0003 | <+0.0001 | +0.0004 | Steady, High |
| 2011-24-0015 | 220 | +0.0002 | <+0.0001 | +0.0004 | Steady, High |
| 2011-08-0004 | 280 | +0.0002 | <+0.0001 | +0.0005 | Noisy, High |

The differences are all small and consistent with the other two cruises.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace any pad values in pressure using linear interpolation based on scan number.

##### Checking Headers

The header check was run.  There were no off-scale fluorescence values and no obvious problems.

Surface check was run and shows an average surface pressure for the cruise was 2.7db. This is a typical value for the Vector.

The cross-reference check was compared with the log book. There was 1 error in station names which was corrected in both bottle and full profile files. In another case the log does not agree with the header station names, but it looks like the log is wrong (GS13 and GS14 should BS13 and BS14).

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. Plots were made of a few casts and the readings fit the plots reasonably well even when there are a lot of spikes near the bottom.

Water Depth readings were also checked. One (cast #15) was significantly different from the log entry and that entry had obviously been changed. Presumably the ship had drifted after the header entry was made. The header was changed to match the log.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. A few casts were examined and the readings look fine.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise the upcast temperature data were quite noisy, making a judgment difficult, but +24 looked best for more casts than any other setting.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of -0.6s to the secondary conductivity. The same settings were found for 2011-08 which used the same equipment.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked and the alignment looks excellent, so no further adjustment is necessary.

##### 14. DELETE

Plots were made of scan # versus pressure to ensure to check for any partial lowerings before the full cast – none were found.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concern files #11 and 85 which contain upcast data only. No action was required.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

These sensors were recalibrated in Jan. 2011 and used a few days before this cruise during 2011-28 and a few days after this one during 2011-08. For 2011-28 the comparison suggested the CTD was low by 0.01, but it was believed the bottle data was untrustworthy, likely because samples sat for 5.7 months before analysis. All bottle data were flagged. For 2011-08 the CTD salinity was low by ~0.004 after exclusion of some outliers. Again, the result was not trusted, though the scatter was not so bad that the bottles were flagged. The analysis occurred after 4.4 months for that cruise.

2. Dissolved Oxygen

This sensor was recalibrated in Dec. 2010 and used for 2011-28 and 2011-08 which bracketed this cruise. The 2011-28 results were very noisy and based on few samples. For 2011-08 there were many bottles and an excellent fit.

3. Pressure

The sensor was recalibrated in Dec. 2009. During 2011-28 and 2011-08 no problems were found with the factory calibration.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were no excursions from the local climatology, but in some areas none was available.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

No post-cruise calibrations were available.

##### DETAILED EDITING

There appears to be slightly more noise in the primary T/C channels than in the secondary, but neither pair had many large spikes. The secondary was selected for 2011-08, so it was also used for this cruise.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as the CTD data do not appear to have equilibrated.

Editing was kept fairly light because this is an area where small instabilities are quite likely to be real. All files required some editing.

##### Initial Recalibration

No recalibration will be applied since:

- The SBE dissolved oxygen data look fine in COMPARE .

- The salinity comparison is very noisy. The only 2 bottles from below 190m have values very close to the CTD salinity. The T and C sensors had been recalibrated recently and the pairs agreed with each other very well. So it is unlikely that there is a significant error in the CTD salinity. There is evidence from other 2011 cruises that having samples sit for 5 months before analysis produces poor results, in particular a lot of outliers.

- Pressure looks fine.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but it is sometimes found appropriate to do a further correction for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure there is a lot of scatter above 15db. The SBE DO values appear to be low if all data are included and high by an average of 0.03mL/L if data above 10db are not included. The average difference is too small to justify recalibration, especially in light of the fact that the standard deviation in that average is 0.12mL/L. During 2011-08 when there was more data available, no calibration was found appropriate.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were not prepared for Dr. Peña because there was no extracted CHL sampling.

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen; there are many unstable features especially near the surface, but no further editing seems justifiable in this region.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following header comments:

Data Processing Notes:

----------------------

Fluorescence and Transmissivity data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±1.5mL/L from 0 to 15db

±0.3mL/L from 15db to1500db

±0.03mL/L below 150db

For details on the processing see processing report: 2011-24-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from ~90% to 142%, with all but one of the values >130% being from Sechelt Inlet. A check of titrated DO values at a few casts with high saturation showed the surface SBE is a little high. This was noted in COMPARE but does not imply that SBE DO values are generally too high below the surface.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was used to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second titrated DO channel (with umol/kg units) was added using Temperature:Draw.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to fix formats and units, fix the agency header, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. One missing DO sample value was found and corrected.

Standards check was run on all files and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

6. Should be station 2.5, not 3.

10/11. Computer crash at 176m of upcast. First file includes bottom bottle, Niskin #1. Then Niskin numbers moved up by one from what is shown on rosette sheet.

68. Computer crash – no upcast, no sampling.

84/85. Computer crash. 3 Niskins on 1st file, 4 on 2nd file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 06Jan11 | Factory |  |  |
| Conductivity | 3579 | 08Jan11 | Factory |  |  |
| Secondary Temp. | 5073 | 06Jan11 | Factory |  |  |
| Secondary Cond. | 3581 | 08Jan11 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 24Dec2010 | Factory |  |  |
| Fluorometer | 2745 |  | IOS |  |  |
| Pressure Sensor | 0941 | 7Dec09 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |