Cruise 2011-26 Line P June 2011

Comparisons of cast #97 (in red) from metre-averaged CTD data and Argo float (in black) first profile before quality checks.

The CTD cast was almost a day earlier.

At depths below 1200db there is a consistent difference with the float having temperatures lower by ~0.015C° and salinity lower by ~0.002 along σt-lines. In the mixed layer the same trend is seen but the differences are a little higher at ~0.03C° and ~0.005. If the CTD data had been recalibrated based on the bottle analysis the differences in salinity would be even higher.

The temperature differences seem surprisingly high. The float may have drifted significantly from the site of Cast #97, but it is surprising if that would produce such a similar affect in the mixed layer and below 1200db.

The 2 CTD temperature and conductivity channels are in good agreement (differences <0.001). A salinity comparison suggests the CTD salinity is slightly low (by ~0.003), though the bottles may have read too high due to a 3-month wait before analysis.