## calibration-1176
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b |  | Summary of Soc Voffset including the original values in the factory calibration
82 | 1001.621 | 2.8856 | 2.8855 | 34.3808 | 34.3811 | 7.4762 | 0.62363 | 9 | 9 | 639 | 8.4449 | 0.047129 | 0.398 |  | 6.1876e-06 | 0.467 | -0.2441
43 | 125.341 | 6.7046 | 6.7035 | 33.747 | 33.7479 | 6.8534 | 1.967 | 18 | 18 | 364 | 6.851 | 0.6745 | 4.621 |  | 7.2609e-06 |  |  |  |  | m | b | Soc | Voffset | R2
82 | 200.12 | 4.1723 | 4.1723 | 33.776 | 33.7763 | 7.2744 | 1.3075 | 15 | 15 | 645 | 7.378 | 0.3665 | 2.704 |  | 1.6699e-05 |  |  |  | factory setting | 0.4453 | -0.2321 | 0.4453 | -0.5212
43 | 3501.548 | 1.5764 | 1.5753 | 34.6636 | 34.6659 | 7.7102 | 0.99529 | 3 | 3 | 349 | 12.1367 | 0.22065 | 2.678 |  | 4.6911e-05 |  |  | 176 pts | all | 0.467 | -0.2441 | 0.467 | -0.5227 | 0.9987
43 | 3501.538 | 1.5765 | 1.5755 | 34.6635 | 34.6659 | 7.7102 | 0.9952 | 2 | 2 | 348 | 12.1366 | 0.22074 | 2.679 |  | 7.8479e-05 |  |  | 171 | no 4 flag | 0.4671 | -0.2445 | 0.4671 | -0.52344 | 0.9987
55 | 124.682 | 6.3153 | 6.3158 | 33.7519 | 33.752 | 6.9152 | 1.8098 | 18 | 18 | 448 | 6.9166 | 0.60116 | 4.158 |  | 9.6156e-05 |  |  | 169 | <.05 | 0.4668 | -0.2443 | 0.4668 | -0.52335 | 0.9994
82 | 1249.223 | 2.5596 | 2.5597 | 34.4526 | 34.453 | 7.5332 | 0.62271 | 8 | 8 | 638 | 8.7977 | 0.046603 | 0.41 |  | 0.00010265 |  |  | 161 | <.02 | 0.4673 | -0.2444 | 0.4673 | -0.523 | 0.9997
43 | 75.552 | 6.9589 | 6.9569 | 32.6519 | 32.6535 | 6.8626 | 2.7068 | 20 | 20 | 366 | 6.8139 | 1.0201 | 6.951 |  | 0.00013554 |  |  | 155 | diff<0.015 | 0.467 | -0.2442 | 0.467 | -0.52291 | 0.9998
55 | 248.547 | 4.8555 | 4.856 | 33.8493 | 33.8497 | 7.1525 | 1.2952 | 14 | 14 | 444 | 7.2891 | 0.36095 | 2.631 |  | 0.00019496 |  |  | 148 | diff<0.01 | 0.4675 | -0.2449 | 0.4675 | -0.52385 | 0.9998
55 | 150.618 | 5.9567 | 5.9577 | 33.8004 | 33.8006 | 6.9711 | 1.6806 | 17 | 17 | 447 | 7 | 0.54057 | 3.784 |  | 0.00019569 |  |  | 134 | diff<0.008 | 0.4674 | -0.245 | 0.4674 | -0.52418 | 0.9999
43 | 9.814 | 9.5473 | 9.5472 | 32.5948 | 32.5954 | 6.4768 | 2.7917 | 23 | 23 | 369 | 6.3578 | 1.0593 | 6.735 |  | 0.0003124 |  |  | 2011-44 | diff<0.015 | 0.4604 | -0.2352 | 0.4604 | -0.51086 | 0.9997
21 | 600.639 | 4.182 | 4.1823 | 34.167 | 34.1678 | 7.2536 | 0.64801 | 6 | 6 | 176 | 7.7493 | 0.058844 | 0.456 |  | 0.00032355
55 | 1999.977 | 1.9384 | 1.9384 | 34.5845 | 34.5854 | 7.6442 | 0.80487 | 6 | 6 | 436 | 9.8697 | 0.13141 | 1.297 |  | 0.00036157
82 | 1999.512 | 1.9214 | 1.9214 | 34.586 | 34.5867 | 7.6474 | 0.80818 | 6 | 6 | 636 | 9.8738 | 0.13379 | 1.321 |  | 0.00046795
82 | 800.313 | 3.1943 | 3.1943 | 34.3054 | 34.3057 | 7.4233 | 0.64735 | 10 | 10 | 640 | 8.1606 | 0.058697 | 0.479 |  | 0.00048446
43 | 200.645 | 5.5933 | 5.5935 | 33.8171 | 33.8175 | 7.03 | 1.6138 | 15 | 15 | 361 | 7.1096 | 0.50889 | 3.618 |  | 0.00066127
43 | 150.572 | 6.2835 | 6.2837 | 33.8126 | 33.813 | 6.9176 | 1.8294 | 17 | 17 | 363 | 6.9424 | 0.60944 | 4.231 |  | 0.00076305
55 | 801.098 | 3.3653 | 3.3657 | 34.2885 | 34.2889 | 7.3931 | 0.63111 | 10 | 10 | 440 | 8.1242 | 0.049851 | 0.405 |  | 0.0007775
43 | 3682.007 | 1.5767 | 1.5755 | 34.6662 | 34.6688 | 7.71 | 0.99256 | 1 | 1 | 347 | 12.4268 | 0.21864 | 2.717 |  | 0.00078435
21 | 299.443 | 5.9089 | 5.9087 | 34.0045 | 34.0052 | 6.9694 | 0.94866 | 8 | 8 | 178 | 7.1346 | 0.20001 | 1.427 |  | 0.0010861
27 | 200.07 | 6.0044 | 6.0046 | 33.8648 | 33.8655 | 6.9603 | 1.3979 | 14 | 14 | 235 | 7.0334 | 0.40763 | 2.867 |  | 0.0011011
27 | 257.3 | 5.5579 | 5.5583 | 33.8894 | 33.8901 | 7.0325 | 1.26 | 13 | 13 | 234 | 7.1648 | 0.3432 | 2.459 |  | 0.0011023
43 | 298.167 | 4.7046 | 4.7052 | 33.8769 | 33.8773 | 7.177 | 1.187 | 13 | 13 | 359 | 7.3636 | 0.30909 | 2.276 |  | 0.0011245
55 | 2499.752 | 1.717 | 1.7168 | 34.6271 | 34.6282 | 7.6848 | 0.91587 | 5 | 5 | 435 | 10.6023 | 0.18241 | 1.934 |  | 0.0011972
21 | 1250.456 | 2.8634 | 2.8635 | 34.4702 | 34.4714 | 7.4758 | 0.64114 | 3 | 3 | 173 | 8.7235 | 0.056514 | 0.493 | 6 | 0.0012014
43 | 50.368 | 7.607 | 7.6074 | 32.613 | 32.6136 | 6.7633 | 2.7754 | 21 | 21 | 367 | 6.6876 | 1.0508 | 7.027 |  | 0.0012514 |  |  |  | 62 | excl. outliers diff>0.02 | 0.4661 | -0.1564 | 0.4661 | -0.33555 | 0.9623
43 | 24.59 | 9.5452 | 9.545 | 32.5941 | 32.5945 | 6.4771 | 2.7839 | 22 | 22 | 368 | 6.3701 | 1.0573 | 6.735 |  | 0.0012829
21 | 100.524 | 7.8003 | 7.8029 | 33.1483 | 33.1478 | 6.7101 | 2.0612 | 14 | 14 | 184 | 6.6761 | 0.71718 | 4.788 |  | 0.0013215
43 | 249.733 | 4.9407 | 4.9411 | 33.8439 | 33.8444 | 7.1382 | 1.3283 | 14 | 14 | 360 | 7.2744 | 0.37487 | 2.727 |  | 0.0013599
82 | 1499.672 | 2.3026 | 2.3026 | 34.507 | 34.5074 | 7.5788 | 0.66725 | 7 | 7 | 637 | 9.1533 | 0.066097 | 0.605 |  | 0.0014092
103 | 501.392 | 3.9257 | 3.9259 | 34.1044 | 34.1048 | 7.3018 | 0.77964 | 2 | 2 | 745 | 7.7058 | 0.12147 | 0.936 |  | 0.0014754
55 | 399.417 | 4.343 | 4.3433 | 34.0227 | 34.0229 | 7.2325 | 0.80707 | 12 | 12 | 442 | 7.5247 | 0.1313 | 0.988 |  | 0.0015011
55 | 201.961 | 5.306 | 5.3058 | 33.8329 | 33.8331 | 7.0771 | 1.4578 | 15 | 15 | 445 | 7.1624 | 0.43519 | 3.117 | 6 | 0.001518
6 | 10.126 | 12.4547 | 12.4777 | 31.8238 | 31.8208 | 6.1152 | 2.8637 | 18 | 18 | 44 | 5.9921 | 1.0916 | 6.541 |  | 0.001625
6 | 101.068 | 8.0326 | 8.0329 | 33.2705 | 33.273 | 6.6697 | 1.8896 | 14 | 14 | 40 | 6.6344 | 0.63668 | 4.224 |  | 0.001667
109 | 124.703 | 6.7841 | 6.7858 | 33.831 | 33.8303 | 6.837 | 0.9383 | 2 | 2 | 776 | 6.8333 | 0.19581 | 1.338 |  | 0.0017193
82 | 600.54 | 3.5596 | 3.5594 | 34.2099 | 34.2102 | 7.3619 | 0.6674 | 11 | 11 | 641 | 7.8774 | 0.069312 | 0.546 | 6 | 0.0017362
55 | 5.104 | 9.2274 | 9.2301 | 32.4892 | 32.4896 | 6.527 | 2.8135 | 24 | 24 | 454 | 6.4051 | 1.0681 | 6.841 |  | 0.0017445
4 | 25.056 | 11.2893 | 11.3208 | 31.7152 | 31.7027 | 6.272 | 2.9931 | 6 | 6 | 22 | 6.1609 | 1.1555 | 7.119 |  | 0.001819
55 | 300.29 | 4.6272 | 4.6263 | 33.9182 | 33.9183 | 7.1883 | 1.0023 | 13 | 13 | 443 | 7.3785 | 0.22579 | 1.666 |  | 0.001839
43 | 174.769 | 5.8637 | 5.8645 | 33.816 | 33.8163 | 6.9855 | 1.7151 | 16 | 16 | 362 | 7.0376 | 0.55871 | 3.932 |  | 0.0018417
43 | 2998.778 | 1.5958 | 1.5953 | 34.6534 | 34.6552 | 7.707 | 0.98671 | 4 | 4 | 350 | 11.3571 | 0.21484 | 2.44 |  | 0.00185
55 | 1001.201 | 3.0048 | 3.0042 | 34.3811 | 34.3818 | 7.4542 | 0.60005 | 9 | 9 | 439 | 8.4166 | 0.034218 | 0.288 |  | 0.0019051
43 | 399.953 | 4.3802 | 4.3806 | 33.9847 | 33.9851 | 7.2279 | 0.91753 | 12 | 12 | 358 | 7.5198 | 0.18245 | 1.372 |  | 0.0019346
82 | 24.826 | 7.8416 | 7.8424 | 32.6853 | 32.6853 | 6.7242 | 2.7882 | 22 | 22 | 652 | 6.6252 | 1.06 | 7.023 | 6 | 0.0020669
21 | 799.394 | 3.694 | 3.6941 | 34.2832 | 34.2842 | 7.3342 | 0.6035 | 5 | 5 | 175 | 8.0505 | 0.03565 | 0.287 |  | 0.0020844
27 | 101.065 | 6.9577 | 6.9577 | 33.2262 | 33.2266 | 6.8369 | 2.1581 | 18 | 18 | 239 | 6.8107 | 0.76145 | 5.186 |  | 0.0022839
21 | 1999.165 | 1.9212 | 1.9202 | 34.593 | 34.5953 | 7.6471 | 0.84362 | 1 | 1 | 171 | 9.873 | 0.15223 | 1.503 |  | 0.0023632
55 | 49.236 | 7.3467 | 7.3455 | 32.5224 | 32.5229 | 6.8076 | 2.7885 | 21 | 21 | 451 | 6.7328 | 1.0556 | 7.107 |  | 0.0025414 |  |  |  | 67 | excl. outliers diff>0.03 | 0.4923 | -0.2152 | 0.4923 | -0.43713 | 0.9302
21 | 1000.796 | 3.3387 | 3.3386 | 34.3824 | 34.3835 | 7.3932 | 0.59948 | 4 | 4 | 174 | 8.339 | 0.033217 | 0.277 |  | 0.0026397 |  |  |  | 71 | excl. outliers diff>0.04 | 0.516 | -0.2697 | 0.516 | -0.52267 | 0.9539
21 | 1500.421 | 2.4143 | 2.4149 | 34.5207 | 34.522 | 7.557 | 0.70475 | 2 | 2 | 172 | 9.1243 | 0.087678 | 0.8 |  | 0.0026598
103 | 9.511 | 8.1141 | 8.1143 | 32.6556 | 32.6554 | 6.6841 | 2.7835 | 8 | 8 | 751 | 6.5707 | 1.0585 | 6.955 | 6 | 0.0027118
55 | 2999.37 | 1.5766 | 1.5759 | 34.6547 | 34.6563 | 7.7106 | 0.99571 | 4 | 4 | 434 | 11.3644 | 0.22368 | 2.542 | 6 | 0.0027853
27 | 124.056 | 6.8708 | 6.8719 | 33.5832 | 33.5828 | 6.8346 | 1.8402 | 17 | 17 | 238 | 6.8293 | 0.61807 | 4.221 | 6 | 0.0028037
55 | 24.295 | 8.9042 | 8.9046 | 32.4982 | 32.4984 | 6.5736 | 2.8061 | 22 | 22 | 452 | 6.4687 | 1.0692 | 6.916 |  | 0.0028175
21 | 399.245 | 4.7728 | 4.7731 | 33.9897 | 33.9904 | 7.1599 | 0.90051 | 7 | 7 | 177 | 7.4418 | 0.17926 | 1.334 |  | 0.0028194
55 | 75.139 | 6.2183 | 6.2182 | 32.7005 | 32.7011 | 6.9793 | 2.5889 | 20 | 20 | 450 | 6.9373 | 0.96204 | 6.674 |  | 0.0028674
43 | 100.925 | 6.7603 | 6.7601 | 33.1617 | 33.1608 | 6.871 | 2.347 | 19 | 19 | 365 | 6.8466 | 0.85517 | 5.855 | 6 | 0.0032019
55 | 100.924 | 6.4216 | 6.421 | 33.315 | 33.3134 | 6.9181 | 2.1064 | 19 | 19 | 449 | 6.8972 | 0.73639 | 5.079 |  | 0.0032178
55 | 1501.012 | 2.318 | 2.318 | 34.5064 | 34.5071 | 7.5759 | 0.67009 | 7 | 7 | 437 | 9.1509 | 0.065567 | 0.6 |  | 0.0032647
27 | 298.919 | 5.2316 | 5.2308 | 33.93 | 33.9308 | 7.085 | 1.0552 | 12 | 12 | 233 | 7.2619 | 0.24539 | 1.782 |  | 0.0032653
43 | 2499.656 | 1.7289 | 1.7283 | 34.6287 | 34.6302 | 7.6824 | 0.91726 | 5 | 5 | 351 | 10.5983 | 0.18097 | 1.918 | 6 | 0.003288
27 | 3268.81 | 1.6763 | 1.6754 | 34.6464 | 34.6489 | 7.6917 | 0.94806 | 1 | 1 | 222 | 11.7385 | 0.19534 | 2.293 |  | 0.0033035
82 | 5.131 | 7.8589 | 7.8588 | 32.6853 | 32.6856 | 6.7215 | 2.7918 | 24 | 24 | 654 | 6.6057 | 1.063 | 7.022 |  | 0.0033588
82 | 2499.938 | 1.6963 | 1.6962 | 34.6301 | 34.631 | 7.6886 | 0.92286 | 5 | 5 | 635 | 10.6087 | 0.19031 | 2.019 |  | 0.003439
55 | 599.177 | 3.8026 | 3.803 | 34.1781 | 34.1784 | 7.32 | 0.66088 | 11 | 11 | 441 | 7.8262 | 0.061077 | 0.478 |  | 0.0034544
21 | 50.844 | 8.6881 | 8.6899 | 32.5446 | 32.5455 | 6.6034 | 2.7391 | 16 | 16 | 186 | 6.5215 | 1.0315 | 6.727 |  | 0.0035239
43 | 600.012 | 3.8402 | 3.8405 | 34.1561 | 34.1566 | 7.3144 | 0.68769 | 11 | 11 | 357 | 7.8204 | 0.073526 | 0.575 |  | 0.0035251
109 | 99.993 | 6.8451 | 6.8454 | 33.7945 | 33.7944 | 6.8291 | 0.97375 | 3 | 3 | 777 | 6.8031 | 0.21417 | 1.457 |  | 0.0035269
55 | 3500.075 | 1.5281 | 1.5273 | 34.6689 | 34.6707 | 7.7194 | 1.0226 | 2 | 2 | 432 | 12.1515 | 0.23701 | 2.88 |  | 0.0035394
82 | 149.445 | 4.5242 | 4.5242 | 33.7219 | 33.7219 | 7.2156 | 1.494 | 17 | 17 | 647 | 7.2648 | 0.45714 | 3.321 |  | 0.0035542
6 | 174.175 | 7.2207 | 7.2208 | 33.8368 | 33.8381 | 6.7686 | 1.4302 | 11 | 11 | 37 | 6.8036 | 0.42022 | 2.859 |  | 0.0035922
27 | 2998.122 | 1.6593 | 1.6586 | 34.6451 | 34.6474 | 7.695 | 0.95761 | 3 | 3 | 224 | 11.3354 | 0.19938 | 2.26 |  | 0.0037277
43 | 1248.809 | 2.664 | 2.6642 | 34.443 | 34.4437 | 7.5142 | 0.61051 | 8 | 8 | 354 | 8.7721 | 0.037163 | 0.326 |  | 0.0038448
27 | 2998.656 | 1.6592 | 1.6583 | 34.6451 | 34.6474 | 7.6951 | 0.95762 | 2 | 2 | 223 | 11.3362 | 0.19918 | 2.258 |  | 0.0039238
82 | 250.171 | 4.0139 | 4.0145 | 33.8612 | 33.8611 | 7.2982 | 1.0987 | 14 | 14 | 644 | 7.4531 | 0.27291 | 2.034 |  | 0.0039288
27 | 50.523 | 8.1566 | 8.1573 | 32.5417 | 32.5419 | 6.6827 | 2.8114 | 20 | 20 | 241 | 6.6035 | 1.0644 | 7.029 |  | 0.0043918
27 | 26.326 | 10.537 | 10.5075 | 32.3305 | 32.3367 | 6.3493 | 2.8191 | 21 | 21 | 242 | 6.2408 | 1.068 | 6.665 |  | 0.0044451
82 | 48.978 | 6.844 | 6.8453 | 32.6945 | 32.6946 | 6.8789 | 2.7646 | 21 | 21 | 651 | 6.8079 | 1.0514 | 7.158 |  | 0.0044563
82 | 99.552 | 4.9258 | 4.9262 | 32.7878 | 32.7878 | 7.1915 | 2.6131 | 19 | 19 | 649 | 7.1878 | 0.98069 | 7.049 |  | 0.0044634
27 | 9.563 | 10.8971 | 10.8967 | 32.2487 | 32.2496 | 6.3035 | 2.8239 | 22 | 22 | 243 | 6.1812 | 1.07 | 6.614 |  | 0.0046178
43 | 1001.058 | 3.0121 | 3.0124 | 34.3695 | 34.3701 | 7.4535 | 0.60308 | 9 | 9 | 355 | 8.4154 | 0.032916 | 0.277 |  | 0.0046224
103 | 4.971 | 8.1165 | 8.1166 | 32.6554 | 32.6553 | 6.6838 | 2.7806 | 9 | 9 | 752 | 6.5665 | 1.0592 | 6.955 |  | 0.0047024
82 | 299.734 | 4.0202 | 4.0205 | 33.9277 | 33.928 | 7.2938 | 0.97286 | 13 | 13 | 643 | 7.4966 | 0.21503 | 1.612 |  | 0.0048058
43 | 2000.252 | 1.9316 | 1.9314 | 34.5888 | 34.5899 | 7.6453 | 0.81728 | 6 | 6 | 352 | 9.8717 | 0.1326 | 1.309 |  | 0.0049685
21 | 199.547 | 6.4857 | 6.4858 | 33.9015 | 33.9023 | 6.8811 | 1.3778 | 10 | 10 | 180 | 6.9472 | 0.40434 | 2.809 |  | 0.0049884
104 | 998.536 | 3.0792 | 3.0795 | 34.3556 | 34.3555 | 7.4418 | 0.59581 | 3 | 3 | 756 | 8.3978 | 0.039177 | 0.329 |  | 0.0050337
4 | 5.276 | 13.1047 | 13.1052 | 31.0883 | 31.0893 | 6.0608 | 3.2719 | 8 | 8 | 24 | 5.9338 | 1.2788 | 7.588 |  | 0.0050867
6 | 300.52 | 6.0939 | 6.093 | 34.009 | 34.0102 | 6.9391 | 0.95845 | 8 | 8 | 34 | 7.1023 | 0.19839 | 1.409 |  | 0.0051089
104 | 800.813 | 3.5377 | 3.5377 | 34.2583 | 34.2582 | 7.3635 | 0.61466 | 4 | 4 | 757 | 8.0875 | 0.048099 | 0.389 |  | 0.0051525
21 | 124.006 | 7.5776 | 7.5785 | 33.6735 | 33.673 | 6.721 | 1.6429 | 13 | 13 | 183 | 6.7091 | 0.51796 | 3.475 |  | 0.0051603
21 | 74.597 | 7.963 | 7.9613 | 32.671 | 32.6743 | 6.7063 | 2.5358 | 15 | 15 | 185 | 6.6489 | 0.9349 | 6.216 |  | 0.0052231
6 | 249.62 | 6.5754 | 6.5754 | 33.9704 | 33.9717 | 6.8637 | 1.052 | 9 | 9 | 35 | 6.9734 | 0.24192 | 1.687 |  | 0.0052799
82 | 400.137 | 3.9463 | 3.9474 | 34.0547 | 34.0544 | 7.3006 | 0.82377 | 12 | 12 | 642 | 7.6035 | 0.14599 | 1.11 |  | 0.0053857
82 | 2999.653 | 1.5676 | 1.5671 | 34.6559 | 34.6571 | 7.7123 | 0.99655 | 4 | 4 | 634 | 11.3677 | 0.22669 | 2.577 | 6 | 0.0054054
82 | 10.123 | 7.8549 | 7.8553 | 32.6853 | 32.6856 | 6.7221 | 2.7884 | 23 | 23 | 653 | 6.6106 | 1.0636 | 7.031 |  | 0.0055018
21 | 174.569 | 6.827 | 6.8309 | 33.873 | 33.8734 | 6.8284 | 1.487 | 11 | 11 | 181 | 6.8681 | 0.45587 | 3.131 |  | 0.0055314
82 | 74.809 | 5.3152 | 5.3161 | 32.7479 | 32.7481 | 7.127 | 2.6585 | 20 | 20 | 650 | 7.0952 | 1.0029 | 7.116 |  | 0.0055372
43 | 799.425 | 3.4083 | 3.4087 | 34.2771 | 34.2775 | 7.3859 | 0.62773 | 10 | 10 | 356 | 8.1136 | 0.043384 | 0.352 |  | 0.0056658
27 | 397.184 | 4.7597 | 4.76 | 34.0024 | 34.0031 | 7.1615 | 0.86561 | 11 | 11 | 232 | 7.4417 | 0.1544 | 1.149 |  | 0.0057402
55 | 4024.718 | 1.544 | 1.5428 | 34.6753 | 34.6778 | 7.7159 | 1.0092 | 1 | 1 | 431 | 13.0096 | 0.23298 | 3.031 |  | 0.0058041
103 | 299.891 | 4.2868 | 4.2871 | 33.9014 | 33.9015 | 7.2482 | 1.0978 | 3 | 3 | 746 | 7.4453 | 0.27454 | 2.044 |  | 0.0059496
6 | 200.09 | 6.9958 | 6.9962 | 33.9054 | 33.9067 | 6.8005 | 1.2671 | 10 | 10 | 36 | 6.8608 | 0.34165 | 2.344 |  | 0.0059972
21 | 148.947 | 7.2605 | 7.2611 | 33.8437 | 33.8444 | 6.7622 | 1.4706 | 12 | 12 | 182 | 6.7748 | 0.44872 | 3.04 |  | 0.0060533
104 | 1499.13 | 2.4247 | 2.4248 | 34.4858 | 34.486 | 7.5568 | 0.63344 | 2 | 2 | 755 | 9.1223 | 0.057771 | 0.527 |  | 0.0060543
6 | 1319.765 | 2.941 | 2.941 | 34.4781 | 34.4802 | 7.461 | 0.65843 | 1 | 1 | 27 | 8.7832 | 0.057268 | 0.503 |  | 0.0061185
103 | 501.646 | 3.9267 | 3.9267 | 34.1038 | 34.1044 | 7.3017 | 0.78002 | 1 | 1 | 744 | 7.7059 | 0.1264 | 0.974 |  | 0.0062279
27 | 150.167 | 6.5732 | 6.573 | 33.772 | 33.7726 | 6.8731 | 1.6535 | 16 | 16 | 237 | 6.8942 | 0.52175 | 3.597 |  | 0.0063145
21 | 252.059 | 6.1465 | 6.1466 | 33.9666 | 33.9677 | 6.9326 | 1.0877 | 9 | 9 | 179 | 7.0507 | 0.27019 | 1.905 |  | 0.0063249
6 | 600.93 | 4.8252 | 4.8252 | 34.1875 | 34.189 | 7.1415 | 0.64709 | 6 | 6 | 32 | 7.6182 | 0.051718 | 0.394 |  | 0.0063728
6 | 24.788 | 9.9713 | 9.9638 | 32.2923 | 32.2943 | 6.4294 | 2.9122 | 17 | 17 | 43 | 6.321 | 1.1095 | 7.013 |  | 0.0063982
27 | 2499.632 | 1.7466 | 1.7461 | 34.6271 | 34.6289 | 7.679 | 0.91738 | 4 | 4 | 225 | 10.5929 | 0.17786 | 1.884 |  | 0.006461
82 | 175.796 | 4.3367 | 4.3321 | 33.7491 | 33.7496 | 7.247 | 1.4046 | 16 | 16 | 646 | 7.3243 | 0.41834 | 3.064 |  | 0.0065067
55 | 1250.517 | 2.6268 | 2.627 | 34.4432 | 34.4437 | 7.5211 | 0.6193 | 8 | 8 | 438 | 8.7832 | 0.038596 | 0.339 |  | 0.0065166
43 | 1499.285 | 2.3888 | 2.3887 | 34.498 | 34.4989 | 7.563 | 0.65375 | 7 | 7 | 353 | 9.131 | 0.054649 | 0.499 |  | 0.006552
82 | 124.804 | 4.5838 | 4.5828 | 33.5622 | 33.5637 | 7.2131 | 1.7051 | 18 | 18 | 648 | 7.2381 | 0.55885 | 4.045 |  | 0.0066725
6 | 1250.62 | 3.1638 | 3.1637 | 34.4507 | 34.4526 | 7.4215 | 0.63595 | 2 | 2 | 28 | 8.6524 | 0.046115 | 0.399 |  | 0.006774
104 | 301.053 | 5.7687 | 5.769 | 33.9589 | 33.9587 | 6.9944 | 0.92333 | 7 | 7 | 760 | 7.1636 | 0.1939 | 1.389 |  | 0.0068009
27 | 598.466 | 4.0741 | 4.0745 | 34.1697 | 34.1703 | 7.2724 | 0.65777 | 10 | 10 | 231 | 7.7693 | 0.056247 | 0.437 |  | 0.0068313
27 | 175.458 | 6.2985 | 6.2986 | 33.8255 | 33.8261 | 6.9146 | 1.5559 | 15 | 15 | 236 | 6.9615 | 0.47547 | 3.31 |  | 0.0070236
4 | 107.087 | 7.1809 | 7.1807 | 33.6182 | 33.6202 | 6.7845 | 1.2653 | 1 | 1 | 17 | 6.7615 | 0.33972 | 2.297 |  | 0.0070593
103 | 75.257 | 5.6248 | 5.6334 | 32.729 | 32.7284 | 7.0758 | 2.6462 | 6 | 6 | 749 | 7.0407 | 0.99877 | 7.032 |  | 0.0070998
27 | 5.469 | 10.8993 | 10.8991 | 32.2479 | 32.2488 | 6.3032 | 2.8299 | 23 | 23 | 244 | 6.1777 | 1.0703 | 6.612 |  | 0.0071705
104 | 400.7 | 5.2829 | 5.285 | 33.9971 | 33.9967 | 7.0732 | 0.83654 | 6 | 6 | 759 | 7.3451 | 0.15384 | 1.13 |  | 0.0072792
6 | 800.246 | 4.2241 | 4.2243 | 34.2848 | 34.2864 | 7.2405 | 0.59965 | 5 | 5 | 31 | 7.9375 | 0.028473 | 0.226 | 6 | 0.007464
82 | 3500.265 | 1.5037 | 1.503 | 34.6715 | 34.673 | 7.724 | 1.0328 | 3 | 3 | 633 | 12.1605 | 0.24571 | 2.988 |  | 0.0074825
104 | 148.623 | 6.3271 | 6.3275 | 33.8543 | 33.8539 | 6.9086 | 1.0863 | 11 | 11 | 764 | 6.9312 | 0.2708 | 1.877 |  | 0.0075971
4 | 100.677 | 7.3743 | 7.3738 | 33.4656 | 33.4679 | 6.7614 | 1.3923 | 2 | 2 | 18 | 6.7311 | 0.39845 | 2.682 |  | 0.0076443
6 | 50.428 | 8.7823 | 8.7945 | 32.501 | 32.5017 | 6.5914 | 2.6677 | 16 | 16 | 42 | 6.5086 | 0.99406 | 6.47 |  | 0.0076722
104 | 200.56 | 6.1409 | 6.1413 | 33.9074 | 33.9071 | 6.9362 | 1.0101 | 9 | 9 | 762 | 7.0078 | 0.23531 | 1.649 |  | 0.0077015
104 | 24.233 | 8.7103 | 8.7134 | 32.5674 | 32.567 | 6.5991 | 2.753 | 16 | 18 | 771 | 6.495 | 1.0493 | 6.815 |  | 0.0077086
103 | 40.812 | 7.9246 | 7.9356 | 32.6645 | 32.6636 | 6.7124 | 2.7659 | 7 | 7 | 750 | 6.6265 | 1.0553 | 6.993 | 6 | 0.0077345
27 | 799.429 | 3.7104 | 3.7106 | 34.3066 | 34.3074 | 7.3301 | 0.60456 | 9 | 9 | 230 | 8.0457 | 0.030451 | 0.245 |  | 0.0077784
27 | 1003.977 | 3.0658 | 3.0659 | 34.3705 | 34.3713 | 7.4436 | 0.60635 | 8 | 8 | 229 | 8.406 | 0.031287 | 0.263 |  | 0.0077784
82 | 4001.058 | 1.4971 | 1.496 | 34.6797 | 34.6816 | 7.7249 | 1.0344 | 2 | 2 | 632 | 12.9873 | 0.24701 | 3.208 |  | 0.0080308
6 | 150.041 | 7.6662 | 7.6568 | 33.7306 | 33.734 | 6.705 | 1.515 | 12 | 12 | 38 | 6.7147 | 0.45527 | 3.057 |  | 0.0081421
27 | 1252.294 | 2.6744 | 2.6745 | 34.4478 | 34.4487 | 7.512 | 0.6249 | 7 | 7 | 228 | 8.7732 | 0.039438 | 0.346 |  | 0.0082901
6 | 1000.655 | 3.7272 | 3.727 | 34.3783 | 34.3802 | 7.3236 | 0.60164 | 4 | 4 | 30 | 8.2512 | 0.028481 | 0.235 |  | 0.0083851
82 | 4317.458 | 1.5225 | 1.5212 | 34.6817 | 34.6842 | 7.7198 | 1.017 | 1 | 1 | 631 | 13.5266 | 0.23953 | 3.24 |  | 0.0086788
27 | 1500.017 | 2.3486 | 2.3488 | 34.5082 | 34.5092 | 7.57 | 0.68496 | 6 | 6 | 227 | 9.1416 | 0.066837 | 0.611 |  | 0.0089392
109 | 51.041 | 7.4989 | 7.4974 | 33.2578 | 33.2594 | 6.7514 | 1.3956 | 5 | 5 | 779 | 6.6773 | 0.41663 | 2.782 |  | 0.008988
104 | 251.051 | 5.9598 | 5.9607 | 33.9361 | 33.9357 | 6.9643 | 0.95859 | 8 | 8 | 761 | 7.0844 | 0.21272 | 1.507 |  | 0.00916
27 | 2002.652 | 1.9236 | 1.9233 | 34.5914 | 34.5928 | 7.6467 | 0.84177 | 5 | 5 | 226 | 9.8769 | 0.13982 | 1.381 |  | 0.0091856
55 | 174.681 | 5.601 | 5.6022 | 33.8159 | 33.8163 | 7.0288 | 1.5319 | 16 | 16 | 446 | 7.0845 | 0.48063 | 3.405 |  | 0.0093342
103 | 121.29 | 5.4932 | 5.4938 | 33.6717 | 33.6706 | 7.0534 | 1.8798 | 5 | 5 | 748 | 7.0619 | 0.64317 | 4.542 | 6 | 0.0094174
103 | 121.028 | 5.4925 | 5.4929 | 33.6718 | 33.6715 | 7.0535 | 1.8774 | 4 | 4 | 747 | 7.0618 | 0.64219 | 4.535 |  | 0.0095578
6 | 5.427 | 13.0021 | 13.0033 | 31.7916 | 31.7922 | 6.0471 | 2.8663 | 19 | 19 | 45 | 5.9206 | 1.0848 | 6.423 |  | 0.0096004
109 | 75.283 | 7.1816 | 7.1853 | 33.5313 | 33.5285 | 6.7883 | 1.2021 | 4 | 4 | 778 | 6.7377 | 0.32712 | 2.204 | 6 | 0.0098538
6 | 76.042 | 8.0396 | 8.0433 | 32.6555 | 32.6495 | 6.6954 | 2.5615 | 15 | 15 | 41 | 6.6386 | 0.94206 | 6.254 | 6 | 0.010055
104 | 48.934 | 6.8098 | 6.8061 | 32.6183 | 32.6184 | 6.8878 | 2.7284 | 15 | 15 | 768 | 6.817 | 1.0406 | 7.094 |  | 0.01057
104 | 38.648 | 8.483 | 8.4861 | 32.5718 | 32.5693 | 6.6326 | 2.7606 | 19 | 17 | 770 | 6.5416 | 1.0338 | 6.763 |  | 0.011238
104 | 124.889 | 6.4464 | 6.4469 | 33.753 | 33.752 | 6.8942 | 1.2319 | 12 | 12 | 765 | 6.8942 | 0.34362 | 2.369 | 6 | 0.012423
4 | 75.724 | 8.0356 | 8.0389 | 32.7604 | 32.7597 | 6.6914 | 2.1598 | 4 | 4 | 20 | 6.6345 | 0.75183 | 4.988 |  | 0.01271
104 | 600.6 | 4.1256 | 4.1261 | 34.1241 | 34.1237 | 7.2656 | 0.66835 | 5 | 5 | 758 | 7.7631 | 0.082312 | 0.639 |  | 0.014293
4 | 75.466 | 8.0358 | 8.0379 | 32.7611 | 32.7599 | 6.6914 | 2.1632 | 3 | 3 | 19 | 6.6342 | 0.75111 | 4.983 |  | 0.014991
4 | 10.327 | 12.4786 | 12.4832 | 31.0801 | 31.0802 | 6.1407 | 3.3273 | 7 | 7 | 23 | 6.0172 | 1.2945 | 7.789 | 6 | 0.015287
104 | 2000.602 | 1.9661 | 1.9656 | 34.5781 | 34.5791 | 7.6393 | 0.78192 | 1 | 1 | 754 | 9.863 | 0.13819 | 1.363 |  | 0.017136
6 | 126.14 | 7.828 | 7.8282 | 33.599 | 33.6001 | 6.6862 | 1.6372 | 13 | 13 | 39 | 6.674 | 0.5024 | 3.353 |  | 0.018064
109 | 30.59 | 9.2121 | 9.2112 | 32.1596 | 32.1598 | 6.5432 | 2.1409 | 6 | 6 | 780 | 6.4419 | 0.77415 | 4.987 |  | 0.018427
104 | 75.4 | 5.9601 | 5.9603 | 32.7533 | 32.7547 | 7.0191 | 2.5785 | 14 | 14 | 767 | 6.9802 | 0.97948 | 6.837 |  | 0.019443
109 | 5.135 | 11.3135 | 11.3132 | 31.5968 | 31.5967 | 6.2735 | 3.6067 | 9 | 9 | 783 | 6.1467 | 1.4601 | 8.975 |  | 0.019907
109 | 20.221 | 9.8119 | 9.7993 | 31.911 | 31.9138 | 6.4676 | 2.395 | 7 | 7 | 781 | 6.3558 | 0.89604 | 5.695 |  | 0.021696
103 | 2.264 | 8.1177 | 8.1179 | 32.6553 | 32.6551 | 6.6836 | 2.7473 | 10 | 10 | 753 | 6.564 | 1.0609 | 6.964 |  | 0.022033
4 | 49.915 | 8.6094 | 8.6105 | 32.3922 | 32.393 | 6.6215 | 2.3872 | 5 | 5 | 21 | 6.5392 | 0.84613 | 5.533 |  | 0.024608
104 | 48.986 | 6.823 | 6.8245 | 32.6174 | 32.6159 | 6.8858 | 2.7328 | 20 | 16 | 769 | 6.8149 | 1.0015 | 6.825 |  | 0.030638
27 | 75.193 | 7.2506 | 7.2506 | 32.7518 | 32.7527 | 6.8124 | 2.5159 | 19 | 19 | 240 | 6.7608 | 0.96556 | 6.528 |  | 0.034746
21 | 25.183 | 11.2958 | 11.3125 | 32.1542 | 32.1473 | 6.2537 | 2.8689 | 19 | 19 | 189 | 6.143 | 1.0571 | 6.494 |  | 0.038559
21 | 43.277 | 9.0621 | 9.0608 | 32.4977 | 32.4987 | 6.5506 | 2.8193 | 20 | 17 | 187 | 6.4607 | 1.0318 | 6.666 | 6 | 0.040712
55 | 3499.66 | 1.528 | 1.5272 | 34.6689 | 34.6708 | 7.7194 | 1.0227 | 3 | 3 | 433 | 12.1509 | 0.18674 | 2.269 |  | 0.046765
109 | 9.942 | 11.0822 | 11.1085 | 31.6461 | 31.6389 | 6.3026 | 3.3044 | 8 | 8 | 782 | 6.1799 | 1.3696 | 8.464 |  | 0.070571
104 | 98.667 | 6.339 | 6.3374 | 33.325 | 33.322 | 6.931 | 1.8789 | 13 | 13 | 766 | 6.9089 | 0.71908 | 4.968 |  | 0.08572
21 | 33.821 | 9.4684 | 9.4679 | 32.4938 | 32.4948 | 6.4922 | 2.8893 | 21 | 18 | 188 | 6.3929 | 1.0149 | 6.488 |  | 0.090344
43 | 4.69 | 9.5493 | 9.5493 | 32.5948 | 32.5955 | 6.4765 | 2.804 | 24 | 24 | 370 | 6.3533 | 1.0631 | 6.754 | 3 | 0.0022815
109 | 129.538 | 6.7261 | 6.7265 | 33.8678 | 33.8678 | 6.8445 | 0.91702 | 1 | 1 | 775 | 6.8457 | 0.18289 | 1.252 | 3 | 0.0012587
6 | 1250.567 | 3.1639 | 3.1638 | 34.4507 | 34.4527 | 7.4215 | 0.63577 | 3 | 3 | 29 | 8.6523 | 0.039643 | 0.343 | 3 | 0.013162
104 | 175.536 | 6.1963 | 6.1966 | 33.8955 | 33.8953 | 6.9278 | 1.0346 | 10 | 10 | 763 | 6.9761 | 0.27164 | 1.895 | 4 | 0.032593

## hysteresis study
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
27 | 200.07 | 6.0044 | 6.0046 | 33.8648 | 33.8655 | 6.9603 | 1.3979 | 14 | 14 | 235 | 7.0435 | 0.40704 | 2.867 |  | 0.0016849
6 | 200.09 | 6.9958 | 6.9962 | 33.9054 | 33.9067 | 6.8005 | 1.2671 | 10 | 10 | 36 | 6.8706 | 0.34116 | 2.344 |  | 0.0064849 |  |  |  | Comparison of concentration found for voltage=1 using different values of E above and below DO minimum
82 | 200.12 | 4.1723 | 4.1723 | 33.776 | 33.7763 | 7.2744 | 1.3075 | 15 | 15 | 645 | 7.3887 | 0.36597 | 2.704 |  | 0.00054525 |  |  |  |  | Data included have voltage range 0.6 to 1.6 above the min and 0.6 to 1.0 below the minimum
104 | 200.56 | 6.1409 | 6.1413 | 33.9074 | 33.9071 | 6.9362 | 1.0101 | 9 | 9 | 762 | 7.0179 | 0.23497 | 1.649 |  | 0.0073638
43 | 200.645 | 5.5933 | 5.5935 | 33.8171 | 33.8175 | 7.03 | 1.6138 | 15 | 15 | 361 | 7.1199 | 0.50816 | 3.618 |  | 0.0013934 |  |  |  |  | above DO min |  |  | below DO min
55 | 201.961 | 5.306 | 5.3058 | 33.8329 | 33.8331 | 7.0771 | 1.4578 | 15 | 15 | 445 | 7.1728 | 0.43456 | 3.117 | 6 | 0.0021488 |  |  |  | Results for voltage = 1
55 | 248.547 | 4.8555 | 4.856 | 33.8493 | 33.8497 | 7.1525 | 1.2952 | 14 | 14 | 444 | 7.3022 | 0.3603 | 2.631 |  | 0.00044987 |  |  |  | E | m | b | m*1+b | m | b | m*1+b | Shallow-Deep | max-min
6 | 249.62 | 6.5754 | 6.5754 | 33.9704 | 33.9717 | 6.8637 | 1.052 | 9 | 9 | 35 | 6.9859 | 0.24149 | 1.687 |  | 0.0057113 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.2235 | 0.4898 | -0.2626 | 0.2272 | -0.0037
43 | 249.733 | 4.9407 | 4.9411 | 33.8439 | 33.8444 | 7.1382 | 1.3283 | 14 | 14 | 360 | 7.2875 | 0.3742 | 2.727 |  | 0.0020326 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.2233 | 0.4826 | -0.2581 | 0.2245 | -0.0012
82 | 250.171 | 4.0139 | 4.0145 | 33.8612 | 33.8611 | 7.2982 | 1.0987 | 14 | 14 | 644 | 7.4665 | 0.27242 | 2.034 |  | 0.0034366 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.2232 | 0.479 | -0.2558 | 0.2232 | 0
104 | 251.051 | 5.9598 | 5.9607 | 33.9361 | 33.9357 | 6.9643 | 0.95859 | 8 | 8 | 761 | 7.0971 | 0.21234 | 1.507 |  | 0.0087777 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.2231 | 0.4754 | -0.2536 | 0.2218 | 0.0013
21 | 252.059 | 6.1465 | 6.1466 | 33.9666 | 33.9677 | 6.9326 | 1.0877 | 9 | 9 | 179 | 7.0635 | 0.2697 | 1.905 |  | 0.0058376 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.223 | 0.4719 | -0.2514 | 0.2205 | 0.0025
27 | 257.3 | 5.5579 | 5.5583 | 33.8894 | 33.8901 | 7.0325 | 1.26 | 13 | 13 | 234 | 7.1781 | 0.34257 | 2.459 |  | 0.0017354 |  |  |  | 0.038 | 0.4674 | -0.2446 | 0.2228 | 0.4684 | -0.2492 | 0.2192 | 0.0036
43 | 298.167 | 4.7046 | 4.7052 | 33.8769 | 33.8773 | 7.177 | 1.187 | 13 | 13 | 359 | 7.3795 | 0.30842 | 2.276 |  | 0.0017871 |  |  |  | 0.0385 | 0.4673 | -0.2446 | 0.2227 | 0.465 | -0.247 | 0.218 | 0.0047
27 | 298.919 | 5.2316 | 5.2308 | 33.93 | 33.9308 | 7.085 | 1.0552 | 12 | 12 | 233 | 7.2775 | 0.24486 | 1.782 |  | 0.0037917 |  |  |  | 0.039 | 0.4673 | -0.2446 | 0.2227 | 0.4615 | -0.2449 | 0.2166 | 0.0061 | 0.0098
21 | 299.443 | 5.9089 | 5.9087 | 34.0045 | 34.0052 | 6.9694 | 0.94866 | 8 | 8 | 178 | 7.15 | 0.19958 | 1.427 |  | 0.00065737 |  |  |  | Results for voltage = 0.9
82 | 299.734 | 4.0202 | 4.0205 | 33.9277 | 33.928 | 7.2938 | 0.97286 | 13 | 13 | 643 | 7.5128 | 0.21457 | 1.612 |  | 0.0043412 |  |  |  | E | m | b | m*0.9+b | m | b | m*0.9+b | Shallow-Deep
103 | 299.891 | 4.2868 | 4.2871 | 33.9014 | 33.9015 | 7.2482 | 1.0978 | 3 | 3 | 746 | 7.4614 | 0.27394 | 2.044 |  | 0.0053567 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.17668 | 0.4898 | -0.2626 | 0.17822 | -0.00154
55 | 300.29 | 4.6272 | 4.6263 | 33.9182 | 33.9183 | 7.1883 | 1.0023 | 13 | 13 | 443 | 7.3945 | 0.2253 | 1.666 |  | 0.0013514 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.17651 | 0.4826 | -0.2581 | 0.17624 | 0.00027
6 | 300.52 | 6.0939 | 6.093 | 34.009 | 34.0102 | 6.9391 | 0.95845 | 8 | 8 | 34 | 7.1176 | 0.19796 | 1.409 |  | 0.0055354 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.17642 | 0.479 | -0.2558 | 0.1753 | 0.00112
104 | 301.053 | 5.7687 | 5.769 | 33.9589 | 33.9587 | 6.9944 | 0.92333 | 7 | 7 | 760 | 7.1791 | 0.19348 | 1.389 |  | 0.0063828 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.17633 | 0.4754 | -0.2536 | 0.17426 | 0.00207
27 | 397.184 | 4.7597 | 4.76 | 34.0024 | 34.0031 | 7.1615 | 0.86561 | 11 | 11 | 232 | 7.463 | 0.15396 | 1.149 |  | 0.0061809 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.17624 | 0.4719 | -0.2514 | 0.17331 | 0.00293
21 | 399.245 | 4.7728 | 4.7731 | 33.9897 | 33.9904 | 7.1599 | 0.90051 | 7 | 7 | 177 | 7.4632 | 0.17874 | 1.334 |  | 0.0023051 |  |  |  | 0.038 | 0.4674 | -0.2446 | 0.17606 | 0.4684 | -0.2492 | 0.17236 | 0.0037
55 | 399.417 | 4.343 | 4.3433 | 34.0227 | 34.0229 | 7.2325 | 0.80707 | 12 | 12 | 442 | 7.5464 | 0.13092 | 0.988 |  | 0.0018786 |  |  |  | 0.0385 | 0.4673 | -0.2446 | 0.17597 | 0.465 | -0.247 | 0.1715 | 0.00447
43 | 399.953 | 4.3802 | 4.3806 | 33.9847 | 33.9851 | 7.2279 | 0.91753 | 12 | 12 | 358 | 7.5415 | 0.18193 | 1.372 |  | 0.0024597 |  |  |  | 0.039 | 0.4673 | -0.2446 | 0.17597 | 0.4615 | -0.2449 | 0.17045 | 0.00552 | 0.00706
82 | 400.137 | 3.9463 | 3.9474 | 34.0547 | 34.0544 | 7.3006 | 0.82377 | 12 | 12 | 642 | 7.6254 | 0.14557 | 1.11 |  | 0.0049647 |  |  |  | Results for voltage = 0.8
104 | 400.7 | 5.2829 | 5.285 | 33.9971 | 33.9967 | 7.0732 | 0.83654 | 6 | 6 | 759 | 7.3663 | 0.1534 | 1.13 |  | 0.006837 |  |  |  | E | m | b | m*0.8+b | m | b | m*0.8+b | Shallow-Deep
103 | 501.392 | 3.9257 | 3.9259 | 34.1044 | 34.1048 | 7.3018 | 0.77964 | 2 | 2 | 745 | 7.7337 | 0.12103 | 0.936 |  | 0.0010366 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.12986 | 0.4898 | -0.2626 | 0.12924 | 0.00062
103 | 501.646 | 3.9267 | 3.9267 | 34.1038 | 34.1044 | 7.3017 | 0.78002 | 1 | 1 | 744 | 7.7338 | 0.12594 | 0.974 |  | 0.0057711 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.12972 | 0.4826 | -0.2581 | 0.12798 | 0.00174
27 | 598.466 | 4.0741 | 4.0745 | 34.1697 | 34.1703 | 7.2724 | 0.65777 | 10 | 10 | 231 | 7.8029 | 0.056005 | 0.437 |  | 0.0070737 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.12964 | 0.479 | -0.2558 | 0.1274 | 0.00224
55 | 599.177 | 3.8026 | 3.803 | 34.1781 | 34.1784 | 7.32 | 0.66088 | 11 | 11 | 441 | 7.8602 | 0.060813 | 0.478 |  | 0.0037181 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.12956 | 0.4754 | -0.2536 | 0.12672 | 0.00284
43 | 600.012 | 3.8402 | 3.8405 | 34.1561 | 34.1566 | 7.3144 | 0.68769 | 11 | 11 | 357 | 7.8543 | 0.073208 | 0.575 |  | 0.003843 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.12948 | 0.4719 | -0.2514 | 0.12612 | 0.00336
82 | 600.54 | 3.5596 | 3.5594 | 34.2099 | 34.2102 | 7.3619 | 0.6674 | 11 | 11 | 641 | 7.9117 | 0.069012 | 0.546 | 6 | 0.001436 |  |  |  | 0.038 | 0.4674 | -0.2446 | 0.12932 | 0.4684 | -0.2492 | 0.12552 | 0.0038
104 | 600.6 | 4.1256 | 4.1261 | 34.1241 | 34.1237 | 7.2656 | 0.66835 | 5 | 5 | 758 | 7.7968 | 0.081956 | 0.639 |  | 0.013937 |  |  |  | 0.0385 | 0.4673 | -0.2446 | 0.12924 | 0.465 | -0.247 | 0.125 | 0.00424
21 | 600.639 | 4.182 | 4.1823 | 34.167 | 34.1678 | 7.2536 | 0.64801 | 6 | 6 | 176 | 7.7829 | 0.05859 | 0.456 |  | 6.9216e-05 |  |  |  | 0.039 | 0.4673 | -0.2446 | 0.12924 | 0.4615 | -0.2449 | 0.1243 | 0.00494 | 0.00432
6 | 600.93 | 4.8252 | 4.8252 | 34.1875 | 34.189 | 7.1415 | 0.64709 | 6 | 6 | 32 | 7.6512 | 0.051495 | 0.394 |  | 0.0065959 |  |  |  | Results for voltage = 0.7
21 | 799.394 | 3.694 | 3.6941 | 34.2832 | 34.2842 | 7.3342 | 0.6035 | 5 | 5 | 175 | 8.0971 | 0.035445 | 0.287 |  | 0.0022897 |  |  |  | E | m | b | m*0.7+b | m | b | m*0.7+b | Shallow-Deep
43 | 799.425 | 3.4083 | 3.4087 | 34.2771 | 34.2775 | 7.3859 | 0.62773 | 10 | 10 | 356 | 8.1606 | 0.043134 | 0.352 |  | 0.0059159 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.08304 | 0.4898 | -0.2626 | 0.08026 | 0.00278
27 | 799.429 | 3.7104 | 3.7106 | 34.3066 | 34.3074 | 7.3301 | 0.60456 | 9 | 9 | 230 | 8.0923 | 0.030276 | 0.245 |  | 0.0079537 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.08293 | 0.4826 | -0.2581 | 0.07972 | 0.00321
6 | 800.246 | 4.2241 | 4.2243 | 34.2848 | 34.2864 | 7.2405 | 0.59965 | 5 | 5 | 31 | 7.9834 | 0.028309 | 0.226 | 6 | 0.0076278 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.08286 | 0.479 | -0.2558 | 0.0795 | 0.00336
82 | 800.313 | 3.1943 | 3.1943 | 34.3054 | 34.3057 | 7.4233 | 0.64735 | 10 | 10 | 640 | 8.208 | 0.058358 | 0.479 |  | 0.00014546 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.08279 | 0.4754 | -0.2536 | 0.07918 | 0.00361
104 | 800.813 | 3.5377 | 3.5377 | 34.2583 | 34.2582 | 7.3635 | 0.61466 | 4 | 4 | 757 | 8.1345 | 0.047821 | 0.389 |  | 0.0048749 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.08272 | 0.4719 | -0.2514 | 0.07893 | 0.00379
55 | 801.098 | 3.3653 | 3.3657 | 34.2885 | 34.2889 | 7.3931 | 0.63111 | 10 | 10 | 440 | 8.1714 | 0.049563 | 0.405 |  | 0.0010655 |  |  |  | 0.038 | 0.4674 | -0.2446 | 0.08258 | 0.4684 | -0.2492 | 0.07868 | 0.0039
[19 empty] | 0.0385 | 0.4673 | -0.2446 | 0.08251 | 0.465 | -0.247 | 0.0785 | 0.00401
[19 empty] | 0.039 | 0.4673 | -0.2446 | 0.08251 | 0.4615 | -0.2449 | 0.07815 | 0.00436 | 0.00158
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b |  | Results for voltage = 0.6
43 | 1248.809 | 2.664 | 2.6642 | 34.443 | 34.4437 | 7.5142 | 0.61051 | 8 | 8 | 354 | 8.892 | 0.036662 | 0.326 |  | 0.0043462 |  |  |  | E | m | b | m*0.6+b | m | b | m*0.6+b | Shallow-Deep
82 | 1249.223 | 2.5596 | 2.5597 | 34.4526 | 34.453 | 7.5332 | 0.62271 | 8 | 8 | 638 | 8.9181 | 0.045974 | 0.41 |  | 0.00073183 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.03622 | 0.4898 | -0.2626 | 0.03128 | 0.00494
21 | 1250.456 | 2.8634 | 2.8635 | 34.4702 | 34.4714 | 7.4758 | 0.64114 | 3 | 3 | 173 | 8.8429 | 0.055751 | 0.493 | 6 | 0.00043853 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.03614 | 0.4826 | -0.2581 | 0.03146 | 0.00468
55 | 1250.517 | 2.6268 | 2.627 | 34.4432 | 34.4437 | 7.5211 | 0.6193 | 8 | 8 | 438 | 8.9035 | 0.038075 | 0.339 |  | 0.0070381 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.03608 | 0.479 | -0.2558 | 0.0316 | 0.00448
6 | 1250.62 | 3.1638 | 3.1637 | 34.4507 | 34.4526 | 7.4215 | 0.63595 | 2 | 2 | 28 | 8.7706 | 0.045493 | 0.399 |  | 0.0073959 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.03602 | 0.4754 | -0.2536 | 0.03164 | 0.00438
27 | 1252.294 | 2.6744 | 2.6745 | 34.4478 | 34.4487 | 7.512 | 0.6249 | 7 | 7 | 228 | 8.8935 | 0.038905 | 0.346 |  | 0.0088236 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.03596 | 0.4719 | -0.2514 | 0.03174 | 0.00422
6 | 1319.765 | 2.941 | 2.941 | 34.4781 | 34.4802 | 7.461 | 0.65843 | 1 | 1 | 27 | 8.9101 | 0.056453 | 0.503 |  | 0.0069339 |  |  |  | 0.038 | 0.4674 | -0.2446 | 0.03584 | 0.4684 | -0.2492 | 0.03184 | 0.004
104 | 1499.13 | 2.4247 | 2.4248 | 34.4858 | 34.486 | 7.5568 | 0.63344 | 2 | 2 | 755 | 9.2724 | 0.056836 | 0.527 |  | 0.0051191 |  |  |  | 0.0385 | 0.4673 | -0.2446 | 0.03578 | 0.465 | -0.247 | 0.032 | 0.00378
43 | 1499.285 | 2.3888 | 2.3887 | 34.498 | 34.4989 | 7.563 | 0.65375 | 7 | 7 | 353 | 9.2812 | 0.053764 | 0.499 |  | 0.0074368 |  |  |  | 0.039 | 0.4673 | -0.2446 | 0.03578 | 0.4615 | -0.2449 | 0.032 | 0.00378 | 0.00116
82 | 1499.672 | 2.3026 | 2.3026 | 34.507 | 34.5074 | 7.5788 | 0.66725 | 7 | 7 | 637 | 9.304 | 0.065026 | 0.605 |  | 0.00248
27 | 1500.017 | 2.3486 | 2.3488 | 34.5082 | 34.5092 | 7.57 | 0.68496 | 6 | 6 | 227 | 9.2922 | 0.065754 | 0.611 |  | 0.010022
21 | 1500.421 | 2.4143 | 2.4149 | 34.5207 | 34.522 | 7.557 | 0.70475 | 2 | 2 | 172 | 9.2746 | 0.086257 | 0.8 |  | 0.0012392
55 | 1501.012 | 2.318 | 2.318 | 34.5064 | 34.5071 | 7.5759 | 0.67009 | 7 | 7 | 437 | 9.3017 | 0.064504 | 0.6 |  | 0.0043278
21 | 1999.165 | 1.9212 | 1.9202 | 34.593 | 34.5953 | 7.6471 | 0.84362 | 1 | 1 | 171 | 10.0906 | 0.14895 | 1.503 |  | 0.00092009 | [5 empty] | E | m | b | effect near surface if E changes - use voltage = 3.5
82 | 1999.512 | 1.9214 | 1.9214 | 34.586 | 34.5867 | 7.6474 | 0.80818 | 6 | 6 | 636 | 10.0915 | 0.1309 | 1.321 |  | 0.002418 | [5 empty] | 0.035 | 0.4682 | -0.2447 | 1.394
55 | 1999.977 | 1.9384 | 1.9384 | 34.5845 | 34.5854 | 7.6442 | 0.80487 | 6 | 6 | 436 | 10.0873 | 0.12858 | 1.297 |  | 0.0031968 | [5 empty] | 0.36 | 0.4679 | -0.2446 | 1.393
43 | 2000.252 | 1.9316 | 1.9314 | 34.5888 | 34.5899 | 7.6453 | 0.81728 | 6 | 6 | 352 | 10.0894 | 0.12974 | 1.309 |  | 0.0078298 | [5 empty] | 0.365 | 0.4678 | -0.2446 | 1.3927
27 | 2002.652 | 1.9236 | 1.9233 | 34.5914 | 34.5928 | 7.6467 | 0.84177 | 5 | 5 | 226 | 10.095 | 0.1368 | 1.381 |  | 0.012206 | [5 empty] | 0.037 | 0.4677 | -0.2446 | 1.3923
27 | 2499.632 | 1.7466 | 1.7461 | 34.6271 | 34.6289 | 7.679 | 0.91738 | 4 | 4 | 225 | 10.8858 | 0.17307 | 1.884 |  | 0.011247 | [5 empty] | 0.0375 | 0.4676 | -0.2446 | 1.392
43 | 2499.656 | 1.7289 | 1.7283 | 34.6287 | 34.6302 | 7.6824 | 0.91726 | 5 | 5 | 351 | 10.8914 | 0.1761 | 1.918 | 6 | 0.0081584 | [5 empty] | 0.038 | 0.4674 | -0.2446 | 1.3913
55 | 2499.752 | 1.717 | 1.7168 | 34.6271 | 34.6282 | 7.6848 | 0.91587 | 5 | 5 | 435 | 10.8955 | 0.1775 | 1.934 |  | 0.0061068 | [5 empty] | 0.0385 | 0.4673 | -0.2446 | 1.3909
82 | 2499.938 | 1.6963 | 1.6962 | 34.6301 | 34.631 | 7.6886 | 0.92286 | 5 | 5 | 635 | 10.9022 | 0.18519 | 2.019 |  | 0.0016839 | [5 empty] | 0.039 | 0.4673 | -0.2446 | 1.3909
27 | 2998.122 | 1.6593 | 1.6586 | 34.6451 | 34.6474 | 7.695 | 0.95761 | 3 | 3 | 224 | 11.7125 | 0.19296 | 2.26 |  | 0.010148 | [8 empty] | 0.00305 | Max-Min
27 | 2998.656 | 1.6592 | 1.6583 | 34.6451 | 34.6474 | 7.6951 | 0.95762 | 2 | 2 | 223 | 11.7134 | 0.19277 | 2.258 |  | 0.010339
43 | 2998.778 | 1.5958 | 1.5953 | 34.6534 | 34.6552 | 7.707 | 0.98671 | 4 | 4 | 350 | 11.7351 | 0.20792 | 2.44 |  | 0.0087709 | [5 empty] | E | m | b | effect at depth with voltage=.6
55 | 2999.37 | 1.5766 | 1.5759 | 34.6547 | 34.6563 | 7.7106 | 0.99571 | 4 | 4 | 434 | 11.7427 | 0.21647 | 2.542 | 6 | 0.0044223 | [5 empty] | 0.035 | 0.4898 | -0.2626 | 0.03128
82 | 2999.653 | 1.5676 | 1.5671 | 34.6559 | 34.6571 | 7.7123 | 0.99655 | 4 | 4 | 634 | 11.7463 | 0.21939 | 2.577 | 6 | 0.0019001 | [5 empty] | 0.36 | 0.4826 | -0.2581 | 0.03146
27 | 3268.81 | 1.6763 | 1.6754 | 34.6464 | 34.6489 | 7.6917 | 0.94806 | 1 | 1 | 222 | 12.1649 | 0.18849 | 2.293 |  | 0.010151 | [5 empty] | 0.365 | 0.479 | -0.2558 | 0.0316
55 | 3500.075 | 1.5281 | 1.5273 | 34.6689 | 34.6707 | 7.7194 | 1.0226 | 2 | 2 | 432 | 12.625 | 0.22812 | 2.88 |  | 0.0053498 | [5 empty] | 0.037 | 0.4754 | -0.2536 | 0.03164
82 | 3500.265 | 1.5037 | 1.503 | 34.6715 | 34.673 | 7.724 | 1.0328 | 3 | 3 | 633 | 12.6344 | 0.2365 | 2.988 |  | 0.0017345 | [5 empty] | 0.0375 | 0.4719 | -0.2514 | 0.03174
43 | 3501.538 | 1.5765 | 1.5755 | 34.6635 | 34.6659 | 7.7102 | 0.9952 | 2 | 2 | 348 | 12.6097 | 0.21246 | 2.679 |  | 0.0082024 | [5 empty] | 0.038 | 0.4684 | -0.2492 | 0.03184
43 | 3501.548 | 1.5764 | 1.5753 | 34.6636 | 34.6659 | 7.7102 | 0.99529 | 3 | 3 | 349 | 12.6097 | 0.21238 | 2.678 |  | 0.0083247 | [5 empty] | 0.0385 | 0.465 | -0.247 | 0.032
43 | 3682.007 | 1.5767 | 1.5755 | 34.6662 | 34.6688 | 7.71 | 0.99256 | 1 | 1 | 347 | 12.9366 | 0.21002 | 2.717 |  | 0.0094009 | [5 empty] | 0.039 | 0.4615 | -0.2449 | 0.032
82 | 4001.058 | 1.4971 | 1.496 | 34.6797 | 34.6816 | 7.7249 | 1.0344 | 2 | 2 | 632 | 13.5675 | 0.23645 | 3.208 |  | 0.002532 | [8 empty] | 0.00072 | Max-Min
55 | 4024.718 | 1.544 | 1.5428 | 34.6753 | 34.6778 | 7.7159 | 1.0092 | 1 | 1 | 431 | 13.5942 | 0.22296 | 3.031 |  | 0.0042148
82 | 4317.458 | 1.5225 | 1.5212 | 34.6817 | 34.6842 | 7.7198 | 1.017 | 1 | 1 | 631 | 14.1798 | 0.22849 | 3.24 |  | 0.0023541 | [5 empty] | E | m | b | effect near surface if E changes - use voltage = 4.5
[11 empty] | 0 | [9 empty] | 0.035 | 0.4682 | -0.2447 | 1.8622
[11 empty] | 0 | [9 empty] | 0.36 | 0.4679 | -0.2446 | 1.861
[11 empty] | 0 | [9 empty] | 0.365 | 0.4678 | -0.2446 | 1.8605
[11 empty] | 0 | [9 empty] | 0.037 | 0.4677 | -0.2446 | 1.86
[11 empty] | 0 | [9 empty] | 0.0375 | 0.4676 | -0.2446 | 1.8596
[18 empty] | 1.0344 |  |  | 0.038 | 0.4674 | -0.2446 | 1.8587
[21 empty] | 0.0385 | 0.4673 | -0.2446 | 1.8582
[21 empty] | 0.039 | 0.4673 | -0.2446 | 1.8582
[24 empty] | 0.00395 | Max-Min
Shallow data
103 | 4.971 | 8.1165 | 8.1166 | 32.6554 | 32.6553 | 6.6838 | 2.7806 | 9 | 9 | 752 | 6.5665 | 1.0592 | 6.955 |  | 0.0047024 | 0.467 | -0.2441
55 | 5.104 | 9.2274 | 9.2301 | 32.4892 | 32.4896 | 6.527 | 2.8135 | 24 | 24 | 454 | 6.4051 | 1.0681 | 6.841 |  | 0.0017445
82 | 5.131 | 7.8589 | 7.8588 | 32.6853 | 32.6856 | 6.7215 | 2.7918 | 24 | 24 | 654 | 6.6057 | 1.063 | 7.022 |  | 0.0033588
4 | 5.276 | 13.1047 | 13.1052 | 31.0883 | 31.0893 | 6.0608 | 3.2719 | 8 | 8 | 24 | 5.9338 | 1.2788 | 7.588 |  | 0.0050867
6 | 5.427 | 13.0021 | 13.0033 | 31.7916 | 31.7922 | 6.0471 | 2.8663 | 19 | 19 | 45 | 5.9206 | 1.0848 | 6.423 |  | 0.0096004
27 | 5.469 | 10.8993 | 10.8991 | 32.2479 | 32.2488 | 6.3032 | 2.8299 | 23 | 23 | 244 | 6.1777 | 1.0703 | 6.612 |  | 0.0071705
103 | 9.511 | 8.1141 | 8.1143 | 32.6556 | 32.6554 | 6.6841 | 2.7835 | 8 | 8 | 751 | 6.5707 | 1.0585 | 6.955 | 6 | 0.0027118
27 | 9.563 | 10.8971 | 10.8967 | 32.2487 | 32.2496 | 6.3035 | 2.8239 | 22 | 22 | 243 | 6.1812 | 1.07 | 6.614 |  | 0.0046178
43 | 9.814 | 9.5473 | 9.5472 | 32.5948 | 32.5954 | 6.4768 | 2.7917 | 23 | 23 | 369 | 6.3578 | 1.0593 | 6.735 |  | 0.0003124
82 | 10.123 | 7.8549 | 7.8553 | 32.6853 | 32.6856 | 6.7221 | 2.7884 | 23 | 23 | 653 | 6.6106 | 1.0636 | 7.031 |  | 0.0055018
6 | 10.126 | 12.4547 | 12.4777 | 31.8238 | 31.8208 | 6.1152 | 2.8637 | 18 | 18 | 44 | 5.9921 | 1.0916 | 6.541 |  | 0.001625
104 | 24.233 | 8.7103 | 8.7134 | 32.5674 | 32.567 | 6.5991 | 2.753 | 16 | 18 | 771 | 6.495 | 1.0493 | 6.815 |  | 0.0077086
55 | 24.295 | 8.9042 | 8.9046 | 32.4982 | 32.4984 | 6.5736 | 2.8061 | 22 | 22 | 452 | 6.4687 | 1.0692 | 6.916 |  | 0.0028175
43 | 24.59 | 9.5452 | 9.545 | 32.5941 | 32.5945 | 6.4771 | 2.7839 | 22 | 22 | 368 | 6.3701 | 1.0573 | 6.735 |  | 0.0012829
6 | 24.788 | 9.9713 | 9.9638 | 32.2923 | 32.2943 | 6.4294 | 2.9122 | 17 | 17 | 43 | 6.321 | 1.1095 | 7.013 |  | 0.0063982
82 | 24.826 | 7.8416 | 7.8424 | 32.6853 | 32.6853 | 6.7242 | 2.7882 | 22 | 22 | 652 | 6.6252 | 1.06 | 7.023 | 6 | 0.0020669
4 | 25.056 | 11.2893 | 11.3208 | 31.7152 | 31.7027 | 6.272 | 2.9931 | 6 | 6 | 22 | 6.1609 | 1.1555 | 7.119 |  | 0.001819
27 | 26.326 | 10.537 | 10.5075 | 32.3305 | 32.3367 | 6.3493 | 2.8191 | 21 | 21 | 242 | 6.2408 | 1.068 | 6.665 |  | 0.0044451
104 | 38.648 | 8.483 | 8.4861 | 32.5718 | 32.5693 | 6.6326 | 2.7606 | 19 | 17 | 770 | 6.5416 | 1.0338 | 6.763 |  | 0.011238
103 | 40.812 | 7.9246 | 7.9356 | 32.6645 | 32.6636 | 6.7124 | 2.7659 | 7 | 7 | 750 | 6.6265 | 1.0553 | 6.993 | 6 | 0.0077345
104 | 48.934 | 6.8098 | 6.8061 | 32.6183 | 32.6184 | 6.8878 | 2.7284 | 15 | 15 | 768 | 6.817 | 1.0406 | 7.094 |  | 0.01057
82 | 48.978 | 6.844 | 6.8453 | 32.6945 | 32.6946 | 6.8789 | 2.7646 | 21 | 21 | 651 | 6.8079 | 1.0514 | 7.158 |  | 0.0044563
55 | 49.236 | 7.3467 | 7.3455 | 32.5224 | 32.5229 | 6.8076 | 2.7885 | 21 | 21 | 451 | 6.7328 | 1.0556 | 7.107 |  | 0.0025414
43 | 50.368 | 7.607 | 7.6074 | 32.613 | 32.6136 | 6.7633 | 2.7754 | 21 | 21 | 367 | 6.6876 | 1.0508 | 7.027 |  | 0.0012514
6 | 50.428 | 8.7823 | 8.7945 | 32.501 | 32.5017 | 6.5914 | 2.6677 | 16 | 16 | 42 | 6.5086 | 0.99406 | 6.47 |  | 0.0076722
27 | 50.523 | 8.1566 | 8.1573 | 32.5417 | 32.5419 | 6.6827 | 2.8114 | 20 | 20 | 241 | 6.6035 | 1.0644 | 7.029 |  | 0.0043918 |  |  |  | 62 | excl. outliers diff>0.02 | 0.4661 | -0.1564 | 0.4661 | -0.33555 | 0.9623
21 | 50.844 | 8.6881 | 8.6899 | 32.5446 | 32.5455 | 6.6034 | 2.7391 | 16 | 16 | 186 | 6.5215 | 1.0315 | 6.727 |  | 0.0035239
109 | 51.041 | 7.4989 | 7.4974 | 33.2578 | 33.2594 | 6.7514 | 1.3956 | 5 | 5 | 779 | 6.6773 | 0.41663 | 2.782 |  | 0.008988
21 | 74.597 | 7.963 | 7.9613 | 32.671 | 32.6743 | 6.7063 | 2.5358 | 15 | 15 | 185 | 6.6489 | 0.9349 | 6.216 |  | 0.0052231
82 | 74.809 | 5.3152 | 5.3161 | 32.7479 | 32.7481 | 7.127 | 2.6585 | 20 | 20 | 650 | 7.0952 | 1.0029 | 7.116 |  | 0.0055372
55 | 75.139 | 6.2183 | 6.2182 | 32.7005 | 32.7011 | 6.9793 | 2.5889 | 20 | 20 | 450 | 6.9373 | 0.96204 | 6.674 |  | 0.0028674
103 | 75.257 | 5.6248 | 5.6334 | 32.729 | 32.7284 | 7.0758 | 2.6462 | 6 | 6 | 749 | 7.0407 | 0.99877 | 7.032 |  | 0.0070998
109 | 75.283 | 7.1816 | 7.1853 | 33.5313 | 33.5285 | 6.7883 | 1.2021 | 4 | 4 | 778 | 6.7377 | 0.32712 | 2.204 | 6 | 0.0098538
4 | 75.466 | 8.0358 | 8.0379 | 32.7611 | 32.7599 | 6.6914 | 2.1632 | 3 | 3 | 19 | 6.6342 | 0.75111 | 4.983 |  | 0.014991
43 | 75.552 | 6.9589 | 6.9569 | 32.6519 | 32.6535 | 6.8626 | 2.7068 | 20 | 20 | 366 | 6.8139 | 1.0201 | 6.951 |  | 0.00013554
4 | 75.724 | 8.0356 | 8.0389 | 32.7604 | 32.7597 | 6.6914 | 2.1598 | 4 | 4 | 20 | 6.6345 | 0.75183 | 4.988 |  | 0.01271
6 | 76.042 | 8.0396 | 8.0433 | 32.6555 | 32.6495 | 6.6954 | 2.5615 | 15 | 15 | 41 | 6.6386 | 0.94206 | 6.254 | 6 | 0.010055
82 | 99.552 | 4.9258 | 4.9262 | 32.7878 | 32.7878 | 7.1915 | 2.6131 | 19 | 19 | 649 | 7.1878 | 0.98069 | 7.049 |  | 0.0044634
109 | 99.993 | 6.8451 | 6.8454 | 33.7945 | 33.7944 | 6.8291 | 0.97375 | 3 | 3 | 777 | 6.8031 | 0.21417 | 1.457 |  | 0.0035269
21 | 100.524 | 7.8003 | 7.8029 | 33.1483 | 33.1478 | 6.7101 | 2.0612 | 14 | 14 | 184 | 6.6761 | 0.71718 | 4.788 |  | 0.0013215
4 | 100.677 | 7.3743 | 7.3738 | 33.4656 | 33.4679 | 6.7614 | 1.3923 | 2 | 2 | 18 | 6.7311 | 0.39845 | 2.682 |  | 0.0076443
55 | 100.924 | 6.4216 | 6.421 | 33.315 | 33.3134 | 6.9181 | 2.1064 | 19 | 19 | 449 | 6.8972 | 0.73639 | 5.079 |  | 0.0032178
43 | 100.925 | 6.7603 | 6.7601 | 33.1617 | 33.1608 | 6.871 | 2.347 | 19 | 19 | 365 | 6.8466 | 0.85517 | 5.855 | 6 | 0.0032019
27 | 101.065 | 6.9577 | 6.9577 | 33.2262 | 33.2266 | 6.8369 | 2.1581 | 18 | 18 | 239 | 6.8107 | 0.76145 | 5.186 |  | 0.0022839
6 | 101.068 | 8.0326 | 8.0329 | 33.2705 | 33.273 | 6.6697 | 1.8896 | 14 | 14 | 40 | 6.6344 | 0.63668 | 4.224 |  | 0.001667
4 | 107.087 | 7.1809 | 7.1807 | 33.6182 | 33.6202 | 6.7845 | 1.2653 | 1 | 1 | 17 | 6.7615 | 0.33972 | 2.297 |  | 0.0070593
103 | 121.028 | 5.4925 | 5.4929 | 33.6718 | 33.6715 | 7.0535 | 1.8774 | 4 | 4 | 747 | 7.0618 | 0.64219 | 4.535 |  | 0.0095578
103 | 121.29 | 5.4932 | 5.4938 | 33.6717 | 33.6706 | 7.0534 | 1.8798 | 5 | 5 | 748 | 7.0619 | 0.64317 | 4.542 | 6 | 0.0094174
21 | 124.006 | 7.5776 | 7.5785 | 33.6735 | 33.673 | 6.721 | 1.6429 | 13 | 13 | 183 | 6.7091 | 0.51796 | 3.475 |  | 0.0051603 |  |  |  | 67 | excl. outliers diff>0.03 | 0.4923 | -0.2152 | 0.4923 | -0.43713 | 0.9302
27 | 124.056 | 6.8708 | 6.8719 | 33.5832 | 33.5828 | 6.8346 | 1.8402 | 17 | 17 | 238 | 6.8293 | 0.61807 | 4.221 | 6 | 0.0028037 |  |  |  | 71 | excl. outliers diff>0.04 | 0.516 | -0.2697 | 0.516 | -0.52267 | 0.9539
55 | 124.682 | 6.3153 | 6.3158 | 33.7519 | 33.752 | 6.9152 | 1.8098 | 18 | 18 | 448 | 6.9166 | 0.60116 | 4.158 |  | 9.6156e-05
109 | 124.703 | 6.7841 | 6.7858 | 33.831 | 33.8303 | 6.837 | 0.9383 | 2 | 2 | 776 | 6.8333 | 0.19581 | 1.338 |  | 0.0017193
82 | 124.804 | 4.5838 | 4.5828 | 33.5622 | 33.5637 | 7.2131 | 1.7051 | 18 | 18 | 648 | 7.2381 | 0.55885 | 4.045 |  | 0.0066725
104 | 124.889 | 6.4464 | 6.4469 | 33.753 | 33.752 | 6.8942 | 1.2319 | 12 | 12 | 765 | 6.8942 | 0.34362 | 2.369 | 6 | 0.012423
43 | 125.341 | 6.7046 | 6.7035 | 33.747 | 33.7479 | 6.8534 | 1.967 | 18 | 18 | 364 | 6.851 | 0.6745 | 4.621 |  | 7.2609e-06
104 | 148.623 | 6.3271 | 6.3275 | 33.8543 | 33.8539 | 6.9086 | 1.0863 | 11 | 11 | 764 | 6.9312 | 0.2708 | 1.877 |  | 0.0075971
21 | 148.947 | 7.2605 | 7.2611 | 33.8437 | 33.8444 | 6.7622 | 1.4706 | 12 | 12 | 182 | 6.7748 | 0.44872 | 3.04 |  | 0.0060533
82 | 149.445 | 4.5242 | 4.5242 | 33.7219 | 33.7219 | 7.2156 | 1.494 | 17 | 17 | 647 | 7.2648 | 0.45714 | 3.321 |  | 0.0035542
6 | 150.041 | 7.6662 | 7.6568 | 33.7306 | 33.734 | 6.705 | 1.515 | 12 | 12 | 38 | 6.7147 | 0.45527 | 3.057 |  | 0.0081421
27 | 150.167 | 6.5732 | 6.573 | 33.772 | 33.7726 | 6.8731 | 1.6535 | 16 | 16 | 237 | 6.8942 | 0.52175 | 3.597 |  | 0.0063145
43 | 150.572 | 6.2835 | 6.2837 | 33.8126 | 33.813 | 6.9176 | 1.8294 | 17 | 17 | 363 | 6.9424 | 0.60944 | 4.231 |  | 0.00076305
55 | 150.618 | 5.9567 | 5.9577 | 33.8004 | 33.8006 | 6.9711 | 1.6806 | 17 | 17 | 447 | 7 | 0.54057 | 3.784 |  | 0.00019569
6 | 174.175 | 7.2207 | 7.2208 | 33.8368 | 33.8381 | 6.7686 | 1.4302 | 11 | 11 | 37 | 6.8036 | 0.42022 | 2.859 |  | 0.0035922
21 | 174.569 | 6.827 | 6.8309 | 33.873 | 33.8734 | 6.8284 | 1.487 | 11 | 11 | 181 | 6.8681 | 0.45587 | 3.131 |  | 0.0055314
55 | 174.681 | 5.601 | 5.6022 | 33.8159 | 33.8163 | 7.0288 | 1.5319 | 16 | 16 | 446 | 7.0845 | 0.48063 | 3.405 |  | 0.0093342
43 | 174.769 | 5.8637 | 5.8645 | 33.816 | 33.8163 | 6.9855 | 1.7151 | 16 | 16 | 362 | 7.0376 | 0.55871 | 3.932 |  | 0.0018417
27 | 175.458 | 6.2985 | 6.2986 | 33.8255 | 33.8261 | 6.9146 | 1.5559 | 15 | 15 | 236 | 6.9615 | 0.47547 | 3.31 |  | 0.0070236
82 | 175.796 | 4.3367 | 4.3321 | 33.7491 | 33.7496 | 7.247 | 1.4046 | 16 | 16 | 646 | 7.3243 | 0.41834 | 3.064 |  | 0.0065067
21 | 199.547 | 6.4857 | 6.4858 | 33.9015 | 33.9023 | 6.8811 | 1.3778 | 10 | 10 | 180 | 6.9472 | 0.40434 | 2.809 |  | 0.0049884
The following case was in an eddy and right at the minimum
104 | 998.536 | 3.0792 | 3.0795 | 34.3556 | 34.3555 | 7.4418 | 0.59581 | 3 | 3 | 756 | 8.4435 | 0.038965 | 0.329 |  | 0.0048218
These are near the min
6 | 1000.655 | 3.7272 | 3.727 | 34.3783 | 34.3802 | 7.3236 | 0.60164 | 4 | 4 | 30 | 8.311 | 0.028276 | 0.235 |  | 0.0085903 |  |  |  | E | m | b | m*0.6+b | m | b | m*0.6+b | Shallow-Deep
21 | 1000.796 | 3.3387 | 3.3386 | 34.3824 | 34.3835 | 7.3932 | 0.59948 | 4 | 4 | 174 | 8.3996 | 0.032978 | 0.277 |  | 0.0028793 |  |  |  | 0.035 | 0.4682 | -0.2447 | 0.03622 | 0.4892 | -0.262 | 0.55552 | -0.5193
43 | 1001.058 | 3.0121 | 3.0124 | 34.3695 | 34.3701 | 7.4535 | 0.60308 | 9 | 9 | 355 | 8.4766 | 0.032678 | 0.277 |  | 0.0048602 |  |  |  | 0.36 | 0.4679 | -0.2446 | 0.03614 | 0.4823 | -0.2578 | 0.54718 | -0.51104
55 | 1001.201 | 3.0048 | 3.0042 | 34.3811 | 34.3818 | 7.4542 | 0.60005 | 9 | 9 | 439 | 8.4778 | 0.033971 | 0.288 |  | 0.0021523 |  |  |  | 0.365 | 0.4678 | -0.2446 | 0.03608 | 0.4789 | -0.2557 | 0.54304 | -0.50696
82 | 1001.621 | 2.8856 | 2.8855 | 34.3808 | 34.3811 | 7.4762 | 0.62363 | 9 | 9 | 639 | 8.5064 | 0.046788 | 0.398 |  | 0.00034697 |  |  |  | 0.037 | 0.4677 | -0.2446 | 0.03602 | 0.4755 | -0.2536 | 0.5389 | -0.50288
27 | 1003.977 | 3.0658 | 3.0659 | 34.3705 | 34.3713 | 7.4436 | 0.60635 | 8 | 8 | 229 | 8.4674 | 0.03106 | 0.263 |  | 0.008005 |  |  |  | 0.0375 | 0.4676 | -0.2446 | 0.03596 | 0.4721 | -0.2515 | 0.53476 | -0.4988

## bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Sample_number | Oxygen:Dissolved | Flag:Oxygen:Dissolved
4 | 107.087 | 7.1809 | 7.1807 | 33.6182 | 33.6202 | 6.7845 | 1.2653 | 1 | 1 | 17 | 17 | 2.297
4 | 100.677 | 7.3743 | 7.3738 | 33.4656 | 33.4679 | 6.7614 | 1.3923 | 2 | 2 | 18 | 18 | 2.682
4 | 75.466 | 8.0358 | 8.0379 | 32.7611 | 32.7599 | 6.6914 | 2.1632 | 3 | 3 | 19 | 19 | 4.983
4 | 75.724 | 8.0356 | 8.0389 | 32.7604 | 32.7597 | 6.6914 | 2.1598 | 4 | 4 | 20 | 20 | 4.988
4 | 49.915 | 8.6094 | 8.6105 | 32.3922 | 32.393 | 6.6215 | 2.3872 | 5 | 5 | 21 | 21 | 5.533
4 | 25.056 | 11.2893 | 11.3208 | 31.7152 | 31.7027 | 6.272 | 2.9931 | 6 | 6 | 22 | 22 | 7.119
4 | 10.327 | 12.4786 | 12.4832 | 31.0801 | 31.0802 | 6.1407 | 3.3273 | 7 | 7 | 23 | 23 | 7.789 | 6
4 | 5.276 | 13.1047 | 13.1052 | 31.0883 | 31.0893 | 6.0608 | 3.2719 | 8 | 8 | 24 | 24 | 7.588
6 | 1319.765 | 2.941 | 2.941 | 34.4781 | 34.4802 | 7.461 | 0.65843 | 1 | 1 | 27 | 27 | 0.503
6 | 1250.62 | 3.1638 | 3.1637 | 34.4507 | 34.4526 | 7.4215 | 0.63595 | 2 | 2 | 28 | 28 | 0.399
6 | 1250.567 | 3.1639 | 3.1638 | 34.4507 | 34.4527 | 7.4215 | 0.63577 | 3 | 3 | 29 | 29 | 0.343 | 3
6 | 1000.655 | 3.7272 | 3.727 | 34.3783 | 34.3802 | 7.3236 | 0.60164 | 4 | 4 | 30 | 30 | 0.235
6 | 800.246 | 4.2241 | 4.2243 | 34.2848 | 34.2864 | 7.2405 | 0.59965 | 5 | 5 | 31 | 31 | 0.226 | 6
6 | 600.93 | 4.8252 | 4.8252 | 34.1875 | 34.189 | 7.1415 | 0.64709 | 6 | 6 | 32 | 32 | 0.394
6 | 400.746 | 5.5437 | 5.5438 | 34.07 | 34.0714 | 7.0263 | 0.78617 | 7 | 7 | 33 | 33 | -99 | 9
6 | 300.52 | 6.0939 | 6.093 | 34.009 | 34.0102 | 6.9391 | 0.95845 | 8 | 8 | 34 | 34 | 1.409
6 | 249.62 | 6.5754 | 6.5754 | 33.9704 | 33.9717 | 6.8637 | 1.052 | 9 | 9 | 35 | 35 | 1.687
6 | 200.09 | 6.9958 | 6.9962 | 33.9054 | 33.9067 | 6.8005 | 1.2671 | 10 | 10 | 36 | 36 | 2.344
6 | 174.175 | 7.2207 | 7.2208 | 33.8368 | 33.8381 | 6.7686 | 1.4302 | 11 | 11 | 37 | 37 | 2.859
6 | 150.041 | 7.6662 | 7.6568 | 33.7306 | 33.734 | 6.705 | 1.515 | 12 | 12 | 38 | 38 | 3.057
6 | 126.14 | 7.828 | 7.8282 | 33.599 | 33.6001 | 6.6862 | 1.6372 | 13 | 13 | 39 | 39 | 3.353
6 | 101.068 | 8.0326 | 8.0329 | 33.2705 | 33.273 | 6.6697 | 1.8896 | 14 | 14 | 40 | 40 | 4.224
6 | 76.042 | 8.0396 | 8.0433 | 32.6555 | 32.6495 | 6.6954 | 2.5615 | 15 | 15 | 41 | 41 | 6.254 | 6
6 | 50.428 | 8.7823 | 8.7945 | 32.501 | 32.5017 | 6.5914 | 2.6677 | 16 | 16 | 42 | 42 | 6.47
6 | 24.788 | 9.9713 | 9.9638 | 32.2923 | 32.2943 | 6.4294 | 2.9122 | 17 | 17 | 43 | 43 | 7.013
6 | 10.126 | 12.4547 | 12.4777 | 31.8238 | 31.8208 | 6.1152 | 2.8637 | 18 | 18 | 44 | 44 | 6.541
6 | 5.427 | 13.0021 | 13.0033 | 31.7916 | 31.7922 | 6.0471 | 2.8663 | 19 | 19 | 45 | 45 | 6.423
21 | 1999.165 | 1.9212 | 1.9202 | 34.593 | 34.5953 | 7.6471 | 0.84362 | 1 | 1 | 171 | 171 | 1.503
21 | 1500.421 | 2.4143 | 2.4149 | 34.5207 | 34.522 | 7.557 | 0.70475 | 2 | 2 | 172 | 172 | 0.8
21 | 1250.456 | 2.8634 | 2.8635 | 34.4702 | 34.4714 | 7.4758 | 0.64114 | 3 | 3 | 173 | 173 | 0.493 | 6
21 | 1000.796 | 3.3387 | 3.3386 | 34.3824 | 34.3835 | 7.3932 | 0.59948 | 4 | 4 | 174 | 174 | 0.277
21 | 799.394 | 3.694 | 3.6941 | 34.2832 | 34.2842 | 7.3342 | 0.6035 | 5 | 5 | 175 | 175 | 0.287
21 | 600.639 | 4.182 | 4.1823 | 34.167 | 34.1678 | 7.2536 | 0.64801 | 6 | 6 | 176 | 176 | 0.456
21 | 399.245 | 4.7728 | 4.7731 | 33.9897 | 33.9904 | 7.1599 | 0.90051 | 7 | 7 | 177 | 177 | 1.334
21 | 299.443 | 5.9089 | 5.9087 | 34.0045 | 34.0052 | 6.9694 | 0.94866 | 8 | 8 | 178 | 178 | 1.427
21 | 252.059 | 6.1465 | 6.1466 | 33.9666 | 33.9677 | 6.9326 | 1.0877 | 9 | 9 | 179 | 179 | 1.905
21 | 199.547 | 6.4857 | 6.4858 | 33.9015 | 33.9023 | 6.8811 | 1.3778 | 10 | 10 | 180 | 180 | 2.809
21 | 174.569 | 6.827 | 6.8309 | 33.873 | 33.8734 | 6.8284 | 1.487 | 11 | 11 | 181 | 181 | 3.131
21 | 148.947 | 7.2605 | 7.2611 | 33.8437 | 33.8444 | 6.7622 | 1.4706 | 12 | 12 | 182 | 182 | 3.04
21 | 124.006 | 7.5776 | 7.5785 | 33.6735 | 33.673 | 6.721 | 1.6429 | 13 | 13 | 183 | 183 | 3.475
21 | 100.524 | 7.8003 | 7.8029 | 33.1483 | 33.1478 | 6.7101 | 2.0612 | 14 | 14 | 184 | 184 | 4.788
21 | 74.597 | 7.963 | 7.9613 | 32.671 | 32.6743 | 6.7063 | 2.5358 | 15 | 15 | 185 | 185 | 6.216
21 | 50.844 | 8.6881 | 8.6899 | 32.5446 | 32.5455 | 6.6034 | 2.7391 | 16 | 16 | 186 | 186 | 6.727
21 | 43.277 | 9.0621 | 9.0608 | 32.4977 | 32.4987 | 6.5506 | 2.8193 | 20 | 17 | 187 | 187 | 6.666 | 6
21 | 33.821 | 9.4684 | 9.4679 | 32.4938 | 32.4948 | 6.4922 | 2.8893 | 21 | 18 | 188 | 188 | 6.488
21 | 25.183 | 11.2958 | 11.3125 | 32.1542 | 32.1473 | 6.2537 | 2.8689 | 19 | 19 | 189 | 189 | 6.494
27 | 3268.81 | 1.6763 | 1.6754 | 34.6464 | 34.6489 | 7.6917 | 0.94806 | 1 | 1 | 222 | 222 | 2.293
27 | 2998.656 | 1.6592 | 1.6583 | 34.6451 | 34.6474 | 7.6951 | 0.95762 | 2 | 2 | 223 | 223 | 2.258
27 | 2998.122 | 1.6593 | 1.6586 | 34.6451 | 34.6474 | 7.695 | 0.95761 | 3 | 3 | 224 | 224 | 2.26
27 | 2499.632 | 1.7466 | 1.7461 | 34.6271 | 34.6289 | 7.679 | 0.91738 | 4 | 4 | 225 | 225 | 1.884
27 | 2002.652 | 1.9236 | 1.9233 | 34.5914 | 34.5928 | 7.6467 | 0.84177 | 5 | 5 | 226 | 226 | 1.381
27 | 1500.017 | 2.3486 | 2.3488 | 34.5082 | 34.5092 | 7.57 | 0.68496 | 6 | 6 | 227 | 227 | 0.611
27 | 1252.294 | 2.6744 | 2.6745 | 34.4478 | 34.4487 | 7.512 | 0.6249 | 7 | 7 | 228 | 228 | 0.346
27 | 1003.977 | 3.0658 | 3.0659 | 34.3705 | 34.3713 | 7.4436 | 0.60635 | 8 | 8 | 229 | 229 | 0.263
27 | 799.429 | 3.7104 | 3.7106 | 34.3066 | 34.3074 | 7.3301 | 0.60456 | 9 | 9 | 230 | 230 | 0.245
27 | 598.466 | 4.0741 | 4.0745 | 34.1697 | 34.1703 | 7.2724 | 0.65777 | 10 | 10 | 231 | 231 | 0.437
27 | 397.184 | 4.7597 | 4.76 | 34.0024 | 34.0031 | 7.1615 | 0.86561 | 11 | 11 | 232 | 232 | 1.149
27 | 298.919 | 5.2316 | 5.2308 | 33.93 | 33.9308 | 7.085 | 1.0552 | 12 | 12 | 233 | 233 | 1.782
27 | 257.3 | 5.5579 | 5.5583 | 33.8894 | 33.8901 | 7.0325 | 1.26 | 13 | 13 | 234 | 234 | 2.459
27 | 200.07 | 6.0044 | 6.0046 | 33.8648 | 33.8655 | 6.9603 | 1.3979 | 14 | 14 | 235 | 235 | 2.867
27 | 175.458 | 6.2985 | 6.2986 | 33.8255 | 33.8261 | 6.9146 | 1.5559 | 15 | 15 | 236 | 236 | 3.31
27 | 150.167 | 6.5732 | 6.573 | 33.772 | 33.7726 | 6.8731 | 1.6535 | 16 | 16 | 237 | 237 | 3.597
27 | 124.056 | 6.8708 | 6.8719 | 33.5832 | 33.5828 | 6.8346 | 1.8402 | 17 | 17 | 238 | 238 | 4.221 | 6
27 | 101.065 | 6.9577 | 6.9577 | 33.2262 | 33.2266 | 6.8369 | 2.1581 | 18 | 18 | 239 | 239 | 5.186
27 | 75.193 | 7.2506 | 7.2506 | 32.7518 | 32.7527 | 6.8124 | 2.5159 | 19 | 19 | 240 | 240 | 6.528
27 | 50.523 | 8.1566 | 8.1573 | 32.5417 | 32.5419 | 6.6827 | 2.8114 | 20 | 20 | 241 | 241 | 7.029
27 | 26.326 | 10.537 | 10.5075 | 32.3305 | 32.3367 | 6.3493 | 2.8191 | 21 | 21 | 242 | 242 | 6.665
27 | 9.563 | 10.8971 | 10.8967 | 32.2487 | 32.2496 | 6.3035 | 2.8239 | 22 | 22 | 243 | 243 | 6.614
27 | 5.469 | 10.8993 | 10.8991 | 32.2479 | 32.2488 | 6.3032 | 2.8299 | 23 | 23 | 244 | 244 | 6.612
43 | 3682.007 | 1.5767 | 1.5755 | 34.6662 | 34.6688 | 7.71 | 0.99256 | 1 | 1 | 347 | 347 | 2.717
43 | 3501.538 | 1.5765 | 1.5755 | 34.6635 | 34.6659 | 7.7102 | 0.9952 | 2 | 2 | 348 | 348 | 2.679
43 | 3501.548 | 1.5764 | 1.5753 | 34.6636 | 34.6659 | 7.7102 | 0.99529 | 3 | 3 | 349 | 349 | 2.678
43 | 2998.778 | 1.5958 | 1.5953 | 34.6534 | 34.6552 | 7.707 | 0.98671 | 4 | 4 | 350 | 350 | 2.44
43 | 2499.656 | 1.7289 | 1.7283 | 34.6287 | 34.6302 | 7.6824 | 0.91726 | 5 | 5 | 351 | 351 | 1.918 | 6
43 | 2000.252 | 1.9316 | 1.9314 | 34.5888 | 34.5899 | 7.6453 | 0.81728 | 6 | 6 | 352 | 352 | 1.309
43 | 1499.285 | 2.3888 | 2.3887 | 34.498 | 34.4989 | 7.563 | 0.65375 | 7 | 7 | 353 | 353 | 0.499
43 | 1248.809 | 2.664 | 2.6642 | 34.443 | 34.4437 | 7.5142 | 0.61051 | 8 | 8 | 354 | 354 | 0.326
43 | 1001.058 | 3.0121 | 3.0124 | 34.3695 | 34.3701 | 7.4535 | 0.60308 | 9 | 9 | 355 | 355 | 0.277
43 | 799.425 | 3.4083 | 3.4087 | 34.2771 | 34.2775 | 7.3859 | 0.62773 | 10 | 10 | 356 | 356 | 0.352
43 | 600.012 | 3.8402 | 3.8405 | 34.1561 | 34.1566 | 7.3144 | 0.68769 | 11 | 11 | 357 | 357 | 0.575
43 | 399.953 | 4.3802 | 4.3806 | 33.9847 | 33.9851 | 7.2279 | 0.91753 | 12 | 12 | 358 | 358 | 1.372
43 | 298.167 | 4.7046 | 4.7052 | 33.8769 | 33.8773 | 7.177 | 1.187 | 13 | 13 | 359 | 359 | 2.276
43 | 249.733 | 4.9407 | 4.9411 | 33.8439 | 33.8444 | 7.1382 | 1.3283 | 14 | 14 | 360 | 360 | 2.727
43 | 200.645 | 5.5933 | 5.5935 | 33.8171 | 33.8175 | 7.03 | 1.6138 | 15 | 15 | 361 | 361 | 3.618
43 | 174.769 | 5.8637 | 5.8645 | 33.816 | 33.8163 | 6.9855 | 1.7151 | 16 | 16 | 362 | 362 | 3.932
43 | 150.572 | 6.2835 | 6.2837 | 33.8126 | 33.813 | 6.9176 | 1.8294 | 17 | 17 | 363 | 363 | 4.231
43 | 125.341 | 6.7046 | 6.7035 | 33.747 | 33.7479 | 6.8534 | 1.967 | 18 | 18 | 364 | 364 | 4.621
43 | 100.925 | 6.7603 | 6.7601 | 33.1617 | 33.1608 | 6.871 | 2.347 | 19 | 19 | 365 | 365 | 5.855 | 6
43 | 75.552 | 6.9589 | 6.9569 | 32.6519 | 32.6535 | 6.8626 | 2.7068 | 20 | 20 | 366 | 366 | 6.951
43 | 50.368 | 7.607 | 7.6074 | 32.613 | 32.6136 | 6.7633 | 2.7754 | 21 | 21 | 367 | 367 | 7.027
43 | 24.59 | 9.5452 | 9.545 | 32.5941 | 32.5945 | 6.4771 | 2.7839 | 22 | 22 | 368 | 368 | 6.735
43 | 9.814 | 9.5473 | 9.5472 | 32.5948 | 32.5954 | 6.4768 | 2.7917 | 23 | 23 | 369 | 369 | 6.735
43 | 4.69 | 9.5493 | 9.5493 | 32.5948 | 32.5955 | 6.4765 | 2.804 | 24 | 24 | 370 | 370 | 6.754 | 3
55 | 4024.718 | 1.544 | 1.5428 | 34.6753 | 34.6778 | 7.7159 | 1.0092 | 1 | 1 | 431 | 431 | 3.031
55 | 3500.075 | 1.5281 | 1.5273 | 34.6689 | 34.6707 | 7.7194 | 1.0226 | 2 | 2 | 432 | 432 | 2.88
55 | 3499.66 | 1.528 | 1.5272 | 34.6689 | 34.6708 | 7.7194 | 1.0227 | 3 | 3 | 433 | 433 | 2.269
55 | 2999.37 | 1.5766 | 1.5759 | 34.6547 | 34.6563 | 7.7106 | 0.99571 | 4 | 4 | 434 | 434 | 2.542 | 6
55 | 2499.752 | 1.717 | 1.7168 | 34.6271 | 34.6282 | 7.6848 | 0.91587 | 5 | 5 | 435 | 435 | 1.934
55 | 1999.977 | 1.9384 | 1.9384 | 34.5845 | 34.5854 | 7.6442 | 0.80487 | 6 | 6 | 436 | 436 | 1.297
55 | 1501.012 | 2.318 | 2.318 | 34.5064 | 34.5071 | 7.5759 | 0.67009 | 7 | 7 | 437 | 437 | 0.6
55 | 1250.517 | 2.6268 | 2.627 | 34.4432 | 34.4437 | 7.5211 | 0.6193 | 8 | 8 | 438 | 438 | 0.339
55 | 1001.201 | 3.0048 | 3.0042 | 34.3811 | 34.3818 | 7.4542 | 0.60005 | 9 | 9 | 439 | 439 | 0.288
55 | 801.098 | 3.3653 | 3.3657 | 34.2885 | 34.2889 | 7.3931 | 0.63111 | 10 | 10 | 440 | 440 | 0.405
55 | 599.177 | 3.8026 | 3.803 | 34.1781 | 34.1784 | 7.32 | 0.66088 | 11 | 11 | 441 | 441 | 0.478
55 | 399.417 | 4.343 | 4.3433 | 34.0227 | 34.0229 | 7.2325 | 0.80707 | 12 | 12 | 442 | 442 | 0.988
55 | 300.29 | 4.6272 | 4.6263 | 33.9182 | 33.9183 | 7.1883 | 1.0023 | 13 | 13 | 443 | 443 | 1.666
55 | 248.547 | 4.8555 | 4.856 | 33.8493 | 33.8497 | 7.1525 | 1.2952 | 14 | 14 | 444 | 444 | 2.631
55 | 201.961 | 5.306 | 5.3058 | 33.8329 | 33.8331 | 7.0771 | 1.4578 | 15 | 15 | 445 | 445 | 3.117 | 6
55 | 174.681 | 5.601 | 5.6022 | 33.8159 | 33.8163 | 7.0288 | 1.5319 | 16 | 16 | 446 | 446 | 3.405
55 | 150.618 | 5.9567 | 5.9577 | 33.8004 | 33.8006 | 6.9711 | 1.6806 | 17 | 17 | 447 | 447 | 3.784
55 | 124.682 | 6.3153 | 6.3158 | 33.7519 | 33.752 | 6.9152 | 1.8098 | 18 | 18 | 448 | 448 | 4.158
55 | 100.924 | 6.4216 | 6.421 | 33.315 | 33.3134 | 6.9181 | 2.1064 | 19 | 19 | 449 | 449 | 5.079
55 | 75.139 | 6.2183 | 6.2182 | 32.7005 | 32.7011 | 6.9793 | 2.5889 | 20 | 20 | 450 | 450 | 6.674
55 | 49.236 | 7.3467 | 7.3455 | 32.5224 | 32.5229 | 6.8076 | 2.7885 | 21 | 21 | 451 | 451 | 7.107
55 | 24.295 | 8.9042 | 8.9046 | 32.4982 | 32.4984 | 6.5736 | 2.8061 | 22 | 22 | 452 | 452 | 6.916
55 | 10.032 | 9.1871 | 9.189 | 32.4916 | 32.4916 | 6.5328 | 2.8063 | 23 | 23 | 453 | 453 | -99 | 9
55 | 5.104 | 9.2274 | 9.2301 | 32.4892 | 32.4896 | 6.527 | 2.8135 | 24 | 24 | 454 | 454 | 6.841
82 | 4317.458 | 1.5225 | 1.5212 | 34.6817 | 34.6842 | 7.7198 | 1.017 | 1 | 1 | 631 | 631 | 3.24
82 | 4001.058 | 1.4971 | 1.496 | 34.6797 | 34.6816 | 7.7249 | 1.0344 | 2 | 2 | 632 | 632 | 3.208
82 | 3500.265 | 1.5037 | 1.503 | 34.6715 | 34.673 | 7.724 | 1.0328 | 3 | 3 | 633 | 633 | 2.988
82 | 2999.653 | 1.5676 | 1.5671 | 34.6559 | 34.6571 | 7.7123 | 0.99655 | 4 | 4 | 634 | 634 | 2.577 | 6
82 | 2499.938 | 1.6963 | 1.6962 | 34.6301 | 34.631 | 7.6886 | 0.92286 | 5 | 5 | 635 | 635 | 2.019
82 | 1999.512 | 1.9214 | 1.9214 | 34.586 | 34.5867 | 7.6474 | 0.80818 | 6 | 6 | 636 | 636 | 1.321
82 | 1499.672 | 2.3026 | 2.3026 | 34.507 | 34.5074 | 7.5788 | 0.66725 | 7 | 7 | 637 | 637 | 0.605
82 | 1249.223 | 2.5596 | 2.5597 | 34.4526 | 34.453 | 7.5332 | 0.62271 | 8 | 8 | 638 | 638 | 0.41
82 | 1001.621 | 2.8856 | 2.8855 | 34.3808 | 34.3811 | 7.4762 | 0.62363 | 9 | 9 | 639 | 639 | 0.398
82 | 800.313 | 3.1943 | 3.1943 | 34.3054 | 34.3057 | 7.4233 | 0.64735 | 10 | 10 | 640 | 640 | 0.479
82 | 600.54 | 3.5596 | 3.5594 | 34.2099 | 34.2102 | 7.3619 | 0.6674 | 11 | 11 | 641 | 641 | 0.546 | 6
82 | 400.137 | 3.9463 | 3.9474 | 34.0547 | 34.0544 | 7.3006 | 0.82377 | 12 | 12 | 642 | 642 | 1.11
82 | 299.734 | 4.0202 | 4.0205 | 33.9277 | 33.928 | 7.2938 | 0.97286 | 13 | 13 | 643 | 643 | 1.612
82 | 250.171 | 4.0139 | 4.0145 | 33.8612 | 33.8611 | 7.2982 | 1.0987 | 14 | 14 | 644 | 644 | 2.034
82 | 200.12 | 4.1723 | 4.1723 | 33.776 | 33.7763 | 7.2744 | 1.3075 | 15 | 15 | 645 | 645 | 2.704
82 | 175.796 | 4.3367 | 4.3321 | 33.7491 | 33.7496 | 7.247 | 1.4046 | 16 | 16 | 646 | 646 | 3.064
82 | 149.445 | 4.5242 | 4.5242 | 33.7219 | 33.7219 | 7.2156 | 1.494 | 17 | 17 | 647 | 647 | 3.321
82 | 124.804 | 4.5838 | 4.5828 | 33.5622 | 33.5637 | 7.2131 | 1.7051 | 18 | 18 | 648 | 648 | 4.045
82 | 99.552 | 4.9258 | 4.9262 | 32.7878 | 32.7878 | 7.1915 | 2.6131 | 19 | 19 | 649 | 649 | 7.049
82 | 74.809 | 5.3152 | 5.3161 | 32.7479 | 32.7481 | 7.127 | 2.6585 | 20 | 20 | 650 | 650 | 7.116
82 | 48.978 | 6.844 | 6.8453 | 32.6945 | 32.6946 | 6.8789 | 2.7646 | 21 | 21 | 651 | 651 | 7.158
82 | 24.826 | 7.8416 | 7.8424 | 32.6853 | 32.6853 | 6.7242 | 2.7882 | 22 | 22 | 652 | 652 | 7.023 | 6
82 | 10.123 | 7.8549 | 7.8553 | 32.6853 | 32.6856 | 6.7221 | 2.7884 | 23 | 23 | 653 | 653 | 7.031
82 | 5.131 | 7.8589 | 7.8588 | 32.6853 | 32.6856 | 6.7215 | 2.7918 | 24 | 24 | 654 | 654 | 7.022
103 | 501.646 | 3.9267 | 3.9267 | 34.1038 | 34.1044 | 7.3017 | 0.78002 | 1 | 1 | 744 | 744 | 0.974
103 | 501.392 | 3.9257 | 3.9259 | 34.1044 | 34.1048 | 7.3018 | 0.77964 | 2 | 2 | 745 | 745 | 0.936
103 | 299.891 | 4.2868 | 4.2871 | 33.9014 | 33.9015 | 7.2482 | 1.0978 | 3 | 3 | 746 | 746 | 2.044
103 | 121.028 | 5.4925 | 5.4929 | 33.6718 | 33.6715 | 7.0535 | 1.8774 | 4 | 4 | 747 | 747 | 4.535
103 | 121.29 | 5.4932 | 5.4938 | 33.6717 | 33.6706 | 7.0534 | 1.8798 | 5 | 5 | 748 | 748 | 4.542 | 6
103 | 75.257 | 5.6248 | 5.6334 | 32.729 | 32.7284 | 7.0758 | 2.6462 | 6 | 6 | 749 | 749 | 7.032
103 | 40.812 | 7.9246 | 7.9356 | 32.6645 | 32.6636 | 6.7124 | 2.7659 | 7 | 7 | 750 | 750 | 6.993 | 6
103 | 9.511 | 8.1141 | 8.1143 | 32.6556 | 32.6554 | 6.6841 | 2.7835 | 8 | 8 | 751 | 751 | 6.955 | 6
103 | 4.971 | 8.1165 | 8.1166 | 32.6554 | 32.6553 | 6.6838 | 2.7806 | 9 | 9 | 752 | 752 | 6.955
103 | 2.264 | 8.1177 | 8.1179 | 32.6553 | 32.6551 | 6.6836 | 2.7473 | 10 | 10 | 753 | 753 | 6.964
104 | 2000.602 | 1.9661 | 1.9656 | 34.5781 | 34.5791 | 7.6393 | 0.78192 | 1 | 1 | 754 | 754 | 1.363
104 | 1499.13 | 2.4247 | 2.4248 | 34.4858 | 34.486 | 7.5568 | 0.63344 | 2 | 2 | 755 | 755 | 0.527
104 | 998.536 | 3.0792 | 3.0795 | 34.3556 | 34.3555 | 7.4418 | 0.59581 | 3 | 3 | 756 | 756 | 0.329
104 | 800.813 | 3.5377 | 3.5377 | 34.2583 | 34.2582 | 7.3635 | 0.61466 | 4 | 4 | 757 | 757 | 0.389
104 | 600.6 | 4.1256 | 4.1261 | 34.1241 | 34.1237 | 7.2656 | 0.66835 | 5 | 5 | 758 | 758 | 0.639
104 | 400.7 | 5.2829 | 5.285 | 33.9971 | 33.9967 | 7.0732 | 0.83654 | 6 | 6 | 759 | 759 | 1.13
104 | 301.053 | 5.7687 | 5.769 | 33.9589 | 33.9587 | 6.9944 | 0.92333 | 7 | 7 | 760 | 760 | 1.389
104 | 251.051 | 5.9598 | 5.9607 | 33.9361 | 33.9357 | 6.9643 | 0.95859 | 8 | 8 | 761 | 761 | 1.507
104 | 200.56 | 6.1409 | 6.1413 | 33.9074 | 33.9071 | 6.9362 | 1.0101 | 9 | 9 | 762 | 762 | 1.649
104 | 175.536 | 6.1963 | 6.1966 | 33.8955 | 33.8953 | 6.9278 | 1.0346 | 10 | 10 | 763 | 763 | 1.895 | 4
104 | 148.623 | 6.3271 | 6.3275 | 33.8543 | 33.8539 | 6.9086 | 1.0863 | 11 | 11 | 764 | 764 | 1.877
104 | 124.889 | 6.4464 | 6.4469 | 33.753 | 33.752 | 6.8942 | 1.2319 | 12 | 12 | 765 | 765 | 2.369 | 6
104 | 98.667 | 6.339 | 6.3374 | 33.325 | 33.322 | 6.931 | 1.8789 | 13 | 13 | 766 | 766 | 4.968
104 | 75.4 | 5.9601 | 5.9603 | 32.7533 | 32.7547 | 7.0191 | 2.5785 | 14 | 14 | 767 | 767 | 6.837
104 | 48.934 | 6.8098 | 6.8061 | 32.6183 | 32.6184 | 6.8878 | 2.7284 | 15 | 15 | 768 | 768 | 7.094
104 | 48.986 | 6.823 | 6.8245 | 32.6174 | 32.6159 | 6.8858 | 2.7328 | 20 | 16 | 769 | 769 | 6.825
104 | 38.648 | 8.483 | 8.4861 | 32.5718 | 32.5693 | 6.6326 | 2.7606 | 19 | 17 | 770 | 770 | 6.763
104 | 24.233 | 8.7103 | 8.7134 | 32.5674 | 32.567 | 6.5991 | 2.753 | 16 | 18 | 771 | 771 | 6.815
109 | 129.538 | 6.7261 | 6.7265 | 33.8678 | 33.8678 | 6.8445 | 0.91702 | 1 | 1 | 775 | 775 | 1.252 | 3
109 | 124.703 | 6.7841 | 6.7858 | 33.831 | 33.8303 | 6.837 | 0.9383 | 2 | 2 | 776 | 776 | 1.338
109 | 99.993 | 6.8451 | 6.8454 | 33.7945 | 33.7944 | 6.8291 | 0.97375 | 3 | 3 | 777 | 777 | 1.457
109 | 75.283 | 7.1816 | 7.1853 | 33.5313 | 33.5285 | 6.7883 | 1.2021 | 4 | 4 | 778 | 778 | 2.204 | 6
109 | 51.041 | 7.4989 | 7.4974 | 33.2578 | 33.2594 | 6.7514 | 1.3956 | 5 | 5 | 779 | 779 | 2.782
109 | 30.59 | 9.2121 | 9.2112 | 32.1596 | 32.1598 | 6.5432 | 2.1409 | 6 | 6 | 780 | 780 | 4.987
109 | 20.221 | 9.8119 | 9.7993 | 31.911 | 31.9138 | 6.4676 | 2.395 | 7 | 7 | 781 | 781 | 5.695
109 | 9.942 | 11.0822 | 11.1085 | 31.6461 | 31.6389 | 6.3026 | 3.3044 | 8 | 8 | 782 | 782 | 8.464
109 | 5.135 | 11.3135 | 11.3132 | 31.5968 | 31.5967 | 6.2735 | 3.6067 | 9 | 9 | 783 | 783 | 8.975

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers