## Notes
Analysis: | Hugh MacLean, OSP onboard CCGS JP Tully, 4-JUN-11 to 21-JUN-11 using IOS Kit #2. Data processing by Kenny Scozzafava
2011-10-26 |  | KS
5 Duplicate Niskin samples underwent a precision analysis but unique values are reported in QF spreasheet.
Casts 21 and 27 were analyzed with the incorrect KIO3 N (0.0101334) but recalculated using correct N of WAKO produced standard (0.0100000). This difference is reflected in the NEW OXY column in the Raw AutoOxy Output Worksheet
Cast 55 did not have a water seal and Hugh questioned whether or not they had been shaken a second time. Values may be off
Germaine, can youplease confirm this? #757 Duplicates flagged as outliers according to Chauvenet's criteria.
Just based on the plot in the 'Raw AutoOxy Output' Worksheet, the first of the two does not seem to track as well, therefore not used and only reported the second of the dups in QF worksheet.
There are quite a few samples that had comments about bubbles on the lens while the titration was occurring but no flags. I left them out of the QF
There were also many poorly chosen endpoints throughout the cruise, but corrections were made on all of those identified.

## QF 2011-26 OXY FINAL
HEADER: | Oxygen samples were analyzed at sea using an automated Winkler titration system
(Brinkmann Dosimat model 765 and a PC 950 Colorimeter controlled by AutoOxy software V3.4)
with modifications based on Carpenter (1965) and adhering to WOCE protocols (Culberson 1991). For details including a duplicate analysis, see document QF2011-26oxy.xls.
Cruise_Number | Event_Number | Station | Sample_number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
2011-26 | 4 | P2 | 17 | 2.2969 |  | 7.4
2011-26 | 4 | P2 | 18 | 2.6817 |  | 7.5
2011-26 | 4 | P2 | 19 | 4.9832 |  | 8
2011-26 | 4 | P2 | 20 | 4.9877 |  | 8.1
2011-26 | 4 | P2 | 21 | 5.533 |  | 8.5
2011-26 | 4 | P2 | 22 | 7.1189 |  | 11.1
2011-26 | 4 | P2 | 23 | 7.7895 | 6 | 12.2
2011-26 | 4 | P2 | 24 | 7.5878 |  | 12.9
2011-26 | 6 | P4 | 27 | 0.50268 |  | 3.9
2011-26 | 6 | P4 | 28 | 0.39944 |  | 4
2011-26 | 6 | P4 | 29 | 0.34293 | 3 | 4 | OXY: Air valve open. Mild outlier in COMPARE
2011-26 | 6 | P4 | 30 | 0.23463 |  | 4.6
2011-26 | 6 | P4 | 31 | 0.22583 | 6 | 5.3
2011-26 | 6 | P4 | 32 | 0.39376 |  | 5.7
2011-26 | 6 | P4 | 33 | -99 | 9 |  | No trip
2011-26 | 6 | P4 | 34 | 1.4089 |  | 6.7
2011-26 | 6 | P4 | 35 | 1.6865 |  | 7
2011-26 | 6 | P4 | 36 | 2.3436 |  | 7.3
2011-26 | 6 | P4 | 37 | 2.8591 |  | 7.5
2011-26 | 6 | P4 | 38 | 3.057 |  | 7.9
2011-26 | 6 | P4 | 39 | 3.3528 |  | 8
2011-26 | 6 | P4 | 40 | 4.2236 |  | 8.2
2011-26 | 6 | P4 | 41 | 6.254 | 6 | 8.25
2011-26 | 6 | P4 | 42 | 6.4697 |  | 8.8
2011-26 | 6 | P4 | 43 | 7.0128 |  | 9.9
2011-26 | 6 | P4 | 44 | 6.5407 |  | 12
2011-26 | 6 | P4 | 45 | 6.4225 |  | 12.7
2011-26 | 21 | P8 | 171 | 1.5027 |  | 2.6
2011-26 | 21 | P8 | 172 | 0.80008 |  | 3
2011-26 | 21 | P8 | 173 | 0.49308 | 6 | 3.45
2011-26 | 21 | P8 | 174 | 0.27736 |  | 3.9
2011-26 | 21 | P8 | 175 | 0.28702 |  | 4.4
2011-26 | 21 | P8 | 176 | 0.45554 |  | 4.9
2011-26 | 21 | P8 | 177 | 1.334 |  | 5.5
2011-26 | 21 | P8 | 178 | 1.4271 |  | 6.3
2011-26 | 21 | P8 | 179 | 1.9051 |  | 6.5
2011-26 | 21 | P8 | 180 | 2.8088 |  | 6.7
2011-26 | 21 | P8 | 181 | 3.1307 |  | 7
2011-26 | 21 | P8 | 182 | 3.0401 |  | 7.3
2011-26 | 21 | P8 | 183 | 3.4752 |  | 7.7
2011-26 | 21 | P8 | 184 | 4.7885 |  | 7.8
2011-26 | 21 | P8 | 185 | 6.2164 |  | 8
2011-26 | 21 | P8 | 186 | 6.7269 |  | 8.6
2011-26 | 21 | P8 | 187 | 6.6658 | 6 | 10.8
2011-26 | 21 | P8 | 188 | 6.4881 |  | 11.2
2011-26 | 21 | P8 | 189 | 6.494 |  | 11.1
2011-26 | 27 | P12 | 222 | 2.2928 |  | 2.1
2011-26 | 27 | P12 | 223 | 2.2579 |  | 2.1
2011-26 | 27 | P12 | 224 | 2.2597 |  | 2.1
2011-26 | 27 | P12 | 225 | 1.8838 |  | 2.3
2011-26 | 27 | P12 | 226 | 1.3813 |  | 2.6
2011-26 | 27 | P12 | 227 | 0.611 |  | 2.9
2011-26 | 27 | P12 | 228 | 0.34587 |  | 3.3
2011-26 | 27 | P12 | 229 | 0.26336 |  | 3.7
2011-26 | 27 | P12 | 230 | 0.24498 |  | 4.4
2011-26 | 27 | P12 | 231 | 0.43711 |  | 4.8
2011-26 | 27 | P12 | 232 | 1.1491 |  | 5.5
2011-26 | 27 | P12 | 233 | 1.7819 |  | 5.9
2011-26 | 27 | P12 | 234 | 2.4588 |  | 6.1
2011-26 | 27 | P12 | 235 | 2.8674 |  | 6.5
2011-26 | 27 | P12 | 236 | 3.3098 |  | 6.7
2011-26 | 27 | P12 | 237 | 3.5967 |  | 6.9
2011-26 | 27 | P12 | 238 | 4.2209 | 6 | 7.15
2011-26 | 27 | P12 | 239 | 5.186 |  | 7.2
2011-26 | 27 | P12 | 240 | 6.5277 |  | 7.4
2011-26 | 27 | P12 | 241 | 7.0285 |  | 8.1
2011-26 | 27 | P12 | 242 | 6.6649 |  | 9.9
2011-26 | 27 | P12 | 243 | 6.6143 |  | 10.1
2011-26 | 27 | P12 | 244 | 6.6121 |  | 10.1
2011-26 | 43 | P16 | 347 | 2.7168 |  | 1.9
2011-26 | 43 | P16 | 348 | 2.679 |  | 1.8
2011-26 | 43 | P16 | 349 | 2.6776 |  | 1.8
2011-26 | 43 | P16 | 350 | 2.4396 |  | 1.9
2011-26 | 43 | P16 | 351 | 1.9182 | 6 | 2.25
2011-26 | 43 | P16 | 352 | 1.3089 |  | 2.4
2011-26 | 43 | P16 | 353 | 0.49946 |  | 2.9
2011-26 | 43 | P16 | 354 | 0.32584 |  | 3.1
2011-26 | 43 | P16 | 355 | 0.27695 |  | 3.5
2011-26 | 43 | P16 | 356 | 0.3517 |  | 4
2011-26 | 43 | P16 | 357 | 0.57543 |  | 4.5
2011-26 | 43 | P16 | 358 | 1.3717 |  | 5.1
2011-26 | 43 | P16 | 359 | 2.2761 |  | 5.7
2011-26 | 43 | P16 | 360 | 2.727 |  | 5.7
2011-26 | 43 | P16 | 361 | 3.618 |  | 6.2
2011-26 | 43 | P16 | 362 | 3.9322 |  | 6.3
2011-26 | 43 | P16 | 363 | 4.2308 |  | 6.7
2011-26 | 43 | P16 | 364 | 4.6212 |  | 6.9
2011-26 | 43 | P16 | 365 | 5.8549 | 6 | 7.15
2011-26 | 43 | P16 | 366 | 6.9507 |  | 7.2
2011-26 | 43 | P16 | 367 | 7.0265 |  | 7.7
2011-26 | 43 | P16 | 368 | 6.7346 |  | 9.2
2011-26 | 43 | P16 | 369 | 6.735 |  | 9.3
2011-26 | 43 | P16 | 370 | 6.7535 | 3 | 9.2 | OXY: Valve open. Bubble on lens. Looks ok in COMPARE
2011-26 | 55 | P20 | 431 | 3.031 |  | 1.7
2011-26 | 55 | P20 | 432 | 2.8803 |  | 1.7
2011-26 | 55 | P20 | 433 | -99 | 5 | 1.8 | OXY: Possible mistrip.
2011-26 | 55 | P20 | 434 | 2.5417 | 6 | 1.8
2011-26 | 55 | P20 | 435 | 1.9339 |  | 2
2011-26 | 55 | P20 | 436 | 1.2966 |  | 2.2
2011-26 | 55 | P20 | 437 | 0.59958 |  | 2.6
2011-26 | 55 | P20 | 438 | 0.33947 |  | 2.9
2011-26 | 55 | P20 | 439 | 0.28773 |  | 3.3
2011-26 | 55 | P20 | 440 | 0.40521 |  | 3.8
2011-26 | 55 | P20 | 441 | 0.47847 |  | 4.2
2011-26 | 55 | P20 | 442 | 0.98836 |  | 4.9
2011-26 | 55 | P20 | 443 | 1.6663 |  | 5.2
2011-26 | 55 | P20 | 444 | 2.6315 |  | 5.5
2011-26 | 55 | P20 | 445 | 3.1166 | 6 | 5.95
2011-26 | 55 | P20 | 446 | 3.4055 |  | 6.1
2011-26 | 55 | P20 | 447 | 3.7837 |  | 6.4
2011-26 | 55 | P20 | 448 | 4.1581 |  | 6.6
2011-26 | 55 | P20 | 449 | 5.0794 |  | 6.8
2011-26 | 55 | P20 | 450 | 6.6742 |  | 6.6
2011-26 | 55 | P20 | 451 | 7.1074 |  | 7.5
2011-26 | 55 | P20 | 452 | 6.9163 |  | 8.6
2011-26 | 55 | P20 | 453 | -99 | 9 |  | No trip
2011-26 | 55 | P20 | 454 | 6.8411 |  | 8.9
2011-26 | 82 | P26 | 631 | 3.2399 |  | 1.4
2011-26 | 82 | P26 | 632 | 3.2079 |  | 1.4
2011-26 | 82 | P26 | 633 | 2.9879 |  | 1.4
2011-26 | 82 | P26 | 634 | 2.5774 | 6 | 1.6
2011-26 | 82 | P26 | 635 | 2.0189 |  | 1.7
2011-26 | 82 | P26 | 636 | 1.3209 |  | 2
2011-26 | 82 | P26 | 637 | 0.60546 |  | 2.4
2011-26 | 82 | P26 | 638 | 0.41045 |  | 2.7
2011-26 | 82 | P26 | 639 | 0.39827 |  | 3.1
2011-26 | 82 | P26 | 640 | 0.47906 |  | 3.3
2011-26 | 82 | P26 | 641 | 0.54624 | 6 | 3.85
2011-26 | 82 | P26 | 642 | 1.1105 |  | 4.3
2011-26 | 82 | P26 | 643 | 1.6121 |  | 4.3
2011-26 | 82 | P26 | 644 | 2.0336 |  | 4.4
2011-26 | 82 | P26 | 645 | 2.7042 |  | 4.6
2011-26 | 82 | P26 | 646 | 3.0643 |  | 4.8
2011-26 | 82 | P26 | 647 | 3.3212 |  | 5.1
2011-26 | 82 | P26 | 648 | 4.0446 |  | 5.2
2011-26 | 82 | P26 | 649 | 7.0491 |  | 5.5
2011-26 | 82 | P26 | 650 | 7.116 |  | 5.8
2011-26 | 82 | P26 | 651 | 7.1577 |  | 7
2011-26 | 82 | P26 | 652 | 7.0233 | 6 | 7.65
2011-26 | 82 | P26 | 653 | 7.0314 |  | 7.7
2011-26 | 82 | P26 | 654 | 7.0216 |  | 7.7
2011-26 | 103 | Uwfloat | 744 | 0.97434 |  | 4.2
2011-26 | 103 | Uwfloat | 745 | 0.93571 |  | 4.1
2011-26 | 103 | Uwfloat | 746 | 2.0443 |  | 4.5
2011-26 | 103 | Uwfloat | 747 | 4.5351 |  | 5.9
2011-26 | 103 | Uwfloat | 748 | 4.5421 | 6 | 5.8
2011-26 | 103 | Uwfloat | 749 | 7.0318 |  | 5.9
2011-26 | 103 | Uwfloat | 750 | 6.9925 | 6 | 7.7
2011-26 | 103 | Uwfloat | 751 | 6.9553 | 6 | 7.9
2011-26 | 103 | Uwfloat | 752 | 6.9551 |  | 7.9
2011-26 | 103 | Uwfloat | 753 | 6.9641 |  | 7.9
2011-26 | 104 | Eddy | 754 | 1.3625 |  | 2.6
2011-26 | 104 | Eddy | 755 | 0.52664 |  | 2.7
2011-26 | 104 | Eddy | 756 | 0.32935 |  | 3.3
2011-26 | 104 | Eddy | 757 | 0.389 |  | 4 | OXY: Duplicate not used
2011-26 | 104 | Eddy | 758 | 0.63897 |  | 4.5
2011-26 | 104 | Eddy | 759 | 1.1303 |  | 5.7
2011-26 | 104 | Eddy | 760 | 1.389 |  | 6.1
2011-26 | 104 | Eddy | 761 | 1.5072 |  | 6.2
2011-26 | 104 | Eddy | 762 | 1.6494 |  | 6.4
2011-26 | 104 | Eddy | 763 | 1.8945 | 4 | 6.5 | OXY: Endpoint overshot due to obstruction of light path. Difficult to know where real endpoint is. Confirmed in COMPARE
2011-26 | 104 | Eddy | 764 | 1.8775 |  | 6.5
2011-26 | 104 | Eddy | 765 | 2.3693 | 6 | 6.6
2011-26 | 104 | Eddy | 766 | 4.9678 | 3 | 6.5 | OXY: outlier when compared to CTD data
2011-26 | 104 | Eddy | 767 | 6.8372 |  | 6.5
2011-26 | 104 | Eddy | 768 | 7.0942 |  | 6.9
2011-26 | 104 | Eddy | 769 | 6.8254 |  | 8.3
2011-26 | 104 | Eddy | 770 | 6.763 |  | 8.4
2011-26 | 104 | Eddy | 771 | 6.8145 |  | 8.4
2011-26 | 109 | LB08 | 775 | 1.2524 | 3 | 6.9 | OXY: sample redrawn in same flask. Looks ok in COMPARE
2011-26 | 109 | LB08 | 776 | 1.3376 |  | 6.9
2011-26 | 109 | LB08 | 777 | 1.4574 |  | 7
2011-26 | 109 | LB08 | 778 | 2.2036 | 6 | 7.3
2011-26 | 109 | LB08 | 779 | 2.7819 |  | 7.6
2011-26 | 109 | LB08 | 780 | 4.9871 |  | 9
2011-26 | 109 | LB08 | 781 | 5.6951 |  | 9.6
2011-26 | 109 | LB08 | 782 | 8.464 |  | 10.8
2011-26 | 109 | LB08 | 783 | 8.9747 |  | 10.9

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Sp for Dissolved Oxygen for the range 0.224 to 7.791 mL/L was 0.009, k = 19 with 1 outlier removed
Precision Statement for Replicate Samples Drawn From two Niskin Bottles Closed at Same Depth:
Sp for Dissolved Oxygen for the range 0.343 to 4.988 mL/L was 0.024, k = 5 with 1 outlier removed
Precision Calculation for Duplicate Samples:
Precision was determined by analyzing replicate samples drawn from one Niskin and from two Niskins closed at the same depth
Pooled standard deviation of pairs of samples (Sp) was calucalated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs
Determination of Outliers
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded
Duplicate Analysis for Samples Drawn From a Single Niskin bottle:
Event # | Sample# | Pressure | DO1 | DO2 | DO Avg | Draw T1 | Draw T2 | Draw T Avg | Difference | Diff^2 | Chauvenet | Reject? | Comment | Critical values for Chauvenet statistic
P2 | 23 | 10 | 7.7878 | 7.7912 | 7.7895 | 12.2 | 12.2 | 12.2 | 0.003419 | 1.1689e-05 | 0.25643 |  |  | n | Chauvenet
P4 | 31 | 800 | 0.22737 | 0.2243 | 0.22583 | 5.2 | 5.4 | 5.3 | -0.0030691 | 9.4192e-06 | 0.20003 |  |  | 3 | 1.37
P4 | 41 | 75 | 6.2527 | 6.2553 | 6.254 | 8.2 | 8.3 | 8.25 | 0.0025672 | 6.5906e-06 | 0.24903 |  |  | 4 | 1.53
P8 | 173 | 1250 | 0.47351 | 0.51264 | 0.49308 | 3.4 | 3.5 | 3.45 | 0.039132 | 0.0015313 | 0.56688 |  |  | 5 | 1.64
P8 | 187 | 25 | 6.6568 | 6.6747 | 6.6658 | 10.8 | 10.8 | 10.8 | 0.017891 | 0.0003201 | 0.38224 |  |  | 6 | 1.73
P12 | 238 | 125 | 4.2206 | 4.2213 | 4.2209 | 7.1 | 7.2 | 7.15 | 0.00070334 | 4.9469e-07 | 0.23283 |  |  | 7 | 1.8
P16 | 351 | 2500 | 1.9195 | 1.917 | 1.9182 | 2.1 | 2.4 | 2.25 | -0.0025291 | 6.3964e-06 | 0.20473 |  |  | 8 | 1.87
P16 | 365 | 100 | 5.8523 | 5.8575 | 5.8549 | 7.1 | 7.2 | 7.15 | 0.0052313 | 2.7366e-05 | 0.27219 |  |  | 9 | 1.91
P20 | 434 | 3000 | 2.5526 | 2.5307 | 2.5417 | 1.7 | 1.9 | 1.8 | -0.021905 | 0.00047984 | 0.036297 |  |  | 10 | 1.97
P20 | 445 | 200 | 3.1179 | 3.1154 | 3.1166 | 5.9 | 6 | 5.95 | -0.0024804 | 6.1522e-06 | 0.20515 |  |  | 20 | 2.24
P26 | 634 | 3000 | 2.576 | 2.5787 | 2.5774 | 1.5 | 1.7 | 1.6 | 0.0027441 | 7.53e-06 | 0.25057 |  |  | 30 | 2.39
P26 | 641 | 600 | 0.55229 | 0.54018 | 0.54624 | 3.8 | 3.9 | 3.85 | -0.012104 | 0.0001465 | 0.1215 |  |  | 50 | 2.57
P26 | 652 | 25 | 7.0258 | 7.0207 | 7.0233 | 7.6 | 7.7 | 7.65 | -0.0051538 | 2.6562e-05 | 0.18191 |  |  | 80 | 2.74
Uwfloat | 748 | 120 | 4.5435 | 4.5406 | 4.5421 | 5.8 | 5.8 | 5.8 | -0.0028788 | 8.2874e-06 | 0.20169 |  |  | 100 | 2.81
Uwfloat | 750 | 40 | 6.9958 | 6.9892 | 6.9925 | 7.7 | 7.7 | 7.7 | -0.0065816 | 4.3318e-05 | 0.1695 |  |  | 150 | 2.93
Uwfloat | 751 | 10 | 6.9575 | 6.9532 | 6.9553 | 7.9 | 7.9 | 7.9 | -0.0043214 | 1.8674e-05 | 0.18915 |  |  | 200 | 3.02
Eddy | 757 | 800 | 0.88726 | 0.38894 | 0.6381 | 3.9 | 4 | 3.95 | -0.49832 | 0.24832 | 4.105 | y
Eddy | 765 | 125 | 2.3671 | 2.3715 | 2.3693 | 6.6 | 6.6 | 6.6 | 0.0043513 | 1.8933e-05 | 0.26454
LB08 | 778 | 75 | 2.2097 | 2.1974 | 2.2036 | 7.3 | 7.3 | 7.3 | -0.012238 | 0.00014976 | 0.12033
[8 empty] | mean diff | -0.026081
[8 empty] | stdev | 0.11504
[8 empty] | count |  | 19
[8 empty] | SUM |  | 0.25114
[8 empty] | Sp |  | 0.081295
[8 empty] | MIN/MAX | 0.2243 | 7.7912
1 Outlier Removed
[8 empty] | count |  | 18
[8 empty] | SUM |  | 0.002819
[8 empty] | Sp |  | 0.008849
Duplicate Samples From two Niskin Bottles Closed at Same Depth
Event # | Sample# | Pressure | DO1 | DO2 | DO Avg | Draw T1 | Draw T2 | Draw T Avg | Difference | Diff^2 | Chauvenet | Reject? | Comment | Critical values for Chauvenet statistic
4 | 19/20 | 75 | 4.9832 | 4.9877 | 4.9855 | 8 | 8.1 | 8.05 | 0.0045286 | 2.0508e-05 | 0.54583 |  |  | n | Chauvenet
6 | 28/29 | 1250 | 0.39944 | 0.34293 | 0.37118 | 4 | 4 | 4 | -0.056506 | 0.0031929 | 0.31548 |  |  | 3 | 1.37
27 | 223/224 | 3000 | 2.2579 | 2.2597 | 2.2588 | 2.1 | 2.1 | 2.1 | 0.0018828 | 3.545e-06 | 0.53585 |  |  | 4 | 1.53
55 | 432/433 | 3500 | 2.8803 | 2.2686 | 2.5744 | 1.7 | 1.8 | 1.75 | -0.61175 | 0.37423 | 1.7801 | y |  | 5 | 1.64
103 | 744/745 | 500 | 0.97434 | 0.93571 | 0.95503 | 4.2 | 4.1 | 4.15 | -0.038631 | 0.0014924 | 0.38294 |  |  | 6 | 1.73
[14 empty] | 7 | 1.8
[8 empty] | mean diff | -0.14009 |  |  |  |  | 8 | 1.87
[8 empty] | stdev | 0.26496 |  |  |  |  | 9 | 1.91
[8 empty] | count |  | 5 |  |  |  | 10 | 1.97
[8 empty] | SUM |  | 0.37894 |  |  |  | 20 | 2.24
[8 empty] | Sp |  | 0.19466 |  |  |  | 30 | 2.39
[8 empty] | MIN/MAX | 0.34293 | 4.9877 |  |  |  | 50 | 2.57
[14 empty] | 80 | 2.74
[8 empty] | 1 Outlier Removed | [5 empty] | 100 | 2.81
[8 empty] | count |  | 4 |  |  |  | 150 | 2.93
[8 empty] | SUM |  | 0.0047094 |  |  |  | 200 | 3.02
[8 empty] | Sp |  | 0.024263

## Flags
Institute of Ocean Sciences -Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## Raw AutoOxy Data
|  |  | Count | 197 | Min/Max | 0.2243 | 8.9747
File Name | STN:STATION | LOC:EVENT_NUMBER | Sample_number | Depth | Oxygen:Dissolved | Thiosulfate_titer | Flask_Number | Flask_Volume | Blank | Standard | KIO3 Std volume ml | KIO3 Normality | THIO N | Draw Temp | NEW OXY | Flag | Comment
2011-26-0004.oxy | P2 | 4 | 17 | 107 | 2.297 | 0.415 | 2068 | 141.98 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7.4 | 2.2969
2011-26-0004.oxy | P2 | 4 | 18 | 100 | 2.682 | 0.466 | 1144 | 136.84 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7.5 | 2.6817
2011-26-0004.oxy | P2 | 4 | 19 | 75 | 4.983 | 0.858 | 631 | 136.38 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8 | 4.9832
2011-26-0004.oxy | P2 | 4 | 20 | 75 | 4.988 | 0.849 | 468 | 134.84 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.1 | 4.9877
2011-26-0004.oxy | P2 | 4 | 21 | 50 | 5.533 | 0.909 | 1475 | 130.27 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.5 | 5.533
2011-26-0004.oxy | P2 | 4 | 22 | 25 | 7.119 | 1.305 | 1248 | 145.41 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 11.1 | 7.1189 |  | Bubble on lens
2011-26-0004.oxy | P2 | 4 | 23 | 10 | 7.788 | 1.391 | 3028 | 141.77 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 12.2 | 7.7878
2011-26-0004.oxy | P2 | 4 | 23 | 10 | 7.791 | 1.323 | 1397 | 134.85 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 12.2 | 7.7912
2011-26-0004.oxy | P2 | 4 | 24 | 5 | 7.588 | 1.292 | 1373 | 135.2 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 12.9 | 7.5878
2011-26-0006.oxy | P4 | 6 | 27 | 1320 | 0.503 | 0.096 | 1298 | 142.91 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 3.9 | 0.50268
2011-26-0006.oxy | P4 | 6 | 28 | 1250 | 0.399 | 0.078 | 2014 | 143.92 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 4 | 0.39944
2011-26-0006.oxy | P4 | 6 | 29 | 1250 | 0.343 | 0.062 | 1408 | 130.64 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 4 | 0.34293 | 3 | valve open
2011-26-0006.oxy | P4 | 6 | 30 | 1000 | 0.235 | 0.045 | 1307 | 133.08 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 4.6 | 0.23463
2011-26-0006.oxy | P4 | 6 | 31 | 800 | 0.227 | 0.044 | 3206 | 133.81 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 5.2 | 0.22737
2011-26-0006.oxy | P4 | 6 | 31 | 800 | 0.224 | 0.043 | 1392 | 132.11 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 5.4 | 0.2243
2011-26-0006.oxy | P4 | 6 | 32 | 600 | 0.394 | 0.072 | 3090 | 133.99 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 5.7 | 0.39376
2011-26-0006.oxy | P4 | 6 | 33 | 400 | [11 empty] | 9 | no trip
2011-26-0006.oxy | P4 | 6 | 34 | 300 | 1.409 | 0.263 | 3126 | 145.42 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 6.7 | 1.4089
2011-26-0006.oxy | P4 | 6 | 35 | 250 | 1.686 | 0.286 | 1378 | 132.53 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7 | 1.6865
2011-26-0006.oxy | P4 | 6 | 36 | 200 | 2.344 | 0.385 | 647 | 129.13 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7.3 | 2.3436
2011-26-0006.oxy | P4 | 6 | 37 | 175 | 2.859 | 0.482 | 3157 | 132.87 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7.5 | 2.8591
2011-26-0006.oxy | P4 | 6 | 38 | 150 | 3.057 | 0.505 | 621 | 130.31 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 7.9 | 3.057
2011-26-0006.oxy | P4 | 6 | 39 | 125 | 3.353 | 0.56 | 1261 | 131.88 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8 | 3.3528
2011-26-0006.oxy | P4 | 6 | 40 | 100 | 4.224 | 0.68 | 1146 | 127.43 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.2 | 4.2236
2011-26-0006.oxy | P4 | 6 | 41 | 75 | 6.253 | 1.024 | 528 | 129.96 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.2 | 6.2527
2011-26-0006.oxy | P4 | 6 | 41 | 75 | 6.255 | 1.028 | 371 | 130.41 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.3 | 6.2553
2011-26-0006.oxy | P4 | 6 | 42 | 50 | 6.47 | 1.114 | 196 | 136.6 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 8.8 | 6.4697
2011-26-0006.oxy | P4 | 6 | 43 | 25 | 7.013 | 1.208 | 569 | 136.71 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 9.9 | 7.0128
2011-26-0006.oxy | P4 | 6 | 44 | 10 | 6.541 | 1.118 | 3182 | 135.62 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 12 | 6.5407
2011-26-0006.oxy | P4 | 6 | 45 | 5 | 6.423 | 1.089 | 1374 | 134.53 | 0.0037 | 0.7267 | 10.0158 | 0.010133 | 0.14038 | 12.7 | 6.4225
2011-26-0021.oxy | P8 | 21 | 171 | 2000 | 1.523 | 0.274 | 1298 | 142.91 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.6 | 1.5027 |  | Casts 21 and 27 were analyzed with the incorrect KIO3 N (0.0101334) but recalculated using correct N of WAKO produced standard (0.0100000)
2011-26-0021.oxy | P8 | 21 | 172 | 1500 | 0.811 | 0.149 | 2014 | 143.92 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3 | 0.80008
2011-26-0021.oxy | P8 | 21 | 173 | 1250 | 0.48 | 0.082 | 1408 | 130.64 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.4 | 0.47351
2011-26-0021.oxy | P8 | 21 | 173 | 1250 | 0.52 | 0.09 | 1307 | 133.08 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.5 | 0.51264
2011-26-0021.oxy | P8 | 21 | 174 | 1000 | 0.281 | 0.051 | 3206 | 133.81 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.9 | 0.27736
2011-26-0021.oxy | P8 | 21 | 175 | 800 | 0.291 | 0.052 | 1392 | 132.11 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.4 | 0.28702
2011-26-0021.oxy | P8 | 21 | 176 | 600 | 0.462 | 0.081 | 3090 | 133.99 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.9 | 0.45554
2011-26-0021.oxy | P8 | 21 | 177 | 400 | 1.352 | 0.249 | 1404 | 146.01 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.5 | 1.334 |  | Bubble on lens
2011-26-0021.oxy | P8 | 21 | 178 | 300 | 1.446 | 0.265 | 3126 | 145.42 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.3 | 1.4271
2011-26-0021.oxy | P8 | 21 | 179 | 250 | 1.931 | 0.321 | 1378 | 132.53 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.5 | 1.9051
2011-26-0021.oxy | P8 | 21 | 180 | 200 | 2.847 | 0.459 | 647 | 129.13 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.7 | 2.8088
2011-26-0021.oxy | P8 | 21 | 181 | 175 | 3.173 | 0.526 | 3157 | 132.87 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7 | 3.1307
2011-26-0021.oxy | P8 | 21 | 182 | 150 | 3.081 | 0.501 | 621 | 130.31 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.3 | 3.0401
2011-26-0021.oxy | P8 | 21 | 183 | 125 | 3.522 | 0.579 | 1261 | 131.88 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.7 | 3.4752
2011-26-0021.oxy | P8 | 21 | 184 | 100 | 4.853 | 0.769 | 1146 | 127.43 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.8 | 4.7885
2011-26-0021.oxy | P8 | 21 | 185 | 75 | 6.299 | 1.017 | 528 | 129.96 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 8 | 6.2164
2011-26-0021.oxy | P8 | 21 | 186 | 50 | 6.817 | 1.104 | 371 | 130.41 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 8.6 | 6.7269
2011-26-0021.oxy | P8 | 21 | 187 | 25 | 6.746 | 1.145 | 196 | 136.6 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 10.8 | 6.6568
2011-26-0021.oxy | P8 | 21 | 187 | 25 | 6.764 | 1.149 | 569 | 136.71 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 10.8 | 6.6747
2011-26-0021.oxy | P8 | 21 | 188 | 10 | 6.575 | 1.108 | 3182 | 135.62 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 11.2 | 6.4881
2011-26-0021.oxy | P8 | 21 | 189 | 5 | 6.581 | 1.1 | 1374 | 134.53 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 11.1 | 6.494
2011-26-0027.oxy | P12 | 27 | 222 | 3270 | 2.307 | 0.413 | 2068 | 141.98 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.1 | 2.2928 |  | Endpoint early at 0.4100. Recalculated using new thio titer of 0.4130
2011-26-0027.oxy | P12 | 27 | 223 | 3000 | 2.288 | 0.392 | 1144 | 136.84 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.1 | 2.2579 |  | Bubble on lens
2011-26-0027.oxy | P12 | 27 | 224 | 3000 | 2.29 | 0.391 | 631 | 136.38 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.1 | 2.2597
2011-26-0027.oxy | P12 | 27 | 225 | 2500 | 1.909 | 0.323 | 468 | 134.84 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.3 | 1.8838
2011-26-0027.oxy | P12 | 27 | 226 | 2000 | 1.4 | 0.23 | 1475 | 130.27 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.6 | 1.3813
2011-26-0027.oxy | P12 | 27 | 227 | 1500 | 0.619 | 0.116 | 1248 | 145.41 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.9 | 0.611
2011-26-0027.oxy | P12 | 27 | 228 | 1250 | 0.351 | 0.066 | 3028 | 141.77 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.3 | 0.34587
2011-26-0027.oxy | P12 | 27 | 229 | 1000 | 0.267 | 0.049 | 1397 | 134.85 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.7 | 0.26336
2011-26-0027.oxy | P12 | 27 | 230 | 800 | 0.248 | 0.046 | 1373 | 135.2 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.4 | 0.24498
2011-26-0027.oxy | P12 | 27 | 231 | 600 | 0.443 | 0.083 | 3151 | 143.15 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.8 | 0.43711
2011-26-0027.oxy | P12 | 27 | 232 | 400 | 1.165 | 0.197 | 1318 | 133.63 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.5 | 1.1491
2011-26-0027.oxy | P12 | 27 | 233 | 300 | 1.806 | 0.295 | 745 | 130.09 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.9 | 1.7819
2011-26-0027.oxy | P12 | 27 | 234 | 250 | 2.492 | 0.455 | 2020 | 145.95 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.1 | 2.4588
2011-26-0027.oxy | P12 | 27 | 235 | 200 | 2.906 | 0.498 | 1300 | 137.22 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.5 | 2.8674
2011-26-0027.oxy | P12 | 27 | 236 | 175 | 3.354 | 0.567 | 1395 | 135.52 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.7 | 3.3098
2011-26-0027.oxy | P12 | 27 | 237 | 150 | 3.645 | 0.619 | 1321 | 136.23 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.9 | 3.5967
2011-26-0027.oxy | P12 | 27 | 238 | 125 | 4.277 | 0.735 | 1406 | 137.98 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.1 | 4.2206
2011-26-0027.oxy | P12 | 27 | 238 | 125 | 4.278 | 0.76 | 1407 | 142.61 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.2 | 4.2213
2011-26-0027.oxy | P12 | 27 | 239 | 100 | 5.256 | 0.866 | 1476 | 132.51 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.2 | 5.186
2011-26-0027.oxy | P12 | 27 | 240 | 75 | 6.615 | 1.08 | 1128 | 131.44 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.4 | 6.5277
2011-26-0027.oxy | P12 | 27 | 241 | 50 | 7.122 | 1.189 | 1330 | 134.4 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 8.1 | 7.0285
2011-26-0027.oxy | P12 | 27 | 242 | 25 | 6.754 | 1.176 | 1411 | 140.09 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 9.9 | 6.6649
2011-26-0027.oxy | P12 | 27 | 243 | 10 | 6.703 | 1.101 | 650 | 132.24 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 10.1 | 6.6143
2011-26-0027.oxy | P12 | 27 | 244 | 5 | 6.701 | 1.205 | 3124 | 144.64 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 10.1 | 6.6121
2011-26-0043.oxy | P16 | 43 | 347 | 3682 | 2.717 | 0.467 | 559 | 135.75 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.9 | 2.7168
2011-26-0043.oxy | P16 | 43 | 348 | 3500 | 2.679 | 0.489 | 2088 | 144.09 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.8 | 2.679
2011-26-0043.oxy | P16 | 43 | 349 | 3500 | 2.678 | 0.477 | 1491 | 140.64 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.8 | 2.6776
2011-26-0043.oxy | P16 | 43 | 350 | 3000 | 2.44 | 0.433 | 3094 | 140 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.9 | 2.4396
2011-26-0043.oxy | P16 | 43 | 351 | 2500 | 1.919 | 0.32 | 1487 | 131.14 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.1 | 1.9195
2011-26-0043.oxy | P16 | 43 | 351 | 2500 | 1.917 | 0.322 | 1465 | 132.13 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.4 | 1.917
2011-26-0043.oxy | P16 | 43 | 352 | 2000 | 1.309 | 0.228 | 632 | 136.17 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.4 | 1.3089
2011-26-0043.oxy | P16 | 43 | 353 | 1500 | 0.499 | 0.087 | 552 | 131.82 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.9 | 0.49946
2011-26-0043.oxy | P16 | 43 | 354 | 1250 | 0.326 | 0.063 | 1175 | 143.13 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.1 | 0.32584
2011-26-0043.oxy | P16 | 43 | 355 | 1000 | 0.277 | 0.053 | 1496 | 139.68 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.5 | 0.27695
2011-26-0043.oxy | P16 | 43 | 356 | 800 | 0.352 | 0.062 | 1158 | 130.52 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4 | 0.3517
2011-26-0043.oxy | P16 | 43 | 357 | 600 | 0.575 | 0.1 | 1483 | 132.43 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.5 | 0.57543
2011-26-0043.oxy | P16 | 43 | 358 | 400 | 1.372 | 0.235 | 1306 | 134.03 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.1 | 1.3717
2011-26-0043.oxy | P16 | 43 | 359 | 300 | 2.276 | 0.386 | 620 | 133.69 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.7 | 2.2761
2011-26-0043.oxy | P16 | 43 | 360 | 250 | 2.727 | 0.461 | 1333 | 133.52 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.7 | 2.727
2011-26-0043.oxy | P16 | 43 | 361 | 200 | 3.618 | 0.583 | 1105 | 127.62 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.2 | 3.618
2011-26-0043.oxy | P16 | 43 | 362 | 175 | 3.932 | 0.656 | 1085 | 132.17 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.3 | 3.9322
2011-26-0043.oxy | P16 | 43 | 363 | 150 | 4.231 | 0.703 | 611 | 131.71 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.7 | 4.2308
2011-26-0043.oxy | P16 | 43 | 364 | 125 | 4.621 | 0.785 | 2090 | 134.69 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.9 | 4.6212
2011-26-0043.oxy | P16 | 43 | 365 | 100 | 5.852 | 1.064 | 3121 | 144.23 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.1 | 5.8523
2011-26-0043.oxy | P16 | 43 | 365 | 100 | 5.857 | 0.988 | 510 | 133.91 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.2 | 5.8575
2011-26-0043.oxy | P16 | 43 | 366 | 75 | 6.951 | 1.261 | 1218 | 144.02 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.2 | 6.9507
2011-26-0043.oxy | P16 | 43 | 367 | 50 | 7.027 | 1.174 | 396 | 132.76 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.7 | 7.0265 |  | Bubble on lens
2011-26-0043.oxy | P16 | 43 | 368 | 25 | 6.735 | 1.108 | 505 | 130.73 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 9.2 | 6.7346
2011-26-0043.oxy | P16 | 43 | 369 | 10 | 6.735 | 1.21 | 2015 | 142.62 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 9.3 | 6.735
2011-26-0043.oxy | P16 | 43 | 370 | 5 | 6.754 | 1.192 | 3119 | 140.14 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 9.2 | 6.7535 | 3 | Valve open. Bubble on lens
2011-26-0055.oxy | P20 | 55 | 431 | 4024 | 3.031 | 0.545 | 2056 | 142.1 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.7 | 3.031 |  | Hugh questioned if cast 55's samples were reshaken since they had no water seal
2011-26-0055.oxy | P20 | 55 | 432 | 3500 | 2.88 | 0.49 | 1386 | 134.43 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.7 | 2.8803
2011-26-0055.oxy | P20 | 55 | 433 | 3500 | 2.269 | 0.412 | 3061 | 143.13 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.8 | 2.2686
2011-26-0055.oxy | P20 | 55 | 434 | 3000 | 2.553 | 0.463 | 3178 | 143.12 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.7 | 2.5526
2011-26-0055.oxy | P20 | 55 | 434 | 3000 | 2.531 | 0.436 | 1201 | 135.96 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 1.9 | 2.5307
2011-26-0055.oxy | P20 | 55 | 435 | 2500 | 1.934 | 0.326 | 1414 | 132.62 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2 | 1.9339
2011-26-0055.oxy | P20 | 55 | 436 | 2000 | 1.297 | 0.223 | 1327 | 134.41 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.2 | 1.2966 |  | Bubble on lens
2011-26-0055.oxy | P20 | 55 | 437 | 1500 | 0.6 | 0.111 | 3107 | 141.59 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.6 | 0.59958
2011-26-0055.oxy | P20 | 55 | 438 | 1250 | 0.339 | 0.062 | 1335 | 135.15 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 2.9 | 0.33947
2011-26-0055.oxy | P20 | 55 | 439 | 1000 | 0.288 | 0.054 | 1134 | 137.25 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.3 | 0.28773
2011-26-0055.oxy | P20 | 55 | 440 | 800 | 0.405 | 0.076 | 570 | 140.69 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 3.8 | 0.40521
2011-26-0055.oxy | P20 | 55 | 441 | 600 | 0.478 | 0.089 | 3118 | 140.8 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.2 | 0.47847
2011-26-0055.oxy | P20 | 55 | 442 | 400 | 0.988 | 0.183 | 2060 | 143.91 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 4.9 | 0.98836
2011-26-0055.oxy | P20 | 55 | 443 | 300 | 1.666 | 0.298 | 3135 | 140.39 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.2 | 1.6663
2011-26-0055.oxy | P20 | 55 | 444 | 250 | 2.632 | 0.443 | 1459 | 132.92 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.5 | 2.6315
2011-26-0055.oxy | P20 | 55 | 445 | 200 | 3.118 | 0.536 | 612 | 135.93 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 5.9 | 3.1179
2011-26-0055.oxy | P20 | 55 | 445 | 200 | 3.115 | 0.556 | 3116 | 141.08 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6 | 3.1154
2011-26-0055.oxy | P20 | 55 | 446 | 175 | 3.405 | 0.611 | 3164 | 141.92 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.1 | 3.4055
2011-26-0055.oxy | P20 | 55 | 447 | 150 | 3.784 | 0.706 | 1460 | 147.66 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.4 | 3.7837
2011-26-0055.oxy | P20 | 55 | 448 | 125 | 4.158 | 0.69 | 2000 | 131.52 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.6 | 4.1581
2011-26-0055.oxy | P20 | 55 | 449 | 100 | 5.079 | 0.849 | 1302 | 132.62 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.8 | 5.0794
2011-26-0055.oxy | P20 | 55 | 450 | 75 | 6.674 | 1.206 | 1260 | 143.43 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 6.6 | 6.6742
2011-26-0055.oxy | P20 | 55 | 451 | 50 | 7.131 | 1.206 | 2075 | 134.81 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 7.5 | 7.1074 |  | Endpoint late at 1.2100. Recalculated using new thio titer of 1.2060
2011-26-0055.oxy | P20 | 55 | 452 | 25 | 6.917 | 1.173 | 3194 | 134.73 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 8.6 | 6.9163
2011-26-0055.oxy | P20 | 55 | 453 | 10 | [11 empty] | 9 | no trip
2011-26-0055.oxy | P20 | 55 | 454 | 5 | 6.841 | 1.173 | 1326 | 136.19 | 0.0023 | 0.716 | 10.0158 | 0.01 | 0.14034 | 8.9 | 6.8411
2011-26-0082.oxy | P26 | 82 | 631 | 4317 | 3.24 | 0.582 | 2056 | 142.1 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.4 | 3.2399
2011-26-0082.oxy | P26 | 82 | 632 | 4000 | 3.208 | 0.545 | 1386 | 134.43 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.4 | 3.2079
2011-26-0082.oxy | P26 | 82 | 633 | 3500 | 2.988 | 0.541 | 3061 | 143.13 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.4 | 2.9879
2011-26-0082.oxy | P26 | 82 | 634 | 3000 | 2.576 | 0.467 | 3178 | 143.12 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.5 | 2.576
2011-26-0082.oxy | P26 | 82 | 634 | 3000 | 2.579 | 0.444 | 1201 | 135.96 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.7 | 2.5787
2011-26-0082.oxy | P26 | 82 | 635 | 2500 | 2.019 | 0.34 | 1414 | 132.62 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 1.7 | 2.0189
2011-26-0082.oxy | P26 | 82 | 636 | 2000 | 1.321 | 0.227 | 1327 | 134.41 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 2 | 1.3209
2011-26-0082.oxy | P26 | 82 | 637 | 1500 | 0.605 | 0.112 | 3107 | 141.59 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 2.4 | 0.60546
2011-26-0082.oxy | P26 | 82 | 638 | 1250 | 0.41 | 0.074 | 1335 | 135.15 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 2.7 | 0.41045
2011-26-0082.oxy | P26 | 82 | 639 | 1000 | 0.398 | 0.073 | 1134 | 137.25 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.1 | 0.39827
2011-26-0082.oxy | P26 | 82 | 640 | 800 | 0.479 | 0.089 | 570 | 140.69 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.3 | 0.47906
2011-26-0082.oxy | P26 | 82 | 641 | 600 | 0.552 | 0.102 | 3118 | 140.8 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.8 | 0.55229
2011-26-0082.oxy | P26 | 82 | 641 | 600 | 0.54 | 0.102 | 2060 | 143.91 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.9 | 0.54018
2011-26-0082.oxy | P26 | 82 | 642 | 400 | 1.11 | 0.2 | 3135 | 140.39 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.3 | 1.1105
2011-26-0082.oxy | P26 | 82 | 643 | 300 | 1.612 | 0.273 | 1459 | 132.92 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.3 | 1.6121
2011-26-0082.oxy | P26 | 82 | 644 | 250 | 2.034 | 0.351 | 612 | 135.93 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.4 | 2.0336
2011-26-0082.oxy | P26 | 82 | 645 | 200 | 2.704 | 0.483 | 3116 | 141.08 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.6 | 2.7042
2011-26-0082.oxy | P26 | 82 | 646 | 175 | 3.064 | 0.55 | 3164 | 141.92 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.8 | 3.0643
2011-26-0082.oxy | P26 | 82 | 647 | 150 | 3.321 | 0.62 | 1460 | 147.66 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.1 | 3.3212
2011-26-0082.oxy | P26 | 82 | 648 | 125 | 4.045 | 0.671 | 2000 | 131.52 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.2 | 4.0446
2011-26-0082.oxy | P26 | 82 | 649 | 100 | 7.049 | 1.176 | 1302 | 132.62 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.5 | 7.0491
2011-26-0082.oxy | P26 | 82 | 650 | 75 | 7.116 | 1.285 | 1260 | 143.43 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.8 | 7.116
2011-26-0082.oxy | P26 | 82 | 651 | 50 | 7.158 | 1.214 | 2075 | 134.81 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7 | 7.1577
2011-26-0082.oxy | P26 | 82 | 652 | 25 | 7.026 | 1.191 | 3194 | 134.73 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.6 | 7.0258
2011-26-0082.oxy | P26 | 82 | 652 | 25 | 7.021 | 1.254 | 2027 | 141.88 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.7 | 7.0207
2011-26-0082.oxy | P26 | 82 | 653 | 10 | 7.032 | 1.205 | 1326 | 136.19 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.7 | 7.0314
2011-26-0082.oxy | P26 | 82 | 654 | 5 | 7.022 | 1.258 | 1142 | 142.31 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.7 | 7.0216
2011-26-0103.oxy | Uwfloat | 103 | 744 | 500 | 0.974 | 0.178 | 2068 | 141.98 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.2 | 0.97434
2011-26-0103.oxy | Uwfloat | 103 | 745 | 500 | 0.936 | 0.165 | 1144 | 136.84 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.1 | 0.93571
2011-26-0103.oxy | Uwfloat | 103 | 746 | 300 | 2.044 | 0.354 | 631 | 136.38 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.5 | 2.0443
2011-26-0103.oxy | Uwfloat | 103 | 747 | 120 | 4.535 | 0.771 | 468 | 134.84 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.9 | 4.5351
2011-26-0103.oxy | Uwfloat | 103 | 748 | 120 | 4.544 | 0.746 | 1475 | 130.27 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.8 | 4.5435
2011-26-0103.oxy | Uwfloat | 103 | 748 | 120 | 4.541 | 0.833 | 1248 | 145.41 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.8 | 4.5406 | 4 | Curve does not flatten. No real endpoint. This point may have been picked arbitrarily
2011-26-0103.oxy | Uwfloat | 103 | 749 | 75 | 7.032 | 1.255 | 3028 | 141.77 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.9 | 7.0318
2011-26-0103.oxy | Uwfloat | 103 | 750 | 40 | 6.996 | 1.187 | 1397 | 134.85 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.7 | 6.9958
2011-26-0103.oxy | Uwfloat | 103 | 750 | 40 | 6.989 | 1.189 | 1373 | 135.2 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.7 | 6.9892
2011-26-0103.oxy | Uwfloat | 103 | 751 | 10 | 6.957 | 1.254 | 3151 | 143.15 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.9 | 6.9575
2011-26-0103.oxy | Uwfloat | 103 | 751 | 10 | 6.953 | 1.169 | 1318 | 133.63 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.9 | 6.9532
2011-26-0103.oxy | Uwfloat | 103 | 752 | 5 | 6.955 | 1.138 | 745 | 130.09 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.9 | 6.9551
2011-26-0103.oxy | Uwfloat | 103 | 753 | 0 | 6.964 | 1.28 | 2020 | 145.95 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.9 | 6.9641
2011-26-0104.oxy | Eddy | 104 | 754 | 2000 | 1.363 | 0.238 | 569 | 136.71 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 2.6 | 1.3625
2011-26-0104.oxy | Eddy | 104 | 755 | 1500 | 0.527 | 0.094 | 3182 | 135.62 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 2.7 | 0.52664
2011-26-0104.oxy | Eddy | 104 | 756 | 1000 | 0.329 | 0.06 | 1374 | 134.53 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.3 | 0.32935
2011-26-0104.oxy | Eddy | 104 | 757 | 800 | 0.887 | 0.154 | 1402 | 134.46 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 3.9 | 0.88726
2011-26-0104.oxy | Eddy | 104 | 757 | 800 | 0.389 | 0.069 | 3196 | 132.41 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4 | 0.38894
2011-26-0104.oxy | Eddy | 104 | 758 | 600 | 0.639 | 0.118 | 1393 | 141.65 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 4.5 | 0.63897
2011-26-0104.oxy | Eddy | 104 | 759 | 400 | 1.13 | 0.189 | 621 | 130.31 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 5.7 | 1.1303
2011-26-0104.oxy | Eddy | 104 | 760 | 300 | 1.389 | 0.234 | 1261 | 131.88 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.1 | 1.389
2011-26-0104.oxy | Eddy | 104 | 761 | 250 | 1.507 | 0.245 | 1146 | 127.43 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.2 | 1.5072
2011-26-0104.oxy | Eddy | 104 | 762 | 200 | 1.649 | 0.273 | 528 | 129.96 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.4 | 1.6494
2011-26-0104.oxy | Eddy | 104 | 763 | 175 | 1.894 | 0.314 | 371 | 130.41 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.5 | 1.8945 | 4 | Endpoint overshot due to obstruction of light path. Difficult to know where real endpoint is.
2011-26-0104.oxy | Eddy | 104 | 764 | 150 | 1.878 | 0.326 | 196 | 136.6 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.5 | 1.8775
2011-26-0104.oxy | Eddy | 104 | 765 | 125 | 2.367 | 0.402 | 3090 | 133.99 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.6 | 2.3671
2011-26-0104.oxy | Eddy | 104 | 765 | 125 | 2.371 | 0.439 | 1404 | 146.01 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.6 | 2.3715
2011-26-0104.oxy | Eddy | 104 | 766 | 100 | 4.968 | 0.911 | 3126 | 145.42 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.5 | 4.9678
2011-26-0104.oxy | Eddy | 104 | 767 | 75 | 6.837 | 1.14 | 1378 | 132.53 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.5 | 6.8372
2011-26-0104.oxy | Eddy | 104 | 768 | 50 | 7.094 | 1.152 | 647 | 129.13 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.9 | 7.0942
2011-26-0104.oxy | Eddy | 104 | 769 | 25 | 6.825 | 1.141 | 3157 | 132.87 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 8.3 | 6.8254
2011-26-0104.oxy | Eddy | 104 | 770 | 10 | 6.763 | 1.217 | 1298 | 142.91 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 8.4 | 6.763
2011-26-0104.oxy | Eddy | 104 | 771 | 5 | 6.815 | 1.235 | 2014 | 143.92 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 8.4 | 6.8145
2011-26-0109.oxy | LB08 | 109 | 775 | 129 | 1.252 | 0.229 | 1298 | 142.91 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.9 | 1.2524 | 3 | sample redrawn in same flask
2011-26-0109.oxy | LB08 | 109 | 776 | 125 | 1.338 | 0.246 | 2014 | 143.92 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 6.9 | 1.3376
2011-26-0109.oxy | LB08 | 109 | 777 | 100 | 1.457 | 0.243 | 1408 | 130.64 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7 | 1.4574
2011-26-0109.oxy | LB08 | 109 | 778 | 75 | 2.21 | 0.373 | 1307 | 133.08 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.3 | 2.2097
2011-26-0109.oxy | LB08 | 109 | 778 | 75 | 2.197 | 0.373 | 3206 | 133.81 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.3 | 2.1974
2011-26-0109.oxy | LB08 | 109 | 779 | 50 | 2.782 | 0.465 | 1392 | 132.11 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 7.6 | 2.7819
2011-26-0109.oxy | LB08 | 109 | 780 | 30 | 4.987 | 0.842 | 3090 | 133.99 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 9 | 4.9871 |  | Bubble on lens
2011-26-0109.oxy | LB08 | 109 | 781 | 20 | 5.695 | 1.048 | 1404 | 146.01 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 9.6 | 5.6951
2011-26-0109.oxy | LB08 | 109 | 782 | 10 | 8.475 | 1.549 | 3126 | 145.42 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 10.8 | 8.464 |  | Endpoint late at 1.5510. Recalculated using new thio titer of 1.5490
2011-26-0109.oxy | LB08 | 109 | 783 | 5 | 8.974 | 1.495 | 1378 | 132.53 | 0.0023 | 0.7157 | 10.0158 | 0.01 | 0.1404 | 10.9 | 8.9747