## comparison-1176
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 199.834 | 7.6094 | 7.6094 | 33.8631 | 33.8629 | 6.7078 | 1.1892 | 11 | 11 | 149 | 6.761 | 0.32007 | 2.164 |  | 0.00024275 | 0.4647 | -0.2328 |  |  |  | Summary of Soc Voffset including the original values in the factory calibration
2 | 4.709 | 13.3395 | 13.3405 | 31.2305 | 31.2305 | 6.0261 | 3.0281 | 2 | 2 | 13 | 5.8989 | 1.1746 | 6.929 |  | 0.00026909 | [7 empty] | m | b | Soc | Voffset | R2
31 | 2499.957 | 1.7586 | 1.758 | 34.6272 | 34.6278 | 7.6767 | 0.87756 | 4 | 4 | 195 | 10.5896 | 0.17442 | 1.847 | 6 | 0.00058545 | [5 empty] | Bottles used | Original | 0.4453 | -0.23209 | 0.4453 | -0.5212 |  | 2.885
12 | 150.378 | 7.4103 | 7.4103 | 33.757 | 33.7571 | 6.7429 | 1.5668 | 12 | 12 | 73 | 6.7553 | 0.4959 | 3.35 |  | 0.00063133 | [5 empty] | 164 | all | 0.4647 | -0.2328 | 0.4647 | -0.50097 | 0.9845 | 3.0201
31 | 200.906 | 5.8658 | 5.8674 | 33.8286 | 33.8282 | 6.9846 | 1.5925 | 15 | 15 | 206 | 7.0604 | 0.50649 | 3.576 |  | 0.00073073 | [5 empty] | 155 | all data except 3 and 4 flags | 0.4698 | -0.2394 | 0.4698 | -0.50958 | 0.9989 | 3.0492
24 | 249.722 | 7.1701 | 7.1706 | 33.9424 | 33.9422 | 6.7718 | 1.0526 | 10 | 10 | 148 | 6.8737 | 0.25561 | 1.757 |  | 0.00073443 | [5 empty] | 144 | no cast 1,no 3 or 4 flags | 0.4705 | -0.239 | 0.4705 | -0.50797 | 0.9995 | 3.0545
12 | 301.316 | 5.9538 | 5.9542 | 33.9553 | 33.9553 | 6.9644 | 1.0639 | 8 | 8 | 69 | 7.1306 | 0.26057 | 1.858 |  | 0.0010511 | [5 empty] | 142 | excl. outliers diff>0.05 & cast #1 | 0.4692 | -0.2377 | 0.4692 | -0.50661 | 0.9998 | 3.0467
12 | 200.112 | 6.8348 | 6.8383 | 33.8887 | 33.8883 | 6.8265 | 1.4051 | 10 | 10 | 71 | 6.8887 | 0.4188 | 2.885 |  | 0.0013704 | [5 empty] | 141 | excl. outliers diff>0.03 & cast #1 | 0.4694 | -0.2382 | 0.4694 | -0.50746 | 0.9998 | 3.0476
12 | 100.895 | 7.6901 | 7.6913 | 33.0805 | 33.0785 | 6.7299 | 2.1145 | 14 | 14 | 75 | 6.6971 | 0.74809 | 5.01 |  | 0.0017036 | [5 empty] | 139 | excl. outliers diff>0.02 & cast #1 | 0.469 | -0.2378 | 0.469 | -0.50704 | 0.9998 | 3.0452
24 | 175.781 | 7.7335 | 7.7336 | 33.8421 | 33.8419 | 6.6899 | 1.2051 | 12 | 12 | 150 | 6.7211 | 0.32524 | 2.186 | 6 | 0.0019698 | [5 empty] | 134 | excl. outliers diff>0.015 & cast #1 | 0.4689 | -0.2378 | 0.4689 | -0.50714 | 0.9999 | 3.0445
12 | 601.42 | 4.6668 | 4.6668 | 34.1784 | 34.1785 | 7.169 | 0.62026 | 6 | 6 | 67 | 7.6508 | 0.053328 | 0.408 |  | 0.0021072
31 | 174.986 | 6.1489 | 6.1512 | 33.816 | 33.8155 | 6.9391 | 1.6803 | 16 | 16 | 207 | 6.9876 | 0.55012 | 3.844 |  | 0.0021074
31 | 251.715 | 5.3447 | 5.3454 | 33.8799 | 33.8796 | 7.0684 | 1.2356 | 14 | 14 | 205 | 7.1992 | 0.34351 | 2.473 |  | 0.0021163 | [7 empty] | m | b | Soc | Voffset | R2
31 | 3000.008 | 1.6679 | 1.6669 | 34.6451 | 34.6464 | 7.6934 | 0.91836 | 2 | 2 | 193 | 11.3353 | 0.19638 | 2.226 |  | 0.0024164 | [5 empty] | casts 4-31 | exc. outliers diff>0.03 | 0.4703 | -0.2412 | 0.4703 | -0.51286 | 0.9997 | 3.0509
24 | 601.24 | 4.8435 | 4.8435 | 34.1593 | 34.1593 | 7.1397 | 0.63927 | 6 | 6 | 144 | 7.6163 | 0.061447 | 0.468 |  | 0.0028216 | [5 empty] | casts 37-70 | exc. outliers diff>0.03 | 0.4688 | -0.236 | 0.4688 | -0.50341 | 0.9999 | 3.0456
24 | 400.115 | 5.998 | 5.9981 | 34.0309 | 34.0307 | 6.9537 | 0.85262 | 8 | 8 | 146 | 7.2104 | 0.16032 | 1.156 |  | 0.0030885
31 | 1999.872 | 1.9142 | 1.9138 | 34.5933 | 34.5936 | 7.6485 | 0.81564 | 5 | 5 | 196 | 9.8759 | 0.14287 | 1.411 |  | 0.0033546
24 | 125.197 | 7.9385 | 7.9394 | 33.8189 | 33.8185 | 6.66 | 1.2172 | 14 | 14 | 152 | 6.646 | 0.32937 | 2.189 |  | 0.0034642
24 | 150.003 | 7.8298 | 7.8296 | 33.8319 | 33.8317 | 6.6758 | 1.2111 | 13 | 13 | 151 | 6.684 | 0.32645 | 2.182 |  | 0.0035403
4 | 10.474 | 12.1618 | 12.1439 | 31.9933 | 32.0004 | 6.1463 | 2.8433 | 7 | 7 | 21 | 6.0236 | 1.092 | 6.578 |  | 0.0035504
12 | 1315.005 | 2.8194 | 2.8191 | 34.4902 | 34.4911 | 7.4829 | 0.64344 | 1 | 1 | 62 | 8.8069 | 0.062565 | 0.551 |  | 0.0036419
12 | 400.403 | 5.3628 | 5.3629 | 34.0219 | 34.022 | 7.0587 | 0.83021 | 7 | 7 | 68 | 7.3286 | 0.14914 | 1.093 |  | 0.0038563
31 | 51.017 | 8.8677 | 8.8685 | 32.566 | 32.5657 | 6.5761 | 2.7946 | 21 | 21 | 212 | 6.4934 | 1.0697 | 6.946 |  | 0.0038614
24 | 50.106 | 8.4249 | 8.4257 | 32.583 | 32.5828 | 6.6407 | 2.5591 | 17 | 17 | 155 | 6.5597 | 0.96026 | 6.299 |  | 0.0038646
31 | 3269.304 | 1.6798 | 1.6786 | 34.6471 | 34.6486 | 7.6909 | 0.9121 | 1 | 1 | 192 | 11.738 | 0.19492 | 2.288 |  | 0.0038702
31 | 400.324 | 4.7331 | 4.7354 | 34.0377 | 34.0374 | 7.1644 | 0.78437 | 12 | 12 | 203 | 7.4482 | 0.12755 | 0.95 | 6 | 0.0041483
24 | 799.882 | 4.1331 | 4.1336 | 34.2903 | 34.2902 | 7.2561 | 0.57363 | 5 | 5 | 143 | 7.9561 | 0.029286 | 0.233 |  | 0.0044801
4 | 50.299 | 7.7661 | 7.7661 | 32.8496 | 32.85 | 6.7285 | 2.3212 | 5 | 5 | 19 | 6.6517 | 0.85076 | 5.659 |  | 0.004896
19 | 4.934 | 16.1084 | 16.1044 | 32.0273 | 32.0268 | 5.6715 | 2.8161 | 2 | 2 | 134 | 5.5494 | 1.0808 | 5.998 |  | 0.005015
24 | 499.592 | 5.3011 | 5.3012 | 34.0857 | 34.0855 | 7.066 | 0.73191 | 7 | 7 | 145 | 7.4318 | 0.10213 | 0.759 |  | 0.0051896
31 | 5.565 | 16.1145 | 16.1136 | 32.4284 | 32.4279 | 5.657 | 2.8229 | 24 | 24 | 215 | 5.5356 | 1.0842 | 6.002 |  | 0.0052439
31 | 150.234 | 6.494 | 6.5009 | 33.7601 | 33.7587 | 6.8862 | 1.7754 | 17 | 17 | 208 | 6.9083 | 0.58698 | 4.055 |  | 0.0052654
12 | 5.265 | 16.3588 | 16.3583 | 32.1851 | 32.1848 | 5.6383 | 2.7737 | 19 | 19 | 80 | 5.5171 | 1.0616 | 5.857 |  | 0.0054801
31 | 299.4 | 5.0165 | 5.0165 | 33.9447 | 33.9444 | 7.1205 | 0.98958 | 13 | 13 | 204 | 7.3021 | 0.22131 | 1.616 |  | 0.0057502
24 | 999.133 | 3.6428 | 3.6426 | 34.3781 | 34.3782 | 7.3386 | 0.57507 | 4 | 4 | 142 | 8.2685 | 0.028663 | 0.237 |  | 0.005772
31 | 1499.272 | 2.3401 | 2.34 | 34.5138 | 34.5137 | 7.5714 | 0.6784 | 6 | 6 | 197 | 9.1426 | 0.076346 | 0.698 |  | 0.0061067
12 | 50.131 | 9.3381 | 9.3469 | 32.5185 | 32.5179 | 6.5099 | 2.7125 | 16 | 16 | 77 | 6.4244 | 1.0339 | 6.642 |  | 0.0061521
12 | 1250.621 | 3.0639 | 3.0638 | 34.4627 | 34.4635 | 7.4392 | 0.61866 | 2 | 2 | 63 | 8.6756 | 0.048527 | 0.421 |  | 0.0061646
31 | 500.543 | 4.4196 | 4.4197 | 34.123 | 34.1226 | 7.2143 | 0.66706 | 11 | 11 | 202 | 7.6035 | 0.070494 | 0.536 |  | 0.0066893
31 | 1001.626 | 3.219 | 3.2188 | 34.4015 | 34.4012 | 7.414 | 0.58425 | 8 | 8 | 199 | 8.3663 | 0.031794 | 0.266 |  | 0.0069066
12 | 76.089 | 8.0908 | 8.0911 | 32.6575 | 32.6574 | 6.6876 | 2.5466 | 15 | 15 | 76 | 6.6305 | 0.95769 | 6.35 | 6 | 0.0070838
31 | 600.037 | 4.2045 | 4.2046 | 34.2139 | 34.2138 | 7.2474 | 0.60272 | 10 | 10 | 201 | 7.7416 | 0.040173 | 0.311 |  | 0.0071114
12 | 1250.736 | 3.0654 | 3.065 | 34.4627 | 34.4634 | 7.4389 | 0.61828 | 3 | 3 | 64 | 8.6754 | 0.04726 | 0.41 |  | 0.0072546
4 | 74.731 | 7.5122 | 7.5145 | 33.1604 | 33.1568 | 6.7537 | 1.9419 | 3 | 3 | 17 | 6.6998 | 0.67689 | 4.535 |  | 0.0072765
12 | 1001.127 | 3.6231 | 3.6227 | 34.3883 | 34.3887 | 7.3416 | 0.58046 | 4 | 4 | 65 | 8.2745 | 0.029609 | 0.245 |  | 0.0073307
21 | 5.666 | 16.2526 | 16.2502 | 32.2581 | 32.257 | 5.6476 | 2.7885 | 2 | 2 | 138 | 5.5264 | 1.0705 | 5.916 |  | 0.0074499
31 | 1250.014 | 2.6937 | 2.6941 | 34.4642 | 34.464 | 7.5076 | 0.61554 | 7 | 7 | 198 | 8.7649 | 0.045523 | 0.399 |  | 0.0077189
4 | 74.807 | 7.5201 | 7.5225 | 33.1454 | 33.1413 | 6.7531 | 1.9665 | 4 | 4 | 18 | 6.6992 | 0.67321 | 4.51 |  | 0.007807
31 | 800.426 | 3.723 | 3.7232 | 34.3255 | 34.325 | 7.3269 | 0.57667 | 9 | 9 | 200 | 8.0429 | 0.026732 | 0.215 |  | 0.008447
24 | 24.498 | 12.9709 | 12.9935 | 32.3603 | 32.3587 | 6.0295 | 3.0473 | 21 | 18 | 156 | 5.9177 | 1.1741 | 6.948 |  | 0.0091735
24 | 99.756 | 8.1154 | 8.1158 | 33.7201 | 33.7197 | 6.6378 | 1.3221 | 15 | 15 | 153 | 6.6009 | 0.37237 | 2.458 |  | 0.0091924
28 | 5.291 | 16.3738 | 16.3722 | 32.473 | 32.4725 | 5.6268 | 2.8693 | 2 | 2 | 168 | 5.5059 | 1.1101 | 6.112 |  | 0.0095325
31 | 100.648 | 6.9771 | 6.9767 | 33.3276 | 33.3285 | 6.8293 | 2.0755 | 19 | 19 | 210 | 6.8026 | 0.74134 | 5.043 |  | 0.009649
31 | 125.837 | 6.8368 | 6.8368 | 33.558 | 33.5577 | 6.841 | 1.9454 | 18 | 18 | 209 | 6.8378 | 0.66045 | 4.516 |  | 0.010781
12 | 10.038 | 16.3193 | 16.3249 | 32.1836 | 32.1829 | 5.6428 | 2.7646 | 18 | 18 | 79 | 5.5247 | 1.0629 | 5.872 |  | 0.010958
12 | 125.402 | 7.7548 | 7.7552 | 33.5329 | 33.5346 | 6.7002 | 1.6801 | 13 | 13 | 74 | 6.6879 | 0.55937 | 3.741 |  | 0.011433
26 | 4.375 | 16.9403 | 16.9338 | 32.0694 | 32.0679 | 5.5788 | 2.796 | 2 | 2 | 164 | 5.4582 | 1.08 | 5.895 |  | 0.013543
31 | 10.797 | 16.08 | 16.0675 | 32.4266 | 32.4261 | 5.6609 | 2.8151 | 23 | 23 | 214 | 5.5431 | 1.0904 | 6.044 |  | 0.015
24 | 74.625 | 8.065 | 8.0659 | 32.9979 | 32.9986 | 6.6766 | 2.0435 | 16 | 16 | 154 | 6.6186 | 0.7012 | 4.641 |  | 0.015622
4 | 25.136 | 8.6286 | 8.6351 | 32.4955 | 32.4936 | 6.6143 | 1.918 | 6 | 6 | 20 | 6.5112 | 0.67468 | 4.393 |  | 0.016166
4 | 99.752 | 7.2461 | 7.2444 | 33.6623 | 33.6635 | 6.7725 | 1.5462 | 2 | 2 | 16 | 6.7425 | 0.50308 | 3.392 |  | 0.017354
4 | 5.282 | 14.9696 | 14.9738 | 31.5045 | 31.5033 | 5.8197 | 2.8851 | 8 | 8 | 22 | 5.6953 | 1.1304 | 6.438 |  | 0.022507
12 | 24.694 | 14.9601 | 14.9615 | 32.255 | 32.253 | 5.7941 | 2.8688 | 17 | 17 | 78 | 5.684 | 1.124 | 6.389 |  | 0.023702
4 | 105.069 | 7.1029 | 7.0994 | 33.7254 | 33.7284 | 6.7919 | 1.0123 | 1 | 1 | 15 | 6.7678 | 0.2701 | 1.828 |  | 0.032499
31 | 25.259 | 13.5157 | 13.5164 | 32.4344 | 32.4332 | 5.9595 | 2.8731 | 22 | 22 | 213 | 5.8485 | 1.1529 | 6.743 |  | 0.050633
10 | 20.298 | 11.8372 | 11.845 | 32.4001 | 32.3989 | 6.1726 | 2.9606 | 4 | 2 | 50 | 6.0578 | 1.2115 | 7.339 |  | 0.06848
42 | 174.945 | 6.1566 | 6.1573 | 33.8318 | 33.8313 | 6.9371 | 1.7171 | 16 | 16 | 280 | 6.9855 | 0.56503 | 3.947 |  | 0.00012413
57 | 2499.998 | 1.6956 | 1.6954 | 34.6319 | 34.6322 | 7.6887 | 0.9013 | 5 | 5 | 373 | 10.6089 | 0.18626 | 1.976 | 6 | 0.00022431
57 | 299.9 | 4.3839 | 4.3857 | 33.9055 | 33.9048 | 7.2311 | 1.0189 | 13 | 13 | 381 | 7.4261 | 0.24104 | 1.79 |  | 0.00034192
70 | 1999.531 | 1.9238 | 1.9235 | 34.586 | 34.586 | 7.647 | 0.79262 | 6 | 6 | 499 | 9.8732 | 0.13481 | 1.331 |  | 0.00072071
70 | 600.49 | 3.6099 | 3.6097 | 34.1978 | 34.1973 | 7.3535 | 0.68117 | 11 | 11 | 504 | 7.8673 | 0.083002 | 0.653 |  | 0.00073784
57 | 24.167 | 12.6205 | 12.6082 | 32.5147 | 32.5129 | 6.0677 | 2.7833 | 22 | 22 | 390 | 5.9557 | 1.0613 | 6.321 |  | 0.00074017
57 | 250.138 | 4.644 | 4.6438 | 33.8604 | 33.8599 | 7.1882 | 1.1373 | 14 | 14 | 382 | 7.3304 | 0.29494 | 2.162 |  | 0.00075641
42 | 399.265 | 4.2658 | 4.2661 | 33.9783 | 33.9778 | 7.2481 | 0.91752 | 12 | 12 | 276 | 7.5422 | 0.19278 | 1.454 |  | 0.00078925
70 | 301.499 | 3.9619 | 3.9621 | 33.9169 | 33.9162 | 7.3047 | 0.95643 | 13 | 13 | 506 | 7.5105 | 0.21051 | 1.581 | 6 | 0.0011478
42 | 2499.125 | 1.7245 | 1.7239 | 34.6302 | 34.6308 | 7.6832 | 0.89118 | 5 | 5 | 269 | 10.5988 | 0.18257 | 1.935 | 6 | 0.0012359
40 | 4.915 | 14.2902 | 14.29 | 32.4036 | 32.4031 | 5.8673 | 2.7897 | 2 | 2 | 255 | 5.7423 | 1.0623 | 6.1 |  | 0.0012873
42 | 125.664 | 6.6292 | 6.6301 | 33.7074 | 33.706 | 6.8671 | 1.8875 | 18 | 18 | 282 | 6.8659 | 0.64566 | 4.433 |  | 0.0013364
70 | 800.324 | 3.2004 | 3.2 | 34.307 | 34.3066 | 7.4221 | 0.6483 | 10 | 10 | 503 | 8.1591 | 0.070105 | 0.572 |  | 0.0016404
57 | 399.365 | 4.0996 | 4.1 | 34.0019 | 34.0013 | 7.2762 | 0.86094 | 12 | 12 | 380 | 7.5744 | 0.16543 | 1.253 |  | 0.0018534
70 | 400.439 | 3.9111 | 3.9113 | 34.0329 | 34.0324 | 7.3079 | 0.81066 | 12 | 12 | 505 | 7.612 | 0.14595 | 1.111 |  | 0.0020397
42 | 150.479 | 6.4011 | 6.4056 | 33.8233 | 33.8225 | 6.8982 | 1.7794 | 17 | 17 | 281 | 6.9215 | 0.59625 | 4.127 |  | 0.0021626
70 | 999.832 | 2.8818 | 2.8818 | 34.3793 | 34.3789 | 7.477 | 0.63213 | 9 | 9 | 502 | 8.4439 | 0.058622 | 0.495 |  | 0.0023286
57 | 74.958 | 6.3908 | 6.3952 | 32.7868 | 32.7841 | 6.9473 | 2.6117 | 20 | 20 | 388 | 6.9035 | 0.98327 | 6.788 | 6 | 0.0023925
57 | 124.366 | 6.2528 | 6.2528 | 33.6201 | 33.6192 | 6.9313 | 1.7926 | 18 | 18 | 386 | 6.9331 | 0.60261 | 4.178 |  | 0.0024114
57 | 1999.763 | 1.9308 | 1.9306 | 34.5858 | 34.5859 | 7.6456 | 0.78759 | 6 | 6 | 374 | 9.8715 | 0.13068 | 1.29 |  | 0.0025138
42 | 300.142 | 4.5655 | 4.5656 | 33.8948 | 33.8944 | 7.2001 | 1.0966 | 13 | 13 | 277 | 7.3915 | 0.27423 | 2.027 |  | 0.0025459
57 | 174.66 | 5.5219 | 5.5254 | 33.8163 | 33.8156 | 7.0419 | 1.4707 | 16 | 16 | 384 | 7.0987 | 0.4536 | 3.22 |  | 0.0029732
42 | 3685.22 | 1.5801 | 1.5786 | 34.6666 | 34.6682 | 7.7093 | 0.96075 | 1 | 1 | 265 | 12.4307 | 0.21672 | 2.694 |  | 0.0030614
42 | 1999.958 | 1.9433 | 1.943 | 34.5881 | 34.5883 | 7.6431 | 0.79006 | 6 | 6 | 270 | 9.868 | 0.13123 | 1.295 |  | 0.0031088
70 | 2500.322 | 1.703 | 1.7025 | 34.6296 | 34.6298 | 7.6874 | 0.89916 | 5 | 5 | 498 | 10.6072 | 0.18836 | 1.998 | 6 | 0.0033224
42 | 200.348 | 5.7078 | 5.707 | 33.8237 | 33.8233 | 7.0108 | 1.579 | 15 | 15 | 279 | 7.0884 | 0.50434 | 3.575 |  | 0.0033969
70 | 1250.434 | 2.5776 | 2.5776 | 34.448 | 34.4478 | 7.5301 | 0.61889 | 8 | 8 | 501 | 8.795 | 0.051393 | 0.452 |  | 0.0034051
42 | 3001.204 | 1.6019 | 1.6011 | 34.6533 | 34.6541 | 7.7058 | 0.95246 | 4 | 4 | 268 | 11.3586 | 0.21367 | 2.427 |  | 0.0038616
70 | 249.078 | 3.9688 | 3.969 | 33.8475 | 33.8469 | 7.3069 | 1.0994 | 14 | 14 | 507 | 7.4617 | 0.28198 | 2.104 |  | 0.0038699
70 | 174.859 | 4.1289 | 4.1294 | 33.738 | 33.7375 | 7.2839 | 1.3767 | 16 | 16 | 509 | 7.3642 | 0.41145 | 3.03 |  | 0.0044754
57 | 2999.278 | 1.5753 | 1.5746 | 34.6558 | 34.6566 | 7.7108 | 0.96689 | 4 | 4 | 372 | 11.3646 | 0.22104 | 2.512 |  | 0.0045239
70 | 1499.325 | 2.3017 | 2.3016 | 34.5072 | 34.507 | 7.579 | 0.65588 | 7 | 7 | 500 | 9.1531 | 0.067409 | 0.617 |  | 0.0045784
42 | 50.502 | 7.4864 | 7.4808 | 32.6281 | 32.6274 | 6.7812 | 2.7618 | 21 | 21 | 285 | 6.7065 | 1.0552 | 7.077 |  | 0.0046478
70 | 201.165 | 4.0056 | 4.006 | 33.7731 | 33.7725 | 7.304 | 1.2638 | 15 | 15 | 508 | 7.4118 | 0.35916 | 2.662 |  | 0.0046492
57 | 149.932 | 5.9367 | 5.9389 | 33.7883 | 33.7876 | 6.9749 | 1.6674 | 17 | 17 | 385 | 7.0035 | 0.54673 | 3.829 |  | 0.0046713
42 | 3500.601 | 1.5722 | 1.571 | 34.6649 | 34.6663 | 7.711 | 0.96691 | 2 | 2 | 266 | 12.1366 | 0.22131 | 2.686 |  | 0.0047912
57 | 600.064 | 3.71 | 3.7102 | 34.1767 | 34.1763 | 7.3366 | 0.66924 | 11 | 11 | 379 | 7.8468 | 0.073279 | 0.575 | 6 | 0.0049173
57 | 98.96 | 6.2524 | 6.2526 | 33.1489 | 33.1494 | 6.9531 | 2.1819 | 19 | 19 | 387 | 6.9321 | 0.78605 | 5.449 |  | 0.0049293
42 | 1000.517 | 3.0275 | 3.0277 | 34.3671 | 34.3668 | 7.4508 | 0.58506 | 9 | 9 | 273 | 8.4113 | 0.034002 | 0.286 |  | 0.0050756
42 | 599.876 | 3.8304 | 3.8309 | 34.1593 | 34.1588 | 7.316 | 0.66675 | 11 | 11 | 275 | 7.8221 | 0.07172 | 0.561 | 6 | 0.0053188
57 | 4.839 | 13.413 | 13.4136 | 32.4719 | 32.4713 | 5.9707 | 2.7578 | 24 | 24 | 392 | 5.8446 | 1.0541 | 6.161 |  | 0.0053672
62 | 5.528 | 13.2136 | 13.2131 | 32.4852 | 32.4843 | 5.9947 | 2.7561 | 2 | 2 | 431 | 5.869 | 1.0533 | 6.182 |  | 0.0053923
42 | 24.983 | 12.6 | 12.598 | 32.4795 | 32.479 | 6.0716 | 2.8297 | 22 | 22 | 286 | 5.9602 | 1.0875 | 6.482 |  | 0.0054017
57 | 1500.095 | 2.3321 | 2.332 | 34.5044 | 34.5043 | 7.5733 | 0.64997 | 7 | 7 | 375 | 9.1463 | 0.063633 | 0.582 |  | 0.0056085
57 | 800.135 | 3.3273 | 3.3273 | 34.2883 | 34.2879 | 7.4 | 0.62037 | 10 | 10 | 378 | 8.1317 | 0.049682 | 0.404 |  | 0.0058037
57 | 3501.493 | 1.5269 | 1.526 | 34.67 | 34.6711 | 7.7195 | 0.99276 | 3 | 3 | 371 | 12.1541 | 0.23441 | 2.849 |  | 0.0058709
57 | 999.566 | 3.0091 | 3.009 | 34.3633 | 34.363 | 7.4543 | 0.60141 | 9 | 9 | 377 | 8.4148 | 0.040762 | 0.343 |  | 0.0059136
70 | 76.088 | 4.6462 | 4.6461 | 32.7944 | 32.7938 | 7.2395 | 2.597 | 20 | 20 | 513 | 7.2179 | 0.98047 | 7.077 | 6 | 0.0064372
70 | 3000.115 | 1.5683 | 1.5678 | 34.6564 | 34.6568 | 7.7122 | 0.9732 | 4 | 4 | 497 | 11.3682 | 0.22598 | 2.569 |  | 0.0065362
42 | 3500.403 | 1.5723 | 1.5711 | 34.6648 | 34.6662 | 7.711 | 0.96686 | 3 | 3 | 267 | 12.1363 | 0.22305 | 2.707 |  | 0.0065509
70 | 4315.976 | 1.5216 | 1.5201 | 34.6824 | 34.6841 | 7.7199 | 0.99571 | 1 | 1 | 494 | 13.5243 | 0.23646 | 3.198 |  | 0.0065563
70 | 149.541 | 4.3393 | 4.3397 | 33.6801 | 33.6796 | 7.2499 | 1.5292 | 17 | 17 | 510 | 7.3022 | 0.48437 | 3.537 |  | 0.0065622
57 | 48.629 | 7.4448 | 7.4437 | 32.662 | 32.6614 | 6.7862 | 2.7973 | 21 | 21 | 389 | 6.7101 | 1.0738 | 7.205 |  | 0.0066315
42 | 800.521 | 3.3926 | 3.3926 | 34.2771 | 34.2768 | 7.3887 | 0.61287 | 10 | 10 | 274 | 8.1182 | 0.04533 | 0.368 |  | 0.0066704
60 | 5.374 | 13.6483 | 13.6475 | 32.4951 | 32.4943 | 5.9411 | 2.7802 | 2 | 2 | 427 | 5.8157 | 1.0659 | 6.199 |  | 0.0067384
57 | 200.057 | 5.1109 | 5.1146 | 33.8301 | 33.8291 | 7.11 | 1.3609 | 15 | 15 | 383 | 7.1968 | 0.40643 | 2.925 |  | 0.0068319
66 | 5.531 | 13.2816 | 13.2806 | 32.5043 | 32.5032 | 5.9856 | 2.7637 | 2 | 2 | 459 | 5.86 | 1.0584 | 6.202 |  | 0.006866
57 | 1250.473 | 2.6308 | 2.6308 | 34.4449 | 34.4446 | 7.5203 | 0.60197 | 8 | 8 | 376 | 8.7821 | 0.039968 | 0.351 |  | 0.0069676
57 | 3500.604 | 1.5265 | 1.5256 | 34.6701 | 34.6711 | 7.7196 | 0.99284 | 2 | 2 | 370 | 12.1528 | 0.23558 | 2.863 |  | 0.0070108
57 | 4022.677 | 1.545 | 1.5436 | 34.6763 | 34.678 | 7.7157 | 0.97997 | 1 | 1 | 369 | 13.0056 | 0.22975 | 2.988 |  | 0.0071548
42 | 1500.454 | 2.387 | 2.387 | 34.5007 | 34.5006 | 7.5632 | 0.64244 | 7 | 7 | 271 | 9.1327 | 0.058581 | 0.535 |  | 0.0071609
42 | 1251.288 | 2.6654 | 2.6654 | 34.444 | 34.4438 | 7.5139 | 0.59492 | 8 | 8 | 272 | 8.7745 | 0.036241 | 0.318 |  | 0.007418
70 | 24.929 | 12.9303 | 12.9302 | 32.5171 | 32.5158 | 6.0287 | 2.7596 | 22 | 22 | 515 | 5.9173 | 1.0572 | 6.256 |  | 0.0076533
49 | 4.629 | 14.0401 | 14.0384 | 32.4157 | 32.415 | 5.8967 | 2.7622 | 2 | 2 | 343 | 5.7711 | 1.0589 | 6.111 |  | 0.0080855
42 | 10 | 14.0898 | 14.0883 | 32.4033 | 32.4024 | 5.8912 | 2.7701 | 23 | 23 | 287 | 5.7696 | 1.0628 | 6.132 |  | 0.0083411
42 | 75.583 | 6.8569 | 6.8599 | 32.7248 | 32.7238 | 6.8755 | 2.6349 | 20 | 20 | 284 | 6.8277 | 1.0002 | 6.829 |  | 0.0085484
52 | 5.184 | 13.7407 | 13.741 | 32.4492 | 32.4483 | 5.9315 | 2.7564 | 2 | 2 | 347 | 5.8061 | 1.0568 | 6.136 |  | 0.0087465
70 | 3997.958 | 1.5006 | 1.4995 | 34.6796 | 34.6809 | 7.7242 | 1.0084 | 2 | 2 | 495 | 12.9807 | 0.24475 | 3.177 |  | 0.0089308
70 | 101.145 | 4.3334 | 4.3344 | 32.8117 | 32.8112 | 7.2934 | 2.5553 | 19 | 19 | 512 | 7.3002 | 0.96408 | 7.038 |  | 0.0094259
70 | 49.345 | 6.3132 | 6.3145 | 32.7508 | 32.7499 | 6.9615 | 2.755 | 21 | 21 | 514 | 6.8957 | 1.0572 | 7.29 |  | 0.0097173
70 | 10.935 | 12.9578 | 12.9574 | 32.5165 | 32.5158 | 6.0253 | 2.7632 | 23 | 23 | 516 | 5.9035 | 1.0611 | 6.264 |  | 0.0097992
42 | 249.977 | 4.9948 | 4.9964 | 33.8569 | 33.8563 | 7.1284 | 1.2233 | 14 | 14 | 278 | 7.2638 | 0.34596 | 2.513 |  | 0.010288
57 | 11.027 | 13.381 | 13.3805 | 32.4709 | 32.4703 | 5.9746 | 2.749 | 23 | 23 | 391 | 5.8531 | 1.0552 | 6.176 |  | 0.010521
37 | 4.829 | 15.5371 | 15.5372 | 32.2835 | 32.2826 | 5.7268 | 2.7893 | 2 | 2 | 251 | 5.6037 | 1.0741 | 6.019 |  | 0.010713
65 | 4.134 | 13.3814 | 13.3804 | 32.5398 | 32.5391 | 5.972 | 2.7729 | 2 | 3 | 457 | 5.8455 | 1.0666 | 6.235 |  | 0.010858
70 | 3500 | 1.5115 | 1.5107 | 34.6712 | 34.672 | 7.7225 | 1.0032 | 3 | 3 | 496 | 12.1572 | 0.24479 | 2.976 |  | 0.011382
42 | 100.797 | 6.658 | 6.656 | 33.2789 | 33.2611 | 6.882 | 2.1688 | 19 | 19 | 283 | 6.8584 | 0.78896 | 5.411 | 6 | 0.013895
70 | 125.246 | 4.407 | 4.4073 | 33.1387 | 33.1376 | 7.2645 | 2.1651 | 18 | 18 | 511 | 7.2929 | 0.79009 | 5.762 |  | 0.016771
1 | 100.625 | 8.894 | 8.8938 | 30.9281 | 30.9284 | 6.6422 | 0.81048 | 6 | 6 | 6 | 6.6003 | 0.14257 | 0.941 | 6 | 0.0012597
1 | 150.489 | 9.1746 | 9.1747 | 31.1832 | 31.1835 | 6.59 | 0.53967 | 4 | 4 | 4 | 6.5882 | 0.010625 | 0.07 |  | 0.0073596
1 | 50.331 | 9.095 | 9.0963 | 29.6533 | 29.6556 | 6.6672 | 1.4754 | 8 | 8 | 8 | 6.5813 | 0.45507 | 2.995 |  | 0.0022608
1 | 125.255 | 9.026 | 9.0253 | 31.0871 | 31.0869 | 6.6159 | 0.71728 | 5 | 5 | 5 | 6.5939 | 0.095695 | 0.631 |  | 0.0048255
1 | 25.368 | 10.0734 | 10.0147 | 29.0545 | 29.0828 | 6.5495 | 1.6572 | 9 | 9 | 9 | 6.439 | 0.53688 | 3.457 |  | 0.00040936
1 | 74.886 | 8.7565 | 8.7565 | 30.5806 | 30.5809 | 6.6777 | 0.90843 | 7 | 7 | 7 | 6.6147 | 0.17703 | 1.171 |  | 0.012317
1 | 175.02 | 9.2492 | 9.2491 | 31.2212 | 31.2216 | 6.5775 | 0.54107 | 3 | 3 | 3 | 6.5958 | 0 | 0 |  | 0.018635
1 | 199.963 | 9.2507 | 9.2508 | 31.2345 | 31.2349 | 6.5767 | 0.54343 | 2 | 2 | 2 | 6.616 | 0 | 0 |  | 0.019732
1 | 215.514 | 9.2114 | 9.2111 | 31.2345 | 31.2351 | 6.5824 | 0.54796 | 1 | 1 | 1 | 6.6352 | 0 | 0 |  | 0.021837
1 | 10.637 | 12.139 | 12.1021 | 28.2322 | 28.2498 | 6.2965 | 1.9851 | 10 | 10 | 10 | 6.171 | 0.66796 | 4.122 |  | 0.02171
1 | 4.836 | 14.891 | 14.5504 | 27.8528 | 27.896 | 5.9614 | 2.6222 | 11 | 11 | 11 | 5.8337 | 0.87045 | 5.078 |  | 0.11528
24 | 2000.389 | 1.965 | 1.9642 | 34.5863 | 34.5873 | 7.6391 | 0.79575 | 1 | 1 | 139 | 9.8625 | 0.1309 | 1.291 | 4 | 0.0060851
24 | 1499.807 | 2.5813 | 2.5809 | 34.5052 | 34.5057 | 7.5264 | 0.65521 | 2 | 2 | 140 | 9.0815 | 0.07697 | 0.699 | 4 | 0.0052935
31 | 2999.79 | 1.668 | 1.667 | 34.6451 | 34.6463 | 7.6933 | 0.91838 | 3 | 3 | 194 | 11.3349 | 0.24949 | 2.828 | 4 | 0.055523
12 | 175.129 | 7.1187 | 7.1273 | 33.8475 | 33.8455 | 6.784 | 1.4842 | 11 | 11 | 72 | 6.8209 | 0.73685 | 5.026 | 3 | 0.27995
24 | 300.497 | 6.8943 | 6.8945 | 33.976 | 33.9757 | 6.8132 | 0.98408 | 9 | 9 | 147 | 6.964 | 0.22186 | 1.545 | 3 | 0.0026461
42 | 5.609 | 14.1132 | 14.1128 | 32.4031 | 32.4027 | 5.8884 | 2.7698 | 24 | 24 | 288 | 5.7636 | 1.062 | 6.121 | 3 | 0.0076524
70 | 5.134 | 12.9604 | 12.9592 | 32.5164 | 32.516 | 6.025 | 2.7737 | 24 | 24 | 517 | 5.8988 | 1.0639 | 6.276 | 3 | 0.007791
24 | 1248.956 | 2.9708 | 2.9707 | 34.4582 | 34.4584 | 7.4566 | 0.60685 | 3 | 3 | 141 | 8.6965 | 0.044501 | 0.387 | 36 | 0.0047024
31 | 74.775 | 7.4436 | 7.4461 | 32.6849 | 32.6836 | 6.7853 | 2.5526 | 20 | 20 | 211 | 6.7319 | 0.40791 | 2.746 | 36 | 0.54546

## bottle data
Event_Number | Station | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
1 | SI03 | 1 | 0 |  | 9.4 | OXY: H2S. Anoxic
1 | SI03 | 2 | 0 |  | 9.4 | OXY: H2S. Anoxic
1 | SI03 | 3 | 0 |  | 9.5 | OXY: H2S. Anoxic
1 | SI03 | 4 | 0.07 |  | 9.4 | OXY: H2S. Titration curve ok.
1 | SI03 | 5 | 0.631 |  | 9.3
1 | SI03 | 6 | 0.941 | 6 | 9.2
1 | SI03 | 7 | 1.171 |  | 9.1
1 | SI03 | 8 | 2.995 |  | 9.3
1 | SI03 | 9 | 3.457 |  | 10
1 | SI03 | 10 | 4.122 |  | 11.9
1 | SI03 | 11 | 5.078 |  | 13.7
2 | P1 | 13 | 6.929 |  | 12.2
4 | P2 | 15 | 1.828 |  | 7.6
4 | P2 | 16 | 3.392 |  | 7.7
4 | P2 | 17 | 4.535 |  | 7.9
4 | P2 | 18 | 4.51 |  | 7.9
4 | P2 | 19 | 5.659 |  | 8.2
4 | P2 | 20 | 4.393 |  | 8.8
4 | P2 | 21 | 6.578 |  | 12.1
4 | P2 | 22 | 6.438 |  | 14.9
9 | P3 | 50 | 7.339 |  | 14.8
12 | P4 | 62 | 0.551 |  | 4
12 | P4 | 63 | 0.421 |  | 4.2
12 | P4 | 64 | 0.41 |  | 4.2
12 | P4 | 65 | 0.245 |  | 4.8
12 | P4 | 66 | -99 | 9 |  | OXY: Niskin empty
12 | P4 | 67 | 0.408 |  | 5.9
12 | P4 | 68 | 1.093 |  | 6.4
12 | P4 | 69 | 1.858 |  | 6.8
12 | P4 | 70 | -99 | 9 |  | OXY: Niskin empty
12 | P4 | 71 | 2.885 |  | 7.5
12 | P4 | 72 | 5.026 | 3 | 8.2 | OXY: Possible misfire
12 | P4 | 73 | 3.35 |  | 8
12 | P4 | 74 | 3.741 |  | 8.3
12 | P4 | 75 | 5.01 |  | 8.3
12 | P4 | 76 | 6.35 | 6 | 8.7
12 | P4 | 77 | 6.642 |  | 9.7
12 | P4 | 78 | 6.389 |  | 14.6
12 | P4 | 79 | 5.872 |  | 16
12 | P4 | 80 | 5.857 |  | 16.3
19 | P5 | 134 | 5.998 |  | 16.4
21 | P7 | 138 | 5.916 |  | 17.2
24 | P8 | 139 | 1.291 | 4 | 3.8 | OXY: Odd curve. No clear endpoint
24 | P8 | 140 | 0.699 | 4 | 4 | OXY: Bubble resulting in an odd curve.
24 | P8 | 141 | 0.387 | 36 | 4.9 | OXY: Poor duplication
24 | P8 | 142 | 0.237 |  | 5.5
24 | P8 | 143 | 0.233 |  | 5.9
24 | P8 | 144 | 0.468 |  | 6.4
24 | P8 | 145 | 0.759 |  | 6.7
24 | P8 | 146 | 1.156 |  | 7.2
24 | P8 | 147 | 1.545 | 3 | 7.9 | OXY: Bubbles in burette
24 | P8 | 148 | 1.757 |  | 8.2
24 | P8 | 149 | 2.164 |  | 8.7
24 | P8 | 150 | 2.186 | 6 | 8.8
24 | P8 | 151 | 2.182 |  | 8.9
24 | P8 | 152 | 2.189 |  | 9.1
24 | P8 | 153 | 2.458 |  | 9.3
24 | P8 | 154 | 4.641 |  | 9.3
24 | P8 | 155 | 6.299 |  | 9.7
24 | P8 | 156 | 6.948 |  | 13.5
24 | P8 | 157 | 5.918 |  | 16.3
24 | P8 | 158 | 5.861 |  | 16.6
26 | P9 | 164 | 5.895 |  | 17
28 | P11 | 168 | 6.112 |  | 17.6
31 | P12 | 192 | 2.288 |  | 3.5
31 | P12 | 193 | 2.226 |  | 2.7
31 | P12 | 194 | 2.828 | 4 | 2.6 | OXY: Bubble, murky. Titration messy and incomplete. Value not trusted.
31 | P12 | 195 | 1.847 | 6 | 3.1
31 | P12 | 196 | 1.411 |  | 3.3
31 | P12 | 197 | 0.698 |  | 3.5
31 | P12 | 198 | 0.399 |  | 4
31 | P12 | 199 | 0.266 |  | 4.5
31 | P12 | 200 | 0.215 |  | 5.1
31 | P12 | 201 | 0.311 |  | 5.6
31 | P12 | 202 | 0.536 |  | 5.9
31 | P12 | 203 | 0.95 | 6 | 6
31 | P12 | 204 | 1.616 |  | 6.2
31 | P12 | 205 | 2.473 |  | 6.3
31 | P12 | 206 | 3.576 |  | 6.5
31 | P12 | 207 | 3.844 |  | 6.9
31 | P12 | 208 | 4.055 |  | 7.2
31 | P12 | 209 | 4.516 |  | 7.5
31 | P12 | 210 | 5.043 |  | 7.7
31 | P12 | 211 | 2.746 | 36 | 8.1 | OXY: Replicate Outliers
31 | P12 | 212 | 6.946 |  | 8.5
31 | P12 | 213 | 6.743 |  | 9.5
31 | P12 | 214 | 6.044 |  | 13.8
31 | P12 | 215 | 6.002 |  | 16.5
37 | P13 | 251 | 6.019 |  | 15.8
40 | P15 | 255 | 6.1 |  | 14.4
42 | P16 | 265 | 2.694 |  | 2.9
42 | P16 | 266 | 2.686 |  | 2.6
42 | P16 | 267 | 2.707 |  | 2.7
42 | P16 | 268 | 2.427 |  | 2.7
42 | P16 | 269 | 1.935 | 6 | 3.3
42 | P16 | 270 | 1.295 |  | 3.3
42 | P16 | 271 | 0.535 |  | 3.8
42 | P16 | 272 | 0.318 |  | 4.2
42 | P16 | 273 | 0.286 |  | 4.9
42 | P16 | 274 | 0.368 |  | 4.9
42 | P16 | 275 | 0.561 | 6 | 5.5
42 | P16 | 276 | 1.454 |  | 5.7
42 | P16 | 277 | 2.027 |  | 6
42 | P16 | 278 | 2.513 |  | 6.3
42 | P16 | 279 | 3.575 |  | 6.9
42 | P16 | 280 | 3.947 |  | 7.2
42 | P16 | 281 | 4.127 |  | 7.4
42 | P16 | 282 | 4.433 |  | 7.6
42 | P16 | 283 | 5.411 | 6 | 7.9
42 | P16 | 284 | 6.829 |  | 7.9
42 | P16 | 285 | 7.077 |  | 8.5
42 | P16 | 286 | 6.482 |  | 12.7
42 | P16 | 287 | 6.132 |  | 14.2
42 | P16 | 288 | 6.121 | 3 | 14.1 | OXY: Leaky spigot
49 | P17 | 343 | 6.111 |  | 14.1
52 | P19 | 347 | 6.136 |  | 13.7
57 | P20 | 369 | 2.988 |  | 2.5
57 | P20 | 370 | 2.863 |  | 2.6
57 | P20 | 371 | 2.849 |  | 2.4
57 | P20 | 372 | 2.512 |  | 2.5
57 | P20 | 373 | 1.976 | 6 | 2.8
57 | P20 | 374 | 1.29 |  | 3
57 | P20 | 375 | 0.582 |  | 3.4
57 | P20 | 376 | 0.351 |  | 3.5
57 | P20 | 377 | 0.343 |  | 4
57 | P20 | 378 | 0.404 |  | 4.5
57 | P20 | 379 | 0.575 | 6 | 5
57 | P20 | 380 | 1.253 |  | 5.4
57 | P20 | 381 | 1.79 |  | 5.7
57 | P20 | 382 | 2.162 |  | 5.9
57 | P20 | 383 | 2.925 |  | 6.3
57 | P20 | 384 | 3.22 |  | 6.5
57 | P20 | 385 | 3.829 |  | 7
57 | P20 | 386 | 4.178 |  | 7.1
57 | P20 | 387 | 5.449 |  | 7.3
57 | P20 | 388 | 6.788 | 6 | 7.4
57 | P20 | 389 | 7.205 |  | 8.4
57 | P20 | 390 | 6.321 |  | 12.7
57 | P20 | 391 | 6.176 |  | 13.5
57 | P20 | 392 | 6.161 |  | 13.4
60 | P21 | 427 | 6.199 |  | 13.9
62 | P23 | 431 | 6.182 |  | 13.1
65 | P25 | 457 | 6.235 |  | 13.3
66 | P35 | 459 | 6.202 |  | 13.1
70 | P26 | 494 | 3.198 |  | 2.5
70 | P26 | 495 | 3.177 |  | 2.4
70 | P26 | 496 | 2.976 |  | 2.4
70 | P26 | 497 | 2.569 |  | 2.4
70 | P26 | 498 | 1.998 | 6 | 2.7
70 | P26 | 499 | 1.331 |  | 2.8
70 | P26 | 500 | 0.617 |  | 3.2
70 | P26 | 501 | 0.452 |  | 3.4 | OXY: Wavy end of curve but end point seems reasonable
70 | P26 | 502 | 0.495 |  | 3.7
70 | P26 | 503 | 0.572 |  | 4.1
70 | P26 | 504 | 0.653 |  | 4.5
70 | P26 | 505 | 1.111 |  | 5
70 | P26 | 506 | 1.581 | 6 | 5.6
70 | P26 | 507 | 2.104 |  | 5.2
70 | P26 | 508 | 2.662 |  | 5.2
70 | P26 | 509 | 3.03 |  | 5.4
70 | P26 | 510 | 3.537 |  | 5.6
70 | P26 | 511 | 5.762 |  | 5.6
70 | P26 | 512 | 7.038 |  | 5.7
70 | P26 | 513 | 7.077 | 6 | 5.9
70 | P26 | 514 | 7.29 |  | 7.2
70 | P26 | 515 | 6.256 |  | 12.6
70 | P26 | 516 | 6.264 |  | 12.8
70 | P26 | 517 | 6.276 | 3 | 12.7 | OXY: Top valve not sealed

## ctd data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Number_of_bin_records
1 | 215.514 | 9.2114 | 9.2111 | 31.2345 | 31.2351 | 6.5824 | 0.54796 | 1 | 1 | 1 | 241
1 | 199.963 | 9.2507 | 9.2508 | 31.2345 | 31.2349 | 6.5767 | 0.54343 | 2 | 2 | 2 | 241
1 | 175.02 | 9.2492 | 9.2491 | 31.2212 | 31.2216 | 6.5775 | 0.54107 | 3 | 3 | 3 | 241
1 | 150.489 | 9.1746 | 9.1747 | 31.1832 | 31.1835 | 6.59 | 0.53967 | 4 | 4 | 4 | 241
1 | 125.255 | 9.026 | 9.0253 | 31.0871 | 31.0869 | 6.6159 | 0.71728 | 5 | 5 | 5 | 241
1 | 100.625 | 8.894 | 8.8938 | 30.9281 | 30.9284 | 6.6422 | 0.81048 | 6 | 6 | 6 | 241
1 | 74.886 | 8.7565 | 8.7565 | 30.5806 | 30.5809 | 6.6777 | 0.90843 | 7 | 7 | 7 | 241
1 | 50.331 | 9.095 | 9.0963 | 29.6533 | 29.6556 | 6.6672 | 1.4754 | 8 | 8 | 8 | 241
1 | 25.368 | 10.0734 | 10.0147 | 29.0545 | 29.0828 | 6.5495 | 1.6572 | 9 | 9 | 9 | 241
1 | 10.637 | 12.139 | 12.1021 | 28.2322 | 28.2498 | 6.2965 | 1.9851 | 10 | 10 | 10 | 241
1 | 4.836 | 14.891 | 14.5504 | 27.8528 | 27.896 | 5.9614 | 2.6222 | 11 | 11 | 11 | 241
2 | 4.709 | 13.3395 | 13.3405 | 31.2305 | 31.2305 | 6.0261 | 3.0281 | 2 | 2 | 13 | 241
4 | 105.069 | 7.1029 | 7.0994 | 33.7254 | 33.7284 | 6.7919 | 1.0123 | 1 | 1 | 15 | 241
4 | 99.752 | 7.2461 | 7.2444 | 33.6623 | 33.6635 | 6.7725 | 1.5462 | 2 | 2 | 16 | 241
4 | 74.807 | 7.5201 | 7.5225 | 33.1454 | 33.1413 | 6.7531 | 1.9665 | 4 | 4 | 18 | 241
4 | 74.731 | 7.5122 | 7.5145 | 33.1604 | 33.1568 | 6.7537 | 1.9419 | 3 | 3 | 17 | 139
4 | 50.299 | 7.7661 | 7.7661 | 32.8496 | 32.85 | 6.7285 | 2.3212 | 5 | 5 | 19 | 241
4 | 25.136 | 8.6286 | 8.6351 | 32.4955 | 32.4936 | 6.6143 | 1.918 | 6 | 6 | 20 | 241
4 | 10.474 | 12.1618 | 12.1439 | 31.9933 | 32.0004 | 6.1463 | 2.8433 | 7 | 7 | 21 | 241
4 | 5.282 | 14.9696 | 14.9738 | 31.5045 | 31.5033 | 5.8197 | 2.8851 | 8 | 8 | 22 | 241
10 | 20.298 | 11.8372 | 11.845 | 32.4001 | 32.3989 | 6.1726 | 2.9606 | 4 | 2 | 50 | 241
12 | 1315.005 | 2.8194 | 2.8191 | 34.4902 | 34.4911 | 7.4829 | 0.64344 | 1 | 1 | 62 | 241
12 | 1250.736 | 3.0654 | 3.065 | 34.4627 | 34.4634 | 7.4389 | 0.61828 | 3 | 3 | 64 | 241
12 | 1250.621 | 3.0639 | 3.0638 | 34.4627 | 34.4635 | 7.4392 | 0.61866 | 2 | 2 | 63 | 241
12 | 1001.127 | 3.6231 | 3.6227 | 34.3883 | 34.3887 | 7.3416 | 0.58046 | 4 | 4 | 65 | 241
12 | 601.42 | 4.6668 | 4.6668 | 34.1784 | 34.1785 | 7.169 | 0.62026 | 6 | 6 | 67 | 241
12 | 400.403 | 5.3628 | 5.3629 | 34.0219 | 34.022 | 7.0587 | 0.83021 | 7 | 7 | 68 | 241
12 | 301.316 | 5.9538 | 5.9542 | 33.9553 | 33.9553 | 6.9644 | 1.0639 | 8 | 8 | 69 | 241
12 | 200.112 | 6.8348 | 6.8383 | 33.8887 | 33.8883 | 6.8265 | 1.4051 | 10 | 10 | 71 | 241
12 | 175.129 | 7.1187 | 7.1273 | 33.8475 | 33.8455 | 6.784 | 1.4842 | 11 | 11 | 72 | 241
12 | 150.378 | 7.4103 | 7.4103 | 33.757 | 33.7571 | 6.7429 | 1.5668 | 12 | 12 | 73 | 241
12 | 125.402 | 7.7548 | 7.7552 | 33.5329 | 33.5346 | 6.7002 | 1.6801 | 13 | 13 | 74 | 241
12 | 100.895 | 7.6901 | 7.6913 | 33.0805 | 33.0785 | 6.7299 | 2.1145 | 14 | 14 | 75 | 241
12 | 76.089 | 8.0908 | 8.0911 | 32.6575 | 32.6574 | 6.6876 | 2.5466 | 15 | 15 | 76 | 241
12 | 50.131 | 9.3381 | 9.3469 | 32.5185 | 32.5179 | 6.5099 | 2.7125 | 16 | 16 | 77 | 241
12 | 24.694 | 14.9601 | 14.9615 | 32.255 | 32.253 | 5.7941 | 2.8688 | 17 | 17 | 78 | 241
12 | 10.038 | 16.3193 | 16.3249 | 32.1836 | 32.1829 | 5.6428 | 2.7646 | 18 | 18 | 79 | 241
12 | 5.265 | 16.3588 | 16.3583 | 32.1851 | 32.1848 | 5.6383 | 2.7737 | 19 | 19 | 80 | 241
19 | 4.934 | 16.1084 | 16.1044 | 32.0273 | 32.0268 | 5.6715 | 2.8161 | 2 | 2 | 134 | 72
21 | 5.666 | 16.2526 | 16.2502 | 32.2581 | 32.257 | 5.6476 | 2.7885 | 2 | 2 | 138 | 241
24 | 2000.389 | 1.965 | 1.9642 | 34.5863 | 34.5873 | 7.6391 | 0.79575 | 1 | 1 | 139 | 241
24 | 1499.807 | 2.5813 | 2.5809 | 34.5052 | 34.5057 | 7.5264 | 0.65521 | 2 | 2 | 140 | 241
24 | 1248.956 | 2.9708 | 2.9707 | 34.4582 | 34.4584 | 7.4566 | 0.60685 | 3 | 3 | 141 | 241
24 | 999.133 | 3.6428 | 3.6426 | 34.3781 | 34.3782 | 7.3386 | 0.57507 | 4 | 4 | 142 | 241
24 | 799.882 | 4.1331 | 4.1336 | 34.2903 | 34.2902 | 7.2561 | 0.57363 | 5 | 5 | 143 | 241
24 | 601.24 | 4.8435 | 4.8435 | 34.1593 | 34.1593 | 7.1397 | 0.63927 | 6 | 6 | 144 | 241
24 | 499.592 | 5.3011 | 5.3012 | 34.0857 | 34.0855 | 7.066 | 0.73191 | 7 | 7 | 145 | 241
24 | 400.115 | 5.998 | 5.9981 | 34.0309 | 34.0307 | 6.9537 | 0.85262 | 8 | 8 | 146 | 241
24 | 300.497 | 6.8943 | 6.8945 | 33.976 | 33.9757 | 6.8132 | 0.98408 | 9 | 9 | 147 | 241
24 | 249.722 | 7.1701 | 7.1706 | 33.9424 | 33.9422 | 6.7718 | 1.0526 | 10 | 10 | 148 | 241
24 | 199.834 | 7.6094 | 7.6094 | 33.8631 | 33.8629 | 6.7078 | 1.1892 | 11 | 11 | 149 | 241
24 | 175.781 | 7.7335 | 7.7336 | 33.8421 | 33.8419 | 6.6899 | 1.2051 | 12 | 12 | 150 | 241
24 | 150.003 | 7.8298 | 7.8296 | 33.8319 | 33.8317 | 6.6758 | 1.2111 | 13 | 13 | 151 | 241
24 | 125.197 | 7.9385 | 7.9394 | 33.8189 | 33.8185 | 6.66 | 1.2172 | 14 | 14 | 152 | 241
24 | 99.756 | 8.1154 | 8.1158 | 33.7201 | 33.7197 | 6.6378 | 1.3221 | 15 | 15 | 153 | 241
24 | 74.625 | 8.065 | 8.0659 | 32.9979 | 32.9986 | 6.6766 | 2.0435 | 16 | 16 | 154 | 241
24 | 50.106 | 8.4249 | 8.4257 | 32.583 | 32.5828 | 6.6407 | 2.5591 | 17 | 17 | 155 | 241
24 | 46.774 | 8.5552 | 8.5543 | 32.5684 | 32.5685 | 6.622 | 2.5923 | 21 | 18 | 156 | 241
24 | 36.701 | 9.0176 | 9.0182 | 32.5211 | 32.5208 | 6.5561 | 2.7364 | 22 | 19 | 157 | 241
24 | 24.498 | 12.9709 | 12.9935 | 32.3603 | 32.3587 | 6.0295 | 3.0473 | 18 | 20 | 158 | 241
26 | 4.375 | 16.9403 | 16.9338 | 32.0694 | 32.0679 | 5.5788 | 2.796 | 2 | 2 | 164 | 241
28 | 5.291 | 16.3738 | 16.3722 | 32.473 | 32.4725 | 5.6268 | 2.8693 | 2 | 2 | 168 | 241
31 | 3269.304 | 1.6798 | 1.6786 | 34.6471 | 34.6486 | 7.6909 | 0.9121 | 1 | 1 | 192 | 241
31 | 3000.008 | 1.6679 | 1.6669 | 34.6451 | 34.6464 | 7.6934 | 0.91836 | 2 | 2 | 193 | 169
31 | 2999.79 | 1.668 | 1.667 | 34.6451 | 34.6463 | 7.6933 | 0.91838 | 3 | 3 | 194 | 241
31 | 2499.957 | 1.7586 | 1.758 | 34.6272 | 34.6278 | 7.6767 | 0.87756 | 4 | 4 | 195 | 241
31 | 1999.872 | 1.9142 | 1.9138 | 34.5933 | 34.5936 | 7.6485 | 0.81564 | 5 | 5 | 196 | 241
31 | 1499.272 | 2.3401 | 2.34 | 34.5138 | 34.5137 | 7.5714 | 0.6784 | 6 | 6 | 197 | 241
31 | 1250.014 | 2.6937 | 2.6941 | 34.4642 | 34.464 | 7.5076 | 0.61554 | 7 | 7 | 198 | 241
31 | 1001.626 | 3.219 | 3.2188 | 34.4015 | 34.4012 | 7.414 | 0.58425 | 8 | 8 | 199 | 241
31 | 800.426 | 3.723 | 3.7232 | 34.3255 | 34.325 | 7.3269 | 0.57667 | 9 | 9 | 200 | 241
31 | 600.037 | 4.2045 | 4.2046 | 34.2139 | 34.2138 | 7.2474 | 0.60272 | 10 | 10 | 201 | 241
31 | 500.543 | 4.4196 | 4.4197 | 34.123 | 34.1226 | 7.2143 | 0.66706 | 11 | 11 | 202 | 241
31 | 400.324 | 4.7331 | 4.7354 | 34.0377 | 34.0374 | 7.1644 | 0.78437 | 12 | 12 | 203 | 241
31 | 299.4 | 5.0165 | 5.0165 | 33.9447 | 33.9444 | 7.1205 | 0.98958 | 13 | 13 | 204 | 241
31 | 251.715 | 5.3447 | 5.3454 | 33.8799 | 33.8796 | 7.0684 | 1.2356 | 14 | 14 | 205 | 241
31 | 200.906 | 5.8658 | 5.8674 | 33.8286 | 33.8282 | 6.9846 | 1.5925 | 15 | 15 | 206 | 241
31 | 174.986 | 6.1489 | 6.1512 | 33.816 | 33.8155 | 6.9391 | 1.6803 | 16 | 16 | 207 | 241
31 | 150.234 | 6.494 | 6.5009 | 33.7601 | 33.7587 | 6.8862 | 1.7754 | 17 | 17 | 208 | 241
31 | 125.837 | 6.8368 | 6.8368 | 33.558 | 33.5577 | 6.841 | 1.9454 | 18 | 18 | 209 | 241
31 | 100.648 | 6.9771 | 6.9767 | 33.3276 | 33.3285 | 6.8293 | 2.0755 | 19 | 19 | 210 | 241
31 | 74.775 | 7.4436 | 7.4461 | 32.6849 | 32.6836 | 6.7853 | 2.5526 | 20 | 20 | 211 | 241
31 | 51.017 | 8.8677 | 8.8685 | 32.566 | 32.5657 | 6.5761 | 2.7946 | 21 | 21 | 212 | 241
31 | 25.259 | 13.5157 | 13.5164 | 32.4344 | 32.4332 | 5.9595 | 2.8731 | 22 | 22 | 213 | 241
31 | 10.797 | 16.08 | 16.0675 | 32.4266 | 32.4261 | 5.6609 | 2.8151 | 23 | 23 | 214 | 241
31 | 5.565 | 16.1145 | 16.1136 | 32.4284 | 32.4279 | 5.657 | 2.8229 | 24 | 24 | 215 | 241
37 | 4.829 | 15.5371 | 15.5372 | 32.2835 | 32.2826 | 5.7268 | 2.7893 | 2 | 2 | 251 | 241
40 | 4.915 | 14.2902 | 14.29 | 32.4036 | 32.4031 | 5.8673 | 2.7897 | 2 | 2 | 255 | 241
42 | 3685.22 | 1.5801 | 1.5786 | 34.6666 | 34.6682 | 7.7093 | 0.96075 | 1 | 1 | 265 | 241
42 | 3500.601 | 1.5722 | 1.571 | 34.6649 | 34.6663 | 7.711 | 0.96691 | 2 | 2 | 266 | 39
42 | 3500.403 | 1.5723 | 1.5711 | 34.6648 | 34.6662 | 7.711 | 0.96686 | 3 | 3 | 267 | 241
42 | 3001.204 | 1.6019 | 1.6011 | 34.6533 | 34.6541 | 7.7058 | 0.95246 | 4 | 4 | 268 | 241
42 | 2499.125 | 1.7245 | 1.7239 | 34.6302 | 34.6308 | 7.6832 | 0.89118 | 5 | 5 | 269 | 241
42 | 1999.958 | 1.9433 | 1.943 | 34.5881 | 34.5883 | 7.6431 | 0.79006 | 6 | 6 | 270 | 241
42 | 1500.454 | 2.387 | 2.387 | 34.5007 | 34.5006 | 7.5632 | 0.64244 | 7 | 7 | 271 | 241
42 | 1251.288 | 2.6654 | 2.6654 | 34.444 | 34.4438 | 7.5139 | 0.59492 | 8 | 8 | 272 | 241
42 | 1000.517 | 3.0275 | 3.0277 | 34.3671 | 34.3668 | 7.4508 | 0.58506 | 9 | 9 | 273 | 241
42 | 800.521 | 3.3926 | 3.3926 | 34.2771 | 34.2768 | 7.3887 | 0.61287 | 10 | 10 | 274 | 241
42 | 599.876 | 3.8304 | 3.8309 | 34.1593 | 34.1588 | 7.316 | 0.66675 | 11 | 11 | 275 | 241
42 | 399.265 | 4.2658 | 4.2661 | 33.9783 | 33.9778 | 7.2481 | 0.91752 | 12 | 12 | 276 | 241
42 | 300.142 | 4.5655 | 4.5656 | 33.8948 | 33.8944 | 7.2001 | 1.0966 | 13 | 13 | 277 | 241
42 | 249.977 | 4.9948 | 4.9964 | 33.8569 | 33.8563 | 7.1284 | 1.2233 | 14 | 14 | 278 | 241
42 | 200.348 | 5.7078 | 5.707 | 33.8237 | 33.8233 | 7.0108 | 1.579 | 15 | 15 | 279 | 241
42 | 174.945 | 6.1566 | 6.1573 | 33.8318 | 33.8313 | 6.9371 | 1.7171 | 16 | 16 | 280 | 241
42 | 150.479 | 6.4011 | 6.4056 | 33.8233 | 33.8225 | 6.8982 | 1.7794 | 17 | 17 | 281 | 241
42 | 125.664 | 6.6292 | 6.6301 | 33.7074 | 33.706 | 6.8671 | 1.8875 | 18 | 18 | 282 | 241
42 | 100.797 | 6.658 | 6.656 | 33.2789 | 33.2611 | 6.882 | 2.1688 | 19 | 19 | 283 | 241
42 | 75.583 | 6.8569 | 6.8599 | 32.7248 | 32.7238 | 6.8755 | 2.6349 | 20 | 20 | 284 | 241
42 | 50.502 | 7.4864 | 7.4808 | 32.6281 | 32.6274 | 6.7812 | 2.7618 | 21 | 21 | 285 | 241
42 | 24.983 | 12.6 | 12.598 | 32.4795 | 32.479 | 6.0716 | 2.8297 | 22 | 22 | 286 | 241
42 | 10 | 14.0898 | 14.0883 | 32.4033 | 32.4024 | 5.8912 | 2.7701 | 23 | 23 | 287 | 241
42 | 5.609 | 14.1132 | 14.1128 | 32.4031 | 32.4027 | 5.8884 | 2.7698 | 24 | 24 | 288 | 241
49 | 4.629 | 14.0401 | 14.0384 | 32.4157 | 32.415 | 5.8967 | 2.7622 | 2 | 2 | 343 | 241
52 | 5.184 | 13.7407 | 13.741 | 32.4492 | 32.4483 | 5.9315 | 2.7564 | 2 | 2 | 347 | 241
57 | 4022.677 | 1.545 | 1.5436 | 34.6763 | 34.678 | 7.7157 | 0.97997 | 1 | 1 | 369 | 241
57 | 3501.493 | 1.5269 | 1.526 | 34.67 | 34.6711 | 7.7195 | 0.99276 | 3 | 3 | 371 | 241
57 | 3500.604 | 1.5265 | 1.5256 | 34.6701 | 34.6711 | 7.7196 | 0.99284 | 2 | 2 | 370 | 124
57 | 2999.278 | 1.5753 | 1.5746 | 34.6558 | 34.6566 | 7.7108 | 0.96689 | 4 | 4 | 372 | 241
57 | 2499.998 | 1.6956 | 1.6954 | 34.6319 | 34.6322 | 7.6887 | 0.9013 | 5 | 5 | 373 | 241
57 | 1999.763 | 1.9308 | 1.9306 | 34.5858 | 34.5859 | 7.6456 | 0.78759 | 6 | 6 | 374 | 241
57 | 1500.095 | 2.3321 | 2.332 | 34.5044 | 34.5043 | 7.5733 | 0.64997 | 7 | 7 | 375 | 241
57 | 1250.473 | 2.6308 | 2.6308 | 34.4449 | 34.4446 | 7.5203 | 0.60197 | 8 | 8 | 376 | 241
57 | 999.566 | 3.0091 | 3.009 | 34.3633 | 34.363 | 7.4543 | 0.60141 | 9 | 9 | 377 | 241
57 | 800.135 | 3.3273 | 3.3273 | 34.2883 | 34.2879 | 7.4 | 0.62037 | 10 | 10 | 378 | 241
57 | 600.064 | 3.71 | 3.7102 | 34.1767 | 34.1763 | 7.3366 | 0.66924 | 11 | 11 | 379 | 241
57 | 399.365 | 4.0996 | 4.1 | 34.0019 | 34.0013 | 7.2762 | 0.86094 | 12 | 12 | 380 | 241
57 | 299.9 | 4.3839 | 4.3857 | 33.9055 | 33.9048 | 7.2311 | 1.0189 | 13 | 13 | 381 | 241
57 | 250.138 | 4.644 | 4.6438 | 33.8604 | 33.8599 | 7.1882 | 1.1373 | 14 | 14 | 382 | 241
57 | 200.057 | 5.1109 | 5.1146 | 33.8301 | 33.8291 | 7.11 | 1.3609 | 15 | 15 | 383 | 241
57 | 174.66 | 5.5219 | 5.5254 | 33.8163 | 33.8156 | 7.0419 | 1.4707 | 16 | 16 | 384 | 241
57 | 149.932 | 5.9367 | 5.9389 | 33.7883 | 33.7876 | 6.9749 | 1.6674 | 17 | 17 | 385 | 241
57 | 124.366 | 6.2528 | 6.2528 | 33.6201 | 33.6192 | 6.9313 | 1.7926 | 18 | 18 | 386 | 241
57 | 98.96 | 6.2524 | 6.2526 | 33.1489 | 33.1494 | 6.9531 | 2.1819 | 19 | 19 | 387 | 241
57 | 74.958 | 6.3908 | 6.3952 | 32.7868 | 32.7841 | 6.9473 | 2.6117 | 20 | 20 | 388 | 241
57 | 48.629 | 7.4448 | 7.4437 | 32.662 | 32.6614 | 6.7862 | 2.7973 | 21 | 21 | 389 | 241
57 | 24.167 | 12.6205 | 12.6082 | 32.5147 | 32.5129 | 6.0677 | 2.7833 | 22 | 22 | 390 | 241
57 | 11.027 | 13.381 | 13.3805 | 32.4709 | 32.4703 | 5.9746 | 2.749 | 23 | 23 | 391 | 241
57 | 4.839 | 13.413 | 13.4136 | 32.4719 | 32.4713 | 5.9707 | 2.7578 | 24 | 24 | 392 | 241
60 | 5.374 | 13.6483 | 13.6475 | 32.4951 | 32.4943 | 5.9411 | 2.7802 | 2 | 2 | 427 | 241
62 | 5.528 | 13.2136 | 13.2131 | 32.4852 | 32.4843 | 5.9947 | 2.7561 | 2 | 2 | 431 | 241
65 | 4.134 | 13.3814 | 13.3804 | 32.5398 | 32.5391 | 5.972 | 2.7729 | 2 | 3 | 457 | 241
66 | 5.531 | 13.2816 | 13.2806 | 32.5043 | 32.5032 | 5.9856 | 2.7637 | 2 | 2 | 459 | 241
70 | 4315.976 | 1.5216 | 1.5201 | 34.6824 | 34.6841 | 7.7199 | 0.99571 | 1 | 1 | 494 | 241
70 | 3997.958 | 1.5006 | 1.4995 | 34.6796 | 34.6809 | 7.7242 | 1.0084 | 2 | 2 | 495 | 241
70 | 3500 | 1.5115 | 1.5107 | 34.6712 | 34.672 | 7.7225 | 1.0032 | 3 | 3 | 496 | 241
70 | 3000.115 | 1.5683 | 1.5678 | 34.6564 | 34.6568 | 7.7122 | 0.9732 | 4 | 4 | 497 | 241
70 | 2500.322 | 1.703 | 1.7025 | 34.6296 | 34.6298 | 7.6874 | 0.89916 | 5 | 5 | 498 | 241
70 | 1999.531 | 1.9238 | 1.9235 | 34.586 | 34.586 | 7.647 | 0.79262 | 6 | 6 | 499 | 241
70 | 1499.325 | 2.3017 | 2.3016 | 34.5072 | 34.507 | 7.579 | 0.65588 | 7 | 7 | 500 | 241
70 | 1250.434 | 2.5776 | 2.5776 | 34.448 | 34.4478 | 7.5301 | 0.61889 | 8 | 8 | 501 | 241
70 | 999.832 | 2.8818 | 2.8818 | 34.3793 | 34.3789 | 7.477 | 0.63213 | 9 | 9 | 502 | 241
70 | 800.324 | 3.2004 | 3.2 | 34.307 | 34.3066 | 7.4221 | 0.6483 | 10 | 10 | 503 | 241
70 | 600.49 | 3.6099 | 3.6097 | 34.1978 | 34.1973 | 7.3535 | 0.68117 | 11 | 11 | 504 | 241
70 | 400.439 | 3.9111 | 3.9113 | 34.0329 | 34.0324 | 7.3079 | 0.81066 | 12 | 12 | 505 | 241
70 | 301.499 | 3.9619 | 3.9621 | 33.9169 | 33.9162 | 7.3047 | 0.95643 | 13 | 13 | 506 | 241
70 | 249.078 | 3.9688 | 3.969 | 33.8475 | 33.8469 | 7.3069 | 1.0994 | 14 | 14 | 507 | 241
70 | 201.165 | 4.0056 | 4.006 | 33.7731 | 33.7725 | 7.304 | 1.2638 | 15 | 15 | 508 | 241
70 | 174.859 | 4.1289 | 4.1294 | 33.738 | 33.7375 | 7.2839 | 1.3767 | 16 | 16 | 509 | 241
70 | 149.541 | 4.3393 | 4.3397 | 33.6801 | 33.6796 | 7.2499 | 1.5292 | 17 | 17 | 510 | 241
70 | 125.246 | 4.407 | 4.4073 | 33.1387 | 33.1376 | 7.2645 | 2.1651 | 18 | 18 | 511 | 241
70 | 101.145 | 4.3334 | 4.3344 | 32.8117 | 32.8112 | 7.2934 | 2.5553 | 19 | 19 | 512 | 241
70 | 76.088 | 4.6462 | 4.6461 | 32.7944 | 32.7938 | 7.2395 | 2.597 | 20 | 20 | 513 | 241
70 | 49.345 | 6.3132 | 6.3145 | 32.7508 | 32.7499 | 6.9615 | 2.755 | 21 | 21 | 514 | 241
70 | 24.929 | 12.9303 | 12.9302 | 32.5171 | 32.5158 | 6.0287 | 2.7596 | 22 | 22 | 515 | 241
70 | 10.935 | 12.9578 | 12.9574 | 32.5165 | 32.5158 | 6.0253 | 2.7632 | 23 | 23 | 516 | 241
70 | 5.134 | 12.9604 | 12.9592 | 32.5164 | 32.516 | 6.025 | 2.7737 | 24 | 24 | 517 | 241

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers