## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |

## PROCESSING NOTES

Cruise: 2011-28

Agency: OSD

Location: Strait of Georgia

Project: SoG Mooring

Party Chief: Johannessen S.

Platform: Vector

Date: April 4, 2011 – April 7, 2011

Processed by: Germaine Gatien

Date of Processing: 27 October 2011 – 1 November 2011

Number of original HEX files: 12   	Number of original ROS files: 12

Number of CTD files:   2	Number of CHE files:  12

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1483), a Seapoint Fluorometer (#2745) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4694), a surface PAR sensor (#16504) and an altimeter (#1252).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had no equipment list. There was a Rinko III Temperature sensor listed in the configuration file, but it was not actually mounted.

There was one error in the configuration file – the pressure offset (-0.9db) had not been entered.

Draw temperatures were not available for this cruise, so titrated Dissolved Oxygen values are only available in mL/L.

There were 7 duplicate salinity samples, but most were from either the surface or the bottom of the cast. It is better to take mid-depth duplicates if the aim is to test the Autosal and/or sampling protocols, but this does give a measure of variability. The average salinity difference was 0.015 and the standard deviation in the differences was 0.013. Even at mid-depths of the deeper casts the differences were on the order of 0.01.

The salinity comparison was very noisy but suggested that both salinity sensors were reading low by approximately 0.01. This seems unlikely in sensors that were recently serviced and which agreed with each other very well. The noise in the comparison and the small number of samples did not support recalibration. There is no evidence of non-linearity in the Autosal. The poor results may be due to the fact that the samples were not analyzed until 5.7 months after the cruise. Cruises 2011-08 and 2011-16 had similar results, but the differences were smaller, ~0.003 and ~0.004, and the delay in analysis a little shorter at 3 months and 4.5 months, respectively. All salinity samples from this cruise were flagged “3”.

Events 2 to 11 were surface casts only, so not suitable for preparation of CTD files. Only CHE files were prepared.

Starting with 2011 cruises a new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel was compared with bottles to find the slope and offset to enter in the configuration files. While a little less convenient than calibrating oxygen concentration against bottles, it is the standard approach and is recommended by SeaBird, and should produce better results. For this cruise there were few samples, so the results of cruise 2011-08 were used since that cruise used the same equipment shortly afterwards and had more sampling.

Starting with 2011, the transmissivity conversion has been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder:

OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, dissolved oxygen and salinity sample data were obtained in spreadsheet format from the analysts. The file creation date was added to the name of the salinity file to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

Only one configuration was used for the cruise. The calibration constants were checked.

- The only error found was in the offset for the pressure sensor which was changed to -0.9039.

- The transmissometer calibration was entered correctly based on the information available at cruise time, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values were entered in the configuration file. The result will be higher values by <1%.

- There is an entry for a Rinko III Temperature sensor, but a test showed no signal in that channel, so it must be an artefact from another cruise. The same conclusion was reached for 2011-08.

After those changes, the file was saved as 2011-28-ctd-original.con.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted including Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT).

The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-28-DO-cal.csv. The titrated DO values and flags were added to that file; there were not entries for every line so those without DO sampling needed to be removed. A calculation was made of Ф using equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P, T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

When all data were included M= 0.4511 and B= -0.2419. R2 =0.9851

When outliers were gradually removed based on the distance from the fit, the following results were found. From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are defined as Soc = M and Voffset = B/M.

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| original |  | 0.4329 | -0.2218 | 0.4329 | -0.5123 |  |
| early April | all | 0.4511 | -0.2419 | -0.4511 | -0.5362 | 0.9851 |
| early April | 1 outlier excluded | 0.4377 | -0.2165 | 0.4377 | -0.4946 | 0.9963 |
| early April | 2 outliers excluded | 0.4323 | -0.2066 | 0.4323 | -0.4779 | 0.9984 |
| early April | 3 outliers excluded | 0.4334 | -0.2087 | 0.4334 | -0.4815 | 0.9983 |
| early April | 7 outliers excluded | 0.4320 | -0.2102 | 0.4320 | -0.4866 | 0.9981 |
| mid-April | 2011-08 best fit | 0.4346 | -0.2135 | 0.4346 | -0.4913 |  |

The most extreme outlier is likely a bad bottle since there was a bubble in the flask, so that one must be excluded. Judging by the R2 value, the fit with 2 outliers excluded looks best. However, when the results of 2011-08 are considered, it seems like the case with 1 bottle excluded is closer to 2011-08, which had more bottles and a larger range of DO values, so is likely more reliable. Moreover, the 2 most severe outliers for this cruise are the 2 highest DO values and only the one with a bubble looked much worse than nearby bottles. A fit based on so little data is not justified when there is more data from a cruise in the same region and only a few days later.

New values of Soc and Voffset were entered in the configuration file (based on the results of 2011-08) and the corrected file was saved as 2011-28-ctd.con.

(See file 2011-28-do-test-xls for details.)

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. Those files were converted to IOS header format. They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files and no significant outliers were found.

A header check was run and shows that there are no off-scale fluorescence values in the BOT files.

The addsamp.csv file prepared in the DO calibration step was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files.

The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-28-bot-hdr.txt.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-28nuts.xls which included a report on precisions. The file was saved as 2011-28-nuts.csv. Headers were changed to standard format. The data were then reordered on event numbers and then sample numbers. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-28oxy.xls. The spreadsheet page with the final data was simplified by removing a few unnecessary columns or rows and the file was then saved as 2011-28-oxy.csv. That file was converted into individual OXY files.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial #68572. This file was saved as 2011-28-sal.xls. There were no comments or quality flags. There were duplicates so a precision study was done on a separate worksheet and a QF page prepared with averaged values for the duplicates and event numbers added. The pooled standard deviation was 0.01 when 1 outlier pair (determined using Chauvenet’s criterion) was excluded. An internal standard was run at the end of the analysis and it has a value within 0.001 of the average value found for this batch in 9 uses over 3 days.

The sheet was then saved as 2011-28-sal.csv and simplified. The station name and sample numbers were added. The file was converted to individual SAL files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only. The files were sorted on bottle number.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. The files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The scatter is very high and almost identical for both primary and secondary channels. Even below 100db there is a lot of scatter and both CTD salinity channels are lower than bottles by an average of ~0.011. Even below 200m the CTD was low by 0.004, 0.014, 0.017 and 0.018. The standard deviations in the CTD salinity are not high for the bottles in the fit. There was an internal standard run at the end of the analysis and it falls within 0.001 of the average of 9 measurements taken over 3 days. So there is no obvious problem with the Autosal. There does not appear to be a linearity problem in the Autosal. The salinity analysis was run almost 6 months after the cruise. The delay seems a more likely cause of the problem than anything else. This might lead to evaporation raising the bottle values and thus making the CTD look low. This would also explain the scatter. It is also possible there were problems with sampling protocols – if the bottles had not been rinsed well or caps were not tight, that would explain such results, but the analyst did not notice any such problem.

If only bottles between 100m and 250db are included the CTD salinity is low by 0.007, or by 0.011 if one outlier is excluded. These are still much higher differences than expected.

The duplicates that were identified as outliers using Chauvenet’s criterion are from the bottom of one of the deep casts. Neither looks close enough to the CTD to justify using a single value instead of the average.

Flags (“3”) were assigned to all the salinity samples since the long wait appears to have caused a problem. The values are likely good to 0.01, but no better.

For details see 2011-28-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. There is a lot of scatter. When the most extreme outlier is excluded the CTD DO is lower than the bottles by an average of 0.027mL/L and a standard deviation of 0.094. When 3 other outliers are removed the CTD is found to be high by an average of 0.003mL/L with a standard deviation of 0.04mL/L. For 2011-08 when 5 outliers were excluded, the average difference between bottles and SBE DO was very small (0.0002mL/L), as expected, and the standard deviation was 0.043mL/L. For 2011-01, using a different sensor, the standard deviation was <0.03mL/L. This comparison for 2011-28 contains a lot of scatter and little data, so does not justify further recalibration.

The most extreme outlier was examined, and the CTD standard deviation does not look out of line. While this is a near-surface sample and the match is often poor there, for this cast all variables show the top 20m to be very well mixed. There was a bubble in the sample, so a “4” flag is appropriate; that was added after consultation with the DO specialist. The three other outliers do not stand out in the original comparison of section 3; for one the standard deviation in the SBE DO data is high, but not for the others. The problem could be due to noise in the CTD DO due to noise in the temperature channel, or because of small-scale variations so that the 1.5m vertical offset between bottles and CTD become significant. No flags will be applied to the bottle data since they look ok compared to SBE DO voltage. (See 2011-28-dox-comp1*.xls.)

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### DO sensor hysteresis study

This sensor has not been fine-tuned for hysteresis parameters, but there is no deep sampling from this cruise so the tests cannot be done. This does not matter for this cruise since sampling was fairly shallow.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-28-ctd.con.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are in good agreement though there are small spikes in both. Upcasts are noisier than downcasts but in reasonable agreement. There is a section of noisy primary data in the upcast for event #15.

The altimetry looks useful near the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run to advance the DO Voltage by 4.4s relative to the pressure, based on the results of 2011-08. Plots of the two deep casts were examined to see if the results were satisfactory and they were.

##### CELLTM

During 2011-08 tests were run comparing a variety of settings for CELLTM. The goal is to make upcasts look closer to downcasts on a T-S surface. There was little temperature variation and for most of the profiles temperature is increasing with depth, so the usual settings often make things worse. No useful setting was found. For this cruise the 2 deep casts were checked and again, no setting was found that improved the data. CELLTM was not run.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The results of the following cruise, 2011-08, are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-28-0012 | 300 | +0.0002 | <+0.0001 | +0.0005 |  |
| 2011-28-0015 | 300 | +0.0002 | <+0.0001 | +0.0007 |  |
| 2011-08-0004 | 280 | +0.0002 | <+0.0001 | +0.0005 | Noisy, High |
| 2011-08-0025 | 380 | +0.0003 | <+0.0001 | +0.0003 | V. noisy, High |
| 2011-08-0028 | 200 | +0.0005 | <+0.0001 | +0.0006 | X Noisy, High |

The differences are very small.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace any pad values in pressure using linear interpolation based on scan number.

##### Checking Headers

The header check was run.  There were no off-scale fluorescence values.

Surface check was run and shows an average surface pressure for the cruise was 1.6 db. This is reasonable for the Vector and especially for this project.

The cross-reference check was compared with the log book. The only discrepancy was the data for the final CTD cast. The header looks right and this could not have occurred on the 5th of April given another CTD was started at the same time. It is obvious that the new date (6th) had just not been entered. No change was made to the header.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. For the first 10 casts the readings are obviously wrong, but since the full files have pressures <3db, no CTD files will be produced, so there is no need to fix these. For casts #12 and 15 plots were checked and the header readings look ok.

Water Depth readings were also examined and all were close to the log entries

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. As usual with surface-only sampling, the altimeter header readings are not accurate, so they were removed from the MRGCLN2 files for casts #2 to 11. For cast #3 either the header reading is wrong, or the CTD drifted into deeper water during the cast; either is possible, but the header entry was removed since it is not clearly correct. The header entries were fine for casts #12 and 15 which did get close to the bottom.

##### 13. Shift

Because there were only 2 casts available and they both had many bottle stops, it is hard to make judgments about shift settings, so the results of 2011-08 were used.

Fluorescence (From 2011-08 report)

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The shift applied is almost always +24 records. For this cruise the upcast temperature data were quite noisy, making a judgment difficult, but +24 looked best for more casts than any other setting.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity (From 2011-08 report)

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of -0.6s to the secondary conductivity.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen (From 2011-08 report)

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

These sensors were recalibrated in Jan. 2011 and were used for 2011-08 which occurred shortly after this cruise. The comparison was very noisy and showed the primary and secondary salinity to be low by 0.0041 and 0.0034. The salinity analysis was done about 4 months after the cruise.

2. Dissolved Oxygen

This sensor was used for 2011-08 which was after this cruise. The initial calibration against bottles worked well, with no need for a second calibration of downcast data.

3. Pressure

The sensor was recalibrated in Dec. 2009; it was used for 2011-08 when no further offset was found necessary.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. All data fell within the local climatology

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

No post-cruise calibrations were available.

##### DETAILED EDITING

As for 20111-08, the choice of primary or secondary channels was not obvious – neither is very noisy. For 2011-08 the secondary was closer to the bottles, but for these data the differences are almost identical. There is more noise in the secondary channels, so the primary T and S sensors were chosen for archiving.

CTDEDIT was used to remove surface records for which the pump was turned off, to clean small spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and a few records corrupted by shed wakes..

##### Initial Recalibration

No recalibration will be applied. Pressure looks ok. The comparisons of dissolved oxygen are noisy and the salinity comparison is not trusted. The same decisions were made for 2011-08 except that for that cruise the pressure offset had been entered incorrectly, so it needed to be recalibrated. If it is later learned that the salinity is really low, then this decision should be revisited.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but it is sometimes found appropriate to do a further correction for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure there is a lot of scatter. The SBE DO values appear to be low by an average of 0.02mL/L but the standard deviation is 0.085mL/L, even when 2 outliers are excluded. When the other 2 outliers identified in the first run of COMPARE are also excluded there is little difference. A plot against DO shows the noise is chiefly in the high-gradient regions. The two casts look quite different, with cast #12 having DO values too low and #15 having them too high. This is likely because there is a lot of structure in the mid-depths of cast #15 that is not seen in #12. This effectively allows the sensor to “catch up”. So applying a correction based on one of these casts would cause trouble for the other and even if a separate one was used for each cast, what suits part of a profile does not suit other parts. This is a correction that is only applied if there seems to be a simple offset against either pressure or DO.

No further recalibration will be applied. None were applied to 2011-08 either.

Error estimates were based on the plot of DO against pressure..

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

There was no CHL sampling, so no special files were prepared for Dr. Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. The casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following header comments:

Data Processing Notes:

----------------------

Fluorescence and Transmissivity data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.2 mL/L from 0 to 100db

±0.1 mL/L below 100db

For details on the processing see processing report: 2011-28-proc.doc.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

T-S plots were examined and no problems were found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 103% to 107%.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was used to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second titrated DO channel (with umol/kg units) could not be added because there was no draw temperature available.

REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data..

Header Check was run and showed that nutrients from cast #19 had been entered in the cast #9 file. That was fixed, and it was found that the salinity comments had not been entered properly, so the SAL files were reconverted, and then all steps starting with the merge steps were rerun.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no further discrepancies were found.

Standards check was run on all files and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 06Jan11 | Factory |  |  |
| Conductivity | 3579 | 08Jan11 | Factory |  |  |
| Secondary Temp. | 5073 | 06Jan11 | Factory |  |  |
| Secondary Cond. | 3581 | 08Jan11 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug10 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 24Dec2010 | Factory |  |  |
| Fluorometer | 2745 |  | IOS |  |  |
| Pressure Sensor | 0941 | 7Dec09 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |