## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2025 | Updated channel names & formats in TOB files. GG |
| 15 Apr 2024 | Fixed Format for Temperature:Draw. SH |

## PROCESSING NOTES

Cruise: 2011-44

Agency: OSD

Location: WCVI/BC Inlets

Project: Aquaculture, Discovery Passage, SARA, La Perouse, Effingham Inlet

Party Chief: Juhasz T.

Platform: John P. Tully

Date: May 14, 2011 – May 24 2011

Processed by: Germaine Gatien

Date of Processing: 7 November 2011 – 12 December 2011

Number of original HEX files: 37	Number of original TSG HEX files: 5

Number of CTD files:  35 (1 file pumps off, 1 file surface data only)

Number of CHE files:  28

Number of TSG files:  3

# INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs (#0585 and #0550) was used for this cruise. Each was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR, a Wet Labs Eco-AFL/FL Fluorometer (#2216) and an altimeter (no serial # available). Sensor #0585 was mounted with SBE 43 DO sensor #47 and CTD #0550 was mounted with 2 SBE 43 DO sensors (#1438 and #1176).

The deck unit was an SBE model 11, serial #0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS is believed to have been a Guildline model 8400B Autosal, serial # 68572, but the analysis sheet is missing.

A thermosalinograph (SeaBird 21 S/N 3363) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter. The sample interval was set to 30s for the first file, and 10s thereafter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but it was missing altimeter and TSG information. The altimeter serial number was also missing from the configuration file. There was also no listing for a 2nd dissolved oxygen sensor on the CTD #550, though a second one was present. The configuration file changed many times during the cruise, but in only 1 case was there a note in the log to explain why. It appears that only one change of equipment actually occurred. It saves a lot of processing time if notes are made about whether such alterations were made because errors were found in the configuration files or because of equipment changes.

The rosette logs were generally in good order though no sampling was indicated for the first cast; that was the only cast with nutrient sampling. That could lead to trouble since the processor would not have been aware of a problem if the nutrient data had been misplaced.

One of the apparent changes of equipment was actually a case of a file being saved with the wrong file number so that an earlier cast file was overwritten. The file for event #2 was lost. There are titrated dissolved oxygen data from that cast and those may be found in file QF2011-44oxy.xls.

The Dissolved Oxygen sensor used on CTD #0585 appears to have malfunctioned during downcasts so that channel was removed for casts #1 to 9. The upcast data are in reasonable agreement with titrated samples, so the SBE DO channel was left in the CHE files.

There is no explanation of why 2 dissolved oxygen sensors were mounted on CTD #0550. The data from sensor #1176 will be archived. Comparisons showed no large differences, but the response time looks slightly lower for sensor #1176 than for #1438.

Errors in format in the entries of station name and water depth prevented that information from appearing in the IOS headers, so the hex files had to be edited. The colon is critical in entries like STATION: B1 and Depth (m): 125. Entries that apply to the whole cruise are easy to fix, but those that vary from cast to cast are not. This happened on at least 3 successive cruises.

The first two Niskin bottles from cast #6 were fired while CTD pumps were turned off, so there are no data available in the CHE file from the pumped channels: temperature, salinity and dissolved oxygen.

There were only 3 salinity bottle samples and those differed from the two CTD salinity channels by from 0.001 to 0.006. No recalibration was applied to the CTD data as they were trusted more than the bottles, which were analyzed more than 3.5 months after collection.

The WetLabs ECO fluorometer was used for the first time at IOS. The sensor appears to have a poor time response so the traces are much smoother and consequently no filtering is needed; however, the alignment needs a larger correction than for the Sea Point fluorometers.

The first TSG file lacked latitude, longitude, intake temperature, fluorometer and flow rate channels. The second TSG file lacks the intake temperature channel. The third TSG file had all expected channels but only a few data points. Files #1 and #3 were not processed.

The TSG intake thermistor malfunctioned during this cruise producing values that are too high. Because it is connected serially, the calibration parameters can only be checked by connecting the instrument to a computer and examining the parameters through SEATERM. However, there was no problem with the intake temperature data on a subsequent cruise, so it is likely that the parameters were entered correctly. The lab temperature was adjusted by subtracting 0.2C° for use as a proxy for the intake temperature.

The fluorometer in use on the TSG has not been recalibrated in over 10 years. It reads higher than CTD fluorescence and extracted chlorophyll. Recalibration is strongly recommended.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated shortly before this cruise.

The calibration constants were checked for all instruments. There are a number of odd things about these files:

- There were 8 changes of configuration files, using 5 different combinations.

Events 1, 3-5, &10: CTD #0585, Fluorometer scale 1, dark value 0.

Event 2: CTD #0550, Fluorometer scale 1, dark value 0.

Events 6-9:  CTD #0585, Fluorometer scale 1, dark value 0, transmissometer in different voltage spot than for cast 1. Times are correct in HDR file.

Events 11-21 & 23-24: CTD #0550, Fluorometer scale 1, dark value 0.0440.

Events 22 &25-37: CTD #0550, Fluorometer scale 25, dark value 0.0270

- The settings for the fluorometer are also puzzling. A scale of 1 does not make sense for a near-shore May cruise. However, this was a new fluorometer and there may have been confusion about how to set it up. The factory setting for the 25 scale setting shows a dark count of 0.027V. The dark setting of 0.044V matches the 13 scale, not 1. The CTD technician believes it was set to scale 25 and in the absence of any information, the factory reading will be assumed correct for the dark value. The 25/0.027V values will be used for conversion of all files and the results studied to see if they look reasonable.

- The moving of the transmissometer at cast #6 is possibly correct but seems odd, given that it went back to the original for cast #10. There is a note that the transmissivity was not working for cast #6, so most likely the change in the configuration was incorrect. So it will be assumed it was in Voltage0 spot for all CTD #0585 casts. However, it does seem to have been in Voltage1 for CTD #0550. Once again, conversion should clarify this.

All parameters were checked and the only gaps were missing calibration dates for fluorometers. The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset had been changed slightly to fit the method recommended by Sea-Bird. So the slope and offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

After these adjustments there are 2 configuration files: 2011-44-ctd585.con and 2011-44-ctd550.con.

Examining the results of conversion using those configurations provides an answer to the mysterious CTD switch before and after event #2. File #2 is not really event #2 – it looks like event #10. There is a file #10 but it contains only garbage, so they must have changed the CTD and run it again. The log note is not clear about when the swap took place. The file must have been inadvertently saved as file 2011-44-0002.hex, thus replacing the real event #2 data. This is sad, but does explain the configuration change.

The converted files lacked “:” after station and “(m):” after depth entries in the headers. This will prevent these entries being recognized when the files are converted to IOS header format, so the hex files were edited to add those. This happened on at least 3 successive cruises.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records. Because the file for cast #2 was overwritten, sample #s 7 to 15 were not used in this file. Samples #172 and #173 were not used.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-44-DO-cal.csv. Two separate spreadsheets were prepared for the 2 different DO sensors following the following method.

The titrated DO values were added to that file and lines removed for which there was no DO sampling. A calculation was made of  Ф  using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

In order to efficiently identify outliers, the values of M and B using all data were used to calculate the differences from the fit as follows.
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. First differences >0.1 were excluded and M & B from that were substituted in the above equation. Then more and more data were removed until it looks like a good fit and a believable result with not too severe an effect on the range of DO Voltage in the fit.

For SBE DO sensor #0047 there are only 76 bottles and the first fit looks poor. Removing just 2 outliers produces a much better result; those outliers are from a cast where there was a problem with CTD pumps. Better data may be available, but for now they will not be included. When more outliers are excluded based on the difference from the fit the range quickly shrinks and the results vary greatly and look hard to believe. We would expect the slope to increase with time, not get smaller. So the fit using all but 3 bottles looks best, though it is not terribly good. The fit excluding 14 bottles is much tighter, but hard to believe. Tests will be done later to see if using the original factory fit may be better for this sensor. This was the first use since recalibration at the factory.

| Summary of Soc Voffset including the original values in the factory calibration # 0047 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.5298 | -0.3273 | 0.5298 | -0.6177 |  |
| 76 | all data | 0.3877 | 0.0265 | 0.3877 | 0.0684 | 0.7152 |
| 74 | excl. outliers diff>0.1 | 0.4970 | -0.2284 | 0.4970 | -0.4596 | 0.9230 |
| 71 | excl. outliers diff>0.05 | 0.5160 | -0.2697 | 0.5160 | -0.5227 | 0.9539 |
| 67 | excl. outliers diff>0.03 | 0.4923 | -0.2152 | 0.4923 | -0.4371 | 0.9302 |
| 62 | excl. outliers diff>0.02 | 0.4673 | -0.1590 | 0.4673 | -0.3403 | 0.9577 |
| 58 | excl. outliers diff>0.015 | 0.4696 | -0.1650 | 0.4696 | -0.3514 | 0.9649 |
| 53 | excl. outliers diff>0.01 | 0.4870 | -0.2047 | 0.4870 | -0.4203 | 0.9646 |

For SBE DO sensors #1176 and 1438 there are 172 bottles available.

SBE DO #1176

When 102 out of 172 points are removed the results is quite similar to when only 5 were excluded and both those choices have similar offsets to the factory setting. The slope does increase a little, as expected. Plots of pressure and DO versus differences from fit show that the near-surface, higher DO values have the largest differences from the fit, as expected. The fit using 102 bottles does have a smaller range.

There is a possibility that the hysteresis parameters are not chosen ideally, which may lead to a few outliers at depth, but most outliers come from near the surface. Fine-tuning the hysteresis factors is not likely to change the calibration significantly.

| Summary of Soc Voffset including the original values in the factory calibration - #1176 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4453 | -0.2321 | 0.4453 | -0.5212 |  |
| 172 | all data | 0.4494 | -0.2206 | 0.4494 | -0.4909 | 0.9750 |
| 167 | excl. outliers diff>0.1 | 0.4566 | -0.2310 | 0.4566 | -0.5059 | 0.9946 |
| 160 | excl. outliers diff>0.07 | 0.4611 | -0.2370 | 0.4611 | -0.5140 | 0.9974 |
| 153 | excl. outliers diff>0.05 | 0.4631 | -0.2399 | 0.4631 | -0.5180 | 0.9987 |
| 147 | excl. outliers diff>0.03 | 0.4618 | -0.2374 | 0.4618 | -0.5141 | 0.9992 |
| 135 | excl. outliers diff>0.02 | 0.4605 | -0.2355 | 0.4605 | -0.5114 | 0.9995 |
| 124 | excl. outliers diff>0.015 | 0.4604 | -0.2352 | 0.4604 | -0.5109 | 0.9997 |
| 102 | excl. outliers diff>0.01 | 0.4588 | -0.2328 | 0.4588 | -0.5074 | 0.9998 |
| 94 | excl. outliers diff>0.008 | 0.4584 | -0.2328 | 0.4584 | -0.5079 | 0.9998 |
| 38 | excl. outliers diff>0.003 | 0.4563 | -0.2307 | 0.4563 | -0.5056 | 0.9999 |
|  | 2011-16 best result | 0.4661 | -0.2471 | 0.4661 | -0.5301 | 0.9997 |

SBEDO #1438:

Removing more and more data bring the results closer to the factory setting.

| Summary of Soc Voffset including the original values in the factory calibration #1438 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4453 | -0.2284 | 0.4538 | -0.5130 |  |
| 172 | all data | 0.4564 | -0.2179 | 0.4564 | -0.4774 | 0.9767 |
| 165 | excl. outliers diff>0.1 | 0.4672 | -0.2347 | 0.4672 | -0.5024 | 0.9965 |
| 154 | excl. outliers diff>0.05 | 0.4681 | -0.2342 | 0.4681 | -0.5003 | 0.9987 |
| 146 | excl. outliers diff>0.03 | 0.4692 | -0.2358 | 0.4692 | -0.5026 | 0.9994 |
| 143 | excl. outliers diff>0.02 | 0.4691 | -0.2357 | 0.4691 | -0.5025 | 0.9995 |
| 131 | excl. outliers diff>0.015 | 0.4687 | -0.2355 | 0.4687 | -0.5025 | 0.9997 |
| 114 | excl. outliers diff>0.01 | 0.4677 | -0.2337 | 0.4677 | -0.4997 | 0.9998 |

For all the DO sensors, these were the first deployments since factory calibration, so COMPARE was first run with the factory setting but the results in COMPARE were very poor. So the values highlighted above were chosen and the results were better, as will be discussed later.

##### Hysteresis Study

Sensor #0047 has been recalibrated since the last hysteresis checks were done, but there is no deep sampling for this cruise, so no tests can be done to fine-tune H1, H3 and E.

Sensor #1438 and #1176 have been recalibrated since the last hysteresis checks were done. There was some deep sampling, but only 1 cast with bottles, and only 3 of those come from below 1500db, so there is not much evidence. As part of the tests a check was made that the deepest bottles did not stand out as outliers, and they did not. A run of COMPARE using only cast #17 did not suggest there was a problem. There is not enough information for the fine-tuning of H1, H3 and E.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters.

The ROS files were converted to IOS header format. They were put through CLEAN to create BOT files.

No bottle file will be prepared for cast #24 since bottles were only fired to get bulk water.

Temperature and salinity were plotted for all BOT files and problems were found in 3 casts. For casts #17 and 37 small spikes in primary salinity were cleaned using CTDEDIT and the output files were then copied to BOT. The other problem is in cast #6. As noted in the log book the cast was restarted because the pumps were not on. Two bottles had already been fired. For those 2 bottles we can use the unpumped data, namely pressure, transmissivity and fluorescence and the altimetry should be fine, but the pumped channels must be removed. A text editor was used to remove temperature, conductivity, salinity and dissolved oxygen channels from the first two bottles, and that file was then copied to *.BOT.

The addsamp.csv file prepared in the DO calibration step was ordered on sample #s and converted to CST files to be used as a framework for the bottle files. It was also used to create SAM files with bottle #s and bottle positions. The SAM files were then bin-averaged.

The analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-44-bot-hdr.txt.

Dates of creation were added to the names of spreadsheets from analysts.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-44nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and the file was saved as 2011-44-nuts.csv. The file was converted to individual NUT files (actually only 1).

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-44oxy.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2011-44-oxy.csv. The duplicates for sample #112 were averaged, but should be checked later to see if either is a clear outlier. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was done at IOS but the samples were run with those of cruise 2011-16 and were not identified as being from 2011-44. The salinity analysis sheets were not available, but the sample numbers and station names make it clear that 3 of the samples are from this cruise. A file was created with those 3 samples and named 2011-44-sal*.xls (* indicates the date of file creation). There were no comments or flags. The file was saved as 2011-44-sal.csv; and for that version a flag channel was added. Some columns were removed and a column was added for the event number – the latter information was found in the log book. The station name and sample numbers were combined in a single column; those had to be separated. The data were ordered on event number and sample number. The file was converted into individual SAL files. There were no duplicates.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but in case the SAMAVG files are ordered on bottle number, the MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel. There were only 3 samples and the samples were analyzed after about 3.7 months. The differences get larger with depth, but with so little data this could be random. The primary CTD salinity is lower than bottles by 0.0036, 0.0019 and 0.0061 (average 0.0039), while the secondary are low by 0.0025, 0.0010 and 0.0042 (average 0.0026). It seems unlikely that these sensors are reading that low right after being serviced and when the CTD readings are within 0.001 of each other. There is a possibility that waiting for analysis for almost 4 months might be the problem. Other recent cruises using different sensors have produced comparisons that were also doubted because of long waits for analysis. For details see 2011-44-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

For all 3 sensors there is a lot of scatter, but overall the CTD looks close to the bottles. Further fine-tuning of Soc and Voffset might improve the results slightly, but the original fits were difficult to judge and the nature of the waters sampled during this cruise probably mean we can never get a fit that looks really satisfying.

Outliers were examined.

For CTD #0585 most outliers were from near the surface where variability was high, but 6 were examined that came from below 20db.

- Cast #1 – Both CTD temperature and DO varying notably during stops at 50 and 70db. Most stops showed high variability. No evidence of trouble with bottles.

- Cast #5 – the CTD DO never equilibrated during the stop at the bottom. Bottle probably ok.

- Cast #6 had 2 outliers – the CTD pumps were not operating, so the bottles are likely ok.

- Cast #9 – the bottom bottle was already flagged due to a bad curve. A note was added about the COMPARE results since they support the flag decision.

For CTD #0550, there were 2 sensors mounted together. There are larger differences than expected, but on analysis most of the large outliers are from the top 15m of Effingham Inlet casts. This is not surprising as the gradients there are hard for SBE DO sensors to resolve.

Even when the Effingham casts are excluded the results are difficult to compare because the choice of outliers varies from one to the other. In neither case is it easy to say which points are outliers and which are not. There was a lot of variability in many of these casts and sensor #1176 had a slightly higher average standard deviation during the 10s around firing time. Is that good or bad? Given the high variability at many of these casts, having a shorter response time would lead to higher standard deviation. In quiet waters, the opposite would be true as the shorter response time would lead to better equilibration. The average difference between bottles and CTD is -0.023 for sensor #1176 and -0.018 for sensor #1438 including all data.

A study was made of how the 2 sensors behaved during a stop at 1130db during cast #13. This stop looks like an outlier for sensor 1438, but for sensor 1176 the CTD is closer to the bottles. A plot of DO voltage versus scan number shows fine-scale noise in the 1176 sensor but not in 1438. The temperature was varying through this period, so this may just be an indication that the 1176 sensor is responding quickly. Which is better? That is unclear.

Five outliers were identified:

- Cast #13 – at 760db. Temperature does not equilibrate during the stop, so DO probably has not either. No flag will be attached to the bottle.

- Cast #15 – 758db. This only stands out in the fit against pressure. The temperature is not as steady as usual for the depth, so this is not likely a bottle problem, but a poor match due to variability. No flag.

- Cast #15 – 503db. The bottle had been flagged by the analyst because replicates were outliers in the precision study. If the lower value were selected instead of the average, this is not an outlier. The OXY file was adjusted and the flag removed; a comment was entered to explain why.

- Cast #17 – 2083db. Temperature and DO seem to have equilibrated. Bottle may be a little high, or there could be some hysteresis in the SBE data. No flag is justified.

- Cast #17 – 749db. Temperature never equilibrated during this stop, so SBE DO not trusted.

(See 2011-44-dox-comp585.xls and 2011-44-dox-comp550.xls.).

The merge process was repeated to capture the new flags and comments.

Once again data were exported to spreadsheet 2011-44-bottles.xls and compared to the rosette sheets and all expected data are present.

The MRG files were then put through CLEAN to remove the SeaBird headers and comments from the secondary file and to adjust the record numbers and maxima and minima in the headers.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-44-ctd585-recal.con or 2011-44-ctd550-recal.con, as appropriate.

A few casts were examined and all expected channels are present.

For both CTDs the two temperature and conductivity channels are close during downcasts, but there are larger differences during upcasts which look like alignment issues – the CTD is probably swinging. There are some spikes with those in the primary channels being larger than in the secondary. The traces look odd with sudden jumps but the sensor pairs are in agreement.

As noted before with the ECO fluorometer there is a vertical offset between downcast and upcast that is larger than seen in the temperature traces. This sensor is unpumped so this is not an issue of transit time and is likely due to slow response.

The transmissivity looks normal with no hysteresis.

The dissolved oxygen sensor data looks normal.

The altimetry looks useful.

The pressure is odd in the CTD #0550 files, with large negative values in many casts.

File 2011-44-0010.cnv contains no useful data.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

WILDEDIT has reduced the negative pressure readings but not entirely removed them. They all seem to be early in the files. The pressure sensor configuration was rechecked and no error was found.

At this point file 2011-44-0010.cnv was deleted and file 2011-44-0002.cnv was renamed 2011-44-0010.cnv.

##### ALIGN DO

There are 3 different DO sensors to be aligned and the upcasts and downcasts are often so different that tests are not very clear. Even when the differences are not huge, the upcasts are very noisy making comparisons almost impossible.

For sensor #0047 a value of +4s looks reasonable.

For sensor #1176 a value of +4.4s looks best, though +4.0s was used for that sensor during 2011-16.

For sensor #1438 a value of +4.2s seems best and makes it line up well with #1176.

ALIGNCTD was used to advance the DO Voltage relative to the pressure using the above values.

##### CELLTM

For CTD #0585 the upcasts are very noisy so the usual tests to determine the best parameters for use in CELLTM are impossible. There is very little variation in temperature, so skipping CELLTM is reasonable. CELLTM was not applied to casts #1-9.

For CTD #0550 the tests are also very difficult thought there are a few casts with larger temperature range. The same equipment was used for 2011-16 when the choice of (α = 0.02, β=7) for the primary and (α = 0.03, β=9) looked best overall. Tests on a few casts from this cruise were hard to interpret due to the noisy upcast data, but overall using the 2011-16 settings do seem to improve some of the deeper casts.

CELLTM was run on casts #10–37 only, using parameters (α = 0.02, β=7) for the primary and (α = 0.03, β=9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

There were 2 DO sensors on SBE #0550 but it was impossible to derive the difference between these channels at this point. It is possible to choose the difference as a channel during the original conversion. But the alignments are different so this proved to be uninformative.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen. Data from 2011-16 are included for comparison and to look for time dependence:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| CTD #0585: 2011-44 |  |  |  |  |  |
| 2011-44-0007 | 240 | ~0 XN | -0.00002XN | -0.0004XN | Steady, high |
| 2011-44-0009 | 300 | +0.0001 N | -0.00003 N | -0.0005 N | F. steady, moderate |
| CTD #0550: 2011-44 |  |  |  |  |  |
| 2011-44-0013 | 800 1000 | -0.0006 -0.0003 | +0.00009 +0.00007 | +0.0013 +0.0013 | Extremely noisy Very noisy |
| 2011-44-0017 | 800 1000 1950 | -0.0005 -0.0005 -0.0008 | +0.00007 +0.00007 +0.00009 | +0.0013 +0.0014 +0.0020 | Extremely noisy Very noisy Very noisy |
| CTD #0550: 2011-16 |  |  |  |  |  |
| 2011-16-0034 | 800 1000 1950 | -0.0003 -0.0004 -0.0007 | +0.00004 +0.00007 +0.00005 | +0.0008 +0.0011 +0.0015 | Extremely noisy |
| 2011-16-0111 | 800 1000 1950 | -0.0001 -0.0003 -0.0007 | +0.00010 +0.00009 +0.00010 | +0.0014 +0.0015 +0.0019 | Noisy |

For CTD #0585 the differences are extremely small.

For CTD #0550 the differences are also small and are in line with the observations during 2011-16 when the same sensors were used.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. There are large negative pressures as noted earlier. The fluorescence maximum looked like it could be off-scale, but there is only a single spike at the surface of 1 cast.

The surface check was run and the average surface value was 0.76db for CTD #0585 and 0.51db for CTD #0550. These are both low, but the surface salinity is also very low, so it is likely that the CTD was started very close to the surface particularly in the inlets.

CLEAN was run on casts #1-9 to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

For casts #10-37 there was a problem with large negative pressures and a few unbelievable positive values nearby in many casts. These problems are at the surface in records that will likely be removed in the DELETE process, but those values might cause some problems in how DELETE works. The same problem was noted with this equipment during 2011-16. As was done then, CLEAN was run first to replace pressures <-0.1db with pad values and then a second time to replace pad values in the pressure channel with interpolated values based on record number. Care will be needed in the editing phase to ensure that this process worked well.

The cross-reference check was compared with the log book and no errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. The values look reasonable; a selection of casts was checked and no problems were noted.

The Water Depth headers were also examined. There were deviations from the log book entries for 2 cases and 1 header entry was missing.

- For cast #10 the log entry looks to be correct and was entered into the file header.

- For cast #13 the log and file headers differ, but both are lower than the depth of maximum sampling, so the entry was removed since we have no reliable value.

- For cast #15 the log differs from the file header, but the log entry is much less than the maximum sampling pressure and the file header looks reasonable.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. A selection of casts was checked and the only problem was the water depth of cast #13 which was removed.

##### 13. Shift

Fluorescence

The ECO fluorometer is not pumped, so alignment is not needed to correct for transit time. However, there is clearly an excess vertical offset when traces are compared to temperature. This is assumed to be due to poor response time.

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset.

For the ECO sensor the estimates give highly variable results, from about 1.5s to 3.5s. This is partly because of the “steppy” nature of the traces and the fact that peaks are not resolved well, and there were few of the casts without stops for bottles that provide the best evidence. During 2011-16 when the same fluorometer was used, a setting of +48 records was found to be the best SHIFT setting, though there was not a lot of confidence in the choice.

SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records. (Output: SHFFL)

Conductivity

Tests were run on the two conductivity channels for each CTD using a variety of shifts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall for CTD #0585 when a shift of -0.5s was applied to the primary and a shift of -0.7s to the secondary conductivity. For CTD #0550 the best results were with a shift of -0.7 applied to the primary and a shift of +0.5s to the secondary conductivity. 
SHIFT was run twice on all casts using those settings.

(When the #0550 sensors were used during 2011-16 the settings were -0.5s and +0.5s for the primary and secondary sensors respectively.)

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. Because there is so much variability in the temperature data, while DO traces are smoother, it is very hard to make a judgment on the alignment, but it looks reasonable. The two DO sensors are close, though again different response times can complicate that judgment. No further alignment will be applied.

##### 14. DELETE

All casts except #1 had an initial drop before the full cast and in the case of Event #6 there were 2 full casts but no pumps on for the first. Plots were made to determine how many initial records needed to be removed from each cast. CLIP was used to remove those initial records.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concern upcast sections of cast #13.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

CTD #0585: The sensors were recalibrated shortly before this cruise, and no other processed cruise since then has used them.

CTD #0550: The sensors were recalibrated shortly before this cruise and were used during 2011-16 after this cruise. No recalibration was applied to them. The bottle comparison suggested both salinity channels were low but the result was not trusted.

2. Dissolved Oxygen

Sensor #0047: The sensor was recalibrated shortly before this cruise, and it has not been used in any other processed cruise since then.

Sensor #1176: The sensor was recalibrated shortly before this cruise, and was used during 2011-16 shortly after this cruise.

Sensor #1438: The sensor was recalibrated shortly before this cruise, and it has not been used in any other processed cruise since then.

3. Pressure

CTD #0585: The sensor was recalibrated shortly before this cruise, and no other processed cruise since then has used it.

CTD #0550: The sensor was recalibrated shortly before this cruise, and was used during 2011-16 shortly after this cruise.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Temperatures were a little above the historic maxima near the bottom of two casts off Brooks Peninsula. Similar excursions were noted for some deep casts off the west coast of Vancouver Island during 2011-16. Salinity was low only near the surface. These excursions look real, not indicative of instrument malfunction.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

At this stage a decision has to be made about whether to edit primary or secondary channels. The two channels are close for both CTDs, but there are a few more spikes in the primary than the secondary, and the secondary were used for 2011-16 when CTD #-550 was used in late May. So the secondary channels were selected for all casts.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. Some surface records were removed as they appear to be corrupted by either ship effects or too short a wait at the top.

All EDU files were copied to EDT.

In the course of editing it was realized that the pumps were never turned on for cast #24 so it was not edited and will not be processed further.

##### Initial Recalibration

The pressure looks ok. COMPARE shows no need to recalibrate dissolved oxygen for any of the sensors used. The salinity comparison for CTD #0550 is not trusted and is based on only 3 samples. There are no salinity samples to match CTD #0585. No recalibration will be applied to pressure, DO or salinity.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between downcast CTD DO data and upcast bottles.

For CTD #0585 the downcast SBE DO values look too high by an average of about 0.2mL/L, but near the surface and the bottom of the cast they look close. The range of values is low. The profiles look odd suggesting that the DO sensor was not behaving well, with the deepest casts showing a swing towards higher values between 75 and 200db that do not appear in the upcast data and are not reflected in the bottle data. Below 200db the values come back down. Unfortunately the change of DO sensors was made between the inlet casts and the offshore casts, so there is no easy way to determine if the oddities are geographical or related to the DO sensor. But it is hard to imagine a geographic explanation. It looks as if there were some problem with the input to the DO sensor that clears at about 200db. There are interesting variations in temperature and salinity between downcast and upcast, but they look quite possible in an area like this – steps in values that are not as distinct in the upcast, and those steps tend to be in the range where the DO looks unbelievable. Cast #5 is interesting in that the bottom two bottles suggest that even the upcast DO are too high and there is a sudden drop at about 75db of the upcast after which bottles and SBE DO agree quite well. The downcast DO data do not look sufficiently reliable to archive. The upcast data are not useful for a profile because there were so many stops for bottles that it is heavily corrupted. So the SBE DO channel should be removed from the downcast files, but left in the bottle files. (Note: Dario Stucchi examined the DO data in detail and agreed to this step.)

For CTD #0550 DO sensor #1176, the fit against pressure is quite flat when near-surface bottles are excluded. The fit against DO shows the SBE sensor is high by an average of +0.004mL/L when a few outliers are excluded. For sensor #1438 the fit against pressures is even flatter. The fit against DO shows the SBE sensor to be low by an average of 0.02mL/L. For both sensors recalibration is not justified, given the noise level; a small change in how outliers are selected would produce different results.

The two SBE DO sensors used on CTD #0550 are in good agreement – the traces for #1176 show a little more detail, so the response may be a little better.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Only the ECO fluorometer was used and it does not require filtering.

There was no extracted CHL sampling, so special files were not prepared for Dr. Peña.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

At this stage it was found that the fluorescence channel had many small negative values, mostly -0.036. This is assumed to be the dark value and will be added to all values.

CALIBRATE was run on all casts to add +0.036 (2011-44-recal.ccf.)

After this step the only remaining negative values were some -0.002 values in cast #17 – these were replaced with 0 values using a text editor.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on casts 1-9 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Transmissivity, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts 10-37 to remove the following channels:

Scan_Number, Temperature:Primary,  Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Voltage:SBE2, Oxygen:Dissolved:SBE2, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added for each of the sensors for casts #10-37.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

For CTD #0585

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

The transmissivity channel was removed from this cast because

there was no signal.

The SBE Dissolved Oxygen data for casts #1-9 were removed from

the CTD files because the profiles did not agree with upcast

bottles and had unusual shapes between 50db and 200db.

For details on the processing see processing report: 2011-44-proc.doc.

For CTD #0550

Data Processing Notes:

----------------------

Fluorescence and Transmissivity data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±2.0mL/L from  0 to 15db

±0.4mL/L from 15db to 275db

±0.1mL/L from 275db to 500db

±0.03mL/L below 500db

For details on the processing see processing report: 2011-24-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

Header check was run and showed that there were some negative values in the fluorescence channel. These are very small, but will be replaced with 0 values using CLEAN.

Profile plots and track plots look ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from ~90% to ~110% at the offshore sites and from 110% to 135% in Effingham Inlet. These values look reasonable.

##### 24. Final Bottle Files

CALIBRATE was run on the MRGCLN2 files to add 0.036 to the fluorescence channel.

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on casts #1-9 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Transmissivity, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with casts #10-37 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and then a second Oxygen:Dissolved channel (using Temperature:Draw). REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and the only problem found was the format for Temperature:Draw; this was done deliberately because only 1 decimal place is appropriate for these thermometers.

##### 25. Thermosalinograph Data

Data were provided in 5 hex files. There was no loop sampling.

a.) Checking calibrations

This is a new TSG system. There were 3 different configurations:
Cast #1: T and C only. – saved as 2011-44-TSG1.con

Cast #2: T, C, FL and Flow - saved as 2011-44-TSG2.con

Casts #3-5: T, C, FL, intake temperature and flow - saved as 2011-44-TSG3.con.

An error in the fluorescence scale factor was fixed in TSG 2 and TSG3.

The calibrations for the remote thermistor could not be confirmed. It can only be checked by connecting the instrument to the computer directly and using SEATERM.

b.) The files were converted to CNV files using the configuration files mentioned above. The names were non-standard, so were changed.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

For cast #1 there are no positions. The record starts on May 15 at 5:53 and lasts for about 14 hours, but conductivity data are bad except for the last 12 minutes. There was no intake temperature, no overlapping CTD casts and no loop sampling. So even if we could locate the measurements, there would only be temperature values. File #1 will not be processed further.

File #3 contains only 9 records, so it will not be processed further.

Only file #1 had the usual interval of 30s. Logging every 10s seems excessive for this sort of data.

Time-series plots were produced for files 2-5. Overall the records look good, with no problems noted except that the final records of file #5 will need to be removed because the flow was turned off. The salinity is quite noisy but that is reasonable for inlet records.

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db and exported to a spreadsheet which was saved as 2011-01-ctd-tsg-comp.xls. All the data came from ~4.5db.

There is no overlap between TSG file #2 and CTD casts.

There is overlap between TSG files 4 and 5 and CTD casts.

The intake temperature was logged, but contains no useful data in File #4.

The 2 TSG files that overlap CTD casts were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 34 matches. Those data were added to 2011-44-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. With one exception the differences in latitude were all <0.0002° and in longitude <0.0003° and the median differences were 0.00003° and 0.00000°. This shows both the times and positions are reliable for both systems. The one exception was cast #11 at LBP2 where the differences were 0.0038° of latitude and 0.0091° of longitude. The time and position taken from the CTD header may not be a good match if the ship was drifting between the time the position was recorded and the time acquisition began. Given that 34 of the positions compared well, it is assumed the TSG clock was working well.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected for TSG file #5 only. The average difference for all casts was 1.12Cº, with the lab temperature higher than the intake temperature. However, there is a clear distinction between the offshore casts and those from inlets. In the inlets the lab temperature is much higher than the CTD values and the differences are very noisy, but offshore the lab temperatures are slightly lower than the intake. The average and mean differences offshore are both 0.20Cº. The offshore temperatures are higher by about 1Cº, so we would expect a little less heating in the loop, but we would not expect cooling! So it is likely that either the lab or the intake thermistor was a little off. No information was available about the calibration of the intake thermistor, so calibrations could not be checked.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

The differences found among the CTD and TSG temperature channels were unusual and varied according to location. The following table summarizes the comparisons. There was no TSG1 sensor for the early casts

| Area | TSG0-CTD | TSG1-CTD | TSG0-TSG1 | TSGsal-CTDsal |
| --- | --- | --- | --- | --- |
| All | 0.55 | 0.39 |  | 0.164 |
| Early inlets | 0.36 |  |  | -0.635 |
| 3 near Brooks Pen. | 0.64 | 0.37 | 0.28 | -0.171 |
| Offshore | 0.25 | 0.41 | -0.2 | -0.007 |
| Effingham Inlet | 2.87 | 0.14 | 2.68 | -5.800 |

The median values differed significantly from the averages for all except the offshore area which gives a good measure of variability. For most cruises the intake temperature is fairly well within 0.1Cº of CTD temperature, but for these data it was generally much higher and only 1 cast was within 0.1Cº. The lab temperature is for the offshore casts it is higher by 0.25Cº which is about the amount of heating we might expect in the loop. This suggests that the intake thermistor did not perform well, while the lab temperature is probably ok. The mixed layer was very shallow through this cruise, so slight mismatches in intake depth and CTD depth can account for a lot of noise. The local gradients in Effingham Inlet are particularly challenging because of high surface gradients. The differences between intake temperature and CTD are very noisy in Effingham, with some too high and some too low. But the difference between lab temperatures and CTD temperatures are all on the high side.

The TSG salinity is lower than the CTD by a median of 0.16 with offshore salinity low by only 0.007 and Effingham Inlet values low by 5.8.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.2 to 5.9 with a median of 2.0 with most values between 1.2 and 3.2.  (See 2011-01-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons There were no loop samples.

- Loop Bottle - Rosette Comparisons

There were no cases of loop bottles being taken during a CTD cast.

- Calibration History

This is a new TSG system and this is believed to have been the first use of it. It was also used for 2011-16, 2011-26 and 2011-09 in May/June 2011.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was steady.

3. An analysis of heating in the loop is pointless given doubts about the intake temperature throughout and the lab temperature late in the cruise.

4. The TSG intake temperature is further from the CTD temperature than usual.

5. The TSG Salinity is lower than the CTD salinity by ~0.007 during the offshore casts. It is much lower for inlet casts. This is likely because the mixed layer was very shallow and the gradients high in the inlets, especially Effingham Inlet.

6. The TSG fluorometer had values higher than the CTD fluorometer by a median factor of 2. This fluorometer has not been recalibrated in many years.

f.) Editing

The ATC files were copied to *.EDT for files #2, 4 and 5.

The ATC files were opened in CTDEDIT. Only file #5 required editing to remove records at the end after the flow was turned off.

The edited file was copied to *.EDT.

The EDT files were thinned so that there are only records every 30s rather than 10s.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

Work on this project was suspended pending an analysis of the TSG data from cruise 2011-26. During that cruise the intake temperature looked excellent but the salinity looked bad. No obvious explanation was found though it is noted that the salinity during 2011-26 looked reasonable near-shore but mostly bad offshore. This is probably just chance, but if a similar result is found on other cruises it will need investigation.

The intake temperature is not useful, so there will need to be a recalibration of the lab temperature to produce a proxy for the intake temperature. The salinity will be increased by 0.007.

First ADD CHANNEL was used to add channel Temperature:Lab after Temperature:Primary.

CALIBRATE was run to set Temperature:Lab = Temperature:Primary

Then CALIBRATE was run a second time to apply file 2011-44-tsg-recal1.ccf to adjust salinity by adding 0.007 and Temperature:Primary by subtracting 0.25Cº. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Record #, Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0, Position:New, Temperature:Secondary and Flag.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data and all look fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

2. Logged as 2, but actually is the repeat of cast #10.

3. Looked like there was a change of configuration, but only because of overwrite of file #2.

6. CTD returned to surface after 2 bottles had been tripped and cast was rerun. New start  @ scan #16600. No CTD data available for pumped channels for the first 2 bottles.

9. Bottle 1 tripped at 320db, but no 300m bottle was on the rosette sheet as that depth was not expected.

10. CTD SBE550 used for this cast and rest of cruise.

24. 2 5m bottles tripped for Cindy Wright. Pumps never turned on. CTD data not processed.

27. Winch wire is very rusty and rust comes off during deployment.

# CRUISE SUMMARY     CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #585 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug2010 | IOS |  |  |
| SBE 43 DO sensor | 0047 | 29Mar2011 | Factory |  |  |
| Eco-AFL Fluorometer | 2216 | 21Mar2011 | Factory |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

| Calibration Information CTD #550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 15Aug2010 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Mar2011 | Factory |  |  |
| SBE 43 DO sensor #2 | 1438 | 2Apr11 | Factory |  |  |
| Eco-AFL Fluorometer | 2216 | 21Mar2011 | Factory |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2011-44

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 23Mar2011 | Factory |  |  |
| Conductivity | 3363 | 23Mar2011 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Jan01 | “ |  |  |
| Temperature:Secondary | 0603 | 3Mar2011 |  |  |  |

-