## comparison-0047
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
7 | 125.241 | 8.5589 | 8.5619 | 29.4954 | 29.4921 | 6.7548 | 2.2538 | 6 | 6 | 55 | 6.3837 | 0.89227 | 5.696 |  | 0.00054514 | 0.497 | -0.2284 |  |  |  | Summary of Soc Voffset including the original values in the factory calibration
8 | 125.087 | 8.5602 | 8.5603 | 29.5339 | 29.5335 | 6.7529 | 2.2825 | 4 | 4 | 66 | 6.3817 | 0.9054 | 5.778 |  | 0.00061611 | [7 empty] | m | b | Soc | Voffset | R2
9 | 125.231 | 8.3808 | 8.3806 | 29.7749 | 29.7747 | 6.7697 | 2.2879 | 6 | 6 | 79 | 6.4064 | 0.90941 | 5.826 |  | 0.00072009 | [5 empty] | Bottles used | Original | 0.5298 | -0.32726 | 0.5298 | -0.6177
8 | 174.779 | 8.509 | 8.5094 | 29.5977 | 29.5968 | 6.7579 | 2.2768 | 2 | 2 | 64 | 6.4296 | 0.9041 | 5.813 |  | 0.00094321 | [5 empty] | 76 | all data | 0.3877 | 0.0265 | 0.3877 | 0.068352 | 0.7152
9 | 150.912 | 8.3821 | 8.3822 | 29.7758 | 29.7751 | 6.7694 | 2.2804 | 5 | 5 | 78 | 6.4271 | 0.906 | 5.823 |  | 0.0010516 | [5 empty] | 74 | excl. outliers diff>0.1 | 0.497 | -0.2284 | 0.497 | -0.45956 | 0.923
6 | 49.896 | 8.7341 | 8.7338 | 29.3075 | 29.3072 | 6.7364 | 2.1764 | 5 | 5 | 45 | 6.2971 | 0.85214 | 5.366 |  | 0.0011344 | [5 empty] | 71 | excl. outliers diff>0.05 | 0.516 | -0.2697 | 0.516 | -0.52267 | 0.9539
8 | 150.413 | 8.5435 | 8.5436 | 29.5538 | 29.5533 | 6.7546 | 2.2794 | 3 | 3 | 65 | 6.4048 | 0.90323 | 5.785 |  | 0.0012231 | [5 empty] | 67 | excl. outliers diff>0.03 | 0.4923 | -0.2152 | 0.4923 | -0.43713 | 0.9302
9 | 74.123 | 8.4147 | 8.4145 | 29.7244 | 29.7244 | 6.7667 | 2.3007 | 8 | 8 | 81 | 6.3602 | 0.91381 | 5.812 |  | 0.0012422 | [5 empty] | 62 | excl. outliers diff>0.02 | 0.4673 | -0.159 | 0.4673 | -0.34025 | 0.9577
7 | 199.979 | 8.5148 | 8.514 | 29.5542 | 29.5549 | 6.7589 | 2.2399 | 3 | 3 | 52 | 6.4511 | 0.88606 | 5.716 |  | 0.0012459 | [5 empty] | 58 | excl. outliers diff>0.015 | 0.4696 | -0.165 | 0.4696 | -0.35136 | 0.9649
7 | 175.18 | 8.5142 | 8.5142 | 29.5549 | 29.5548 | 6.759 | 2.2514 | 4 | 4 | 53 | 6.4307 | 0.89212 | 5.737 |  | 0.0015659 | [5 empty] | 53 | excl. outliers diff>0.01 | 0.487 | -0.2047 | 0.487 | -0.42033 | 0.9646
8 | 100.389 | 8.574 | 8.5743 | 29.521 | 29.5205 | 6.7514 | 2.2885 | 5 | 5 | 67 | 6.3595 | 0.90714 | 5.769 |  | 0.0018502
7 | 75.152 | 8.6526 | 8.6527 | 29.3786 | 29.3783 | 6.7456 | 2.2078 | 8 | 8 | 57 | 6.3299 | 0.86699 | 5.488 |  | 0.0018916
7 | 50.062 | 8.7007 | 8.701 | 29.3235 | 29.3224 | 6.7407 | 2.1894 | 9 | 9 | 58 | 6.3029 | 0.86167 | 5.431 |  | 0.0019406
5 | 50.651 | 8.564 | 8.5633 | 29.3348 | 29.3351 | 6.761 | 2.2258 | 3 | 3 | 36 | 6.3287 | 0.87995 | 5.569 |  | 0.0021151
9 | 100.545 | 8.4343 | 8.435 | 29.7026 | 29.7015 | 6.7647 | 2.2965 | 7 | 7 | 80 | 6.3789 | 0.91082 | 5.81 |  | 0.0021675
9 | 174.023 | 8.356 | 8.3562 | 29.8151 | 29.8145 | 6.7717 | 2.2746 | 4 | 4 | 77 | 6.4496 | 0.90455 | 5.834 |  | 0.0024603
6 | 10.133 | 8.8127 | 8.8127 | 29.2478 | 29.2475 | 6.7271 | 2.1726 | 8 | 8 | 48 | 6.2529 | 0.85384 | 5.339 |  | 0.0024763
9 | 200.803 | 8.3561 | 8.3568 | 29.8184 | 29.8171 | 6.7716 | 2.2683 | 3 | 3 | 76 | 6.4716 | 0.90164 | 5.835 |  | 0.0026817
8 | 200.019 | 8.4833 | 8.4834 | 29.6316 | 29.6314 | 6.7603 | 2.2747 | 1 | 1 | 63 | 6.454 | 0.89929 | 5.804 |  | 0.0028444
3 | 102.024 | 8.1484 | 8.1485 | 29.8928 | 29.8931 | 6.8002 | 2.2949 | 7 | 7 | 22 | 6.4275 | 0.90922 | 5.844 |  | 0.002961
8 | 75.162 | 8.5936 | 8.5936 | 29.4963 | 29.4959 | 6.7495 | 2.2859 | 6 | 6 | 68 | 6.3363 | 0.91094 | 5.772 |  | 0.003234
3 | 126.505 | 8.1426 | 8.1429 | 29.9277 | 29.9271 | 6.7995 | 2.2905 | 6 | 6 | 21 | 6.4473 | 0.90674 | 5.846 |  | 0.0032353
7 | 30.24 | 8.7278 | 8.7265 | 29.3091 | 29.3089 | 6.7372 | 2.1709 | 10 | 10 | 59 | 6.2824 | 0.85397 | 5.365 |  | 0.0034154
3 | 202.269 | 8.1418 | 8.142 | 29.9703 | 29.9701 | 6.7978 | 2.2734 | 3 | 3 | 18 | 6.5085 | 0.90513 | 5.891 |  | 0.00366
7 | 250.237 | 8.5243 | 8.5243 | 29.5508 | 29.5507 | 6.7576 | 2.2273 | 2 | 2 | 51 | 6.4909 | 0.88247 | 5.728 | 6 | 0.0039025
3 | 151.53 | 8.1384 | 8.1384 | 29.9557 | 29.9553 | 6.799 | 2.2864 | 5 | 5 | 20 | 6.4676 | 0.90389 | 5.846 |  | 0.0040288
7 | 100.098 | 8.6168 | 8.6163 | 29.4198 | 29.4198 | 6.7493 | 2.2178 | 7 | 7 | 56 | 6.3553 | 0.86967 | 5.527 |  | 0.0041657
6 | 75.052 | 8.7252 | 8.7247 | 29.3168 | 29.3168 | 6.7373 | 2.1848 | 4 | 4 | 44 | 6.3186 | 0.85319 | 5.391 |  | 0.0042719
9 | 251.011 | 8.3538 | 8.3535 | 29.8292 | 29.829 | 6.7714 | 2.2561 | 2 | 2 | 75 | 6.5133 | 0.8974 | 5.845 |  | 0.0045136
9 | 50.203 | 8.4824 | 8.4827 | 29.6338 | 29.6331 | 6.7604 | 2.2961 | 9 | 9 | 82 | 6.3316 | 0.91746 | 5.809 | 6 | 0.0047104
7 | 3.047 | 8.7915 | 8.7914 | 29.1961 | 29.1967 | 6.7325 | 2.1721 | 13 | 13 | 62 | 6.2533 | 0.85603 | 5.353 |  | 0.0048947
7 | 303.146 | 8.5016 | 8.5011 | 29.5879 | 29.5885 | 6.7595 | 2.2214 | 1 | 1 | 50 | 6.5379 | 0.88072 | 5.758 |  | 0.0050655
3 | 238.009 | 8.1347 | 8.1348 | 29.9869 | 29.9869 | 6.7982 | 2.2692 | 1 | 1 | 16 | 6.539 | 0.89433 | 5.848 |  | 0.0050919
1 | 20.549 | 7.6952 | 7.6966 | 29.3585 | 29.3598 | 6.8948 | 2.0673 | 4 | 4 | 4 | 6.4715 | 0.79394 | 5.138 |  | 0.0051196
5 | 9.921 | 8.5772 | 8.5773 | 29.0255 | 29.0252 | 6.7726 | 2.2955 | 6 | 6 | 39 | 6.306 | 0.91769 | 5.787 |  | 0.0052231
6 | 3.186 | 8.9501 | 8.9518 | 29.0265 | 29.0232 | 6.716 | 2.1886 | 9 | 9 | 49 | 6.2307 | 0.86458 | 5.387 |  | 0.0052249
4 | 20.388 | 8.5166 | 8.518 | 29.1606 | 29.1597 | 6.7759 | 2.3124 | 3 | 3 | 31 | 6.3204 | 0.91528 | 5.785 |  | 0.0055599
3 | 238.043 | 8.1321 | 8.133 | 29.9893 | 29.9886 | 6.7984 | 2.2689 | 2 | 2 | 17 | 6.5394 | 0.89366 | 5.844 | 6 | 0.0055683
6 | 100.533 | 8.7248 | 8.7249 | 29.3242 | 29.3234 | 6.737 | 2.1854 | 3 | 3 | 43 | 6.339 | 0.85203 | 5.401 | 3 | 0.0056982
5 | 20.26 | 8.5641 | 8.564 | 29.2075 | 29.2068 | 6.7666 | 2.2416 | 5 | 5 | 38 | 6.3094 | 0.89168 | 5.626 |  | 0.0060178
3 | 176.325 | 8.142 | 8.1422 | 29.9582 | 29.9577 | 6.7983 | 2.2776 | 4 | 4 | 19 | 6.4874 | 0.90961 | 5.901 |  | 0.0060278
6 | 30.469 | 8.7654 | 8.7656 | 29.2718 | 29.2714 | 6.7332 | 2.1511 | 6 | 6 | 46 | 6.2771 | 0.84689 | 5.316 | 6 | 0.0061929
7 | 10.438 | 8.7898 | 8.7898 | 29.1964 | 29.1979 | 6.7328 | 2.1617 | 12 | 12 | 61 | 6.2595 | 0.8523 | 5.335 |  | 0.0063429
5 | 76.551 | 8.5018 | 8.5019 | 29.4508 | 29.4506 | 6.7654 | 2.2628 | 2 | 2 | 35 | 6.3568 | 0.88928 | 5.653 |  | 0.0069072
8 | 50.125 | 8.6076 | 8.6075 | 29.4552 | 29.455 | 6.7491 | 2.259 | 7 | 7 | 69 | 6.3151 | 0.9018 | 5.695 | 6 | 0.0074892
3 | 75.93 | 8.1537 | 8.1537 | 29.8493 | 29.8493 | 6.8013 | 2.2936 | 8 | 8 | 23 | 6.4068 | 0.91965 | 5.892 |  | 0.0081437
9 | 30.39 | 8.5131 | 8.5133 | 29.594 | 29.5934 | 6.7574 | 2.3004 | 10 | 10 | 83 | 6.3114 | 0.92309 | 5.826 |  | 0.0081947
9 | 20.39 | 8.491 | 8.4913 | 29.6196 | 29.619 | 6.7597 | 2.2923 | 11 | 11 | 84 | 6.3065 | 0.91937 | 5.798 |  | 0.008506
9 | 10.397 | 8.501 | 8.5006 | 29.606 | 29.6062 | 6.7588 | 2.2984 | 12 | 12 | 85 | 6.2971 | 0.92249 | 5.809 |  | 0.0085753
4 | 29.657 | 8.509 | 8.5091 | 29.2188 | 29.2187 | 6.7745 | 2.2751 | 2 | 2 | 30 | 6.327 | 0.89363 | 5.654 | 6 | 0.0087081
8 | 10.315 | 8.8556 | 8.8582 | 28.9568 | 28.9533 | 6.7333 | 2.3004 | 10 | 10 | 72 | 6.2568 | 0.90541 | 5.665 |  | 0.0095027
6 | 20.499 | 8.7758 | 8.7764 | 29.2648 | 29.2643 | 6.7319 | 2.1512 | 7 | 7 | 47 | 6.2674 | 0.85027 | 5.329 |  | 0.0095055
8 | 20.345 | 8.6686 | 8.6642 | 29.3298 | 29.3375 | 6.7453 | 2.2135 | 9 | 9 | 71 | 6.2848 | 0.88197 | 5.543 |  | 0.010265
5 | 30.63 | 8.5645 | 8.5647 | 29.2888 | 29.2882 | 6.763 | 2.2151 | 4 | 4 | 37 | 6.3144 | 0.88338 | 5.578 | 6 | 0.010883
7 | 19.955 | 8.7584 | 8.76 | 29.2811 | 29.2789 | 6.7338 | 2.1522 | 11 | 11 | 60 | 6.2696 | 0.85365 | 5.352 |  | 0.012395
5 | 2.532 | 8.6626 | 8.6609 | 28.6887 | 28.6928 | 6.7743 | 2.5158 | 7 | 7 | 40 | 6.2977 | 1.0091 | 6.355 |  | 0.012829
9 | 1.882 | 8.4967 | 8.4967 | 29.6112 | 29.6109 | 6.7592 | 2.2929 | 13 | 13 | 86 | 6.2909 | 0.92483 | 5.818 |  | 0.013676
3 | 30.723 | 8.1593 | 8.1595 | 29.6405 | 29.6403 | 6.8097 | 2.2306 | 10 | 10 | 25 | 6.3774 | 0.89487 | 5.707 |  | 0.01466
8 | 3.709 | 8.8698 | 8.8704 | 28.917 | 28.9155 | 6.7329 | 2.3279 | 11 | 11 | 73 | 6.2505 | 0.94376 | 5.899 |  | 0.015212
1 | 30.248 | 7.6952 | 7.7016 | 29.4961 | 29.4995 | 6.8886 | 2.0619 | 3 | 3 | 3 | 6.4738 | 0.81236 | 5.259 |  | 0.016002
8 | 30.322 | 8.6351 | 8.6347 | 29.3904 | 29.3912 | 6.7478 | 2.2067 | 8 | 8 | 70 | 6.2966 | 0.8846 | 5.57 |  | 0.016292
3 | 50.708 | 8.1508 | 8.1508 | 29.7208 | 29.7202 | 6.8075 | 2.267 | 9 | 9 | 24 | 6.3921 | 0.91801 | 5.868 |  | 0.019694
3 | 9.711 | 8.2923 | 8.2941 | 29.3468 | 29.3429 | 6.8022 | 2.2656 | 12 | 12 | 27 | 6.3469 | 0.91887 | 5.832 |  | 0.021273
3 | 2.272 | 8.8032 | 8.8217 | 27.327 | 27.2405 | 6.8127 | 2.4279 | 13 | 13 | 28 | 6.3266 | 1.0002 | 6.328 |  | 0.021943
4 | 41.159 | 8.5105 | 8.5108 | 29.2576 | 29.2572 | 6.7726 | 2.2895 | 1 | 1 | 29 | 6.3344 | 0.88595 | 5.612 |  | 0.023518
3 | 19.812 | 8.1884 | 8.1896 | 29.5799 | 29.5779 | 6.8079 | 2.2401 | 11 | 11 | 26 | 6.3654 | 0.91164 | 5.803 |  | 0.026712
1 | 70.674 | 8.1816 | 8.1815 | 30.2038 | 30.2044 | 6.7813 | 2.121 | 1 | 1 | 1 | 6.3824 | 0.79594 | 5.08 |  | 0.029779
4 | 9.774 | 8.5702 | 8.5701 | 29.035 | 29.0346 | 6.7732 | 2.4676 | 4 | 4 | 32 | 6.3069 | 0.96784 | 6.104 |  | 0.030142
9 | 316.785 | 8.3554 | 8.3554 | 29.8358 | 29.8354 | 6.7709 | 2.2421 | 1 | 1 | 74 | 6.5677 | 0.85525 | 5.617 | 4 | 0.030658
4 | 2.043 | 8.7337 | 8.7323 | 28.6277 | 28.6421 | 6.7661 | 2.9415 | 5 | 5 | 33 | 6.2864 | 1.267 | 7.965 |  | 0.033481
1 | 50.401 | 8.1778 | 8.1788 | 29.9779 | 29.9809 | 6.7919 | 2.2014 | 2 | 2 | 2 | 6.3759 | 0.82169 | 5.239 |  | 0.04403
1 | 9.96 | 8.25 | 8.2727 | 29.3162 | 29.318 | 6.81 | 2.4282 | 5 | 5 | 5 | 6.3565 | 0.92032 | 5.85 |  | 0.058081
5 | 95.934 | 8.4986 | 8.4986 | 29.5235 | 29.5245 | 6.7627 | 2.3684 | 1 | 1 | 34 | 6.3702 | 0.88977 | 5.668 |  | 0.058943
1 | 2.032 | 9.5219 | 9.1269 | 29.0635 | 29.1426 | 6.6295 | 3.0346 | 6 | 6 | 6 | 6.124 | 1.3764 | 8.429 |  | 0.096612
6 | 157.004 | 8.7042 | 8.7089 | 29.3679 | 29.3391 | 6.7383 | 1.6898 | 1 | 1 | 41 | 6.387 | 0.8558 | 5.466 |  | 0.24438
6 | 126.401 | 8.7275 | 8.7218 | 29.3505 | 29.3368 | 6.7355 | 1.616 | 2 | 2 | 42 | 6.3584 | 0.85305 | 5.424 |  | 0.27827

## comparison-1176
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | M/H | Titrated DO Flag | Sample # | Diff in samp #
15 | 758.615 | 4.3093 | 4.3093 | 34.2359 | 34.2373 | 7.228 | 0.60743 | 1 | 1 | 111 | 7.8794 | 0.046323 | 0.365 |  | 2.9166e-05 | 0.4566 | -0.231 | 0.076261 |  |  | Summary of Soc Voffset including the original values in the factory calibration - #1176
13 | 506.822 | 5.0146 | 5.0148 | 34.0893 | 34.0901 | 7.114 | 0.73463 | 4 | 4 | 97 | 7.4938 | 0.10449 | 0.783 |  | 5.3829e-05 |  |  | 0.14223 |  |  |  |  | m | b | Soc | Voffset | R2
34 | 19.979 | 8.5479 | 8.5489 | 31.619 | 31.6184 | 6.664 | 1.9892 | 5 | 5 | 239 | 6.5565 | 0.67735 | 4.441 |  | 8.7974e-05 |  |  | 0.34052 |  |  | Bottles used | Original | 0.4453 | -0.23209 | 0.4453 | -0.5212 |  | 2.1506
22 | 150.646 | 6.6144 | 6.6141 | 33.919 | 33.92 | 6.8599 | 1.2325 | 4 | 4 | 163 | 6.8809 | 0.33193 | 2.284 |  | 0.0001934 |  |  | 0.26933 |  |  | 172 | all data | 0.4494 | -0.2206 | 0.4494 | -0.49088 | 0.975 | 2.2055
17 | 299.482 | 5.7297 | 5.73 | 33.956 | 33.9568 | 7.001 | 1.0344 | 10 | 10 | 135 | 7.1694 | 0.24158 | 1.732 |  | 0.00029344 |  |  | 0.23356 |  |  | 167 | excl. outliers diff>0.1 | 0.4566 | -0.231 | 0.4566 | -0.50591 | 0.9946 | 2.2337
27 | 29.42 | 8.7056 | 8.7059 | 31.8094 | 31.8093 | 6.6323 | 1.7468 | 4 | 4 | 187 | 6.5321 | 0.56628 | 3.699 | 6 | 0.00031349 |  |  | 0.32418 |  |  | 160 | excl. outliers diff>0.07 | 0.4611 | -0.237 | 0.4611 | -0.51399 | 0.9974 | 2.2526
20 | 119.131 | 7.1761 | 7.176 | 33.6833 | 33.6844 | 6.7824 | 1.5393 | 1 | 1 | 153 | 6.7699 | 0.4715 | 3.192 |  | 0.00034908 |  |  | 0.30631 |  |  | 153 | excl. outliers diff>0.05 | 0.4631 | -0.2399 | 0.4631 | -0.51803 | 0.9987 | 2.2606
30 | 20.526 | 8.6444 | 8.6446 | 31.6156 | 31.6168 | 6.6498 | 1.8676 | 7 | 7 | 213 | 6.5422 | 0.62135 | 4.065 |  | 0.00041078 |  |  | 0.33269 |  |  | 147 | excl. outliers diff>0.03 | 0.4618 | -0.2374 | 0.4618 | -0.51408 | 0.9992 | 2.2567
13 | 305.9 | 5.6029 | 5.6031 | 34.0126 | 34.0133 | 7.0192 | 0.8947 | 5 | 5 | 98 | 7.1958 | 0.17705 | 1.274 | 6 | 0.00047188 |  |  | 0.19789 |  |  | 135 | excl. outliers diff>0.02 | 0.4605 | -0.2355 | 0.4605 | -0.5114 | 0.9995 | 2.2516
13 | 1152.059 | 3.408 | 3.4076 | 34.3869 | 34.3884 | 7.3804 | 0.60523 | 1 | 1 | 94 | 8.4884 | 0.044649 | 0.379 |  | 0.00069892 |  |  | 0.073772 |  |  | 124 | excl. outliers diff>0.015 | 0.4604 | -0.2352 | 0.4604 | -0.51086 | 0.9997 | 2.2515
13 | 127.108 | 7.0502 | 7.0493 | 33.8244 | 33.825 | 6.7957 | 1.5313 | 10 | 10 | 103 | 6.7913 | 0.46898 | 3.185 |  | 0.00077986 |  |  | 0.30626 |  |  | 102 | excl. outliers diff>0.01 | 0.4588 | -0.2328 | 0.4588 | -0.50741 | 0.9998 | 2.2454
17 | 749.962 | 4.1201 | 4.1205 | 34.2569 | 34.2577 | 7.26 | 0.5855 | 7 | 7 | 132 | 7.9092 | 0.037298 | 0.295 |  | 0.00095888 |  |  | 0.063703 |  |  | 94 | excl. outliers diff>0.008 | 0.4584 | -0.2328 | 0.4584 | -0.50785 | 0.9998 | 2.2425
22 | 175.255 | 6.3658 | 6.3658 | 33.9576 | 33.9584 | 6.8977 | 1.1204 | 3 | 3 | 162 | 6.9436 | 0.28155 | 1.955 | 6 | 0.00097075 |  |  | 0.25129 |  |  | 38 | excl. outliers diff>0.003 | 0.4563 | -0.2307 | 0.4563 | -0.50559 | 0.9999 | 2.2322
31 | 75.263 | 8.4549 | 8.4551 | 32.2915 | 32.2918 | 6.6488 | 1.844 | 3 | 3 | 218 | 6.5886 | 0.60999 | 4.019 |  | 0.00097711 |  |  | 0.3308 |  |  |  | 2011-16 best result | 0.4661 | -0.2471 | 0.4661 | -0.53014 | 0.9997 | 2.2665
15 | 126.292 | 7.2364 | 7.2362 | 33.7304 | 33.7309 | 6.7709 | 1.5389 | 8 | 8 | 118 | 6.7641 | 0.47057 | 3.183 |  | 0.0010748 |  |  | 0.30579
15 | 305.476 | 5.8722 | 5.8725 | 34.036 | 34.0367 | 6.9739 | 0.85351 | 3 | 3 | 113 | 7.1453 | 0.1601 | 1.144 |  | 0.0013921 |  |  | 0.18758
17 | 1748.799 | 2.2491 | 2.2482 | 34.5486 | 34.5509 | 7.5868 | 0.73677 | 3 | 3 | 128 | 9.468 | 0.10382 | 0.983 |  | 0.0015855 |  |  | 0.14092
11 | 96.136 | 7.1012 | 7.1017 | 33.727 | 33.7279 | 6.7921 | 1.4519 | 1 | 1 | 87 | 6.7603 | 0.43031 | 2.909 |  | 0.0016255 |  |  | 0.29638
22 | 200.063 | 6.2078 | 6.2084 | 33.9767 | 33.9774 | 6.9222 | 1.0544 | 2 | 2 | 161 | 6.9924 | 0.25213 | 1.763 |  | 0.0016956 |  |  | 0.23912
17 | 2083.101 | 1.846 | 1.8453 | 34.6098 | 34.6119 | 7.6607 | 0.84655 | 1 | 1 | 126 | 10.0027 | 0.15726 | 1.573 |  | 0.0017228 |  |  | 0.18576
28 | 20.469 | 8.6762 | 8.6762 | 31.6121 | 31.6079 | 6.6452 | 1.8073 | 3 | 3 | 193 | 6.5374 | 0.59595 | 3.896 |  | 0.0017486 |  |  | 0.32975
15 | 176.008 | 6.5398 | 6.54 | 33.8871 | 33.8877 | 6.8732 | 1.2396 | 6 | 6 | 116 | 6.9176 | 0.33682 | 2.33 |  | 0.0018034 |  |  | 0.27171
13 | 760.488 | 4.3008 | 4.3005 | 34.2273 | 34.2284 | 7.2299 | 0.60773 | 3 | 3 | 96 | 7.8836 | 0.048328 | 0.381 |  | 0.0018389 |  |  | 0.079523
37 | 96.14 | 7.9086 | 7.9199 | 32.6427 | 32.6405 | 6.7158 | 1.6912 | 1 | 1 | 255 | 6.6771 | 0.5393 | 3.601 |  | 0.0019128 |  |  | 0.31888
31 | 20.032 | 8.611 | 8.6112 | 31.5553 | 31.551 | 6.6573 | 1.927 | 6 | 6 | 221 | 6.5495 | 0.64692 | 4.237 |  | 0.0019653 |  |  | 0.33571
15 | 202.872 | 6.3882 | 6.388 | 33.9199 | 33.9206 | 6.8959 | 1.2192 | 5 | 5 | 115 | 6.9662 | 0.32773 | 2.283 |  | 0.0020165 |  |  | 0.26879
27 | 20.385 | 8.7086 | 8.7091 | 31.6222 | 31.6222 | 6.6399 | 1.7674 | 5 | 5 | 188 | 6.532 | 0.57379 | 3.748 |  | 0.0022203 |  |  | 0.32465
18 | 48.411 | 8.5334 | 8.5304 | 32.3121 | 32.317 | 6.6363 | 2.2631 | 1 | 1 | 148 | 6.553 | 0.80009 | 5.243 |  | 0.0022625 |  |  | 0.35353
17 | 2000.077 | 1.8934 | 1.893 | 34.6002 | 34.6021 | 7.6521 | 0.83218 | 2 | 2 | 127 | 9.8816 | 0.14664 | 1.449 |  | 0.0023376 |  |  | 0.17621
30 | 50.348 | 8.5466 | 8.5455 | 32.1589 | 32.1627 | 6.6409 | 1.8641 | 5 | 5 | 211 | 6.5591 | 0.61776 | 4.052 |  | 0.0023847 |  |  | 0.3314
26 | 9.598 | 8.7819 | 8.7743 | 31.2518 | 31.291 | 6.645 | 1.5916 | 5 | 5 | 182 | 6.5274 | 0.4933 | 3.22 |  | 0.0024229 |  |  | 0.30994
32 | 30.459 | 8.5775 | 8.578 | 31.7982 | 31.7937 | 6.6518 | 1.9405 | 3 | 3 | 226 | 6.5531 | 0.65756 | 4.309 | 6 | 0.0025101 |  |  | 0.33885
13 | 1012.795 | 3.6559 | 3.6561 | 34.3477 | 34.349 | 7.3378 | 0.59689 | 2 | 2 | 95 | 8.282 | 0.039 | 0.323 |  | 0.0025395 |  |  | 0.065339
17 | 1499.92 | 2.548 | 2.5482 | 34.5161 | 34.5175 | 7.5321 | 0.68547 | 4 | 4 | 129 | 9.0895 | 0.079432 | 0.722 |  | 0.0025533 |  |  | 0.11588
31 | 50.43 | 8.5509 | 8.5525 | 32.1465 | 32.1433 | 6.6408 | 1.8729 | 4 | 4 | 219 | 6.5591 | 0.62677 | 4.111 |  | 0.0026 |  |  | 0.33465
32 | 50.368 | 8.5326 | 8.5332 | 32.1581 | 32.157 | 6.643 | 1.8843 | 2 | 2 | 225 | 6.5613 | 0.63204 | 4.147 |  | 0.0026415 |  |  | 0.33541
36 | 20.124 | 8.4356 | 8.4309 | 31.6262 | 31.6184 | 6.6805 | 2.0591 | 5 | 5 | 252 | 6.5736 | 0.70646 | 4.644 |  | 0.0027444 |  |  | 0.34308
17 | 249.723 | 6.0473 | 6.0482 | 33.9061 | 33.9068 | 6.9514 | 1.2831 | 11 | 11 | 136 | 7.069 | 0.35776 | 2.529 |  | 0.0028866 |  |  | 0.27882
13 | 175.571 | 6.4911 | 6.4915 | 33.8993 | 33.8999 | 6.8804 | 1.3299 | 8 | 8 | 101 | 6.925 | 0.37921 | 2.626 |  | 0.0029886 |  |  | 0.28515
13 | 51.62 | 7.7942 | 7.7966 | 32.6607 | 32.6603 | 6.7325 | 2.4531 | 13 | 13 | 106 | 6.6566 | 0.88604 | 5.898 | 6 | 0.0030334 |  |  | 0.3612
30 | 30.16 | 8.6506 | 8.6508 | 31.7889 | 31.789 | 6.6414 | 1.8661 | 6 | 6 | 212 | 6.542 | 0.62412 | 4.083 |  | 0.003054 |  |  | 0.33445
11 | 21.039 | 8.4559 | 8.4561 | 32.2424 | 32.2431 | 6.6508 | 2.4781 | 5 | 5 | 91 | 6.545 | 0.90374 | 5.915 |  | 0.0032499 |  |  | 0.36469
29 | 150.285 | 8.2495 | 8.2497 | 32.8188 | 32.8191 | 6.6567 | 0.58679 | 3 | 3 | 198 | 6.6616 | 0.04023 | 0.268 |  | 0.0033022 |  |  | 0.06856
17 | 999.523 | 3.7027 | 3.7026 | 34.3668 | 34.3679 | 7.3285 | 0.57827 | 6 | 6 | 131 | 8.2561 | 0.029554 | 0.244 |  | 0.0034842 |  |  | 0.051107
28 | 29.901 | 8.6626 | 8.6629 | 31.88 | 31.8807 | 6.6357 | 1.8443 | 2 | 2 | 192 | 6.536 | 0.61459 | 4.017 |  | 0.0035031 |  |  | 0.33325
13 | 256.433 | 5.7595 | 5.7592 | 33.9812 | 33.982 | 6.9949 | 0.98797 | 6 | 6 | 99 | 7.1231 | 0.22364 | 1.593 |  | 0.0035313 |  |  | 0.22636
20 | 10.732 | 10.2851 | 10.2847 | 31.2678 | 31.2694 | 6.4276 | 2.9175 | 6 | 6 | 158 | 6.3065 | 1.1047 | 6.967 |  | 0.0035915 |  |  | 0.37865
15 | 19.121 | 8.4897 | 8.4911 | 32.3799 | 32.382 | 6.6398 | 2.6947 | 13 | 13 | 123 | 6.5324 | 1.003 | 6.552 | 6 | 0.0035997 |  |  | 0.37221
28 | 53.005 | 8.5421 | 8.5425 | 32.2224 | 32.2229 | 6.6388 | 1.7911 | 1 | 1 | 191 | 6.5594 | 0.59045 | 3.873 |  | 0.0036605 |  |  | 0.32967
33 | 30.03 | 8.564 | 8.564 | 31.7524 | 31.7528 | 6.6558 | 1.9689 | 2 | 2 | 231 | 6.5567 | 0.67168 | 4.404 |  | 0.0036831 |  |  | 0.34115
35 | 19.969 | 8.5481 | 8.5457 | 31.6869 | 31.6696 | 6.661 | 1.9908 | 4 | 4 | 245 | 6.5535 | 0.67429 | 4.419 |  | 0.0037021 |  |  | 0.33871
22 | 125.224 | 7.1328 | 7.1331 | 33.8457 | 33.8464 | 6.7819 | 1.5008 | 5 | 5 | 164 | 6.7751 | 0.458 | 3.103 |  | 0.0037505 |  |  | 0.30518
13 | 104.407 | 7.3496 | 7.3495 | 33.7094 | 33.71 | 6.7544 | 1.5973 | 11 | 11 | 104 | 6.7276 | 0.50211 | 3.378 |  | 0.0037776 |  |  | 0.31435
17 | 500.01 | 4.7678 | 4.7679 | 34.0958 | 34.0965 | 7.1557 | 0.70148 | 8 | 8 | 133 | 7.5352 | 0.085333 | 0.643 | 6 | 0.0039629 |  |  | 0.12165
34 | 75.095 | 8.1697 | 8.1692 | 32.5816 | 32.5832 | 6.679 | 1.6291 | 2 | 2 | 236 | 6.6205 | 0.50887 | 3.369 |  | 0.0039753 |  |  | 0.31236
29 | 29.166 | 8.6516 | 8.649 | 31.8981 | 31.9077 | 6.6365 | 1.8532 | 8 | 8 | 203 | 6.5364 | 0.61931 | 4.048 | 6 | 0.0041164 |  |  | 0.33417
15 | 9.651 | 9.6996 | 9.7 | 31.2954 | 31.2955 | 6.5093 | 2.8191 | 14 | 14 | 124 | 6.3887 | 1.0603 | 6.774 |  | 0.0041297 |  |  | 0.37612
36 | 30.004 | 8.4873 | 8.4887 | 31.8223 | 31.8276 | 6.6643 | 2.0656 | 4 | 4 | 251 | 6.5656 | 0.70794 | 4.648 |  | 0.0042066 |  |  | 0.34273
22 | 266.655 | 6.1402 | 6.1405 | 34.0034 | 34.0045 | 6.9319 | 0.96648 | 1 | 1 | 160 | 7.0634 | 0.21463 | 1.516 |  | 0.0043332 |  |  | 0.22207
22 | 100.289 | 7.4071 | 7.4079 | 33.6928 | 33.6931 | 6.7463 | 1.6099 | 6 | 6 | 165 | 6.7154 | 0.50853 | 3.415 |  | 0.0044533 |  |  | 0.31588
17 | 400.41 | 5.268 | 5.2687 | 34.0402 | 34.041 | 7.0736 | 0.79759 | 9 | 9 | 134 | 7.3455 | 0.13763 | 1.011 |  | 0.0044548 |  |  | 0.17256
29 | 124.583 | 8.3595 | 8.3597 | 32.7582 | 32.7585 | 6.6429 | 0.6406 | 4 | 4 | 199 | 6.6251 | 0.065961 | 0.437 |  | 0.0044634 |  |  | 0.10297
22 | 29.719 | 8.7651 | 8.7641 | 32.3788 | 32.3785 | 6.5991 | 2.7412 | 9 | 9 | 168 | 6.4992 | 1.0252 | 6.663 |  | 0.004561 |  |  | 0.37399
17 | 74.949 | 7.3564 | 7.3565 | 32.8381 | 32.8389 | 6.792 | 2.4616 | 17 | 17 | 142 | 6.7395 | 0.89755 | 6.049 |  | 0.0045882 |  |  | 0.36462
35 | 79.461 | 8.1145 | 8.1136 | 32.6066 | 32.6078 | 6.6862 | 1.689 | 1 | 1 | 242 | 6.6318 | 0.5448 | 3.613 | 6 | 0.0045934 |  |  | 0.32255
17 | 1249.266 | 3.1439 | 3.1438 | 34.4424 | 34.4436 | 7.4256 | 0.60911 | 5 | 5 | 130 | 8.6561 | 0.042513 | 0.368 |  | 0.0046062 |  |  | 0.069796
22 | 9.82 | 10.1767 | 10.2142 | 31.6317 | 31.5928 | 6.4278 | 2.8898 | 11 | 11 | 170 | 6.3064 | 1.0932 | 6.894 |  | 0.0047105 |  |  | 0.37829
30 | 74.791 | 8.4969 | 8.4975 | 32.2877 | 32.2879 | 6.6427 | 1.8006 | 4 | 4 | 210 | 6.5819 | 0.59588 | 3.922 |  | 0.0047176 |  |  | 0.33093
15 | 151.015 | 6.8913 | 6.8913 | 33.828 | 33.8287 | 6.8203 | 1.4 | 7 | 7 | 117 | 6.8386 | 0.4131 | 2.825 |  | 0.0048753 |  |  | 0.29508
11 | 10.491 | 9.0162 | 9.0208 | 31.6954 | 31.6892 | 6.5914 | 2.5993 | 6 | 6 | 92 | 6.474 | 0.96076 | 6.22 |  | 0.0049313 |  |  | 0.36963
17 | 150.645 | 6.9485 | 6.949 | 33.7734 | 33.7741 | 6.8138 | 1.7824 | 14 | 14 | 139 | 6.8311 | 0.5879 | 4.016 |  | 0.0050472 |  |  | 0.32983
15 | 2.559 | 10.0624 | 10.0626 | 31.0924 | 31.0916 | 6.4661 | 2.8522 | 15 | 15 | 125 | 6.3387 | 1.0766 | 6.824 |  | 0.0052518 |  |  | 0.37745
17 | 125.685 | 7.1734 | 7.1735 | 33.6209 | 33.6223 | 6.7856 | 1.8626 | 15 | 15 | 140 | 6.7788 | 0.62489 | 4.236 |  | 0.00544 |  |  | 0.3355
11 | 50.556 | 7.7379 | 7.7365 | 32.9406 | 32.9426 | 6.7288 | 2.1024 | 3 | 3 | 89 | 6.6525 | 0.73476 | 4.888 |  | 0.0057904 |  |  | 0.34948
29 | 50.062 | 8.5396 | 8.5396 | 32.1799 | 32.1808 | 6.641 | 1.8419 | 7 | 7 | 202 | 6.5591 | 0.6042 | 3.963 |  | 0.0058196 |  |  | 0.32803
20 | 76.366 | 7.8398 | 7.8411 | 32.9242 | 32.9236 | 6.7139 | 2.1497 | 2 | 2 | 154 | 6.659 | 0.75642 | 5.037 |  | 0.0058898 |  |  | 0.35188
29 | 19.896 | 8.6583 | 8.6586 | 31.5649 | 31.5653 | 6.6499 | 1.8261 | 9 | 9 | 204 | 6.5417 | 0.60871 | 3.982 |  | 0.0059038 |  |  | 0.33334
11 | 30.095 | 8.3946 | 8.396 | 32.3146 | 32.3139 | 6.6568 | 2.491 | 4 | 4 | 90 | 6.559 | 0.91234 | 5.984 | 6 | 0.0059639 | 0.465 | -0.2355 | 0.36626
36 | 50.583 | 8.443 | 8.4451 | 32.3237 | 32.32 | 6.6492 | 1.9664 | 3 | 3 | 250 | 6.5683 | 0.66044 | 4.338 |  | 0.0063966 |  |  | 0.33587
17 | 201.061 | 6.4737 | 6.4732 | 33.8578 | 33.8586 | 6.885 | 1.568 | 12 | 12 | 137 | 6.9526 | 0.49147 | 3.417 |  | 0.0064957 |  |  | 0.31343
32 | 19.886 | 8.579 | 8.5794 | 31.505 | 31.5045 | 6.6643 | 2.0117 | 4 | 4 | 227 | 6.5564 | 0.68101 | 4.465 |  | 0.0065227 |  |  | 0.33853
20 | 30.527 | 8.6484 | 8.6487 | 32.3528 | 32.3535 | 6.6174 | 2.5311 | 4 | 4 | 156 | 6.5187 | 0.93147 | 6.072 |  | 0.0067886 |  |  | 0.36802
18 | 30.342 | 8.6779 | 8.6782 | 32.0106 | 32.014 | 6.6278 | 2.2537 | 2 | 2 | 149 | 6.5285 | 0.80493 | 5.255 |  | 0.0069027 |  |  | 0.35716
31 | 30.436 | 8.5764 | 8.5758 | 31.7925 | 31.7885 | 6.6523 | 1.9444 | 5 | 5 | 220 | 6.5534 | 0.64989 | 4.259 |  | 0.0069442 |  |  | 0.33423
34 | 30.02 | 8.5567 | 8.5569 | 31.868 | 31.8622 | 6.6519 | 1.9635 | 4 | 4 | 238 | 6.5529 | 0.65848 | 4.315 |  | 0.0070493 |  |  | 0.33536
26 | 19.773 | 8.7552 | 8.7554 | 31.6765 | 31.6841 | 6.6307 | 1.6309 | 4 | 4 | 181 | 6.5221 | 0.50659 | 3.304 |  | 0.0070894 |  |  | 0.31062
11 | 76.519 | 7.3555 | 7.3557 | 33.5042 | 33.505 | 6.7626 | 1.6239 | 2 | 2 | 88 | 6.7116 | 0.51761 | 3.474 |  | 0.007158 |  |  | 0.31875
25 | 32.411 | 8.7666 | 8.7683 | 31.8559 | 31.854 | 6.6213 | 1.4093 | 1 | 1 | 174 | 6.5233 | 0.41988 | 2.739 |  | 0.0074001 |  |  | 0.29794
13 | 153.036 | 6.7436 | 6.743 | 33.8678 | 33.8685 | 6.8418 | 1.4432 | 9 | 9 | 102 | 6.8634 | 0.4355 | 2.989 |  | 0.0075232 |  |  | 0.30175
11 | 2.334 | 9.6184 | 9.6246 | 31.1604 | 31.156 | 6.5266 | 2.7327 | 7 | 7 | 93 | 6.4002 | 1.0244 | 6.556 |  | 0.0075906 |  |  | 0.37485
17 | 9.747 | 10.0151 | 10.0222 | 31.2387 | 31.2044 | 6.4667 | 2.857 | 21 | 21 | 146 | 6.3454 | 1.0811 | 6.86 |  | 0.0076146 |  |  | 0.37841
17 | 174.866 | 6.7715 | 6.7719 | 33.8302 | 33.8307 | 6.8391 | 1.6399 | 13 | 13 | 138 | 6.8797 | 0.52546 | 3.615 | 6 | 0.00769 |  |  | 0.32043
32 | 82.233 | 8.3686 | 8.3681 | 32.3933 | 32.3949 | 6.6573 | 1.8099 | 1 | 1 | 224 | 6.6035 | 0.60331 | 3.984 |  | 0.0079169 |  |  | 0.33334
15 | 29.712 | 8.2428 | 8.2429 | 32.4322 | 32.4334 | 6.6745 | 2.6183 | 12 | 12 | 122 | 6.5772 | 0.9726 | 6.397 |  | 0.0080959 |  |  | 0.37147
29 | 74.757 | 8.5238 | 8.5242 | 32.3115 | 32.3104 | 6.6377 | 1.7049 | 6 | 6 | 201 | 6.5766 | 0.5556 | 3.654 |  | 0.0081634 |  |  | 0.32589
13 | 203.95 | 6.1927 | 6.1912 | 33.9363 | 33.9371 | 6.9265 | 1.1722 | 7 | 7 | 100 | 7.0004 | 0.31255 | 2.188 |  | 0.0083073 |  |  | 0.26663
17 | 2.183 | 10.0039 | 10.0038 | 31.1105 | 31.1114 | 6.4736 | 2.8519 | 22 | 22 | 147 | 6.3461 | 1.0796 | 6.851 |  | 0.008384 |  |  | 0.37854
13 | 76.789 | 7.5895 | 7.5895 | 33.2853 | 33.2858 | 6.7363 | 1.9618 | 12 | 12 | 105 | 6.6836 | 0.67344 | 4.501 |  | 0.008697 |  |  | 0.34328
22 | 75.652 | 7.725 | 7.7253 | 33.206 | 33.2064 | 6.7191 | 1.9534 | 7 | 7 | 166 | 6.6644 | 0.66968 | 4.463 |  | 0.0087742 |  |  | 0.34283
34 | 95.639 | 8.1404 | 8.1407 | 32.605 | 32.6057 | 6.6824 | 1.5828 | 1 | 1 | 235 | 6.6416 | 0.50063 | 3.325 |  | 0.0089327 |  |  | 0.3163
30 | 140.607 | 8.2872 | 8.281 | 32.7958 | 32.801 | 6.6521 | 0.60695 | 1 | 1 | 207 | 6.6484 | 0.0367 | 0.244 | 6 | 0.0094329 |  |  | 0.060467
13 | 1.73 | 10.1173 | 10.1173 | 30.6944 | 30.6953 | 6.4748 | 2.8705 | 17 | 17 | 110 | 6.3463 | 1.0896 | 6.915 |  | 0.0099533 |  |  | 0.37959
13 | 30.04 | 8.6904 | 8.6841 | 32.0207 | 32.0272 | 6.6255 | 2.5082 | 14 | 14 | 107 | 6.526 | 0.92431 | 6.032 |  | 0.010047 |  |  | 0.36851
22 | 20.253 | 9.0014 | 9.0036 | 32.096 | 32.0963 | 6.5765 | 2.7945 | 10 | 10 | 169 | 6.4676 | 1.0551 | 6.824 | 6 | 0.010134 |  |  | 0.37756
17 | 100.019 | 7.3866 | 7.3869 | 33.3718 | 33.374 | 6.7636 | 1.9856 | 16 | 16 | 141 | 6.7326 | 0.68606 | 4.619 |  | 0.010449 |  |  | 0.34552
15 | 251.925 | 5.9881 | 5.9885 | 33.935 | 33.9358 | 6.9597 | 1.1648 | 4 | 4 | 114 | 7.0802 | 0.31129 | 2.204 |  | 0.010458 |  |  | 0.26726
17 | 29.562 | 8.6969 | 8.6919 | 32.4821 | 32.4828 | 6.6048 | 2.828 | 19 | 19 | 144 | 6.5051 | 1.0713 | 6.969 |  | 0.011046 |  |  | 0.37882
36 | 75.352 | 8.1357 | 8.1388 | 32.5375 | 32.5367 | 6.686 | 1.7799 | 2 | 2 | 249 | 6.628 | 0.57061 | 3.782 | 6 | 0.011079 |  |  | 0.32059
35 | 29.464 | 8.529 | 8.5291 | 31.8694 | 31.8559 | 6.656 | 2.0039 | 3 | 3 | 244 | 6.5567 | 0.6729 | 4.412 |  | 0.011082 |  |  | 0.3358
15 | 75.915 | 7.6946 | 7.6942 | 33.0195 | 33.0219 | 6.7319 | 2.0506 | 10 | 10 | 120 | 6.6776 | 0.71642 | 4.784 |  | 0.011119 |  |  | 0.34937
13 | 9.904 | 10.0029 | 10.0053 | 30.962 | 30.9599 | 6.4799 | 2.8596 | 16 | 16 | 109 | 6.3585 | 1.0859 | 6.905 |  | 0.011269 |  |  | 0.37976
29 | 175.079 | 8.2321 | 8.2322 | 32.8295 | 32.8299 | 6.6589 | 0.58572 | 2 | 2 | 197 | 6.6851 | 0.025131 | 0.168 |  | 0.011309 |  |  | 0.042906
29 | 195.577 | 8.2184 | 8.2187 | 32.8426 | 32.843 | 6.6603 | 0.55826 | 1 | 1 | 196 | 6.7042 | 0.012529 | 0.084 |  | 0.011372 |  |  | 0.022444
27 | 50.121 | 8.6634 | 8.6626 | 31.9703 | 31.9686 | 6.6316 | 1.6801 | 3 | 3 | 186 | 6.549 | 0.52405 | 3.432 |  | 0.012079 |  |  | 0.31192
17 | 20.388 | 9.3644 | 9.3658 | 32.3198 | 32.3207 | 6.5144 | 2.8688 | 20 | 20 | 145 | 6.4044 | 1.091 | 6.987 |  | 0.012083 |  |  | 0.38029
22 | 49.312 | 7.8738 | 7.8713 | 32.5804 | 32.5785 | 6.7238 | 2.5457 | 8 | 8 | 167 | 6.6454 | 0.94351 | 6.27 |  | 0.012163 |  |  | 0.37063
18 | 1.743 | 11.0185 | 11.0264 | 29.8535 | 29.8518 | 6.3835 | 3.2458 | 5 | 5 | 152 | 6.2529 | 1.2634 | 7.9 |  | 0.012377 |  |  | 0.38924
22 | 1.801 | 11.4499 | 11.4497 | 30.7154 | 30.7157 | 6.2902 | 2.8783 | 12 | 12 | 171 | 6.1601 | 1.0958 | 6.75 |  | 0.012552 |  |  | 0.3807
26 | 49.437 | 8.7226 | 8.7236 | 31.9826 | 31.9834 | 6.6224 | 1.4335 | 2 | 2 | 179 | 6.5388 | 0.43616 | 2.852 |  | 0.012646 |  |  | 0.30427
36 | 85.961 | 8.0032 | 8.0024 | 32.6518 | 32.651 | 6.7011 | 1.6349 | 1 | 1 | 248 | 6.653 | 0.52893 | 3.519 |  | 0.013439 |  |  | 0.32353
37 | 1.563 | 13.1632 | 13.2078 | 21.667 | 21.5836 | 6.4199 | 3.0352 | 7 | 7 | 261 | 6.2822 | 1.1687 | 7.342 |  | 0.01381 |  |  | 0.38504
18 | 20.562 | 8.9528 | 8.9576 | 31.5184 | 31.5161 | 6.6083 | 2.3623 | 3 | 3 | 150 | 6.4994 | 0.86162 | 5.6 |  | 0.013993 |  |  | 0.36474
20 | 2.446 | 11.3801 | 11.3663 | 30.6737 | 30.682 | 6.3013 | 2.952 | 7 | 7 | 159 | 6.1716 | 1.131 | 6.98 |  | 0.014119 |  |  | 0.38313
37 | 20.491 | 8.4507 | 8.4555 | 31.583 | 31.5818 | 6.6801 | 2.1224 | 5 | 5 | 259 | 6.5734 | 0.72382 | 4.758 |  | 0.014248 |  |  | 0.34105
15 | 100.871 | 7.4643 | 7.4636 | 33.4486 | 33.4512 | 6.7483 | 1.6812 | 9 | 9 | 119 | 6.7174 | 0.55185 | 3.707 |  | 0.015199 |  |  | 0.32824
37 | 75.04 | 8.1089 | 8.1089 | 32.5342 | 32.5336 | 6.6902 | 1.7713 | 2 | 2 | 256 | 6.6321 | 0.56257 | 3.731 |  | 0.015204 |  |  | 0.3176
15 | 503.637 | 4.8372 | 4.8376 | 34.1381 | 34.1391 | 7.1418 | 0.66829 | 2 | 2 | 112 | 7.5229 | 0.089725 | 0.675 | 3 | 0.015584 |  |  | 0.13426
36 | 1.473 | 13.4518 | 13.4826 | 21.4115 | 21.2949 | 6.391 | 3.1612 | 7 | 7 | 254 | 6.2533 | 1.2289 | 7.685 |  | 0.016557 |  |  | 0.38876
17 | 50.599 | 7.6571 | 7.6564 | 32.5912 | 32.5914 | 6.7565 | 2.7045 | 18 | 18 | 143 | 6.6807 | 1.0215 | 6.824 | 4 | 0.017585 |  |  | 0.37769
13 | 20.017 | 8.9129 | 8.9091 | 31.8195 | 31.823 | 6.6013 | 2.5543 | 15 | 15 | 108 | 6.4923 | 0.95328 | 6.189 |  | 0.017986 |  |  | 0.37321
30 | 124.703 | 8.3592 | 8.3585 | 32.7431 | 32.7434 | 6.6436 | 0.63899 | 2 | 2 | 208 | 6.6259 | 0.042258 | 0.28 |  | 0.018504 |  |  | 0.066133
34 | 1.507 | 13.6213 | 13.5676 | 18.5884 | 18.7547 | 6.4808 | 3.05 | 7 | 7 | 241 | 6.341 | 1.1807 | 7.487 | 6 | 0.019102 |  |  | 0.38713
35 | 50.646 | 8.4268 | 8.4271 | 32.348 | 32.3485 | 6.6506 | 1.9314 | 2 | 2 | 243 | 6.5698 | 0.63137 | 4.148 |  | 0.019525 |  |  | 0.32689
34 | 50.338 | 8.3484 | 8.3469 | 32.4316 | 32.433 | 6.6587 | 1.8545 | 3 | 3 | 237 | 6.5782 | 0.59606 | 3.921 |  | 0.019711 |  |  | 0.32141
31 | 1.575 | 13.4468 | 13.5551 | 16.9974 | 16.6902 | 6.5695 | 3.1067 | 8 | 8 | 223 | 6.4281 | 1.1675 | 7.505 |  | 0.019997 |  |  | 0.37581
20 | 21.078 | 9.1224 | 9.118 | 32.1173 | 32.119 | 6.558 | 2.6693 | 5 | 5 | 157 | 6.4492 | 1.0079 | 6.5 |  | 0.020083 |  |  | 0.37758
27 | 74.661 | 8.5231 | 8.5273 | 32.4852 | 32.4836 | 6.6303 | 0.55092 | 2 | 2 | 185 | 6.5693 | 0 | 0 |  | 0.02055 |  |  | 0
28 | 10.132 | 8.7373 | 8.7333 | 31.1651 | 31.168 | 6.6553 | 2.1331 | 4 | 4 | 194 | 6.5384 | 0.72189 | 4.72 | 6 | 0.021074 |  |  | 0.33843
31 | 100.436 | 8.4036 | 8.405 | 32.3691 | 32.3654 | 6.6531 | 1.7672 | 2 | 2 | 217 | 6.6145 | 0.55333 | 3.66 | 6 | 0.022583 |  |  | 0.31311
20 | 51.124 | 8.2841 | 8.2807 | 32.5 | 32.502 | 6.6654 | 2.5253 | 3 | 3 | 155 | 6.5859 | 0.94505 | 6.224 |  | 0.022997 |  |  | 0.37423
27 | 115.198 | 8.145 | 8.1443 | 32.5718 | 32.5728 | 6.6831 | 0.55955 | 1 | 1 | 184 | 6.6589 | 0 | 0 |  | 0.024491 |  |  | 0
33 | 59.856 | 8.2852 | 8.2769 | 32.4894 | 32.4953 | 6.6657 | 1.8019 | 1 | 1 | 230 | 6.5935 | 0.61742 | 4.071 |  | 0.025675 |  |  | 0.34265
15 | 49.982 | 7.8524 | 7.8526 | 32.6293 | 32.63 | 6.725 | 2.4664 | 11 | 11 | 121 | 6.6473 | 0.92083 | 6.121 |  | 0.025684 |  |  | 0.37335
31 | 113.721 | 8.3959 | 8.3967 | 32.4084 | 32.4099 | 6.6526 | 1.6264 | 1 | 1 | 216 | 6.6253 | 0.5387 | 3.569 |  | 0.027081 |  |  | 0.33122
30 | 10.611 | 8.639 | 8.6409 | 31.3223 | 31.3152 | 6.6632 | 1.9793 | 8 | 8 | 214 | 6.5472 | 0.69999 | 4.583 |  | 0.027255 |  |  | 0.35366
32 | 1.76 | 13.9418 | 13.9793 | 20.9459 | 20.8335 | 6.3438 | 3.2637 | 6 | 6 | 229 | 6.2066 | 1.2865 | 7.985 |  | 0.027328 |  |  | 0.39419
26 | 30.624 | 8.7497 | 8.7536 | 31.8371 | 31.8355 | 6.6246 | 1.5748 | 3 | 3 | 180 | 6.5251 | 0.51646 | 3.37 | 6 | 0.028424 |  |  | 0.32796
37 | 49.987 | 8.5012 | 8.4973 | 32.2483 | 32.25 | 6.6438 | 2.0369 | 3 | 3 | 257 | 6.562 | 0.66671 | 4.375 | 6 | 0.03232 |  |  | 0.32732
26 | 1.405 | 11.2622 | 11.2741 | 12.6875 | 12.2494 | 7.0792 | 2.9151 | 6 | 6 | 183 | 6.9331 | 1.1383 | 7.892 |  | 0.038263 |  |  | 0.39048
25 | 20.038 | 8.8069 | 8.8069 | 31.6315 | 31.6272 | 6.6249 | 1.3427 | 2 | 2 | 175 | 6.5163 | 0.34253 | 2.232 | 6 | 0.03956 |  |  | 0.2551
35 | 1.609 | 13.2721 | 13.268 | 17.5061 | 17.5065 | 6.5731 | 3.0159 | 6 | 6 | 247 | 6.432 | 1.1901 | 7.655 |  | 0.044064 |  |  | 0.39462
33 | 10.015 | 8.649 | 8.6521 | 31.2914 | 31.288 | 6.6631 | 2.2456 | 4 | 4 | 233 | 6.5465 | 0.74788 | 4.896 | 6 | 0.04648 |  |  | 0.33304
18 | 10.295 | 9.6423 | 9.6493 | 31.0521 | 31.042 | 6.5277 | 2.8345 | 4 | 4 | 151 | 6.4076 | 1.111 | 7.119 | 6 | 0.047787 |  |  | 0.39196
28 | 1.537 | 12.1412 | 12.1875 | 22.4079 | 22.1523 | 6.5312 | 3.2312 | 5 | 5 | 195 | 6.3936 | 1.1915 | 7.618 |  | 0.052878 |  |  | 0.36875
30 | 100.252 | 8.4302 | 8.4285 | 32.3594 | 32.3658 | 6.6496 | 1.6804 | 3 | 3 | 209 | 6.6106 | 0.58981 | 3.899 |  | 0.05354 |  |  | 0.35099
27 | 1.493 | 13.3661 | 13.2571 | 21.1019 | 21.7329 | 6.4149 | 3.3254 | 7 | 7 | 190 | 6.2769 | 1.2328 | 7.738 |  | 0.054598 |  |  | 0.37072
25 | 1.542 | 9.9349 | 9.9594 | 15.8868 | 15.9539 | 7.1504 | 2.9993 | 4 | 4 | 177 | 7.0093 | 1.1968 | 8.389 |  | 0.058355 |  |  | 0.39904
33 | 1.921 | 14.0122 | 14.028 | 18.9032 | 18.728 | 6.4151 | 3.0406 | 5 | 5 | 234 | 6.2764 | 1.2193 | 7.653 |  | 0.062013 |  |  | 0.40102
37 | 9.635 | 8.5377 | 8.5414 | 31.185 | 31.1841 | 6.6843 | 2.4704 | 6 | 6 | 260 | 6.5678 | 0.83239 | 5.467 |  | 0.064605 |  |  | 0.33694
34 | 9.516 | 8.7123 | 8.7118 | 31.1572 | 31.1593 | 6.6594 | 2.4729 | 6 | 6 | 240 | 6.542 | 0.82543 | 5.4 |  | 0.072677 |  |  | 0.3338
29 | 1.77 | 10.1391 | 10.058 | 28.0408 | 28.2109 | 6.5826 | 3.263 | 11 | 11 | 206 | 6.4519 | 1.1854 | 7.648 |  | 0.073518 |  |  | 0.36328
29 | 99.717 | 8.5174 | 8.5221 | 32.3933 | 32.395 | 6.6351 | 1.4569 | 5 | 5 | 200 | 6.5951 | 0.36042 | 2.377 |  | 0.073821 |  |  | 0.24738
29 | 10.595 | 8.746 | 8.7379 | 31.1524 | 31.1718 | 6.6546 | 2.102 | 10 | 10 | 205 | 6.538 | 0.65204 | 4.263 |  | 0.076718 |  |  | 0.3102
35 | 10.089 | 8.6841 | 8.6872 | 31.2112 | 31.2046 | 6.6613 | 2.4312 | 5 | 5 | 246 | 6.5446 | 0.79837 | 5.225 |  | 0.080694 |  |  | 0.32839
27 | 10.236 | 8.7337 | 8.7325 | 31.1678 | 31.1724 | 6.6558 | 1.9834 | 6 | 6 | 189 | 6.5389 | 0.58771 | 3.843 |  | 0.086921 |  |  | 0.29631
26 | 66.199 | 8.7417 | 8.7416 | 32.1022 | 32.1036 | 6.6144 | 1.006 | 1 | 1 | 178 | 6.5448 | 0.33034 | 2.162 |  | 0.102 |  |  | 0.32837
30 | 1.79 | 12.93 | 12.9545 | 23.2695 | 23.3439 | 6.3875 | 3.3952 | 9 | 9 | 215 | 6.2512 | 1.2068 | 7.544 |  | 0.11242 |  |  | 0.35545
31 | 10.118 | 8.7604 | 8.7636 | 31.0948 | 31.0902 | 6.6549 | 2.4251 | 7 | 7 | 222 | 6.5378 | 0.72027 | 4.709 |  | 0.15602 |  |  | 0.29701
25 | 10.267 | 8.9426 | 8.9458 | 31.0174 | 31.0108 | 6.6312 | 1.9673 | 3 | 3 | 176 | 6.5134 | 0.50219 | 3.271 |  | 0.16506 |  |  | 0.25527
32 | 10.263 | 8.7976 | 8.8064 | 31.0534 | 31.0393 | 6.6512 | 2.7156 | 5 | 5 | 228 | 6.534 | 0.79676 | 5.206 |  | 0.2122 |  |  | 0.2934
36 | 10.118 | 8.7116 | 8.7143 | 31.1482 | 31.1487 | 6.6599 | 3.169 | 6 | 6 | 253 | 6.543 | 0.82806 | 5.418 |  | 0.38789 |  |  | 0.2613
37 | 29.846 | 8.4687 | 8.4695 | 31.8171 | 31.824 | 6.6673 | 2.0757 | 4 | 4 | 258 | 6.5685 | 1.1115 | 7.301 | 4 | 0.39474 |  |  | 0.53548

## comparison-1438
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE2 | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
15 | 19.121 | 8.4897 | 8.4911 | 32.3799 | 32.382 | 6.6398 | 2.631 | 13 | 13 | 123 | 6.5877 | 0.99457 | 6.552 | 6 | 6.6339e-05 | 0.4672 | -0.2347
29 | 195.577 | 8.2184 | 8.2187 | 32.8426 | 32.843 | 6.6603 | 0.52914 | 1 | 1 | 196 | 6.7605 | 0.012425 | 0.084 |  | 8.9049e-05 |  |  | Summary of Soc Voffset including the original values in the factory calibration
30 | 140.607 | 8.2872 | 8.281 | 32.7958 | 32.801 | 6.6521 | 0.58045 | 1 | 1 | 207 | 6.7044 | 0.036394 | 0.244 | 6 | 9.1985e-05 |  |  |  |  | m | b | Soc | Voffset | R2
17 | 999.523 | 3.7027 | 3.7026 | 34.3668 | 34.3679 | 7.3285 | 0.56564 | 6 | 6 | 131 | 8.3 | 0.029398 | 0.244 |  | 0.00016932 |  |  | Bottles used | Original | 0.4453 | -0.22844 | 0.4538 | -0.513 |  | 2.2098
17 | 1499.92 | 2.548 | 2.5482 | 34.5161 | 34.5175 | 7.5321 | 0.67127 | 4 | 4 | 129 | 9.1254 | 0.07912 | 0.722 |  | 0.00020226 |  |  | 172 | all data | 0.4564 | -0.2179 | 0.4564 | -0.47743 | 0.9767 | 2.261
17 | 74.949 | 7.3564 | 7.3565 | 32.8381 | 32.8389 | 6.792 | 2.4087 | 17 | 17 | 142 | 6.7937 | 0.89038 | 6.049 |  | 0.00026637 |  |  | 165 | excl. outliers diff>0.1 | 0.4672 | -0.2347 | 0.4672 | -0.50235 | 0.9965 | 2.3008
18 | 48.411 | 8.5334 | 8.5304 | 32.3121 | 32.317 | 6.6363 | 2.1998 | 1 | 1 | 148 | 6.6087 | 0.79335 | 5.243 |  | 0.0003061 |  |  | 154 | excl. outliers diff>0.05 | 0.4681 | -0.2342 | 0.4681 | -0.50032 | 0.9987 | 2.3083
13 | 305.9 | 5.6029 | 5.6031 | 34.0126 | 34.0133 | 7.0192 | 0.87787 | 5 | 5 | 98 | 7.2463 | 0.17581 | 1.274 | 6 | 0.00037311 |  |  | 146 | excl. outliers diff>0.03 | 0.4692 | -0.2358 | 0.4692 | -0.50256 | 0.9994 | 2.3126
31 | 20.032 | 8.611 | 8.6112 | 31.5553 | 31.551 | 6.6573 | 1.8765 | 6 | 6 | 221 | 6.6052 | 0.64146 | 4.237 |  | 0.00053874 |  |  | 143 | excl. outliers diff>0.02 | 0.4691 | -0.2357 | 0.4691 | -0.50245 | 0.9995 | 2.3121
17 | 299.482 | 5.7297 | 5.73 | 33.956 | 33.9568 | 7.001 | 1.0146 | 10 | 10 | 135 | 7.2204 | 0.23988 | 1.732 |  | 0.00055528 |  |  | 131 | excl. outliers diff>0.015 | 0.4687 | -0.2355 | 0.4687 | -0.50245 | 0.9997 | 2.3097
36 | 20.124 | 8.4356 | 8.4309 | 31.6262 | 31.6184 | 6.6805 | 2.0005 | 5 | 5 | 252 | 6.6292 | 0.70053 | 4.644 |  | 0.00061741 |  |  | 114 | excl. outliers diff>0.01 | 0.4677 | -0.2337 | 0.4677 | -0.49968 | 0.9998 | 2.3065
13 | 9.904 | 10.0029 | 10.0053 | 30.962 | 30.9599 | 6.4799 | 2.8051 | 16 | 16 | 109 | 6.4143 | 1.0765 | 6.905 |  | 0.0006815
13 | 175.571 | 6.4911 | 6.4915 | 33.8993 | 33.8999 | 6.8804 | 1.3064 | 8 | 8 | 101 | 6.9776 | 0.37635 | 2.626 |  | 0.00069037
17 | 2000.077 | 1.8934 | 1.893 | 34.6002 | 34.6021 | 7.6521 | 0.81703 | 2 | 2 | 127 | 9.9119 | 0.14619 | 1.449 |  | 0.00082862
13 | 30.04 | 8.6904 | 8.6841 | 32.0207 | 32.0272 | 6.6255 | 2.4658 | 14 | 14 | 107 | 6.5816 | 0.91649 | 6.032 |  | 0.00084628
13 | 1012.795 | 3.6559 | 3.6561 | 34.3477 | 34.349 | 7.3378 | 0.58338 | 2 | 2 | 95 | 8.3255 | 0.038796 | 0.323 |  | 0.00094113
17 | 1748.799 | 2.2491 | 2.2482 | 34.5486 | 34.5509 | 7.5868 | 0.72176 | 3 | 3 | 128 | 9.5017 | 0.10346 | 0.983 |  | 0.00094939
13 | 104.407 | 7.3496 | 7.3495 | 33.7094 | 33.71 | 6.7544 | 1.5665 | 11 | 11 | 104 | 6.7817 | 0.4981 | 3.378 |  | 0.00095318
15 | 29.712 | 8.2428 | 8.2429 | 32.4322 | 32.4334 | 6.6745 | 2.5646 | 12 | 12 | 122 | 6.6324 | 0.9645 | 6.397 |  | 0.00099641
15 | 176.008 | 6.5398 | 6.54 | 33.8871 | 33.8877 | 6.8732 | 1.2157 | 6 | 6 | 116 | 6.9704 | 0.33427 | 2.33 |  | 0.001001
30 | 20.526 | 8.6444 | 8.6446 | 31.6156 | 31.6168 | 6.6498 | 1.8188 | 7 | 7 | 213 | 6.598 | 0.6161 | 4.065 |  | 0.0010661
22 | 150.646 | 6.6144 | 6.6141 | 33.919 | 33.92 | 6.8599 | 1.2051 | 4 | 4 | 163 | 6.9337 | 0.32941 | 2.284 |  | 0.0010693
22 | 125.224 | 7.1328 | 7.1331 | 33.8457 | 33.8464 | 6.7819 | 1.4725 | 5 | 5 | 164 | 6.8289 | 0.45439 | 3.103 |  | 0.0011191
17 | 1249.266 | 3.1439 | 3.1438 | 34.4424 | 34.4436 | 7.4256 | 0.59584 | 5 | 5 | 130 | 8.6967 | 0.042315 | 0.368 |  | 0.0013613
15 | 202.872 | 6.3882 | 6.388 | 33.9199 | 33.9206 | 6.8959 | 1.1954 | 5 | 5 | 115 | 7.0187 | 0.32527 | 2.283 |  | 0.0014673
13 | 1.73 | 10.1173 | 10.1173 | 30.6944 | 30.6953 | 6.4748 | 2.811 | 17 | 17 | 110 | 6.402 | 1.0801 | 6.915 |  | 0.0015156
31 | 50.43 | 8.5509 | 8.5525 | 32.1465 | 32.1433 | 6.6408 | 1.8288 | 4 | 4 | 219 | 6.6148 | 0.62149 | 4.111 |  | 0.0017767
17 | 500.01 | 4.7678 | 4.7679 | 34.0958 | 34.0965 | 7.1557 | 0.68768 | 8 | 8 | 133 | 7.583 | 0.084795 | 0.643 | 6 | 0.0017895
22 | 29.719 | 8.7651 | 8.7641 | 32.3788 | 32.3785 | 6.5991 | 2.682 | 9 | 9 | 168 | 6.5548 | 1.0165 | 6.663 |  | 0.0017974
15 | 75.915 | 7.6946 | 7.6942 | 33.0195 | 33.0219 | 6.7319 | 2.0194 | 10 | 10 | 120 | 6.7324 | 0.7106 | 4.784 |  | 0.0018421
32 | 50.368 | 8.5326 | 8.5332 | 32.1581 | 32.157 | 6.643 | 1.8398 | 2 | 2 | 225 | 6.617 | 0.62672 | 4.147 |  | 0.0018473
13 | 506.822 | 5.0146 | 5.0148 | 34.0893 | 34.0901 | 7.114 | 0.72056 | 4 | 4 | 97 | 7.543 | 0.1038 | 0.783 |  | 0.0018593
17 | 249.723 | 6.0473 | 6.0482 | 33.9061 | 33.9068 | 6.9514 | 1.2584 | 11 | 11 | 136 | 7.1208 | 0.35516 | 2.529 |  | 0.0019225
13 | 127.108 | 7.0502 | 7.0493 | 33.8244 | 33.825 | 6.7957 | 1.5025 | 10 | 10 | 103 | 6.845 | 0.4653 | 3.185 |  | 0.0019534
30 | 74.791 | 8.4969 | 8.4975 | 32.2877 | 32.2879 | 6.6427 | 1.7629 | 4 | 4 | 210 | 6.6377 | 0.59087 | 3.922 |  | 0.0019583
31 | 75.263 | 8.4549 | 8.4551 | 32.2915 | 32.2918 | 6.6488 | 1.8014 | 3 | 3 | 218 | 6.6444 | 0.60487 | 4.019 |  | 0.0020416
15 | 151.015 | 6.8913 | 6.8913 | 33.828 | 33.8287 | 6.8203 | 1.3753 | 7 | 7 | 117 | 6.8921 | 0.40989 | 2.825 |  | 0.0020578
27 | 20.385 | 8.7086 | 8.7091 | 31.6222 | 31.6222 | 6.6399 | 1.7246 | 5 | 5 | 188 | 6.5877 | 0.56894 | 3.748 |  | 0.0020738
20 | 76.366 | 7.8398 | 7.8411 | 32.9242 | 32.9236 | 6.7139 | 2.1037 | 2 | 2 | 154 | 6.7139 | 0.75023 | 5.037 |  | 0.0020999
11 | 30.095 | 8.3946 | 8.396 | 32.3146 | 32.3139 | 6.6568 | 2.434 | 4 | 4 | 90 | 6.6144 | 0.90469 | 5.984 | 6 | 0.0022429
11 | 21.039 | 8.4559 | 8.4561 | 32.2424 | 32.2431 | 6.6508 | 2.4157 | 5 | 5 | 91 | 6.6004 | 0.89615 | 5.915 |  | 0.0022486
22 | 100.289 | 7.4071 | 7.4079 | 33.6928 | 33.6931 | 6.7463 | 1.577 | 6 | 6 | 165 | 6.7696 | 0.50446 | 3.415 |  | 0.0023844
29 | 29.166 | 8.6516 | 8.649 | 31.8981 | 31.9077 | 6.6365 | 1.8115 | 8 | 8 | 203 | 6.5921 | 0.61407 | 4.048 | 6 | 0.0024261
29 | 175.079 | 8.2321 | 8.2322 | 32.8295 | 32.8299 | 6.6589 | 0.561 | 2 | 2 | 197 | 6.7412 | 0.024921 | 0.168 |  | 0.0024778
17 | 2.183 | 10.0039 | 10.0038 | 31.1105 | 31.1114 | 6.4736 | 2.7875 | 22 | 22 | 147 | 6.4017 | 1.0702 | 6.851 |  | 0.0025587
34 | 75.095 | 8.1697 | 8.1692 | 32.5816 | 32.5832 | 6.679 | 1.5881 | 2 | 2 | 236 | 6.676 | 0.50464 | 3.369 |  | 0.0026153
11 | 50.556 | 7.7379 | 7.7365 | 32.9406 | 32.9426 | 6.7288 | 2.0566 | 3 | 3 | 89 | 6.7071 | 0.72878 | 4.888 |  | 0.0026511
35 | 19.969 | 8.5481 | 8.5457 | 31.6869 | 31.6696 | 6.661 | 1.9277 | 4 | 4 | 245 | 6.6092 | 0.66861 | 4.419 |  | 0.0026927
17 | 150.645 | 6.9485 | 6.949 | 33.7734 | 33.7741 | 6.8138 | 1.745 | 14 | 14 | 139 | 6.8848 | 0.58331 | 4.016 |  | 0.0027391
17 | 2083.101 | 1.846 | 1.8453 | 34.6098 | 34.6119 | 7.6607 | 0.83167 | 1 | 1 | 126 | 10.0327 | 0.15679 | 1.573 |  | 0.0029314
22 | 175.255 | 6.3658 | 6.3658 | 33.9576 | 33.9584 | 6.8977 | 1.094 | 3 | 3 | 162 | 6.9959 | 0.27945 | 1.955 | 6 | 0.0030287
15 | 305.476 | 5.8722 | 5.8725 | 34.036 | 34.0367 | 6.9739 | 0.83603 | 3 | 3 | 113 | 7.1968 | 0.15896 | 1.144 |  | 0.0030654
13 | 1152.059 | 3.408 | 3.4076 | 34.3869 | 34.3884 | 7.3804 | 0.59072 | 1 | 1 | 94 | 8.5308 | 0.044427 | 0.379 |  | 0.003143
17 | 29.562 | 8.6969 | 8.6919 | 32.4821 | 32.4828 | 6.6048 | 2.7692 | 19 | 19 | 144 | 6.5606 | 1.0623 | 6.969 |  | 0.0031944
37 | 96.14 | 7.9086 | 7.9199 | 32.6427 | 32.6405 | 6.7158 | 1.6403 | 1 | 1 | 255 | 6.7324 | 0.53487 | 3.601 |  | 0.0032243
13 | 76.789 | 7.5895 | 7.5895 | 33.2853 | 33.2858 | 6.7363 | 1.925 | 12 | 12 | 105 | 6.7381 | 0.66799 | 4.501 |  | 0.0033234
32 | 30.459 | 8.5775 | 8.578 | 31.7982 | 31.7937 | 6.6518 | 1.8907 | 3 | 3 | 226 | 6.6088 | 0.65201 | 4.309 | 6 | 0.0033781
17 | 125.685 | 7.1734 | 7.1735 | 33.6209 | 33.6223 | 6.7856 | 1.822 | 15 | 15 | 140 | 6.8328 | 0.61995 | 4.236 |  | 0.0034105
30 | 30.16 | 8.6506 | 8.6508 | 31.7889 | 31.789 | 6.6414 | 1.8194 | 6 | 6 | 212 | 6.5977 | 0.61885 | 4.083 |  | 0.003546
22 | 75.652 | 7.725 | 7.7253 | 33.206 | 33.2064 | 6.7191 | 1.9165 | 7 | 7 | 166 | 6.7191 | 0.66422 | 4.463 |  | 0.0035489
20 | 119.131 | 7.1761 | 7.176 | 33.6833 | 33.6844 | 6.7824 | 1.4955 | 1 | 1 | 153 | 6.8238 | 0.46777 | 3.192 |  | 0.0037785
13 | 256.433 | 5.7595 | 5.7592 | 33.9812 | 33.982 | 6.9949 | 0.96939 | 6 | 6 | 99 | 7.1739 | 0.22205 | 1.593 |  | 0.0038556
34 | 19.979 | 8.5479 | 8.5489 | 31.619 | 31.6184 | 6.664 | 1.9317 | 5 | 5 | 239 | 6.6121 | 0.67164 | 4.441 |  | 0.0038578
28 | 20.469 | 8.6762 | 8.6762 | 31.6121 | 31.6079 | 6.6452 | 1.7588 | 3 | 3 | 193 | 6.5932 | 0.59091 | 3.896 |  | 0.0039252
29 | 19.896 | 8.6583 | 8.6586 | 31.5649 | 31.5653 | 6.6499 | 1.7857 | 9 | 9 | 204 | 6.5975 | 0.60357 | 3.982 |  | 0.0039727
17 | 749.962 | 4.1201 | 4.1205 | 34.2569 | 34.2577 | 7.26 | 0.57306 | 7 | 7 | 132 | 7.9547 | 0.037085 | 0.295 |  | 0.0040515
20 | 30.527 | 8.6484 | 8.6487 | 32.3528 | 32.3535 | 6.6174 | 2.4704 | 4 | 4 | 156 | 6.5743 | 0.9236 | 6.072 |  | 0.0041207
17 | 9.747 | 10.0151 | 10.0222 | 31.2387 | 31.2044 | 6.4667 | 2.7873 | 21 | 21 | 146 | 6.401 | 1.0717 | 6.86 |  | 0.0041599
22 | 49.312 | 7.8738 | 7.8713 | 32.5804 | 32.5785 | 6.7238 | 2.4963 | 8 | 8 | 167 | 6.7004 | 0.93577 | 6.27 |  | 0.0041826
22 | 20.253 | 9.0014 | 9.0036 | 32.096 | 32.0963 | 6.5765 | 2.7321 | 10 | 10 | 169 | 6.5232 | 1.0461 | 6.824 | 6 | 0.0043603 | 0.4453 | -0.2284
27 | 29.42 | 8.7056 | 8.7059 | 31.8094 | 31.8093 | 6.6323 | 1.7136 | 4 | 4 | 187 | 6.5878 | 0.56149 | 3.699 | 6 | 0.0043976
15 | 758.615 | 4.3093 | 4.3093 | 34.2359 | 34.2373 | 7.228 | 0.59135 | 1 | 1 | 111 | 7.9261 | 0.04605 | 0.365 |  | 0.0044716
33 | 30.03 | 8.564 | 8.564 | 31.7524 | 31.7528 | 6.6558 | 1.918 | 2 | 2 | 231 | 6.6124 | 0.66602 | 4.404 |  | 0.0046152
26 | 19.773 | 8.7552 | 8.7554 | 31.6765 | 31.6841 | 6.6307 | 1.5877 | 4 | 4 | 181 | 6.5778 | 0.50229 | 3.304 |  | 0.0047601
15 | 126.292 | 7.2364 | 7.2362 | 33.7304 | 33.7309 | 6.7709 | 1.5118 | 8 | 8 | 118 | 6.8182 | 0.46684 | 3.183 |  | 0.0047648
22 | 200.063 | 6.2078 | 6.2084 | 33.9767 | 33.9774 | 6.9222 | 1.0278 | 2 | 2 | 161 | 7.0444 | 0.25027 | 1.763 |  | 0.0047818
29 | 74.757 | 8.5238 | 8.5242 | 32.3115 | 32.3104 | 6.6377 | 1.6712 | 6 | 6 | 201 | 6.6325 | 0.55093 | 3.654 |  | 0.004838
34 | 30.02 | 8.5567 | 8.5569 | 31.868 | 31.8622 | 6.6519 | 1.9104 | 4 | 4 | 238 | 6.6086 | 0.65294 | 4.315 |  | 0.0048823
13 | 760.488 | 4.3008 | 4.3005 | 34.2273 | 34.2284 | 7.2299 | 0.5947 | 3 | 3 | 96 | 7.9302 | 0.048044 | 0.381 |  | 0.0049002
30 | 50.348 | 8.5466 | 8.5455 | 32.1589 | 32.1627 | 6.6409 | 1.8241 | 5 | 5 | 211 | 6.6148 | 0.61256 | 4.052 |  | 0.0049443
17 | 201.061 | 6.4737 | 6.4732 | 33.8578 | 33.8586 | 6.885 | 1.5357 | 12 | 12 | 137 | 7.0055 | 0.48776 | 3.417 |  | 0.0049838
11 | 96.136 | 7.1012 | 7.1017 | 33.727 | 33.7279 | 6.7921 | 1.405 | 1 | 1 | 87 | 6.8139 | 0.42692 | 2.909 |  | 0.0051995
13 | 153.036 | 6.7436 | 6.743 | 33.8678 | 33.8685 | 6.8418 | 1.4159 | 9 | 9 | 102 | 6.9166 | 0.43215 | 2.989 |  | 0.0053453
36 | 30.004 | 8.4873 | 8.4887 | 31.8223 | 31.8276 | 6.6643 | 2.0165 | 4 | 4 | 251 | 6.6212 | 0.70199 | 4.648 |  | 0.0054025
28 | 29.901 | 8.6626 | 8.6629 | 31.88 | 31.8807 | 6.6357 | 1.7951 | 2 | 2 | 192 | 6.5918 | 0.6094 | 4.017 |  | 0.005422
11 | 2.334 | 9.6184 | 9.6246 | 31.1604 | 31.156 | 6.5266 | 2.6642 | 7 | 7 | 93 | 6.456 | 1.0155 | 6.556 |  | 0.0054755
26 | 9.598 | 8.7819 | 8.7743 | 31.2518 | 31.291 | 6.645 | 1.5611 | 5 | 5 | 182 | 6.5833 | 0.48912 | 3.22 |  | 0.0055195
11 | 10.491 | 9.0162 | 9.0208 | 31.6954 | 31.6892 | 6.5914 | 2.5292 | 6 | 6 | 92 | 6.5298 | 0.95256 | 6.22 |  | 0.0056131
17 | 174.866 | 6.7715 | 6.7719 | 33.8302 | 33.8307 | 6.8391 | 1.6063 | 13 | 13 | 138 | 6.9331 | 0.52141 | 3.615 | 6 | 0.0056665
15 | 2.559 | 10.0624 | 10.0626 | 31.0924 | 31.0916 | 6.4661 | 2.7743 | 15 | 15 | 125 | 6.3943 | 1.0672 | 6.824 |  | 0.005751
15 | 9.651 | 9.6996 | 9.7 | 31.2954 | 31.2955 | 6.5093 | 2.7645 | 14 | 14 | 124 | 6.4445 | 1.0511 | 6.774 |  | 0.0057593
31 | 30.436 | 8.5764 | 8.5758 | 31.7925 | 31.7885 | 6.6523 | 1.8944 | 5 | 5 | 220 | 6.6092 | 0.64441 | 4.259 |  | 0.0059453
17 | 400.41 | 5.268 | 5.2687 | 34.0402 | 34.041 | 7.0736 | 0.78213 | 9 | 9 | 134 | 7.3952 | 0.13671 | 1.011 |  | 0.0059984
17 | 20.388 | 9.3644 | 9.3658 | 32.3198 | 32.3207 | 6.5144 | 2.8033 | 20 | 20 | 145 | 6.46 | 1.0816 | 6.987 |  | 0.0065687
13 | 203.95 | 6.1927 | 6.1912 | 33.9363 | 33.9371 | 6.9265 | 1.1517 | 7 | 7 | 100 | 7.0524 | 0.31025 | 2.188 |  | 0.0068983
37 | 1.563 | 13.1632 | 13.2078 | 21.667 | 21.5836 | 6.4199 | 2.9675 | 7 | 7 | 261 | 6.3359 | 1.1588 | 7.342 |  | 0.0070742
37 | 20.491 | 8.4507 | 8.4555 | 31.583 | 31.5818 | 6.6801 | 2.054 | 5 | 5 | 259 | 6.6291 | 0.71775 | 4.758 |  | 0.0071736
28 | 53.005 | 8.5421 | 8.5425 | 32.2224 | 32.2229 | 6.6388 | 1.74 | 1 | 1 | 191 | 6.6151 | 0.58548 | 3.873 |  | 0.0072498
35 | 29.464 | 8.529 | 8.5291 | 31.8694 | 31.8559 | 6.656 | 1.9461 | 3 | 3 | 244 | 6.6123 | 0.66724 | 4.412 |  | 0.0072685
22 | 1.801 | 11.4499 | 11.4497 | 30.7154 | 30.7157 | 6.2902 | 2.8115 | 12 | 12 | 171 | 6.2142 | 1.0862 | 6.75 |  | 0.0073803
32 | 19.886 | 8.579 | 8.5794 | 31.505 | 31.5045 | 6.6643 | 1.9637 | 4 | 4 | 227 | 6.6122 | 0.67527 | 4.465 |  | 0.0074851
17 | 100.019 | 7.3866 | 7.3869 | 33.3718 | 33.374 | 6.7636 | 1.9428 | 16 | 16 | 141 | 6.787 | 0.68057 | 4.619 |  | 0.0076071
36 | 50.583 | 8.443 | 8.4451 | 32.3237 | 32.32 | 6.6492 | 1.9206 | 3 | 3 | 250 | 6.6239 | 0.6549 | 4.338 |  | 0.0076889
13 | 20.017 | 8.9129 | 8.9091 | 31.8195 | 31.823 | 6.6013 | 2.5087 | 15 | 15 | 108 | 6.548 | 0.94517 | 6.189 |  | 0.007789
22 | 9.82 | 10.1767 | 10.2142 | 31.6317 | 31.5928 | 6.4278 | 2.8051 | 11 | 11 | 170 | 6.3618 | 1.0837 | 6.894 |  | 0.0078195
20 | 10.732 | 10.2851 | 10.2847 | 31.2678 | 31.2694 | 6.4276 | 2.8286 | 6 | 6 | 158 | 6.3619 | 1.0951 | 6.967 |  | 0.0082826
29 | 50.062 | 8.5396 | 8.5396 | 32.1799 | 32.1808 | 6.641 | 1.8034 | 7 | 7 | 202 | 6.6148 | 0.59911 | 3.963 |  | 0.0087202
35 | 79.461 | 8.1145 | 8.1136 | 32.6066 | 32.6078 | 6.6862 | 1.64 | 1 | 1 | 242 | 6.6873 | 0.54028 | 3.613 | 6 | 0.0087583
15 | 251.925 | 5.9881 | 5.9885 | 33.935 | 33.9358 | 6.9597 | 1.145 | 4 | 4 | 114 | 7.1318 | 0.30904 | 2.204 |  | 0.008814
32 | 82.233 | 8.3686 | 8.3681 | 32.3933 | 32.3949 | 6.6573 | 1.7635 | 1 | 1 | 224 | 6.6593 | 0.59826 | 3.984 |  | 0.0090516
18 | 30.342 | 8.6779 | 8.6782 | 32.0106 | 32.014 | 6.6278 | 2.1912 | 2 | 2 | 149 | 6.5842 | 0.79812 | 5.255 |  | 0.0091069
36 | 75.352 | 8.1357 | 8.1388 | 32.5375 | 32.5367 | 6.686 | 1.7343 | 2 | 2 | 249 | 6.6834 | 0.56588 | 3.782 | 6 | 0.0096702
15 | 100.871 | 7.4643 | 7.4636 | 33.4486 | 33.4512 | 6.7483 | 1.6533 | 9 | 9 | 119 | 6.7718 | 0.54742 | 3.707 |  | 0.0096912
22 | 266.655 | 6.1402 | 6.1405 | 34.0034 | 34.0045 | 6.9319 | 0.93716 | 1 | 1 | 160 | 7.1156 | 0.21305 | 1.516 |  | 0.0099125
30 | 124.703 | 8.3592 | 8.3585 | 32.7431 | 32.7434 | 6.6436 | 0.61356 | 2 | 2 | 208 | 6.6818 | 0.041905 | 0.28 |  | 0.01005
18 | 1.743 | 11.0185 | 11.0264 | 29.8535 | 29.8518 | 6.3835 | 3.1607 | 5 | 5 | 152 | 6.308 | 1.2524 | 7.9 |  | 0.010396
11 | 76.519 | 7.3555 | 7.3557 | 33.5042 | 33.505 | 6.7626 | 1.579 | 2 | 2 | 88 | 6.7656 | 0.51348 | 3.474 |  | 0.010463
17 | 50.599 | 7.6571 | 7.6564 | 32.5912 | 32.5914 | 6.7565 | 2.6483 | 18 | 18 | 143 | 6.7353 | 1.0132 | 6.824 | 4 | 0.010581
28 | 10.132 | 8.7373 | 8.7333 | 31.1651 | 31.168 | 6.6553 | 2.0579 | 4 | 4 | 194 | 6.5943 | 0.71577 | 4.72 | 6 | 0.010958
29 | 150.285 | 8.2495 | 8.2497 | 32.8188 | 32.8191 | 6.6567 | 0.56423 | 3 | 3 | 198 | 6.7176 | 0.039895 | 0.268 |  | 0.010987
36 | 1.473 | 13.4518 | 13.4826 | 21.4115 | 21.2949 | 6.391 | 3.0864 | 7 | 7 | 254 | 6.3063 | 1.2186 | 7.685 |  | 0.011352
27 | 74.661 | 8.5231 | 8.5273 | 32.4852 | 32.4836 | 6.6303 | 0.52717 | 2 | 2 | 185 | 6.625 | 0 | 0 |  | 0.011594
29 | 124.583 | 8.3595 | 8.3597 | 32.7582 | 32.7585 | 6.6429 | 0.61677 | 4 | 4 | 199 | 6.681 | 0.06541 | 0.437 |  | 0.011955
34 | 95.639 | 8.1404 | 8.1407 | 32.605 | 32.6057 | 6.6824 | 1.5393 | 1 | 1 | 235 | 6.6971 | 0.49648 | 3.325 |  | 0.01202
13 | 51.62 | 7.7942 | 7.7966 | 32.6607 | 32.6603 | 6.7325 | 2.4093 | 13 | 13 | 106 | 6.7114 | 0.8788 | 5.898 | 6 | 0.012114
27 | 115.198 | 8.145 | 8.1443 | 32.5718 | 32.5728 | 6.6831 | 0.52977 | 1 | 1 | 184 | 6.7146 | 0 | 0 |  | 0.012809
20 | 2.446 | 11.3801 | 11.3663 | 30.6737 | 30.682 | 6.3013 | 2.8735 | 7 | 7 | 159 | 6.2259 | 1.1211 | 6.98 |  | 0.013298
37 | 75.04 | 8.1089 | 8.1089 | 32.5342 | 32.5336 | 6.6902 | 1.7253 | 2 | 2 | 256 | 6.6875 | 0.55791 | 3.731 |  | 0.013454
27 | 50.121 | 8.6634 | 8.6626 | 31.9703 | 31.9686 | 6.6316 | 1.644 | 3 | 3 | 186 | 6.6048 | 0.51962 | 3.432 |  | 0.013737
20 | 21.078 | 9.1224 | 9.118 | 32.1173 | 32.119 | 6.558 | 2.6107 | 5 | 5 | 157 | 6.5049 | 0.99924 | 6.5 |  | 0.014202
15 | 49.982 | 7.8524 | 7.8526 | 32.6293 | 32.63 | 6.725 | 2.4255 | 11 | 11 | 121 | 6.7022 | 0.91328 | 6.121 |  | 0.014794
32 | 1.76 | 13.9418 | 13.9793 | 20.9459 | 20.8335 | 6.3438 | 3.2008 | 6 | 6 | 229 | 6.2584 | 1.2759 | 7.985 |  | 0.015184
25 | 32.411 | 8.7666 | 8.7683 | 31.8559 | 31.854 | 6.6213 | 1.3606 | 1 | 1 | 174 | 6.579 | 0.41632 | 2.739 |  | 0.015335
36 | 85.961 | 8.0032 | 8.0024 | 32.6518 | 32.651 | 6.7011 | 1.5916 | 1 | 1 | 248 | 6.7083 | 0.52457 | 3.519 |  | 0.01568
20 | 51.124 | 8.2841 | 8.2807 | 32.5 | 32.502 | 6.6654 | 2.4733 | 3 | 3 | 155 | 6.6413 | 0.93716 | 6.224 |  | 0.016357
26 | 49.437 | 8.7226 | 8.7236 | 31.9826 | 31.9834 | 6.6224 | 1.3913 | 2 | 2 | 179 | 6.5947 | 0.43247 | 2.852 |  | 0.017159
26 | 1.405 | 11.2622 | 11.2741 | 12.6875 | 12.2494 | 7.0792 | 2.8804 | 6 | 6 | 183 | 6.9941 | 1.1284 | 7.892 |  | 0.017382
31 | 100.436 | 8.4036 | 8.405 | 32.3691 | 32.3654 | 6.6531 | 1.714 | 2 | 2 | 217 | 6.6704 | 0.54869 | 3.66 | 6 | 0.017384
18 | 20.562 | 8.9528 | 8.9576 | 31.5184 | 31.5161 | 6.6083 | 2.2926 | 3 | 3 | 150 | 6.5552 | 0.85428 | 5.6 |  | 0.017886
34 | 50.338 | 8.3484 | 8.3469 | 32.4316 | 32.433 | 6.6587 | 1.8064 | 3 | 3 | 237 | 6.6336 | 0.59108 | 3.921 |  | 0.018196
34 | 1.507 | 13.6213 | 13.5676 | 18.5884 | 18.7547 | 6.4808 | 2.9695 | 7 | 7 | 241 | 6.3944 | 1.1709 | 7.487 | 6 | 0.018215
15 | 503.637 | 4.8372 | 4.8376 | 34.1381 | 34.1391 | 7.1418 | 0.65337 | 2 | 2 | 112 | 7.5712 | 0.089154 | 0.675 | 3 | 0.0186
35 | 50.646 | 8.4268 | 8.4271 | 32.348 | 32.3485 | 6.6506 | 1.8845 | 2 | 2 | 243 | 6.6254 | 0.62608 | 4.148 |  | 0.019663
25 | 20.038 | 8.8069 | 8.8069 | 31.6315 | 31.6272 | 6.6249 | 1.2838 | 2 | 2 | 175 | 6.5721 | 0.33962 | 2.232 | 6 | 0.025478
33 | 59.856 | 8.2852 | 8.2769 | 32.4894 | 32.4953 | 6.6657 | 1.7566 | 1 | 1 | 230 | 6.649 | 0.61227 | 4.071 |  | 0.026283
26 | 30.624 | 8.7497 | 8.7536 | 31.8371 | 31.8355 | 6.6246 | 1.5421 | 3 | 3 | 180 | 6.5809 | 0.51209 | 3.37 | 6 | 0.026304
31 | 113.721 | 8.3959 | 8.3967 | 32.4084 | 32.4099 | 6.6526 | 1.5804 | 1 | 1 | 216 | 6.6812 | 0.53418 | 3.569 |  | 0.030501
30 | 10.611 | 8.639 | 8.6409 | 31.3223 | 31.3152 | 6.6632 | 1.919 | 8 | 8 | 214 | 6.603 | 0.69408 | 4.583 |  | 0.032239
30 | 100.252 | 8.4302 | 8.4285 | 32.3594 | 32.3658 | 6.6496 | 1.6825 | 3 | 3 | 209 | 6.6665 | 0.58486 | 3.899 |  | 0.03349
37 | 49.987 | 8.5012 | 8.4973 | 32.2483 | 32.25 | 6.6438 | 1.9919 | 3 | 3 | 257 | 6.6177 | 0.66111 | 4.375 | 6 | 0.034807
33 | 10.015 | 8.649 | 8.6521 | 31.2914 | 31.288 | 6.6631 | 2.1695 | 4 | 4 | 233 | 6.6023 | 0.74156 | 4.896 | 6 | 0.037311
35 | 1.609 | 13.2721 | 13.268 | 17.5061 | 17.5065 | 6.5731 | 2.9444 | 6 | 6 | 247 | 6.4868 | 1.1801 | 7.655 |  | 0.039156
31 | 1.575 | 13.4468 | 13.5551 | 16.9974 | 16.6902 | 6.5695 | 3.0786 | 8 | 8 | 223 | 6.4825 | 1.1577 | 7.505 |  | 0.045906
37 | 9.635 | 8.5377 | 8.5414 | 31.185 | 31.1841 | 6.6843 | 2.3726 | 6 | 6 | 260 | 6.6236 | 0.82538 | 5.467 |  | 0.048397
27 | 1.493 | 13.3661 | 13.2571 | 21.1019 | 21.7329 | 6.4149 | 3.2262 | 7 | 7 | 190 | 6.3302 | 1.2224 | 7.738 |  | 0.050206
18 | 10.295 | 9.6423 | 9.6493 | 31.0521 | 31.042 | 6.5277 | 2.7494 | 4 | 4 | 151 | 6.4635 | 1.1014 | 7.119 | 6 | 0.051591
25 | 1.542 | 9.9349 | 9.9594 | 15.8868 | 15.9539 | 7.1504 | 2.9298 | 4 | 4 | 177 | 7.0708 | 1.1864 | 8.389 |  | 0.052339
33 | 1.921 | 14.0122 | 14.028 | 18.9032 | 18.728 | 6.4151 | 2.9742 | 5 | 5 | 234 | 6.3286 | 1.2093 | 7.653 |  | 0.054426
28 | 1.537 | 12.1412 | 12.1875 | 22.4079 | 22.1523 | 6.5312 | 3.1626 | 5 | 5 | 195 | 6.4494 | 1.1812 | 7.618 |  | 0.061658
35 | 10.089 | 8.6841 | 8.6872 | 31.2112 | 31.2046 | 6.6613 | 2.3378 | 5 | 5 | 246 | 6.6004 | 0.79162 | 5.225 |  | 0.065894
34 | 9.516 | 8.7123 | 8.7118 | 31.1572 | 31.1593 | 6.6594 | 2.4 | 6 | 6 | 240 | 6.5979 | 0.81844 | 5.4 |  | 0.068118
29 | 99.717 | 8.5174 | 8.5221 | 32.3933 | 32.395 | 6.6351 | 1.4385 | 5 | 5 | 200 | 6.6511 | 0.35739 | 2.377 |  | 0.079968
29 | 10.595 | 8.746 | 8.7379 | 31.1524 | 31.1718 | 6.6546 | 2.0626 | 10 | 10 | 205 | 6.5938 | 0.64651 | 4.263 |  | 0.082416
27 | 10.236 | 8.7337 | 8.7325 | 31.1678 | 31.1724 | 6.6558 | 1.9319 | 6 | 6 | 189 | 6.5948 | 0.58273 | 3.843 |  | 0.08513
29 | 1.77 | 10.1391 | 10.058 | 28.0408 | 28.2109 | 6.5826 | 3.2293 | 11 | 11 | 206 | 6.5085 | 1.1751 | 7.648 |  | 0.09895
26 | 66.199 | 8.7417 | 8.7416 | 32.1022 | 32.1036 | 6.6144 | 0.95982 | 1 | 1 | 178 | 6.6008 | 0.32754 | 2.162 |  | 0.11381
30 | 1.79 | 12.93 | 12.9545 | 23.2695 | 23.3439 | 6.3875 | 3.3132 | 9 | 9 | 215 | 6.3049 | 1.1965 | 7.544 |  | 0.11668
25 | 10.267 | 8.9426 | 8.9458 | 31.0174 | 31.0108 | 6.6312 | 1.8761 | 3 | 3 | 176 | 6.5694 | 0.49791 | 3.271 |  | 0.1439
31 | 10.118 | 8.7604 | 8.7636 | 31.0948 | 31.0902 | 6.6549 | 2.3434 | 7 | 7 | 222 | 6.5937 | 0.71417 | 4.709 |  | 0.14596
32 | 10.263 | 8.7976 | 8.8064 | 31.0534 | 31.0393 | 6.6512 | 2.6077 | 5 | 5 | 228 | 6.5899 | 0.78999 | 5.206 |  | 0.19363
36 | 10.118 | 8.7116 | 8.7143 | 31.1482 | 31.1487 | 6.6599 | 3.0447 | 6 | 6 | 253 | 6.5989 | 0.82105 | 5.418 |  | 0.36673
37 | 29.846 | 8.4687 | 8.4695 | 31.8171 | 31.824 | 6.6673 | 2.0146 | 4 | 4 | 258 | 6.6242 | 1.1022 | 7.301 | 4 | 0.39564

## bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number
1 | 70.674 | 8.1816 | 8.1815 | 30.2038 | 30.2044 | 6.7813 | 2.121 | 1 | 1 | 1 | 5.08
1 | 50.401 | 8.1778 | 8.1788 | 29.9779 | 29.9809 | 6.7919 | 2.2014 | 2 | 2 | 2 | 5.239
1 | 30.248 | 7.6952 | 7.7016 | 29.4961 | 29.4995 | 6.8886 | 2.0619 | 3 | 3 | 3 | 5.259
1 | 20.549 | 7.6952 | 7.6966 | 29.3585 | 29.3598 | 6.8948 | 2.0673 | 4 | 4 | 4 | 5.138
1 | 9.96 | 8.25 | 8.2727 | 29.3162 | 29.318 | 6.81 | 2.4282 | 5 | 5 | 5 | 5.85
1 | 2.032 | 9.5219 | 9.1269 | 29.0635 | 29.1426 | 6.6295 | 3.0346 | 6 | 6 | 6 | 8.429
3 | 238.009 | 8.1347 | 8.1348 | 29.9869 | 29.9869 | 6.7982 | 2.2692 | 1 | 1 | 16 | 5.848
3 | 238.043 | 8.1321 | 8.133 | 29.9893 | 29.9886 | 6.7984 | 2.2689 | 2 | 2 | 17 | 5.844 | 6
3 | 202.269 | 8.1418 | 8.142 | 29.9703 | 29.9701 | 6.7978 | 2.2734 | 3 | 3 | 18 | 5.891
3 | 176.325 | 8.142 | 8.1422 | 29.9582 | 29.9577 | 6.7983 | 2.2776 | 4 | 4 | 19 | 5.901
3 | 151.53 | 8.1384 | 8.1384 | 29.9557 | 29.9553 | 6.799 | 2.2864 | 5 | 5 | 20 | 5.846
3 | 126.505 | 8.1426 | 8.1429 | 29.9277 | 29.9271 | 6.7995 | 2.2905 | 6 | 6 | 21 | 5.846
3 | 102.024 | 8.1484 | 8.1485 | 29.8928 | 29.8931 | 6.8002 | 2.2949 | 7 | 7 | 22 | 5.844
3 | 75.93 | 8.1537 | 8.1537 | 29.8493 | 29.8493 | 6.8013 | 2.2936 | 8 | 8 | 23 | 5.892
3 | 50.708 | 8.1508 | 8.1508 | 29.7208 | 29.7202 | 6.8075 | 2.267 | 9 | 9 | 24 | 5.868
3 | 30.723 | 8.1593 | 8.1595 | 29.6405 | 29.6403 | 6.8097 | 2.2306 | 10 | 10 | 25 | 5.707
3 | 19.812 | 8.1884 | 8.1896 | 29.5799 | 29.5779 | 6.8079 | 2.2401 | 11 | 11 | 26 | 5.803
3 | 9.711 | 8.2923 | 8.2941 | 29.3468 | 29.3429 | 6.8022 | 2.2656 | 12 | 12 | 27 | 5.832
3 | 2.272 | 8.8032 | 8.8217 | 27.327 | 27.2405 | 6.8127 | 2.4279 | 13 | 13 | 28 | 6.328
4 | 41.159 | 8.5105 | 8.5108 | 29.2576 | 29.2572 | 6.7726 | 2.2895 | 1 | 1 | 29 | 5.612
4 | 29.657 | 8.509 | 8.5091 | 29.2188 | 29.2187 | 6.7745 | 2.2751 | 2 | 2 | 30 | 5.654 | 6
4 | 20.388 | 8.5166 | 8.518 | 29.1606 | 29.1597 | 6.7759 | 2.3124 | 3 | 3 | 31 | 5.785
4 | 9.774 | 8.5702 | 8.5701 | 29.035 | 29.0346 | 6.7732 | 2.4676 | 4 | 4 | 32 | 6.104
4 | 2.043 | 8.7337 | 8.7323 | 28.6277 | 28.6421 | 6.7661 | 2.9415 | 5 | 5 | 33 | 7.965
5 | 95.934 | 8.4986 | 8.4986 | 29.5235 | 29.5245 | 6.7627 | 2.3684 | 1 | 1 | 34 | 5.668
5 | 76.551 | 8.5018 | 8.5019 | 29.4508 | 29.4506 | 6.7654 | 2.2628 | 2 | 2 | 35 | 5.653
5 | 50.651 | 8.564 | 8.5633 | 29.3348 | 29.3351 | 6.761 | 2.2258 | 3 | 3 | 36 | 5.569
5 | 30.63 | 8.5645 | 8.5647 | 29.2888 | 29.2882 | 6.763 | 2.2151 | 4 | 4 | 37 | 5.578 | 6
5 | 20.26 | 8.5641 | 8.564 | 29.2075 | 29.2068 | 6.7666 | 2.2416 | 5 | 5 | 38 | 5.626
5 | 9.921 | 8.5772 | 8.5773 | 29.0255 | 29.0252 | 6.7726 | 2.2955 | 6 | 6 | 39 | 5.787
5 | 2.532 | 8.6626 | 8.6609 | 28.6887 | 28.6928 | 6.7743 | 2.5158 | 7 | 7 | 40 | 6.355
6 | 157.004 | 8.7042 | 8.7089 | 29.3679 | 29.3391 | 6.7383 | 1.6898 | 1 | 1 | 41 | 5.466
6 | 126.401 | 8.7275 | 8.7218 | 29.3505 | 29.3368 | 6.7355 | 1.616 | 2 | 2 | 42 | 5.424
6 | 100.533 | 8.7248 | 8.7249 | 29.3242 | 29.3234 | 6.737 | 2.1854 | 3 | 3 | 43 | 5.401 | 3
6 | 75.052 | 8.7252 | 8.7247 | 29.3168 | 29.3168 | 6.7373 | 2.1848 | 4 | 4 | 44 | 5.391
6 | 49.896 | 8.7341 | 8.7338 | 29.3075 | 29.3072 | 6.7364 | 2.1764 | 5 | 5 | 45 | 5.366
6 | 30.469 | 8.7654 | 8.7656 | 29.2718 | 29.2714 | 6.7332 | 2.1511 | 6 | 6 | 46 | 5.316 | 6
6 | 20.499 | 8.7758 | 8.7764 | 29.2648 | 29.2643 | 6.7319 | 2.1512 | 7 | 7 | 47 | 5.329
6 | 10.133 | 8.8127 | 8.8127 | 29.2478 | 29.2475 | 6.7271 | 2.1726 | 8 | 8 | 48 | 5.339
6 | 3.186 | 8.9501 | 8.9518 | 29.0265 | 29.0232 | 6.716 | 2.1886 | 9 | 9 | 49 | 5.387
7 | 303.146 | 8.5016 | 8.5011 | 29.5879 | 29.5885 | 6.7595 | 2.2214 | 1 | 1 | 50 | 5.758
7 | 250.237 | 8.5243 | 8.5243 | 29.5508 | 29.5507 | 6.7576 | 2.2273 | 2 | 2 | 51 | 5.728 | 6
7 | 199.979 | 8.5148 | 8.514 | 29.5542 | 29.5549 | 6.7589 | 2.2399 | 3 | 3 | 52 | 5.716
7 | 175.18 | 8.5142 | 8.5142 | 29.5549 | 29.5548 | 6.759 | 2.2514 | 4 | 4 | 53 | 5.737
7 | 150.153 | 8.5008 | 8.4978 | 29.5652 | 29.5676 | 6.7606 | 2.2609 | 5 | 5 | 54 | -99 | 9
7 | 125.241 | 8.5589 | 8.5619 | 29.4954 | 29.4921 | 6.7548 | 2.2538 | 6 | 6 | 55 | 5.696
7 | 100.098 | 8.6168 | 8.6163 | 29.4198 | 29.4198 | 6.7493 | 2.2178 | 7 | 7 | 56 | 5.527
7 | 75.152 | 8.6526 | 8.6527 | 29.3786 | 29.3783 | 6.7456 | 2.2078 | 8 | 8 | 57 | 5.488
7 | 50.062 | 8.7007 | 8.701 | 29.3235 | 29.3224 | 6.7407 | 2.1894 | 9 | 9 | 58 | 5.431
7 | 30.24 | 8.7278 | 8.7265 | 29.3091 | 29.3089 | 6.7372 | 2.1709 | 10 | 10 | 59 | 5.365
7 | 19.955 | 8.7584 | 8.76 | 29.2811 | 29.2789 | 6.7338 | 2.1522 | 11 | 11 | 60 | 5.352
7 | 10.438 | 8.7898 | 8.7898 | 29.1964 | 29.1979 | 6.7328 | 2.1617 | 12 | 12 | 61 | 5.335
7 | 3.047 | 8.7915 | 8.7914 | 29.1961 | 29.1967 | 6.7325 | 2.1721 | 13 | 13 | 62 | 5.353
8 | 200.019 | 8.4833 | 8.4834 | 29.6316 | 29.6314 | 6.7603 | 2.2747 | 1 | 1 | 63 | 5.804
8 | 174.779 | 8.509 | 8.5094 | 29.5977 | 29.5968 | 6.7579 | 2.2768 | 2 | 2 | 64 | 5.813
8 | 150.413 | 8.5435 | 8.5436 | 29.5538 | 29.5533 | 6.7546 | 2.2794 | 3 | 3 | 65 | 5.785
8 | 125.087 | 8.5602 | 8.5603 | 29.5339 | 29.5335 | 6.7529 | 2.2825 | 4 | 4 | 66 | 5.778
8 | 100.389 | 8.574 | 8.5743 | 29.521 | 29.5205 | 6.7514 | 2.2885 | 5 | 5 | 67 | 5.769
8 | 75.162 | 8.5936 | 8.5936 | 29.4963 | 29.4959 | 6.7495 | 2.2859 | 6 | 6 | 68 | 5.772
8 | 50.125 | 8.6076 | 8.6075 | 29.4552 | 29.455 | 6.7491 | 2.259 | 7 | 7 | 69 | 5.695 | 6
8 | 30.322 | 8.6351 | 8.6347 | 29.3904 | 29.3912 | 6.7478 | 2.2067 | 8 | 8 | 70 | 5.57
8 | 20.345 | 8.6686 | 8.6642 | 29.3298 | 29.3375 | 6.7453 | 2.2135 | 9 | 9 | 71 | 5.543
8 | 10.315 | 8.8556 | 8.8582 | 28.9568 | 28.9533 | 6.7333 | 2.3004 | 10 | 10 | 72 | 5.665
8 | 3.709 | 8.8698 | 8.8704 | 28.917 | 28.9155 | 6.7329 | 2.3279 | 11 | 11 | 73 | 5.899
9 | 316.785 | 8.3554 | 8.3554 | 29.8358 | 29.8354 | 6.7709 | 2.2421 | 1 | 1 | 74 | 5.617 | 4
9 | 251.011 | 8.3538 | 8.3535 | 29.8292 | 29.829 | 6.7714 | 2.2561 | 2 | 2 | 75 | 5.845
9 | 200.803 | 8.3561 | 8.3568 | 29.8184 | 29.8171 | 6.7716 | 2.2683 | 3 | 3 | 76 | 5.835
9 | 174.023 | 8.356 | 8.3562 | 29.8151 | 29.8145 | 6.7717 | 2.2746 | 4 | 4 | 77 | 5.834
9 | 150.912 | 8.3821 | 8.3822 | 29.7758 | 29.7751 | 6.7694 | 2.2804 | 5 | 5 | 78 | 5.823
9 | 125.231 | 8.3808 | 8.3806 | 29.7749 | 29.7747 | 6.7697 | 2.2879 | 6 | 6 | 79 | 5.826
9 | 100.545 | 8.4343 | 8.435 | 29.7026 | 29.7015 | 6.7647 | 2.2965 | 7 | 7 | 80 | 5.81
9 | 74.123 | 8.4147 | 8.4145 | 29.7244 | 29.7244 | 6.7667 | 2.3007 | 8 | 8 | 81 | 5.812
9 | 50.203 | 8.4824 | 8.4827 | 29.6338 | 29.6331 | 6.7604 | 2.2961 | 9 | 9 | 82 | 5.809 | 6
9 | 30.39 | 8.5131 | 8.5133 | 29.594 | 29.5934 | 6.7574 | 2.3004 | 10 | 10 | 83 | 5.826
9 | 20.39 | 8.491 | 8.4913 | 29.6196 | 29.619 | 6.7597 | 2.2923 | 11 | 11 | 84 | 5.798
9 | 10.397 | 8.501 | 8.5006 | 29.606 | 29.6062 | 6.7588 | 2.2984 | 12 | 12 | 85 | 5.809
9 | 1.882 | 8.4967 | 8.4967 | 29.6112 | 29.6109 | 6.7592 | 2.2929 | 13 | 13 | 86 | 5.818
11 | 96.136 | 7.1012 | 7.1017 | 33.727 | 33.7279 | 6.7921 | 1.4519 | 1 | 1 | 87 | 2.909
11 | 76.519 | 7.3555 | 7.3557 | 33.5042 | 33.505 | 6.7626 | 1.6239 | 2 | 2 | 88 | 3.474
11 | 50.556 | 7.7379 | 7.7365 | 32.9406 | 32.9426 | 6.7288 | 2.1024 | 3 | 3 | 89 | 4.888
11 | 30.095 | 8.3946 | 8.396 | 32.3146 | 32.3139 | 6.6568 | 2.491 | 4 | 4 | 90 | 5.984 | 6
11 | 21.039 | 8.4559 | 8.4561 | 32.2424 | 32.2431 | 6.6508 | 2.4781 | 5 | 5 | 91 | 5.915
11 | 10.491 | 9.0162 | 9.0208 | 31.6954 | 31.6892 | 6.5914 | 2.5993 | 6 | 6 | 92 | 6.22
11 | 2.334 | 9.6184 | 9.6246 | 31.1604 | 31.156 | 6.5266 | 2.7327 | 7 | 7 | 93 | 6.556
13 | 1152.059 | 3.408 | 3.4076 | 34.3869 | 34.3884 | 7.3804 | 0.60523 | 1 | 1 | 94 | 0.379
13 | 1012.795 | 3.6559 | 3.6561 | 34.3477 | 34.349 | 7.3378 | 0.59689 | 2 | 2 | 95 | 0.323
13 | 760.488 | 4.3008 | 4.3005 | 34.2273 | 34.2284 | 7.2299 | 0.60773 | 3 | 3 | 96 | 0.381
13 | 506.822 | 5.0146 | 5.0148 | 34.0893 | 34.0901 | 7.114 | 0.73463 | 4 | 4 | 97 | 0.783
13 | 305.9 | 5.6029 | 5.6031 | 34.0126 | 34.0133 | 7.0192 | 0.8947 | 5 | 5 | 98 | 1.274 | 6
13 | 256.433 | 5.7595 | 5.7592 | 33.9812 | 33.982 | 6.9949 | 0.98797 | 6 | 6 | 99 | 1.593
13 | 203.95 | 6.1927 | 6.1912 | 33.9363 | 33.9371 | 6.9265 | 1.1722 | 7 | 7 | 100 | 2.188
13 | 175.571 | 6.4911 | 6.4915 | 33.8993 | 33.8999 | 6.8804 | 1.3299 | 8 | 8 | 101 | 2.626
13 | 153.036 | 6.7436 | 6.743 | 33.8678 | 33.8685 | 6.8418 | 1.4432 | 9 | 9 | 102 | 2.989
13 | 127.108 | 7.0502 | 7.0493 | 33.8244 | 33.825 | 6.7957 | 1.5313 | 10 | 10 | 103 | 3.185
13 | 104.407 | 7.3496 | 7.3495 | 33.7094 | 33.71 | 6.7544 | 1.5973 | 11 | 11 | 104 | 3.378
13 | 76.789 | 7.5895 | 7.5895 | 33.2853 | 33.2858 | 6.7363 | 1.9618 | 12 | 12 | 105 | 4.501
13 | 51.62 | 7.7942 | 7.7966 | 32.6607 | 32.6603 | 6.7325 | 2.4531 | 13 | 13 | 106 | 5.898 | 6
13 | 30.04 | 8.6904 | 8.6841 | 32.0207 | 32.0272 | 6.6255 | 2.5082 | 14 | 14 | 107 | 6.032
13 | 20.017 | 8.9129 | 8.9091 | 31.8195 | 31.823 | 6.6013 | 2.5543 | 15 | 15 | 108 | 6.189
13 | 9.904 | 10.0029 | 10.0053 | 30.962 | 30.9599 | 6.4799 | 2.8596 | 16 | 16 | 109 | 6.905
13 | 1.73 | 10.1173 | 10.1173 | 30.6944 | 30.6953 | 6.4748 | 2.8705 | 17 | 17 | 110 | 6.915
15 | 758.615 | 4.3093 | 4.3093 | 34.2359 | 34.2373 | 7.228 | 0.60743 | 1 | 1 | 111 | 0.365
15 | 503.637 | 4.8372 | 4.8376 | 34.1381 | 34.1391 | 7.1418 | 0.66829 | 2 | 2 | 112 | 0.675 | 3
15 | 305.476 | 5.8722 | 5.8725 | 34.036 | 34.0367 | 6.9739 | 0.85351 | 3 | 3 | 113 | 1.144
15 | 251.925 | 5.9881 | 5.9885 | 33.935 | 33.9358 | 6.9597 | 1.1648 | 4 | 4 | 114 | 2.204
15 | 202.872 | 6.3882 | 6.388 | 33.9199 | 33.9206 | 6.8959 | 1.2192 | 5 | 5 | 115 | 2.283
15 | 176.008 | 6.5398 | 6.54 | 33.8871 | 33.8877 | 6.8732 | 1.2396 | 6 | 6 | 116 | 2.33
15 | 151.015 | 6.8913 | 6.8913 | 33.828 | 33.8287 | 6.8203 | 1.4 | 7 | 7 | 117 | 2.825
15 | 126.292 | 7.2364 | 7.2362 | 33.7304 | 33.7309 | 6.7709 | 1.5389 | 8 | 8 | 118 | 3.183
15 | 100.871 | 7.4643 | 7.4636 | 33.4486 | 33.4512 | 6.7483 | 1.6812 | 9 | 9 | 119 | 3.707
15 | 75.915 | 7.6946 | 7.6942 | 33.0195 | 33.0219 | 6.7319 | 2.0506 | 10 | 10 | 120 | 4.784
15 | 49.982 | 7.8524 | 7.8526 | 32.6293 | 32.63 | 6.725 | 2.4664 | 11 | 11 | 121 | 6.121
15 | 29.712 | 8.2428 | 8.2429 | 32.4322 | 32.4334 | 6.6745 | 2.6183 | 12 | 12 | 122 | 6.397
15 | 19.121 | 8.4897 | 8.4911 | 32.3799 | 32.382 | 6.6398 | 2.6947 | 13 | 13 | 123 | 6.552 | 6
15 | 9.651 | 9.6996 | 9.7 | 31.2954 | 31.2955 | 6.5093 | 2.8191 | 14 | 14 | 124 | 6.774
15 | 2.559 | 10.0624 | 10.0626 | 31.0924 | 31.0916 | 6.4661 | 2.8522 | 15 | 15 | 125 | 6.824
17 | 2083.101 | 1.846 | 1.8453 | 34.6098 | 34.6119 | 7.6607 | 0.84655 | 1 | 1 | 126 | 1.573
17 | 2000.077 | 1.8934 | 1.893 | 34.6002 | 34.6021 | 7.6521 | 0.83218 | 2 | 2 | 127 | 1.449
17 | 1748.799 | 2.2491 | 2.2482 | 34.5486 | 34.5509 | 7.5868 | 0.73677 | 3 | 3 | 128 | 0.983
17 | 1499.92 | 2.548 | 2.5482 | 34.5161 | 34.5175 | 7.5321 | 0.68547 | 4 | 4 | 129 | 0.722
17 | 1249.266 | 3.1439 | 3.1438 | 34.4424 | 34.4436 | 7.4256 | 0.60911 | 5 | 5 | 130 | 0.368
17 | 999.523 | 3.7027 | 3.7026 | 34.3668 | 34.3679 | 7.3285 | 0.57827 | 6 | 6 | 131 | 0.244
17 | 749.962 | 4.1201 | 4.1205 | 34.2569 | 34.2577 | 7.26 | 0.5855 | 7 | 7 | 132 | 0.295
17 | 500.01 | 4.7678 | 4.7679 | 34.0958 | 34.0965 | 7.1557 | 0.70148 | 8 | 8 | 133 | 0.643 | 6
17 | 400.41 | 5.268 | 5.2687 | 34.0402 | 34.041 | 7.0736 | 0.79759 | 9 | 9 | 134 | 1.011
17 | 299.482 | 5.7297 | 5.73 | 33.956 | 33.9568 | 7.001 | 1.0344 | 10 | 10 | 135 | 1.732
17 | 249.723 | 6.0473 | 6.0482 | 33.9061 | 33.9068 | 6.9514 | 1.2831 | 11 | 11 | 136 | 2.529
17 | 201.061 | 6.4737 | 6.4732 | 33.8578 | 33.8586 | 6.885 | 1.568 | 12 | 12 | 137 | 3.417
17 | 174.866 | 6.7715 | 6.7719 | 33.8302 | 33.8307 | 6.8391 | 1.6399 | 13 | 13 | 138 | 3.615 | 6
17 | 150.645 | 6.9485 | 6.949 | 33.7734 | 33.7741 | 6.8138 | 1.7824 | 14 | 14 | 139 | 4.016
17 | 125.685 | 7.1734 | 7.1735 | 33.6209 | 33.6223 | 6.7856 | 1.8626 | 15 | 15 | 140 | 4.236
17 | 100.019 | 7.3866 | 7.3869 | 33.3718 | 33.374 | 6.7636 | 1.9856 | 16 | 16 | 141 | 4.619
17 | 74.949 | 7.3564 | 7.3565 | 32.8381 | 32.8389 | 6.792 | 2.4616 | 17 | 17 | 142 | 6.049
17 | 50.599 | 7.6571 | 7.6564 | 32.5912 | 32.5914 | 6.7565 | 2.7045 | 18 | 18 | 143 | 6.824 | 4
17 | 29.562 | 8.6969 | 8.6919 | 32.4821 | 32.4828 | 6.6048 | 2.828 | 19 | 19 | 144 | 6.969
17 | 20.388 | 9.3644 | 9.3658 | 32.3198 | 32.3207 | 6.5144 | 2.8688 | 20 | 20 | 145 | 6.987
17 | 9.747 | 10.0151 | 10.0222 | 31.2387 | 31.2044 | 6.4667 | 2.857 | 21 | 21 | 146 | 6.86
17 | 2.183 | 10.0039 | 10.0038 | 31.1105 | 31.1114 | 6.4736 | 2.8519 | 22 | 22 | 147 | 6.851
18 | 48.411 | 8.5334 | 8.5304 | 32.3121 | 32.317 | 6.6363 | 2.2631 | 1 | 1 | 148 | 5.243
18 | 30.342 | 8.6779 | 8.6782 | 32.0106 | 32.014 | 6.6278 | 2.2537 | 2 | 2 | 149 | 5.255
18 | 20.562 | 8.9528 | 8.9576 | 31.5184 | 31.5161 | 6.6083 | 2.3623 | 3 | 3 | 150 | 5.6
18 | 10.295 | 9.6423 | 9.6493 | 31.0521 | 31.042 | 6.5277 | 2.8345 | 4 | 4 | 151 | 7.119 | 6
18 | 1.743 | 11.0185 | 11.0264 | 29.8535 | 29.8518 | 6.3835 | 3.2458 | 5 | 5 | 152 | 7.9
20 | 119.131 | 7.1761 | 7.176 | 33.6833 | 33.6844 | 6.7824 | 1.5393 | 1 | 1 | 153 | 3.192
20 | 76.366 | 7.8398 | 7.8411 | 32.9242 | 32.9236 | 6.7139 | 2.1497 | 2 | 2 | 154 | 5.037
20 | 51.124 | 8.2841 | 8.2807 | 32.5 | 32.502 | 6.6654 | 2.5253 | 3 | 3 | 155 | 6.224
20 | 30.527 | 8.6484 | 8.6487 | 32.3528 | 32.3535 | 6.6174 | 2.5311 | 4 | 4 | 156 | 6.072
20 | 21.078 | 9.1224 | 9.118 | 32.1173 | 32.119 | 6.558 | 2.6693 | 5 | 5 | 157 | 6.5
20 | 10.732 | 10.2851 | 10.2847 | 31.2678 | 31.2694 | 6.4276 | 2.9175 | 6 | 6 | 158 | 6.967
20 | 2.446 | 11.3801 | 11.3663 | 30.6737 | 30.682 | 6.3013 | 2.952 | 7 | 7 | 159 | 6.98
22 | 266.655 | 6.1402 | 6.1405 | 34.0034 | 34.0045 | 6.9319 | 0.96648 | 1 | 1 | 160 | 1.516
22 | 200.063 | 6.2078 | 6.2084 | 33.9767 | 33.9774 | 6.9222 | 1.0544 | 2 | 2 | 161 | 1.763
22 | 175.255 | 6.3658 | 6.3658 | 33.9576 | 33.9584 | 6.8977 | 1.1204 | 3 | 3 | 162 | 1.955 | 6
22 | 150.646 | 6.6144 | 6.6141 | 33.919 | 33.92 | 6.8599 | 1.2325 | 4 | 4 | 163 | 2.284
22 | 125.224 | 7.1328 | 7.1331 | 33.8457 | 33.8464 | 6.7819 | 1.5008 | 5 | 5 | 164 | 3.103
22 | 100.289 | 7.4071 | 7.4079 | 33.6928 | 33.6931 | 6.7463 | 1.6099 | 6 | 6 | 165 | 3.415
22 | 75.652 | 7.725 | 7.7253 | 33.206 | 33.2064 | 6.7191 | 1.9534 | 7 | 7 | 166 | 4.463
22 | 49.312 | 7.8738 | 7.8713 | 32.5804 | 32.5785 | 6.7238 | 2.5457 | 8 | 8 | 167 | 6.27
22 | 29.719 | 8.7651 | 8.7641 | 32.3788 | 32.3785 | 6.5991 | 2.7412 | 9 | 9 | 168 | 6.663
22 | 20.253 | 9.0014 | 9.0036 | 32.096 | 32.0963 | 6.5765 | 2.7945 | 10 | 10 | 169 | 6.824 | 6
22 | 9.82 | 10.1767 | 10.2142 | 31.6317 | 31.5928 | 6.4278 | 2.8898 | 11 | 11 | 170 | 6.894
22 | 1.801 | 11.4499 | 11.4497 | 30.7154 | 30.7157 | 6.2902 | 2.8783 | 12 | 12 | 171 | 6.75
25 | 32.411 | 8.7666 | 8.7683 | 31.8559 | 31.854 | 6.6213 | 1.4093 | 1 | 1 | 174 | 2.739
25 | 20.038 | 8.8069 | 8.8069 | 31.6315 | 31.6272 | 6.6249 | 1.3427 | 2 | 2 | 175 | 2.232 | 6
25 | 10.267 | 8.9426 | 8.9458 | 31.0174 | 31.0108 | 6.6312 | 1.9673 | 3 | 3 | 176 | 3.271
25 | 1.542 | 9.9349 | 9.9594 | 15.8868 | 15.9539 | 7.1504 | 2.9993 | 4 | 4 | 177 | 8.389
26 | 66.199 | 8.7417 | 8.7416 | 32.1022 | 32.1036 | 6.6144 | 1.006 | 1 | 1 | 178 | 2.162
26 | 49.437 | 8.7226 | 8.7236 | 31.9826 | 31.9834 | 6.6224 | 1.4335 | 2 | 2 | 179 | 2.852
26 | 30.624 | 8.7497 | 8.7536 | 31.8371 | 31.8355 | 6.6246 | 1.5748 | 3 | 3 | 180 | 3.37 | 6
26 | 19.773 | 8.7552 | 8.7554 | 31.6765 | 31.6841 | 6.6307 | 1.6309 | 4 | 4 | 181 | 3.304
26 | 9.598 | 8.7819 | 8.7743 | 31.2518 | 31.291 | 6.645 | 1.5916 | 5 | 5 | 182 | 3.22
26 | 1.405 | 11.2622 | 11.2741 | 12.6875 | 12.2494 | 7.0792 | 2.9151 | 6 | 6 | 183 | 7.892
27 | 115.198 | 8.145 | 8.1443 | 32.5718 | 32.5728 | 6.6831 | 0.55955 | 1 | 1 | 184 | 0
27 | 74.661 | 8.5231 | 8.5273 | 32.4852 | 32.4836 | 6.6303 | 0.55092 | 2 | 2 | 185 | 0
27 | 50.121 | 8.6634 | 8.6626 | 31.9703 | 31.9686 | 6.6316 | 1.6801 | 3 | 3 | 186 | 3.432
27 | 29.42 | 8.7056 | 8.7059 | 31.8094 | 31.8093 | 6.6323 | 1.7468 | 4 | 4 | 187 | 3.699 | 6
27 | 20.385 | 8.7086 | 8.7091 | 31.6222 | 31.6222 | 6.6399 | 1.7674 | 5 | 5 | 188 | 3.748
27 | 10.236 | 8.7337 | 8.7325 | 31.1678 | 31.1724 | 6.6558 | 1.9834 | 6 | 6 | 189 | 3.843
27 | 1.493 | 13.3661 | 13.2571 | 21.1019 | 21.7329 | 6.4149 | 3.3254 | 7 | 7 | 190 | 7.738
28 | 53.005 | 8.5421 | 8.5425 | 32.2224 | 32.2229 | 6.6388 | 1.7911 | 1 | 1 | 191 | 3.873
28 | 29.901 | 8.6626 | 8.6629 | 31.88 | 31.8807 | 6.6357 | 1.8443 | 2 | 2 | 192 | 4.017
28 | 20.469 | 8.6762 | 8.6762 | 31.6121 | 31.6079 | 6.6452 | 1.8073 | 3 | 3 | 193 | 3.896
28 | 10.132 | 8.7373 | 8.7333 | 31.1651 | 31.168 | 6.6553 | 2.1331 | 4 | 4 | 194 | 4.72 | 6
28 | 1.537 | 12.1412 | 12.1875 | 22.4079 | 22.1523 | 6.5312 | 3.2312 | 5 | 5 | 195 | 7.618
29 | 195.577 | 8.2184 | 8.2187 | 32.8426 | 32.843 | 6.6603 | 0.55826 | 1 | 1 | 196 | 0.084
29 | 175.079 | 8.2321 | 8.2322 | 32.8295 | 32.8299 | 6.6589 | 0.58572 | 2 | 2 | 197 | 0.168
29 | 150.285 | 8.2495 | 8.2497 | 32.8188 | 32.8191 | 6.6567 | 0.58679 | 3 | 3 | 198 | 0.268
29 | 124.583 | 8.3595 | 8.3597 | 32.7582 | 32.7585 | 6.6429 | 0.6406 | 4 | 4 | 199 | 0.437
29 | 99.717 | 8.5174 | 8.5221 | 32.3933 | 32.395 | 6.6351 | 1.4569 | 5 | 5 | 200 | 2.377
29 | 74.757 | 8.5238 | 8.5242 | 32.3115 | 32.3104 | 6.6377 | 1.7049 | 6 | 6 | 201 | 3.654
29 | 50.062 | 8.5396 | 8.5396 | 32.1799 | 32.1808 | 6.641 | 1.8419 | 7 | 7 | 202 | 3.963
29 | 29.166 | 8.6516 | 8.649 | 31.8981 | 31.9077 | 6.6365 | 1.8532 | 8 | 8 | 203 | 4.048 | 6
29 | 19.896 | 8.6583 | 8.6586 | 31.5649 | 31.5653 | 6.6499 | 1.8261 | 9 | 9 | 204 | 3.982
29 | 10.595 | 8.746 | 8.7379 | 31.1524 | 31.1718 | 6.6546 | 2.102 | 10 | 10 | 205 | 4.263
29 | 1.77 | 10.1391 | 10.058 | 28.0408 | 28.2109 | 6.5826 | 3.263 | 11 | 11 | 206 | 7.648
30 | 140.607 | 8.2872 | 8.281 | 32.7958 | 32.801 | 6.6521 | 0.60695 | 1 | 1 | 207 | 0.244 | 6
30 | 124.703 | 8.3592 | 8.3585 | 32.7431 | 32.7434 | 6.6436 | 0.63899 | 2 | 2 | 208 | 0.28
30 | 100.252 | 8.4302 | 8.4285 | 32.3594 | 32.3658 | 6.6496 | 1.6804 | 3 | 3 | 209 | 3.899
30 | 74.791 | 8.4969 | 8.4975 | 32.2877 | 32.2879 | 6.6427 | 1.8006 | 4 | 4 | 210 | 3.922
30 | 50.348 | 8.5466 | 8.5455 | 32.1589 | 32.1627 | 6.6409 | 1.8641 | 5 | 5 | 211 | 4.052
30 | 30.16 | 8.6506 | 8.6508 | 31.7889 | 31.789 | 6.6414 | 1.8661 | 6 | 6 | 212 | 4.083
30 | 20.526 | 8.6444 | 8.6446 | 31.6156 | 31.6168 | 6.6498 | 1.8676 | 7 | 7 | 213 | 4.065
30 | 10.611 | 8.639 | 8.6409 | 31.3223 | 31.3152 | 6.6632 | 1.9793 | 8 | 8 | 214 | 4.583
30 | 1.79 | 12.93 | 12.9545 | 23.2695 | 23.3439 | 6.3875 | 3.3952 | 9 | 9 | 215 | 7.544
31 | 113.721 | 8.3959 | 8.3967 | 32.4084 | 32.4099 | 6.6526 | 1.6264 | 1 | 1 | 216 | 3.569
31 | 100.436 | 8.4036 | 8.405 | 32.3691 | 32.3654 | 6.6531 | 1.7672 | 2 | 2 | 217 | 3.66 | 6
31 | 75.263 | 8.4549 | 8.4551 | 32.2915 | 32.2918 | 6.6488 | 1.844 | 3 | 3 | 218 | 4.019
31 | 50.43 | 8.5509 | 8.5525 | 32.1465 | 32.1433 | 6.6408 | 1.8729 | 4 | 4 | 219 | 4.111
31 | 30.436 | 8.5764 | 8.5758 | 31.7925 | 31.7885 | 6.6523 | 1.9444 | 5 | 5 | 220 | 4.259
31 | 20.032 | 8.611 | 8.6112 | 31.5553 | 31.551 | 6.6573 | 1.927 | 6 | 6 | 221 | 4.237
31 | 10.118 | 8.7604 | 8.7636 | 31.0948 | 31.0902 | 6.6549 | 2.4251 | 7 | 7 | 222 | 4.709
31 | 1.575 | 13.4468 | 13.5551 | 16.9974 | 16.6902 | 6.5695 | 3.1067 | 8 | 8 | 223 | 7.505
32 | 82.233 | 8.3686 | 8.3681 | 32.3933 | 32.3949 | 6.6573 | 1.8099 | 1 | 1 | 224 | 3.984
32 | 50.368 | 8.5326 | 8.5332 | 32.1581 | 32.157 | 6.643 | 1.8843 | 2 | 2 | 225 | 4.147
32 | 30.459 | 8.5775 | 8.578 | 31.7982 | 31.7937 | 6.6518 | 1.9405 | 3 | 3 | 226 | 4.309 | 6
32 | 19.886 | 8.579 | 8.5794 | 31.505 | 31.5045 | 6.6643 | 2.0117 | 4 | 4 | 227 | 4.465
32 | 10.263 | 8.7976 | 8.8064 | 31.0534 | 31.0393 | 6.6512 | 2.7156 | 5 | 5 | 228 | 5.206
32 | 1.76 | 13.9418 | 13.9793 | 20.9459 | 20.8335 | 6.3438 | 3.2637 | 6 | 6 | 229 | 7.985
33 | 59.856 | 8.2852 | 8.2769 | 32.4894 | 32.4953 | 6.6657 | 1.8019 | 1 | 1 | 230 | 4.071
33 | 30.03 | 8.564 | 8.564 | 31.7524 | 31.7528 | 6.6558 | 1.9689 | 2 | 2 | 231 | 4.404
33 | 20.193 | 8.5716 | 8.5722 | 31.6151 | 31.6092 | 6.6606 | 1.9953 | 3 | 3 | 232 | -99 | 9
33 | 10.015 | 8.649 | 8.6521 | 31.2914 | 31.288 | 6.6631 | 2.2456 | 4 | 4 | 233 | 4.896 | 6
33 | 1.921 | 14.0122 | 14.028 | 18.9032 | 18.728 | 6.4151 | 3.0406 | 5 | 5 | 234 | 7.653
34 | 95.639 | 8.1404 | 8.1407 | 32.605 | 32.6057 | 6.6824 | 1.5828 | 1 | 1 | 235 | 3.325
34 | 75.095 | 8.1697 | 8.1692 | 32.5816 | 32.5832 | 6.679 | 1.6291 | 2 | 2 | 236 | 3.369
34 | 50.338 | 8.3484 | 8.3469 | 32.4316 | 32.433 | 6.6587 | 1.8545 | 3 | 3 | 237 | 3.921
34 | 30.02 | 8.5567 | 8.5569 | 31.868 | 31.8622 | 6.6519 | 1.9635 | 4 | 4 | 238 | 4.315
34 | 19.979 | 8.5479 | 8.5489 | 31.619 | 31.6184 | 6.664 | 1.9892 | 5 | 5 | 239 | 4.441
34 | 9.516 | 8.7123 | 8.7118 | 31.1572 | 31.1593 | 6.6594 | 2.4729 | 6 | 6 | 240 | 5.4
34 | 1.507 | 13.6213 | 13.5676 | 18.5884 | 18.7547 | 6.4808 | 3.05 | 7 | 7 | 241 | 7.487 | 6
35 | 79.461 | 8.1145 | 8.1136 | 32.6066 | 32.6078 | 6.6862 | 1.689 | 1 | 1 | 242 | 3.613 | 6
35 | 50.646 | 8.4268 | 8.4271 | 32.348 | 32.3485 | 6.6506 | 1.9314 | 2 | 2 | 243 | 4.148
35 | 29.464 | 8.529 | 8.5291 | 31.8694 | 31.8559 | 6.656 | 2.0039 | 3 | 3 | 244 | 4.412
35 | 19.969 | 8.5481 | 8.5457 | 31.6869 | 31.6696 | 6.661 | 1.9908 | 4 | 4 | 245 | 4.419
35 | 10.089 | 8.6841 | 8.6872 | 31.2112 | 31.2046 | 6.6613 | 2.4312 | 5 | 5 | 246 | 5.225
35 | 1.609 | 13.2721 | 13.268 | 17.5061 | 17.5065 | 6.5731 | 3.0159 | 6 | 6 | 247 | 7.655
36 | 85.961 | 8.0032 | 8.0024 | 32.6518 | 32.651 | 6.7011 | 1.6349 | 1 | 1 | 248 | 3.519
36 | 75.352 | 8.1357 | 8.1388 | 32.5375 | 32.5367 | 6.686 | 1.7799 | 2 | 2 | 249 | 3.782 | 6
36 | 50.583 | 8.443 | 8.4451 | 32.3237 | 32.32 | 6.6492 | 1.9664 | 3 | 3 | 250 | 4.338
36 | 30.004 | 8.4873 | 8.4887 | 31.8223 | 31.8276 | 6.6643 | 2.0656 | 4 | 4 | 251 | 4.648
36 | 20.124 | 8.4356 | 8.4309 | 31.6262 | 31.6184 | 6.6805 | 2.0591 | 5 | 5 | 252 | 4.644
36 | 10.118 | 8.7116 | 8.7143 | 31.1482 | 31.1487 | 6.6599 | 3.169 | 6 | 6 | 253 | 5.418
36 | 1.473 | 13.4518 | 13.4826 | 21.4115 | 21.2949 | 6.391 | 3.1612 | 7 | 7 | 254 | 7.685
37 | 96.14 | 7.9086 | 7.9199 | 32.6427 | 32.6405 | 6.7158 | 1.6912 | 1 | 1 | 255 | 3.601
37 | 75.04 | 8.1089 | 8.1089 | 32.5342 | 32.5336 | 6.6902 | 1.7713 | 2 | 2 | 256 | 3.731
37 | 49.987 | 8.5012 | 8.4973 | 32.2483 | 32.25 | 6.6438 | 2.0369 | 3 | 3 | 257 | 4.375 | 6
37 | 29.846 | 8.4687 | 8.4695 | 31.8171 | 31.824 | 6.6673 | 2.0757 | 4 | 4 | 258 | 7.301 | 4
37 | 20.491 | 8.4507 | 8.4555 | 31.583 | 31.5818 | 6.6801 | 2.1224 | 5 | 5 | 259 | 4.758
37 | 9.635 | 8.5377 | 8.5414 | 31.185 | 31.1841 | 6.6843 | 2.4704 | 6 | 6 | 260 | 5.467
37 | 1.563 | 13.1632 | 13.2078 | 21.667 | 21.5836 | 6.4199 | 3.0352 | 7 | 7 | 261 | 7.342

## hysteresis
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE2 | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
27 | 74.661 | 8.5231 | 8.5273 | 32.4852 | 32.4836 | 6.6303 | 0.52717 | 2 | 2 | 185 | 6.9456 | 0 | 0 |  | 0
29 | 195.577 | 8.2184 | 8.2187 | 32.8426 | 32.843 | 6.6603 | 0.52914 | 1 | 1 | 196 | 7.0646 | 0.01189 | 0.084 |  | 0.01189
27 | 115.198 | 8.145 | 8.1443 | 32.5718 | 32.5728 | 6.6831 | 0.52977 | 1 | 1 | 184 | 7.0113 | 0 | 0 |  | 0
29 | 175.079 | 8.2321 | 8.2322 | 32.8295 | 32.8299 | 6.6589 | 0.561 | 2 | 2 | 197 | 7.0455 | 0.023845 | 0.168 |  | 0.023845
29 | 150.285 | 8.2495 | 8.2497 | 32.8188 | 32.8191 | 6.6567 | 0.56423 | 3 | 3 | 198 | 7.0221 | 0.038165 | 0.268 |  | 0.038165
17 | 749.962 | 4.1201 | 4.1205 | 34.2569 | 34.2577 | 7.26 | 0.57306 | 7 | 7 | 132 | 8.0442 | 0.036672 | 0.295 |  | 0.036672
30 | 140.607 | 8.2872 | 8.281 | 32.7958 | 32.801 | 6.6521 | 0.58045 | 1 | 1 | 207 | 7.0111 | 0.034802 | 0.244 | 6 | 0.034802
15 | 758.615 | 4.3093 | 4.3093 | 34.2359 | 34.2373 | 7.228 | 0.59135 | 1 | 1 | 111 | 8.0238 | 0.04549 | 0.365 |  | 0.04549
13 | 760.488 | 4.3008 | 4.3005 | 34.2273 | 34.2284 | 7.2299 | 0.5947 | 3 | 3 | 96 | 8.0276 | 0.047462 | 0.381 |  | 0.047462
30 | 124.703 | 8.3592 | 8.3585 | 32.7431 | 32.7434 | 6.6436 | 0.61356 | 2 | 2 | 208 | 6.9928 | 0.040041 | 0.28 |  | 0.040041
29 | 124.583 | 8.3595 | 8.3597 | 32.7582 | 32.7585 | 6.6429 | 0.61677 | 4 | 4 | 199 | 6.992 | 0.0625 | 0.437 |  | 0.0625
15 | 503.637 | 4.8372 | 4.8376 | 34.1381 | 34.1391 | 7.1418 | 0.65337 | 2 | 2 | 112 | 7.6888 | 0.08779 | 0.675 | 3 | 0.08779
17 | 500.01 | 4.7678 | 4.7679 | 34.0958 | 34.0965 | 7.1557 | 0.68768 | 8 | 8 | 133 | 7.6975 | 0.083534 | 0.643 | 6 | 0.083534
13 | 506.822 | 5.0146 | 5.0148 | 34.0893 | 34.0901 | 7.114 | 0.72056 | 4 | 4 | 97 | 7.6689 | 0.1021 | 0.783 |  | 0.1021
17 | 400.41 | 5.268 | 5.2687 | 34.0402 | 34.041 | 7.0736 | 0.78213 | 9 | 9 | 134 | 7.5316 | 0.13424 | 1.011 |  | 0.13424
15 | 305.476 | 5.8722 | 5.8725 | 34.036 | 34.0367 | 6.9739 | 0.83603 | 3 | 3 | 113 | 7.3618 | 0.1554 | 1.144 |  | 0.1554
13 | 305.9 | 5.6029 | 5.6031 | 34.0126 | 34.0133 | 7.0192 | 0.87787 | 5 | 5 | 98 | 7.3974 | 0.17222 | 1.274 | 6 | 0.17222
17 | 999.523 | 3.7027 | 3.7026 | 34.3668 | 34.3679 | 7.3285 | 0.56564 | 6 | 6 | 131 | 8.3754 | 0.029133 | 0.244 |  | 0.029133
13 | 1012.795 | 3.6559 | 3.6561 | 34.3477 | 34.349 | 7.3378 | 0.58338 | 2 | 2 | 95 | 8.3993 | 0.038456 | 0.323 |  | 0.038456
13 | 1152.059 | 3.408 | 3.4076 | 34.3869 | 34.3884 | 7.3804 | 0.59072 | 1 | 1 | 94 | 8.5964 | 0.044088 | 0.379 |  | 0.044088
17 | 1249.266 | 3.1439 | 3.1438 | 34.4424 | 34.4436 | 7.4256 | 0.59584 | 5 | 5 | 130 | 8.7536 | 0.04204 | 0.368 |  | 0.04204
17 | 1499.92 | 2.548 | 2.5482 | 34.5161 | 34.5175 | 7.5321 | 0.67127 | 4 | 4 | 129 | 9.1646 | 0.078781 | 0.722 |  | 0.078781
17 | 1748.799 | 2.2491 | 2.2482 | 34.5486 | 34.5509 | 7.5868 | 0.72176 | 3 | 3 | 128 | 9.5335 | 0.10311 | 0.983 |  | 0.10311
17 | 2000.077 | 1.8934 | 1.893 | 34.6002 | 34.6021 | 7.6521 | 0.81703 | 2 | 2 | 127 | 9.9354 | 0.14584 | 1.449 |  | 0.14584
17 | 2083.101 | 1.846 | 1.8453 | 34.6098 | 34.6119 | 7.6607 | 0.83167 | 1 | 1 | 126 | 10.0553 | 0.15644 | 1.573 |  | 0.15644
[15 empty] | DO | .
DO sensor 1176
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
27 | 74.661 | 8.5231 | 8.5273 | 32.4852 | 32.4836 | 6.6303 | 0.55092 | 2 | 2 | 185 | 6.5693 | 0 | 0 |  | 0
27 | 115.198 | 8.145 | 8.1443 | 32.5718 | 32.5728 | 6.6831 | 0.55955 | 1 | 1 | 184 | 6.6589 | 0 | 0 |  | 0
29 | 124.583 | 8.3595 | 8.3597 | 32.7582 | 32.7585 | 6.6429 | 0.6406 | 4 | 4 | 199 | 6.6251 | 0.065961 | 0.437 |  | 0.065961
30 | 124.703 | 8.3592 | 8.3585 | 32.7431 | 32.7434 | 6.6436 | 0.63899 | 2 | 2 | 208 | 6.6259 | 0.042258 | 0.28 |  | 0.042258
30 | 140.607 | 8.2872 | 8.281 | 32.7958 | 32.801 | 6.6521 | 0.60695 | 1 | 1 | 207 | 6.6484 | 0.0367 | 0.244 | 6 | 0.0367
29 | 150.285 | 8.2495 | 8.2497 | 32.8188 | 32.8191 | 6.6567 | 0.58679 | 3 | 3 | 198 | 6.6616 | 0.04023 | 0.268 |  | 0.04023
29 | 175.079 | 8.2321 | 8.2322 | 32.8295 | 32.8299 | 6.6589 | 0.58572 | 2 | 2 | 197 | 6.6851 | 0.025131 | 0.168 |  | 0.025131
29 | 195.577 | 8.2184 | 8.2187 | 32.8426 | 32.843 | 6.6603 | 0.55826 | 1 | 1 | 196 | 6.7042 | 0.012529 | 0.084 |  | 0.012529
15 | 305.476 | 5.8722 | 5.8725 | 34.036 | 34.0367 | 6.9739 | 0.85351 | 3 | 3 | 113 | 7.1453 | 0.1601 | 1.144 |  | 0.1601
13 | 305.9 | 5.6029 | 5.6031 | 34.0126 | 34.0133 | 7.0192 | 0.8947 | 5 | 5 | 98 | 7.1958 | 0.17705 | 1.274 | 6 | 0.17705
17 | 400.41 | 5.268 | 5.2687 | 34.0402 | 34.041 | 7.0736 | 0.79759 | 9 | 9 | 134 | 7.3455 | 0.13763 | 1.011 |  | 0.13763
17 | 500.01 | 4.7678 | 4.7679 | 34.0958 | 34.0965 | 7.1557 | 0.70148 | 8 | 8 | 133 | 7.5352 | 0.085333 | 0.643 | 6 | 0.085333
15 | 503.637 | 4.8372 | 4.8376 | 34.1381 | 34.1391 | 7.1418 | 0.66829 | 2 | 2 | 112 | 7.5229 | 0.089725 | 0.675 | 3 | 0.089725
13 | 506.822 | 5.0146 | 5.0148 | 34.0893 | 34.0901 | 7.114 | 0.73463 | 4 | 4 | 97 | 7.4938 | 0.10449 | 0.783 |  | 0.10449
17 | 749.962 | 4.1201 | 4.1205 | 34.2569 | 34.2577 | 7.26 | 0.5855 | 7 | 7 | 132 | 7.9092 | 0.037298 | 0.295 |  | 0.037298
15 | 758.615 | 4.3093 | 4.3093 | 34.2359 | 34.2373 | 7.228 | 0.60743 | 1 | 1 | 111 | 7.8794 | 0.046323 | 0.365 |  | 0.046323 | 0.4566 | -0.231
13 | 760.488 | 4.3008 | 4.3005 | 34.2273 | 34.2284 | 7.2299 | 0.60773 | 3 | 3 | 96 | 7.8836 | 0.048328 | 0.381 |  | 0.048328
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
17 | 999.523 | 3.7027 | 3.7026 | 34.3668 | 34.3679 | 7.3285 | 0.57827 | 6 | 6 | 131 | 8.2561 | 0.029554 | 0.244 |  | 0.029554
13 | 1012.795 | 3.6559 | 3.6561 | 34.3477 | 34.349 | 7.3378 | 0.59689 | 2 | 2 | 95 | 8.282 | 0.039 | 0.323 |  | 0.039
13 | 1152.059 | 3.408 | 3.4076 | 34.3869 | 34.3884 | 7.3804 | 0.60523 | 1 | 1 | 94 | 8.4884 | 0.044649 | 0.379 |  | 0.044649
17 | 1249.266 | 3.1439 | 3.1438 | 34.4424 | 34.4436 | 7.4256 | 0.60911 | 5 | 5 | 130 | 8.6561 | 0.042513 | 0.368 |  | 0.042513
17 | 1499.92 | 2.548 | 2.5482 | 34.5161 | 34.5175 | 7.5321 | 0.68547 | 4 | 4 | 129 | 9.0895 | 0.079432 | 0.722 |  | 0.079432
17 | 1748.799 | 2.2491 | 2.2482 | 34.5486 | 34.5509 | 7.5868 | 0.73677 | 3 | 3 | 128 | 9.468 | 0.10382 | 0.983 |  | 0.10382
17 | 2000.077 | 1.8934 | 1.893 | 34.6002 | 34.6021 | 7.6521 | 0.83218 | 2 | 2 | 127 | 9.8816 | 0.14664 | 1.449 |  | 0.14664
17 | 2083.101 | 1.846 | 1.8453 | 34.6098 | 34.6119 | 7.6607 | 0.84655 | 1 | 1 | 126 | 10.0027 | 0.15726 | 1.573 |  | 0.15726

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers