## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |
| 4 Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |

# PROCESSING NOTES

Cruise: 2011-45

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: West Coast Vancouver Island / BC Inlets

Project: High Seas Salmon

Party Chief: Trudel M. & Morris J.

Platform: W.E. Ricker

Date: October 1, 2011 – November 6, 2011

Processed by: Germaine Gatien

Date of Processing: March 21, 2012 – 2 April 2012

Number of original hex files: 	153		Number of CTD casts processed: 153

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with an SBE43 Dissolved Oxygen sensor (#1119), an SBE pH sensor #0851 and a WetLabs ECO-FL Fluorometer #2214. The log indicates that the fluorometer was mounted on the primary pump, but this is likely a mistake – it probably should refer to the DO sensor. The salinometer used was a model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The log books were in good order with an equipment list, a record of sampling done and useful notes about problems encountered.

There was no deep salinity calibration sampling, but many samples from ~10m. The comparison with the CTD data suggests that both CTD salinity channels were low but the salinity gradients at 10m were relatively high and the depth of sampling used for the comparison was an estimate since the actual depth was variable and impossible to establish precisely since there was a lot of vertical motion during the bottle stops. The secondary salinity was closer to the bottles, on average.

A special set of files with extensions CTDpH were prepared for the High Seas Salmon program staff; they contain the pH data that were removed from files to be archived.

There was no calibration sampling for Dissolved Oxygen, so those data are considered nominal.

The WetLabs fluorescence data are considered nominal; values were much lower than extracted chlorophyll values.

# PROCESSING SUMMARY

##### 1. Seasave

This step was completed at sea; the raw data files have extension HEX.

##### 2. Preliminary Steps

The 3 Log Books were obtained. A few errors in file names and station names were noted and the fact that one cast was not recorded. There were also details of bottle sampling and a few cases of pumps being turned on late.

Bottle data were provided in 2 spreadsheets – QF2011-45Nuts+Chl*.xls and QF2011-45SAL*.xls. The * stands for the date of the spreadsheet version.

There were many problems with the CHL analysis; fortunately the duplicates were run on separate days and the analysis problems were sporadic, so there is a sample for most casts. However, there are no duplicates. There were no salinity duplicates. There were Autosal analysis log sheets available; the only comment concerns a value lost due to a computer problem. A number of salinity samples required 3 readings.

To enable preparation of a document combining CTD data plus bottle data, the 2 analysis files were edited and combined. Comments for each variable were combined in a single column with prefixes indicating to which variable each pertained. Columns were added to accommodate the CTD data that will be added to the file later. A document file (Data Analysis Comments.doc) was prepared with notes about analysis and quality flag definitions. A note about that file was entered in the first row of the spreadsheet file.

The cruise summary sheet was completed.

The histories of the dissolved oxygen, conductivity and pressure sensors were obtained.

There were no changes in the configuration files through the cruise. A configuration file was saved as 2011-45-ctd.xmlcon. Calibration constants were checked and no errors were found.

There was no DO sampling.

Salinity, chlorophyll and nutrient samples were taken at ~10m. Sometimes the CTD stopped at ~15m on the way up in the main cast, but when pH samples were taken at depth, the CTD was returned to the surface and then went down to about 15m so the Niskin could be fired ~10m.

##### 3. Conversion of Raw Data

All data were converted using configuration file 2011-45-ctd.xmlcon.

Plots were made for a few casts. The CTD typically stopped at 14db and sank a little to 14.5 or 15db, occasionally a little deeper. The Niskin bottle is usually mounted about 5m above the CTD. A few casts were examined in detail:

- The temperature channels track well on the downcast but the upcasts show larger differences. It is not clear which sensor provides better data, though there are more spikes in the primary. There are odd excursions in both. There is also a clear vertical offset that is seen only in the upcasts. These upcast features may be related to pump or plumbing problems.

- The two conductivity channels are similar during the downcasts, but on the upcasts they are noisy with the primary a little noisier than the secondary.

- Fluorescence looks normal, with a larger offset between down and upcasts than is seen in temperature or conductivity. So alignment will be necessary.

- Dissolved Oxygen data has the usual vertical offset between downcast and upcast so will require alignment.

- The descent rate was kept high most of the time. The noise level varies greatly though the cruise, as is expected because some casts are offshore and others in inlets.

- The pH signal looks normal with an offset between downcast and upcast; it will need alignment.

##### 4. WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### 5. ALIGNCTD

##### Tests were run on 3 casts to see if the offset of +2s that was applied to the dissolved oxygen channel for cruise 2011-58 was appropriate for this cruise too. The noisy upcast temperature data makes this a rough judgment, but the results look satisfactory.

##### ALIGNCTD was used to advance the DO signal relative to temperature by 2s.

##### 6. CELLTM

Tests were run on 3 casts using a variety of settings. The noisiness of the upcast temperature and salinity make the judgment a little difficult, but the best choice in each case was (α=0.02, 1/β=7) for both primary and secondary channels.

CELLTM was run using (0.02, 7) for the primary and the secondary conductivity for all casts.

##### 7. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 8. Test Plots and Channel Check

There were few casts deeper than 200m. The differences between the sensors were often noisy. The same sensors were used for 2011-58 so two casts from that cruise are included in the table.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-58-0007 | 470 | +0.0001 | +0.00028 | +0.0030 | VNoisy, fairly High |
| 2011-58-0013 | 470 | +0.0002 | +0.00028 | +0.0027 | VNoisy, fairly High |
| 2011-45-0067 | 220 | +0.0005 | +0.00050 | +0.0050 | XNoisy, fairly High |
| 2011-45-0184 | 240 | +0.0008 | +0.00045 | +0.0045 | XNoisy, High |

The differences are in conductivity and salinity are higher than for the previous cruise, but the shallowness and noisy descent rate may account for this.

##### 9. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### 10. Checking Headers

The header check was run. The fluorescence did not go off-scale; there are some slightly negative values. This should be checked again after DELETE and CTDEDIT. If negative values remain a recalibration will be done to remove them. There are some slightly negative pressure values.

Log records were checked against the cross-reference report and 1 error in station name was found and corrected.

The track plots look reasonable and were added to the end of the report.

The surface report indicates that the average surface pressure was 2.3db which is a little low for Ricker work. A few casts with some negative pressure readings were examined; while they occur with “in-water” salinity values, they mostly look like noise or cases of the CTD bobbing in and out of water at the end of a cast. It is possible that the pressure sensor is reading slightly low, but there is insufficient evidence to justify recalibration.

##### 11. SHIFT

Fluorescence

The WetLabs ECO-AFL fluorometer has been found to need a shift of +48 records in other cruises and tests on a few casts confirm that setting produces good alignment between fluorescence and temperature traces for these data as well.

A shift of +48 records (2s) was applied.

Conductivity

Tests were run on both conductivity channels using 4 casts. For the primary conductivity the best setting was -0.7 records and for the secondary the best varies from feature to feature, but -0.9s was best overall.

All casts were put through SHIFT using -0.7 records for the primary and -0.9s for the secondary conductivity.

Dissolved Oxygen

The SBE DO data have already been aligned, but a few casts were examined to see if further adjustments are required. The vertical offset between upcast and downcast for DO is sometimes closer and sometimes further apart than for temperature, but overall the two look very close.

SHIFT will not be run on the SBE DO channel.

pH

The vertical offset between upcast and downcast for pH is larger than for temperature. Alignment tests were run on a few casts trying different shifts. The best results were with a setting of +72 records.

##### 12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: All warnings concerned the upcast portion of casts except for cast #10 which had noisy pressure in one section. DELETE selected data for that cast appropriately.

##### 13. COMPARISON WITH BOTTLES

There were no deep bottle samples, but there were near-surface samples for most casts. The CTD stopped at between 13 and 16db during upcasts, often dropping during the stop with 14.5db looking like the most common level when the bottle would likely have been closed. So CTD data from about 9.5db was chosen for comparison with the bottle data. Unfortunately, the local gradients at that level were high for most casts. The following steps were taken to obtain useful data for a comparison:

- REVERSE was run on the files after the final alignment step.

- CLIP was run to reduce the data to the top 12db of the final upcast section.

- DELETE was also run on the clipped files to produced *.REVDEL

- Bin-Average was run on the REVDEL files using 0.5db for the bins to produce REVAVG. Standard deviations were included to help reduce the data to casts with low salinity gradient.

- Thin was run to reduce the REVAVG files to a single value at about 9.5db.

- Data from the REVTHN files were exported to a spreadsheet. The bottle data were added.

Casts were removed if there was no salinity sample available or no CTD data available.

The data were then ordered on standard deviation in the salinity standard deviation. When only casts for which the primary salinity had a standard deviation <0.001, the primary salinity was lower than bottles by an average of 0.0122, based on 56 casts. When the same approach was used for the secondary salinity the average was lower than the bottles by 0.0049, based on 54 casts. When more casts were removed by picking a smaller standard deviation limits, the averages were -0.0112 and -0.0068 for standard deviation <0.0005 and -0.01220 and -0.0069 for standard deviation <0.0004. The differences between the two averages for the 2 pairs were 0.007, 0.004 and 0.005 for the standard deviations <0.001, 0.0005 and 0.0004, respectively. The latter two fit the results reported in section 8. The data used in these comparisons can be found in file 2011-45-ctd-bottle-comp.xls. (Note that there are separate tabs for fluorescence and salinity comparisons.)

The ascent rates of the CTD did not vary a lot; sorting on that variable had little effect on fits.

Whether this comparison is sufficient to justify recalibration or to decide which channels to archive is not a clear matter. The actual depths at which bottles were fired are variable and the estimate of 9.5db may be poor. If the values included in the comparison came from cases where bottles closed at 10 or 11db, then the differences would be smaller. There is usually a lot of vertical motion during the stops, so this judgment is not clear. Salinity samples were analyzed soon after collection, so there is no suspicion of evaporation as an issue.

Examination of downcast data shows that both salinity traces have small spikes, but the primary has some larger spikes not seen in the secondary. Upcasts are different with secondary being noisier than the primary. For the profiles, the secondary will be selected since it is closer to the bottles and quieter during downcasts. Recalibration will not be applied. If evidence from later cruises or post-cruise calibration suggests it should have been recalibrated, it can be done later.

A comparison of the fluorometer versus extracted chlorophyll was also done and saved on a separate worksheet in 2011-45-ctd-bottle-comp.xls. There were only 5 bottles with CHL >1.5ug/L. The ratio of FL to CHL ranged from 0.7 for the lowest CHL readings and fell steeply as CHL rose, with a value of 0.12 when CHL reached 0.09ug/L and being approximately constant at 0.05 when CHL>0.5ug/L. A few downcasts were examined to see if values were closer to bottles and they were not. Many bottles were rejected from this comparison because of analysis problems or because there was no CTD fluorescence data at the relevant depth.

File 2011-45-CTD-bottles-summary.xls containing CTD data from 9.5db data plus all bottle sample data.

##### 14. DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to clean the secondary temperature and salinity data for downcasts.

All casts required some editing.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

The edited files were copied to *.EDT files.

##### 15. Other Comparisons

Previous experience with these sensors – all sensors had factory calibrations from March or April 2011.

- Both conductivity sensors have been used on only 1 other processed cruise since the last factory service.  During 2011-58 the bottle comparison was rough, but both sensors were reasonably close to bottles, likely within ±0.002.

- No recalibration was applied to the pressure sensor during 2011-58.

- There was no calibration sampling for the dissolved oxygen sensor during 2011-58.

Historic ranges – There were many excursions from the 3-standard deviation plots of temperature and salinity ranges, but none that were not explained by being close to shore or in inlets, where larger variations are expected and at sites that may not be well represented in the climatology. Others appear to be cases of deep mixing. Most excursions were in the Queen Charlotte Strait with water near the bottom tending to have higher temperature and lower salinity than usual. The excursions are not dramatic enough to suggest instrumental problems. Local climatology data were unavailable for many casts.

##### 16. Initial Recalibration

No recalibration was applied. All indications are that the salinity and pressure data are good, and we have no evidence on which to base a judgment on dissolved oxygen. The most recent cruise that included calibration sampling was before the last factory calibration, so those results are not applicable.

##### 17. Fluorescence Processing

The data from the ECO fluorometer does not require filtering.

##### 18.  BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Though small unstable features remain these may be real, so no further editing was considered necessary.

Profile plots were made to check all channels. The only problem noted is with the fluorescence from Johnstone Strait where it is not clear if there was a malfunction or just extremely low chlorophyll values. Extracted CHL values for the first 2 casts are <0.1 and in Queen Charlotte Strait where there is clearly a signal the values remain extremely low. So it is assumed that the sensor is just not able to register values that low.

##### 19. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  SBE:pH, Status:Pump, Descent_Rate and Flag.

A second set of files were prepared for the Chief Scientist with the pH channel. From this point on both sets are treated in the same way but the 2nd set have “pH” added to the extension in the file names.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. There was a lot of variability with the lowest values (<65%) in Queen Charlotte Strait and some inlets, while the only values >95% were 3 casts in inlets on the west coast of Vancouver Island. Offshore and Queen Charlotte Sound values were between 85% and 95%. These values are not unreasonable, but may be a little low. Calibration drift is generally towards lower values.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and Dissolved Oxygen data are nominal and unedited, except

that some records were removed in editing temperature and salinity.

All sensors were recalibrated in March or April 2011.

There was no calibration sampling for dissolved oxygen.

There was salinity calibration sampling; there are problems with the

comparison between CTD and bottles, but it appears that the

CTD salinity calibration drift is small.

The pH:SBE data were not considered ready for archiving, so that channel was

removed.

For details on the processing see processing report: 2011-45-proc.doc.

The special files for the chief scientist were given the following header comment:

Data Processing Notes:

----------------------

Fluorescence and Dissolved Oxygen data are nominal and unedited, except

that some records were removed in editing temperature and salinity.

All sensors were recalibrated in March or April 2011.

There was no calibration sampling for dissolved oxygen.

There was salinity calibration sampling; there are problems with the

comparison between CTD and bottles, but it appears that the

CTD salinity calibration drift is small.

For details on the processing see processing report: 2011-45-proc.doc.

NOTE: This file was prepared for the use of the Chief Scientist. It includes

the SBE pH channel which is uncalibrated. It is not considered suitable

for archiving.

##### 20. Producing final files

The Standards Check routine was run and no problems were found. The final files were named CTD.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

The special set of files with extensions CTDpH were prepared for the High Seas Salmon program staff; they contain the pH data that were removed from files to be archived.

Particulars:

13. Cast not recorded

142. Pumps not on until 6db down. Affected data removed.

160. Pumps not on until ~76db down. Affected data removed.

169. Pumps not on until ~57db down. Affected data removed.

232. File saved as 231. Fixed.

298. Acquisition did not start until 8db down.

343. Station name wrong in file – should be EP07. Fixed.

361. Very low DO.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 13Apr11 | Factory |  |  |
| Conductivity | 3394 | 29Mar11 | Factory |  |  |
| Secondary Temp. | 4883 | 31Mar11 | Factory |  |  |
| Secondary Cond. | 1763 | 29Mar11 | Factory |  |  |
| pH | 0851 | 21Apr11 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 29Mar2011 | Factory |  |  |
| WetLabs Eco Fluorometer | 2214 |  | IOS |  |  |
| Pressure Sensor | 69698 | 15Apr2011 | Factory |  |  |