## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |

## PROCESSING NOTES

Cruise: 2011-60

Agency: OSD

Location: Strait of Georgia

Project: Moorings / Fraser River Sampling

Party Chief: Johannessen S.

Platform: Vector

Date: August 5, 2011 – August 8, 2011

Processed by: Germaine Gatien

Date of Processing: 10 January 2012 – 17 January 2011

Number of original HEX files: 13

Number of CTD files:   13

Number of bottle casts:	  2 (Casts with maximum pressure <2db were not processed.)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0997), a Wet Labs Eco-AFL/FL Fluorometer (#2215), an SBE18 pH sensor (#0692) and an altimeter (no serial # available).

The deck unit was an SBE 11+ (#0425).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The Rosette and CTD logs were in excellent order with a list of all equipment (except the altimeter) and clear notes about problems encountered. The TSG section included a note to say it was not used.

The SBE data are considered of lower quality than normal because 24 scans were averaged in acquisition due to an improper setting in the configuration file. These averages may contain spikes that would normally be removed in processing and might be systematic in nature. Fine-tuning of alignment settings and CELLTM correction are also less sensitive with smoothed data. Unfortunately the same error was made on several other cruises using the same equipment, so we can not rely on settings determined from those data sets. Given that sampling was in an area that is not usually affected much by shed wakes and the CTD is believed to have worked well with few spikes, it is likely the errors caused by this are small.

The SBE Dissolved Oxygen data was corrected based on bottles, but there are problems in both data sets. As already noted there may be spikes included in the averaged SBE DO data. There are also concerns about sampling protocols. The analyst noted a skew in the duplicates so that the second sample had higher DO values than the first, except for one pair in supersaturated surface waters where the opposite was the case. This suggests that air was drawn into the Niskin between samples. One surface pair does not fit this explanation, though DO variability was high near the surface for that one. It was judged wisest to select the first sample for each pair. No flags were attached, but based on the precision study only 2 decimal places are reported for the titrated DO samples.

For full details see tabs “Notes” and “Precision” in file QF2011-60oxy*.xls.

The CTD salinity appears to be reasonably close to the bottles, but high gradients and limited data in the comparisons make the judgment a rough one. The sensors were freshly recalibrated so were likely working well.

The pH:SBE data are not considered ready archiving, but were supplied to the Chief Scientist.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was a cruise report available which mentioned only one major problem – an announcement by the CTD computer that it had an unauthorized version of Microsoft Office. That was resolved with help via phone. Bubbles were noted in the line for Niskin #17.

A note in the log mentions a problem with the naming of files early in the cruise. Some were corrected at sea. Those for the test cast (#1) were missed, so those were corrected before converting the data.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. None had been used since the latest factory calibrations were done.

The configuration files were identical as far as sensor parameters went, but the first and last files did not have the NMEA box ticked. The calibration constants were checked for all instruments. There were a number of errors in the con files:

- The transmissometer calibration was entered correctly, but starting with 2011 cruises the algorithm for determining the slope and offset has been changed slightly to fit the method recommended by Sea-Bird. So the slope/offset were recalculated and those values entered in the configuration file. The result will be higher values by <1%.

- The calibration for the Wetlabs Eco fluorometer could not be confirmed. The parameters entered are close to those on the characterization sheet from the factory, but a field test was done just before the cruise and no record is available for that, but it looks reasonable

- The most serious error is that “Scans to average” was set to 24. This should be set to 1. There is no way to fix this and acquiring metre-averaged data only.  Single bad points will be included in the average since there is no opportunity to remove them first and fine-tuning of settings will be limited or impossible.

A configuration file used at sea (with NMEA box ticked) was corrected and saved as 2011-60-ctd.xmlcon. Cast #1 was a test only and will not be processed. Positions will be added to cast #15 later.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-60-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. One sample was flagged “5” but not replaced with a pad value, so it was removed. (That sample was later replaced with a pad value.) The analyst expressed doubts about the quality of these data. A calculation was made of Ф using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees. Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.4769 and B = -0.2546 with an R2 value of 0.9971. The initial estimates of M and B were used to calculate a difference between each point and the fit:
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. Each time data were removed the M and B were updated in the difference calculation. When removing a little more data has little effect on the fit, it is judged that a reasonable value has been found unless, in so doing, a whole class of points has been removed such as all high values, or all values from late in the cruise. The following table shows that results:

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4348 | -0.2110 | 0.4348 | -0.4853 |  |
| 23 | all data | 0.4443 | -0.2137 | 0.4443 | -0.4810 | 0.9977 |
| 21 | excl. outliers diff>0.05 | 0.4493 | -0.2214 | 0.4493 | -0.4928 | 0.9991 |
| 18 | excl. outliers diff>0.03 | 0.4410 | -0.2071 | 0.4410 | -0.4696 | 0.9986 |
| 15 | excl. outliers diff>0.02 | 0.4483 | -0.2186 | 0.4483 | -0.4876 | 0.9993 |

With so few bottles to use, a cutoff difference of 0.05 looks reasonable. Removing more data reduces the range of DO values significantly and the R2 value starts to go down. The 2 outliers that were removed were both from near the surface, but a few near-surface bottles are left in the comparison. The outliers were from the last station, SOGS.

The configuration files were updated with the new parameters Soc and Voffset and saved with name 2011-60-ctd-new.xmlcon.

##### Hysteresis Study

This sensor has been recalibrated since the last hysteresis checks were done. There was no deep sampling during this cruise, so hysteresis is not going to be a problem and there is not enough deep data to enable reliable tests to be done to fine-tune H1, H3 and E.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters.

The ROS files were converted to IOS header format.

There was no colon between the name and data for the depth and station name entries so that information was not converted (except for cast #15 which was correct.). This was corrected in the ROS files and conversion was rerun.

There was no position information in file #15 so lines were added to the IOS header for that and the positions entered to match the log book.

The files were put through CLEAN to create BOT files.

A track plot and header check turned up no further problems.

Temperature and salinity were plotted for all BOT files. Most of the samples were very shallow and salinity was very low, so there is a lot of scatter, but no obvious outliers.

The addsamp.csv file prepared in the DO calibration step was ordered on sample #s and converted to CST files to be used as a framework for the bottle files. It was also used to create SAM files with bottle #s and bottle positions. The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets was examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-60-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts. In future references to these files the * stands for the date.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-60nuts*.xls which included a report on precisions.  The file was simplified, reordered on sample numbers, header names were changed to standard format and the file was saved as 2011-60-nuts.csv. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-60oxy*.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified; the file was then saved as 2011-60oxy.csv. That file was converted into individual *.ADD files.

SALINITY

Salinity analysis was done at IOS and saved in file 2011-60sal*.xls. No comments or flags were entered and there is no note of problems on the Autosal analysis log sheets. The file was simplified, flag and event number columns were added; event # information was filled in based on the log book entries.

Station P standards were used, so that information was added to a document tracing the performance of standard samples. Two of those standards were outliers in the comparison of 12 standard samples. This may be a sign that the Autosal does not perform well with such large variations in salinity. The surface salinity values were near zero, so much lower than usually encountered.

Duplicates had not been averaged, so the information was copied to a separate spreadsheet 2011-60-sal-duplicates.xls. A precision study was done and 2 of the 8 samples were found to be Chauvenet outliers.

The average values were substituted in file 2011-60-sal.csv along with “6” flags for the duplicates or “36” for the outlier duplicates.

The data were ordered on event number and sample number.

They were then converted to individual SAL files.

The SAL, OXY and NUT files were merged with CST files in 3 steps. After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

Bottle data were exported to a spreadsheet and compared with the rosette sheets – no problems were found.

##### Compare

Salinity

Compare was run with pressure as reference channel. Given that the salinity varied from 3 to 31 and most bottles were from the surface, it is not surprising that the scatter is large. When only bottles from below 50m are included, the primary and secondary salinity channels are low by an average of 0.0020 and 0.0029 respectively, with a standard deviation of 0.012 for both. The median differences for pressures >50m were -0.0001 and -0.0006. When only the 6 bottles below 200db are included the differences are +0.0016 and +0.0008, but one of those 6 is an outlier. When only 5 are included the differences are 0 and -0.0008.

We cannot expect a good comparison from these data. There are likely problems in analysis due to large changes in salinity and large vertical salinity gradients make the 1.5m between bottles and CTD sensors significant, and not just at the surface. For cast #15 there was a difference of 0.005 over 1.5db around 300db. Moreover, the CTD data is not as reliable as usual due to the averaging of scans over 1s.

There were many outliers, but most are from above 2db with high standard deviation in the CTD data and  very high vertical salinity gradients (changes up to 3 in 1db). Investigations were done on 2 duplicates identified as outliers in the precision study and 1 non-surface outlier identified in COMPARE.

- Cast #12, sample 20: The duplicates were Chauvenet outliers. One differs from the CTD salinity by ~0.18 and the other by 0.38. Given the local gradients either could be accurate. The samples came from <2db. Given the variability the water in the bottle may not be well-mixed, so the Chauvenet criterion is too severe. The “3” flag was removed.

- Cast #15, sample 44: The duplicates were Chauvenet outliers. Both differ from the CTD salinity by >3, but they came from 1.6db in an area with a strong vertical gradient and temporal variation. The bottles would have closed very close to the surface. Given the variability the water in the bottle may not be well-mixed, so the Chauvenet criterion is too severe. The “3” flag will be removed.

- Cast #15, sample #24: This is an outlier in COMPARE. While it is from 250db and the CTD salinity looks quiet during the 10s around firing, there was higher variability during the 20s before that and somewhat lower values, so most of the difference may be due to temporal variations and the vertical salinity gradient. A flag is not justified.

There is insufficient evidence to comment on temporal variations in the comparison.

For details see 2011-60-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When 3 near-surface outliers were excluded, the CTD values were higher by 0.002mL/L with a standard deviation of 0.02mL/L. This is as good a result as we are likely to achieve with this limited data set.

The analyst did not use the Chauvenet criterion because of an apparent skew in the data. The replicates were all considered poor. In 5 pairs the first sample had lower value than the 2nd. Only the surface samples from cast #15 had the 2nd lower.

Nina Nemcek noted the following about these data:

Positively skewed duplicates indicate that this dataset may not have been sampled correctly.  As air enters the top of the Niskin, it creates convective flow and introduces atmospheric oxygen that can quickly contaminate samples with low DO (below saturation).  Sampler(s) likely had too high of a flow rate and probably let a lot of water run through between replicates thus introducing more air than necessary into the top of the Niskin. If sampling is done correctly, there should be no evidence of this kind of positive skew which is exactly why we take replicates, it is more important as a test of sampling proficiency rather than analysis. (As a side note, water samples that are supersaturated would show the opposite trend, with the second sample showing lower DO if flow rate is too vigorous and oxygen is allowed to escape.)

Of the 6 pairs available, the 4 from below the surface had the 2nd value higher than the first (though 1 of the second values was rejected due to analysis problems). For the 2 surface pairs, cast #13 has a lower 1st value, and cast #15 has a higher 1st value. Both are supersaturated during the bottle stop (107% and 112%, respectively) but cast #13 has a lot of variability in the % saturation around 2db (90%-107%) based on the downcast date.  So 4 of the 5 duplicates fit the explanation of poor sampling protocol.

The COMPARE results were examined to see if it supports this interpretation of the precision study. The outliers and replicate samples were investigated in COMPARE:

- Cast #13, sample #45: The 1st duplicate was closer to the CTD value, and since this is a 329db sample, it is probably justified to say the 1st should be used instead of the average.

- Cast #13, sample #51: The 1st duplicate was closer to the CTD value, and since this is a 150db sample, it is probably justified to say the 1st should be used instead of the average.

- Cast #13, sample #67: The 2nd duplicate was closer to the CTD value, but this is a surface sample where the CTD data is noisier and generally less reliable, so the choice of sample should not be based on the CTD value.

- Cast #15, sample #21: The analyst rejected the second value, and the 1st is closer to the CTD value. This was a 310db sample.

- Cast #15, sample #27: The CTD value is roughly half-way between the duplicates, but the standard deviation in the CTD data is higher than usual for 150db sample, so it is not as reliable as usual.

- Cast #15, sample #41: This value is bad, a severe outlier and flagged “5” by analyst; the value was replaced with a pad value.

- Cast #15, sample #43: The 2nd duplicate was slightly closer to the bottle, but since this is a surface sample a comparison with the CTD is not reliable. The standard deviation in the CTD data is fairly high and the sensor does not usually perform well right at the surface.

The CTD comparison supports the use of the first of the duplicates below the surface. The noise in the SBE DO data is too high near the surface to comment on those 2 samples. Since 4 out of 5 pairs support the explanation of sampling problems, it is wise to use the 1st of each pair.

(See 2011-60-dox-comp1*.xls.).

The first sample was selected for each of the replicates. There are no flags assigned since there is no trouble with the analysis and no known problem with the 1st samples. However, the precision study does imply that these data are only good to 2 decimal places, rather than the usual 2.

The merge process was repeated to capture the new flags and comments.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-60-ctd-new.con.

A few casts were examined and all expected channels are present. The conductivity channels track well during downcasts and are close – they are a bit noisier on the upcasts, but not bad. The two temperature traces track well especially during the downcasts though the differences are a little larger than expected in newly calibrated sensors.

The fluorescence looks reasonable with the usual vertical offset. It will be hard to judge the right shift to apply because there are upcast stops for the 2 useful casts.

The DO voltage looks as expected with a vertical offset with some detail to help alignment.

The pH is also offset vertically but will be harder to judge.

Transmissivity looks normal.

The altimetry is noisy at the bottom, but there is a clear signal just above bottom.

##### WILDEDIT

Program WILDEDIT is usually run to remove spikes from the pressure, conductivity & temperature channels. Because the data are already averaged, this step will be skipped.

##### ALIGN DO

Tests were done on 2 casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge in these data because of the lack of detail due to averaging during acquisition, but a setting of +4.5s looks reasonable.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

Tests were run comparing a variety of settings for CELLTM using the only 2 casts that sampled below the surface. The goal is to make upcasts look closer to downcasts on a T-S surface. The test did not work very well since there is so little data, but all choices looked better than no application, with the choice of (α = 0.0245, β=9.5) seeming best overall for both channels.

CELLTM was run on all casts using setting (α = 0.0245, β=9.5) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on 2 casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The two deep casts were plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 13 | 320 | +0.0005 | -0.00004 | -0.0008 | High, F. Steady |
| 15 | 300 | +0.0007 | -0.00003 | -0.0009 | High, F. Steady |

The differences in conductivity are very small while those in temperature are a little bigger than might be expected so soon after recalibration, but not bad. The differences in salinity are in good agreement with the results of COMPARE when only the 5 bottles were included that led to flat fits. For those the primary salinity was closest to the bottles.

##### Conversion to IOS Headers

From this point onwards only casts #13 and 15 will be processed. The other casts sampled only the top 2db and only bottle files will be prepared for those.

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. One bad DO value had to replaced in the upcast section of each cast before conversion worked.

Position information was added to the header of cast #15.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header checks were run on the bottle files since most casts have no full profile file. The minimum fluorescence value was -0.03mg/m^3, so CALIBRATE will be used to add that amount so there are no negative values.

Surface check was run and shows an average surface pressure for the cruise was 1.7db using the bottle files and 1.3db for the 2 deep casts. Given this is an average value the actual minima must have been lower. A low value is expected from this project where near-surface samples are required, though the shallowest values would have been in the upcast in most cases. There is no way to check the pressure accuracy from averaged data.

The cross-reference check was compared with the log book and no problems were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. The values for the 2 deep casts look reasonable. The Water Depth headers look fine.

The altimeter readings from the headers of the BOT files were exported to a spreadsheet. The algorithm failed on 4 of the shallow casts for which the CTD did not sample within 15db of the bottom; those header entries were removed. For 2 other casts it is not absolutely clear the headers are wrong, but they do not agree well with the log records. Surface altimetry tends to be unreliable and these are averaged values, so it is hard to judge what was happening. The altimetry headers were removed for all the shallow casts.

##### 13. Shift

Fluorescence

Fluorometers usually require alignment, either to remove the effects of pumping for SeaPoint sensors or to correct for slow response time in ECO sensors. Data from only one ECO sensor (#2216) has been processed up to this time, and it required an advance of ~2s. The slow response of that sensor surprised the manufacturer who feels it should be fast. These data are from a different sensor (#2215).

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. Usually an advance of +48 records is found appropriate. Dividing by 24 to allow for the averaging, this is equivalent to +2 records. Tests comparing -5 to +5 showed +2 records to be reasonable. SHIFT was applied with a +2 records setting.

Conductivity

The pre-averaging makes the alignment more difficult to asses. Tests were run using values from -0.1 records to +0.1 and none look any better than the original. No shift was applied.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked and no further alignment was found necessary. SHIFT was not run on DO.

pH

When this sensor has been used unpumped in the past a shift of +90 records was found to be the best setting for alignment. Tests of these data show an advance of +4 records provides good results. That is equivalent to +96 records before 1s averaging. SHIFT was run using +4 records.

##### 14. DELETE

The following DELETE parameters were used (adjustments were made to the usual settings to allow for the fact the data are already averaged):

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

**Pressure filtered over 15 points**NOT APPLIED

Swells deleted. Warning message if pressure difference of 2.00

**Drop rates <   0.30m/s (calculated over 11 points) **NOT APPLIED.

**Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval:  taken from header

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  The sensors were both recalibrated in late March 2011 and this was the first known use since that time.

2. Dissolved Oxygen: The DO sensor was repaired and recalibrated in April 2011. This was the first known use after that date.

3. Pressure: The sensor was recalibrated in April 2011 and this was the first known use since that time.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data fell within the ranges, though salinity was near the minimum at station SOGS above 35db. This is likely due to the high run-off from the Fraser River in spring 2011.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

At this stage a decision has to be made about whether to edit primary or secondary channels. The primary is slightly closer to the bottles and given only averaged data it is unknown which was cleaner. The primary channels were selected.

CTDEDIT was used to lightly clean a few points in salinity in both casts.

##### Initial Recalibration

File 2011-60-recal.ccf was prepared to add 0.092mg/m^3 to the fluorescence. While no values that small were found in the profiles, there were some negative values of that size in the bottle files. Given that averaging had been applied and to be consistent with the bottle files, a choice of +0.092 was made.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response-time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are compared to bottle data from the same pressure. For this cruise the twin problems of data smoothed during acquisition and less reliable DO samples mean the comparison is unlikely to be sensitive enough to justify further recalibration, but is still worth doing to get some idea of the accuracy of the DO:SBE data.

The downcast files are usually averaged in 0.5m bins, but in this case the full files were thinned to the levels at which bottles are usually fired. COMPARE was then run to compare the CTD and bottle data. When differences are plotted against pressure the CTD is low by an average of 0.085mL/L with a standard deviation of 0.16. When only data below 15db are included the CTD is low by 0.006mL/L with a standard deviation of 0.038. Given the problems with both data sets that is as good as we can expect. No further recalibration is justified.

##### Special Fluorometer Processing

There was no CHL sampling, so special files were not prepared for Dr. Peña.

No filtering is required for the ECO fluorescence.

##### BIN AVERAGE of CTD files

While the data were average over 24 scans, the descent rate is often less than 1m/s so Bin Average was run on the COR1 files using the following settings:

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

Special files were prepared for the use of the Chief Scientist with the pH channel included. For those files REMOVE was used to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, and to add the following comments:

Data Processing Notes:

----------------------

The SBE data are considered of lower quality than normal because 24 scans

were averaged in acquisition. These averages may contain spikes that

would normally be removed in processing and might be systematic in nature.

Fine-tuning of alignment settings and CELLTM correction are also less

sensitive with smoothed data. Unfortunately the same error was made on

several other cruises using the same equipment, so we can not rely on

settings determined from those data sets. Given the area is in protected

waters and the CTD is believed to have worked well with few spikes, it is

likely the errors caused by this problem are small.

Fluorescence and Transmissivity data are nominal and unedited.

The pH:SBE data were not considered ready for archiving.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The tests for accuracy of Oxygen:Dissolved:SBE data are less reliable than

usual, but the data are considered, very roughly, to be:

±0.6mL/L from 0 to 25db

±0.1mL/L below 25db

For details on the processing see processing report: 2011-60-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found except for formats that were deliberately changed to reflect quality concerns.

The final files were named CTD.

Profile and T-S plots were examined and no problems found.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were both ~110%.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second set of files was created that include pH:SBE. For those REMOVE run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together. (This was also done to the special pH files.)

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data. Those files were named *.CHE.

The special files for the chief scientist were put through HEADER EDIT and saved as *.CHEX.

A header check was run on the final files and no problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Data from the CHE files were exported to spreadsheet 2011-60-bottles-final.xls; no errors were found.

Standards check was run on all files and no errors were found except some formats that were changed deliberately to reflect data quality.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (Notes from log book and sampling notes)

3. Late starting pump

15. Windows says it is not authentic. CTD programme settings lost.

Niskin #17 – when venting gurgles and whistles a lot – seems to produce a lot of small bubbles in the line.

Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 4Aug2011 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 23Apr2011 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2215 | 4Aug2011 |  |  |  |
| SBE pH | 0692 | 29Dec2010 |  |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |