## comparison-0997
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
7 | 9.67 | 11.5781 | 11.5787 | 27.2885 | 27.2919 | 6.4104 | 1.8324 | 6 | 6 | 30 | 6.5402 | 0.59693 | 3.904 |  | 0.00045084 | 0.4398 | -0.2085 |  |  |  | Summary of Soc Voffset including the original values in the factory calibration
18 | 125.173 | 8.1747 | 8.1757 | 32.6424 | 32.6386 | 6.6756 | 1.0004 | 4 | 4 | 98 | 6.8723 | 0.23224 | 1.596 |  | 0.00077051 | [7 empty] | m | b | Soc | Voffset | R2
3 | 9.469 | 11.6346 | 11.6352 | 27.32 | 27.32 | 6.4014 | 1.9665 | 5 | 5 | 9 | 6.5314 | 0.6556 | 4.282 |  | 0.00077063 | [5 empty] | Bottles used | Original | 0.4348 | -0.21101 | 0.4348 | -0.4853 |  | 9.1691 | 2.956
6 | 31.132 | 12.4042 | 12.4053 | 27.9854 | 27.9862 | 6.271 | 1.8399 | 6 | 6 | 21 | 6.4242 | 0.60147 | 3.864 |  | 0.00079344 | [5 empty] | 149 | all unflagged bottles | 0.4395 | -0.2026 | 0.4395 | -0.46098 | 0.9404 | 9.3324 | 3.0596
30 | 74.452 | 7.9712 | 7.9659 | 32.8616 | 32.8706 | 6.6967 | 1.1117 | 2 | 2 | 161 | 6.8471 | 0.28129 | 1.926 |  | 0.00087421 | [5 empty] | 138 | excl. outliers diff>0.1 | 0.4392 | -0.2069 | 0.4392 | -0.47108 | 0.9937 | 9.2994 | 3.0277
25 | 83.245 | 8.2102 | 8.2151 | 32.6115 | 32.6093 | 6.6716 | 1.0678 | 1 | 1 | 126 | 6.8318 | 0.25996 | 1.776 |  | 0.0011666 | [5 empty] | 128 | excl. outliers diff>0.05 | 0.4398 | -0.2085 | 0.4398 | -0.47408 | 0.9966 | 9.3042 | 3.023
27 | 95.051 | 7.9687 | 7.9705 | 32.867 | 32.8678 | 6.6969 | 1.0006 | 1 | 1 | 139 | 6.8653 | 0.23014 | 1.58 |  | 0.0014392 | [5 empty] | 117 | excl. outliers diff>0.04 | 0.445 | -0.2188 | 0.445 | -0.49169 | 0.9978 | 9.3672 | 3.0063
24 | 74.906 | 8.2145 | 8.2149 | 32.591 | 32.5913 | 6.6719 | 1.0932 | 3 | 3 | 120 | 6.8248 | 0.27078 | 1.848 |  | 0.0014907 | [5 empty] | 109 | excl. outliers diff>0.03 | 0.4451 | -0.2191 | 0.4451 | -0.49225 | 0.9985 | 9.3678 | 3.0053
24 | 99.207 | 8.1873 | 8.1877 | 32.6151 | 32.6158 | 6.6749 | 1.0619 | 2 | 2 | 119 | 6.8489 | 0.25698 | 1.76 |  | 0.0015424 | [5 empty] | 93 | excl. outliers diff>0.02 | 0.4474 | -0.2213 | 0.4474 | -0.49464 | 0.9991 | 9.4098 | 3.0137
23 | 99.446 | 8.1823 | 8.1829 | 32.6152 | 32.6154 | 6.6756 | 1.067 | 3 | 3 | 111 | 6.8498 | 0.26264 | 1.799 |  | 0.0018692 | [5 empty] | 79 | excl. outliers diff>0.015 | 0.4457 | -0.2177 | 0.4457 | -0.48845 | 0.9992 | 9.3906 | 3.0207
30 | 19.732 | 10.0906 | 10.0757 | 31.5886 | 31.5948 | 6.4416 | 1.8712 | 5 | 5 | 164 | 6.5638 | 0.61641 | 4.046 |  | 0.0019642 | [5 empty] | 54 | excl. outliers diff>0.01 | 0.4443 | -0.2159 | 0.4443 | -0.48593 | 0.9996 | 9.3678 | 3.0187
28 | 19.625 | 9.8848 | 9.8871 | 31.621 | 31.6207 | 6.4693 | 1.7613 | 5 | 5 | 150 | 6.5896 | 0.56832 | 3.745 |  | 0.0021941
6 | 120.092 | 13.9613 | 13.9617 | 29.4804 | 29.481 | 6.0147 | 0.51342 | 2 | 2 | 17 | 6.247 | 0.014727 | 0.092 | 6 | 0.0025749
29 | 85.664 | 7.9599 | 7.973 | 32.8891 | 32.8783 | 6.6973 | 1.1151 | 1 | 1 | 153 | 6.8574 | 0.27897 | 1.913 |  | 0.0029594
23 | 74.612 | 8.2014 | 8.2019 | 32.5817 | 32.5818 | 6.6742 | 1.1034 | 4 | 4 | 112 | 6.8268 | 0.27363 | 1.868 |  | 0.0031403
30 | 96.816 | 7.8137 | 7.8177 | 33.0759 | 33.0748 | 6.7113 | 1.1568 | 1 | 1 | 160 | 6.8798 | 0.2971 | 2.044 |  | 0.0031434
23 | 9.591 | 10.8141 | 10.7903 | 31.353 | 31.3563 | 6.3509 | 2.1329 | 8 | 8 | 116 | 6.4709 | 0.72586 | 4.697 | 6 | 0.0036759 | [6 empty] | Results from 3 cruises with the same sensor
6 | 19.524 | 12.2711 | 12.2734 | 27.861 | 27.8622 | 6.2936 | 1.8957 | 7 | 7 | 22 | 6.4365 | 0.62145 | 4 |  | 0.003784 | [7 empty] | m | b | Soc | Voffset | R2
29 | 2.49 | 17.3668 | 17.5889 | 27.8874 | 27.7839 | 5.6745 | 3.5747 | 7 | 7 | 159 | 5.8396 | 1.3595 | 7.939 |  | 0.0041262 | [6 empty] | 2011-60 | 0.4493 | -0.2214 | -0.4493 | -0.4928 | 0.9991 | 9.4548
23 | 124.637 | 8.1816 | 8.1821 | 32.6387 | 32.6389 | 6.6747 | 1.0293 | 2 | 2 | 110 | 6.871 | 0.24 | 1.649 |  | 0.0041866 | [6 empty] | 2011-63 | 0.4474 | -0.2213 | 0.4474 | -0.4946 | 0.9991 |  | -3.032
28 | 29.879 | 9.3868 | 9.3907 | 31.7844 | 31.7825 | 6.5337 | 1.6333 | 4 | 4 | 149 | 6.6583 | 0.50539 | 3.365 |  | 0.0044261 | [6 empty] | 2011-10 | 0.4417 | -0.1979 | 0.4417 | -0.448 | 0.9989 | 9.4134 | 3.1133
16 | 1.958 | 18.0887 | 18.0907 | 23.4182 | 23.9527 | 5.7479 | 3.7685 | 8 | 8 | 88 | 5.9214 | 1.4534 | 8.606 |  | 0.0044748
9 | 65.122 | 13.1048 | 13.109 | 28.3303 | 28.3363 | 6.1661 | 2.1452 | 1 | 1 | 38 | 6.3513 | 0.73969 | 4.698 |  | 0.0047392
7 | 19.078 | 12.1857 | 12.1801 | 27.8008 | 27.7986 | 6.3074 | 1.855 | 5 | 5 | 29 | 6.4493 | 0.61216 | 3.948 |  | 0.0048131
11 | 1.881 | 14.1572 | 14.0797 | 28.2973 | 28.2958 | 6.0347 | 2.4411 | 6 | 6 | 54 | 6.1776 | 0.86021 | 5.314 | 6 | 0.0049031 | [5 empty] | Summary of Soc Voffset including the original values in the factory calibration
23 | 49.959 | 8.5811 | 8.5806 | 32.2019 | 32.2021 | 6.6339 | 1.299 | 5 | 5 | 113 | 6.7685 | 0.35724 | 2.418 |  | 0.005554 | [7 empty] | m | b | Soc | Voffset | R2
8 | 49.731 | 13.0502 | 13.0505 | 28.2538 | 28.2545 | 6.1761 | 2.1414 | 2 | 2 | 33 | 6.3487 | 0.73937 | 4.694 |  | 0.0061007
23 | 19.965 | 9.8844 | 9.8902 | 31.5943 | 31.5926 | 6.4705 | 1.8512 | 7 | 7 | 115 | 6.5911 | 0.61189 | 4.033 |  | 0.0062113
13 | 19.277 | 13.1547 | 13.1574 | 28.955 | 28.9518 | 6.1357 | 2.2487 | 5 | 5 | 68 | 6.2842 | 0.78674 | 4.944 |  | 0.0062729
27 | 74.683 | 7.9927 | 7.9897 | 32.8333 | 32.8379 | 6.6947 | 1.0345 | 2 | 2 | 140 | 6.8455 | 0.24016 | 1.644 |  | 0.0062933
17 | 50.217 | 8.578 | 8.5789 | 32.207 | 32.2078 | 6.6341 | 1.2965 | 2 | 2 | 90 | 6.769 | 0.3553 | 2.405 | 6 | 0.0064028
29 | 74.598 | 8.0013 | 8.0022 | 32.8124 | 32.8129 | 6.6943 | 1.1214 | 2 | 2 | 154 | 6.8451 | 0.27801 | 1.903 |  | 0.006666
9 | 49.473 | 13.0379 | 13.0396 | 28.244 | 28.2456 | 6.178 | 2.0877 | 2 | 2 | 39 | 6.3504 | 0.7165 | 4.55 |  | 0.0068334
12 | 147.475 | 13.0117 | 13.012 | 29.1486 | 29.1493 | 6.1466 | 2.1923 | 1 | 1 | 55 | 6.3961 | 0.76265 | 4.878 |  | 0.006952
24 | 19.859 | 9.8116 | 9.8103 | 31.6347 | 31.6349 | 6.4791 | 1.7829 | 6 | 6 | 123 | 6.599 | 0.58266 | 3.845 |  | 0.0070449
11 | 9.543 | 13.1583 | 13.1583 | 28.6013 | 28.6023 | 6.1488 | 2.2525 | 5 | 5 | 53 | 6.2899 | 0.78936 | 4.965 |  | 0.0072002
15 | 30.601 | 8.9604 | 8.9635 | 31.749 | 31.7489 | 6.5973 | 1.2685 | 3 | 3 | 77 | 6.7189 | 0.34217 | 2.299 |  | 0.0072323
12 | 1.569 | 13.6436 | 13.6362 | 28.3835 | 28.3864 | 6.0955 | 2.3347 | 9 | 9 | 63 | 6.2342 | 0.82577 | 5.148 |  | 0.0074854
7 | 49.343 | 13 | 13.001 | 28.2035 | 28.2044 | 6.1845 | 2.0929 | 3 | 3 | 27 | 6.3565 | 0.71958 | 4.574 |  | 0.0076161
6 | 51.288 | 12.8517 | 12.8522 | 28.167 | 28.1677 | 6.2051 | 1.914 | 5 | 5 | 20 | 6.3777 | 0.64098 | 4.088 |  | 0.0077057
12 | 74.742 | 13.0231 | 13.0231 | 29.1119 | 29.1128 | 6.1465 | 2.2107 | 4 | 4 | 58 | 6.3379 | 0.77154 | 4.89 |  | 0.0077595
11 | 50.132 | 13.0121 | 13.0124 | 29.024 | 29.025 | 6.1513 | 2.209 | 2 | 2 | 50 | 6.3232 | 0.77098 | 4.875 |  | 0.0079576
25 | 29.42 | 9.3166 | 9.3136 | 31.8132 | 31.8143 | 6.5426 | 1.6283 | 4 | 4 | 129 | 6.6662 | 0.49894 | 3.326 |  | 0.0086726
11 | 19.955 | 13.0842 | 13.0846 | 28.7475 | 28.7483 | 6.1527 | 2.2155 | 4 | 4 | 52 | 6.3014 | 0.77459 | 4.881 |  | 0.0086957
12 | 10.048 | 13.1144 | 13.1147 | 28.7868 | 28.7913 | 6.1473 | 2.237 | 8 | 8 | 62 | 6.2883 | 0.78415 | 4.931 |  | 0.0088316
27 | 20.001 | 9.8262 | 9.8383 | 31.6408 | 31.6374 | 6.4768 | 1.7577 | 5 | 5 | 143 | 6.5969 | 0.55556 | 3.665 |  | 0.0089554
23 | 30.171 | 9.3752 | 9.3761 | 31.7709 | 31.7708 | 6.5359 | 1.6635 | 6 | 6 | 114 | 6.6607 | 0.51406 | 3.424 |  | 0.0090541
3 | 49.73 | 13.0573 | 13.0589 | 28.2528 | 28.2542 | 6.1752 | 2.1371 | 2 | 2 | 6 | 6.3478 | 0.74057 | 4.701 |  | 0.0091561
11 | 73.687 | 12.9739 | 12.9743 | 29.1127 | 29.1135 | 6.1528 | 2.1821 | 1 | 1 | 49 | 6.3431 | 0.76036 | 4.823 |  | 0.009167
6 | 99.38 | 13.9546 | 13.9549 | 29.4363 | 29.437 | 6.0171 | 0.5084 | 3 | 3 | 18 | 6.2332 | 0.0056151 | 0.035 |  | 0.0094793
26 | 30.126 | 9.3134 | 9.32 | 31.8262 | 31.8238 | 6.5426 | 1.6146 | 3 | 3 | 135 | 6.6667 | 0.49185 | 3.279 |  | 0.0097738
12 | 99.812 | 13.0075 | 13.0078 | 29.1455 | 29.1463 | 6.1472 | 2.2021 | 3 | 3 | 57 | 6.3585 | 0.76983 | 4.895 |  | 0.0098308
24 | 9.744 | 10.8842 | 10.8809 | 31.3385 | 31.3404 | 6.3419 | 2.1754 | 7 | 7 | 124 | 6.4627 | 0.7384 | 4.772 | 6 | 0.0098538
12 | 30.006 | 13.13 | 13.1302 | 28.9579 | 28.9578 | 6.1388 | 2.2431 | 6 | 6 | 60 | 6.2955 | 0.78786 | 4.96 |  | 0.0098542
12 | 124.697 | 13.0046 | 13.0048 | 29.1437 | 29.1445 | 6.1477 | 2.1918 | 2 | 2 | 56 | 6.3789 | 0.76549 | 4.883 |  | 0.010044
12 | 19.67 | 13.1213 | 13.1218 | 28.8275 | 28.8288 | 6.1449 | 2.2393 | 7 | 7 | 61 | 6.2935 | 0.78652 | 4.95 |  | 0.010188
24 | 29.999 | 9.3638 | 9.363 | 31.7728 | 31.7731 | 6.5375 | 1.6618 | 5 | 5 | 122 | 6.662 | 0.51201 | 3.411 |  | 0.010338
6 | 9.311 | 11.58 | 11.5801 | 27.3371 | 27.3303 | 6.4082 | 1.735 | 8 | 8 | 23 | 6.5376 | 0.56488 | 3.693 |  | 0.010339
12 | 49.441 | 13.0652 | 13.0652 | 29.0626 | 29.0636 | 6.143 | 2.2254 | 5 | 5 | 59 | 6.3146 | 0.78073 | 4.93 |  | 0.010477
10 | 53.182 | 13.1304 | 13.1308 | 28.3687 | 28.3692 | 6.1613 | 2.1837 | 1 | 1 | 44 | 6.3371 | 0.76281 | 4.834 |  | 0.010895
11 | 29.926 | 13.0733 | 13.0736 | 28.8054 | 28.8066 | 6.1519 | 2.2178 | 3 | 3 | 51 | 6.3083 | 0.77786 | 4.907 |  | 0.010952
13 | 9.647 | 13.1379 | 13.1379 | 28.9852 | 28.9869 | 6.1367 | 2.2537 | 6 | 6 | 69 | 6.2774 | 0.79364 | 4.982 |  | 0.010971
13 | 74.898 | 12.7325 | 12.7321 | 29.6316 | 29.6333 | 6.1639 | 2.1764 | 2 | 2 | 65 | 6.3529 | 0.75997 | 4.828 |  | 0.011267
13 | 30.178 | 12.9694 | 12.9728 | 29.2623 | 29.2589 | 6.1476 | 2.2288 | 4 | 4 | 67 | 6.3031 | 0.78311 | 4.936 | 6 | 0.01137
7 | 29.645 | 12.4263 | 12.426 | 27.978 | 27.9786 | 6.2684 | 1.8597 | 4 | 4 | 28 | 6.4206 | 0.62082 | 3.986 |  | 0.011402
25 | 1.655 | 18.4717 | 18.6184 | 25.3271 | 25.0615 | 5.6407 | 3.5261 | 7 | 7 | 132 | 5.8142 | 1.3309 | 7.738 |  | 0.011412
13 | 49.988 | 12.9708 | 12.9746 | 29.2613 | 29.257 | 6.1475 | 2.2226 | 3 | 3 | 66 | 6.3187 | 0.78054 | 4.932 |  | 0.011521
13 | 88.205 | 12.676 | 12.6761 | 29.7323 | 29.7332 | 6.1673 | 2.1463 | 1 | 1 | 64 | 6.3665 | 0.74719 | 4.757 |  | 0.011751
25 | 19.965 | 9.7968 | 9.7965 | 31.6384 | 31.6394 | 6.4811 | 1.7762 | 5 | 5 | 130 | 6.6009 | 0.56038 | 3.699 |  | 0.012281
9 | 19.692 | 12.2595 | 12.2592 | 27.8349 | 27.8355 | 6.2962 | 1.9044 | 4 | 4 | 41 | 6.4392 | 0.64139 | 4.13 |  | 0.01231
9 | 29.594 | 12.7574 | 12.7558 | 28.0446 | 28.0449 | 6.2222 | 2.0258 | 3 | 3 | 40 | 6.3768 | 0.69486 | 4.431 |  | 0.0124
3 | 68.614 | 13.1002 | 13.101 | 28.3228 | 28.3238 | 6.167 | 2.1209 | 1 | 1 | 5 | 6.3549 | 0.73722 | 4.685 |  | 0.012953
13 | 1.535 | 13.1422 | 13.1433 | 28.9798 | 28.9796 | 6.1364 | 2.2521 | 7 | 7 | 70 | 6.2707 | 0.79513 | 4.986 |  | 0.013144
3 | 20.027 | 12.2496 | 12.2488 | 27.8147 | 27.8143 | 6.2983 | 1.9416 | 4 | 4 | 8 | 6.4415 | 0.6587 | 4.243 | 6 | 0.013272
30 | 49.978 | 8.7709 | 8.7631 | 32.1431 | 32.1481 | 6.6083 | 1.4547 | 3 | 3 | 162 | 6.7446 | 0.41752 | 2.816 |  | 0.01378
27 | 29.695 | 9.3773 | 9.378 | 31.8014 | 31.8014 | 6.5344 | 1.63 | 4 | 4 | 142 | 6.6587 | 0.49439 | 3.292 |  | 0.013982
16 | 99.648 | 8.2149 | 8.2154 | 32.5594 | 32.559 | 6.6732 | 0.50664 | 2 | 2 | 82 | 6.8478 | 0 | 0 |  | 0.01432
7 | 74.576 | 13.0725 | 13.0738 | 28.2767 | 28.2777 | 6.1723 | 2.061 | 2 | 2 | 26 | 6.3649 | 0.68359 | 4.351 | 6 | 0.014335
1 | 34.893 | 13.5876 | 13.588 | 29.4328 | 29.4324 | 6.0629 | 0.5076 | 3 | 3 | 3 | 6.2262 | 0 | 0 |  | 0.014742
27 | 49.897 | 8.77 | 8.7701 | 32.124 | 32.1245 | 6.6093 | 1.4051 | 3 | 3 | 141 | 6.7455 | 0.39374 | 2.656 |  | 0.015715
16 | 113.973 | 8.1682 | 8.1689 | 32.5702 | 32.5703 | 6.6797 | 0.50991 | 1 | 1 | 81 | 6.8666 | 0 | 0 |  | 0.015758
1 | 40.001 | 13.5879 | 13.5884 | 29.4328 | 29.4326 | 6.0628 | 0.51004 | 2 | 2 | 2 | 6.2302 | 0 | 0 |  | 0.015816
4 | 39.856 | 13.657 | 13.6576 | 29.41 | 29.4106 | 6.055 | 0.5101 | 2 | 2 | 12 | 6.2228 | 0 | 0 |  | 0.015842
29 | 49.984 | 8.75 | 8.7506 | 32.1641 | 32.1641 | 6.6105 | 1.4434 | 3 | 3 | 155 | 6.7466 | 0.41043 | 2.769 |  | 0.015892
18 | 149.99 | 8.2665 | 8.2683 | 32.7754 | 32.7749 | 6.6561 | 0.56417 | 3 | 3 | 97 | 6.875 | 0.022982 | 0.158 |  | 0.01664
17 | 61.039 | 8.2675 | 8.2604 | 32.5152 | 32.5217 | 6.6672 | 1.1894 | 1 | 1 | 89 | 6.8086 | 0.33135 | 2.256 |  | 0.016745
8 | 19.457 | 12.1819 | 12.1872 | 27.8005 | 27.7995 | 6.3079 | 1.867 | 4 | 4 | 35 | 6.4501 | 0.6296 | 4.061 |  | 0.016989
15 | 19.843 | 9.4694 | 9.4795 | 31.5234 | 31.5235 | 6.5328 | 1.8244 | 4 | 4 | 78 | 6.6497 | 0.57671 | 3.835 |  | 0.017161
3 | 29.818 | 12.7926 | 12.7905 | 28.0412 | 28.0401 | 6.2177 | 2.0623 | 3 | 3 | 7 | 6.3728 | 0.7157 | 4.561 |  | 0.017203
6 | 120.04 | 13.9613 | 13.9616 | 29.4799 | 29.4808 | 6.0147 | 0.5144 | 1 | 1 | 16 | 6.2469 | 0 | 0 | 6 | 0.017733
15 | 49.478 | 8.5935 | 8.5945 | 32.2111 | 32.2114 | 6.6317 | 0.91137 | 2 | 2 | 76 | 6.766 | 0.17455 | 1.181 | 6 | 0.017771
4 | 47.367 | 13.5537 | 13.5543 | 29.4216 | 29.4221 | 6.0675 | 0.51555 | 1 | 1 | 11 | 6.2404 | 0 | 0 |  | 0.018239
24 | 117.886 | 8.1909 | 8.1912 | 32.6488 | 32.6491 | 6.6729 | 0.97053 | 1 | 1 | 118 | 6.8633 | 0.23793 | 1.633 |  | 0.019595
7 | 101.876 | 13.9255 | 13.9263 | 29.4249 | 29.4268 | 6.0211 | 0.52032 | 1 | 1 | 25 | 6.2391 | 0 | 0 |  | 0.020337
10 | 19.535 | 12.2858 | 12.2491 | 27.7585 | 27.7452 | 6.2957 | 1.9674 | 3 | 3 | 46 | 6.4388 | 0.67714 | 4.36 |  | 0.020377
10 | 9.79 | 12.6576 | 12.6592 | 27.6952 | 27.6878 | 6.2489 | 2.0876 | 4 | 4 | 47 | 6.3872 | 0.68873 | 4.399 |  | 0.020909
24 | 49.842 | 8.5888 | 8.5913 | 32.2002 | 32.1989 | 6.6328 | 1.3132 | 4 | 4 | 121 | 6.7674 | 0.34769 | 2.353 |  | 0.021364
15 | 66.974 | 8.4878 | 8.4904 | 32.3286 | 32.3301 | 6.6423 | 0.90177 | 1 | 1 | 75 | 6.7908 | 0.16625 | 1.129 |  | 0.021845
18 | 197.637 | 8.2432 | 8.2638 | 32.8303 | 32.8164 | 6.6572 | 0.53263 | 1 | 1 | 95 | 6.9179 | 0.0030356 | 0.021 |  | 0.022715
16 | 30.006 | 9.0699 | 9.0702 | 31.7458 | 31.746 | 6.5814 | 1.5237 | 5 | 5 | 85 | 6.7034 | 0.43874 | 2.941 |  | 0.022879
29 | 19.889 | 9.9241 | 9.9267 | 31.6206 | 31.6194 | 6.4638 | 1.7734 | 5 | 5 | 157 | 6.5846 | 0.59472 | 3.916 |  | 0.023277
16 | 19.927 | 9.6231 | 9.6472 | 31.5563 | 31.5455 | 6.5093 | 1.9494 | 6 | 6 | 86 | 6.6277 | 0.62556 | 4.146 |  | 0.023304
1 | 49.473 | 13.5876 | 13.5701 | 29.4294 | 29.411 | 6.063 | 0.52935 | 1 | 1 | 1 | 6.2378 | 0 | 0 |  | 0.024308
10 | 30.112 | 13.002 | 13.0124 | 28.1415 | 28.1443 | 6.1866 | 2.17 | 2 | 2 | 45 | 6.3433 | 0.77073 | 4.889 |  | 0.024858
18 | 175.347 | 8.2486 | 8.2498 | 32.8194 | 32.8185 | 6.6568 | 0.56079 | 2 | 2 | 96 | 6.8979 | 0.012178 | 0.084 |  | 0.025958
28 | 2.209 | 17.394 | 17.1697 | 27.228 | 27.358 | 5.6953 | 3.616 | 7 | 7 | 152 | 5.861 | 1.3554 | 7.944 |  | 0.026409
23 | 1.871 | 17.1391 | 17.1531 | 27.1652 | 27.1164 | 5.7247 | 3.7554 | 9 | 9 | 117 | 5.8887 | 1.4164 | 8.341 |  | 0.026658
18 | 100.146 | 8.172 | 8.173 | 32.6155 | 32.6143 | 6.6772 | 1.0205 | 5 | 5 | 99 | 6.8519 | 0.26752 | 1.833 |  | 0.027202
10 | 1.743 | 15.1137 | 15.1903 | 25.0036 | 24.8707 | 6.0398 | 2.7466 | 5 | 5 | 48 | 6.1924 | 0.97215 | 6.02 |  | 0.027318
27 | 1.751 | 17.6719 | 17.4116 | 26.1086 | 26.3901 | 5.7022 | 3.4514 | 7 | 7 | 145 | 5.8703 | 1.34 | 7.866 |  | 0.030553
25 | 49.639 | 8.6057 | 8.6039 | 32.2095 | 32.2121 | 6.6299 | 1.3198 | 3 | 3 | 128 | 6.7645 | 0.34105 | 2.307 |  | 0.03092
16 | 9.795 | 10.4286 | 10.4805 | 31.3544 | 31.3508 | 6.404 | 2.3438 | 7 | 7 | 87 | 6.521 | 0.85524 | 5.577 |  | 0.032957
18 | 74.962 | 8.188 | 8.189 | 32.5906 | 32.5902 | 6.6758 | 1.038 | 6 | 6 | 100 | 6.8287 | 0.28117 | 1.92 |  | 0.033156
8 | 9.64 | 11.6645 | 11.668 | 27.337 | 27.3399 | 6.3966 | 2.0049 | 5 | 5 | 36 | 6.527 | 0.63689 | 4.157 |  | 0.036362
26 | 59.789 | 8.2338 | 8.2346 | 32.5849 | 32.5862 | 6.6692 | 1.1225 | 1 | 1 | 133 | 6.8092 | 0.32265 | 2.197 |  | 0.037487
6 | 74.381 | 12.9475 | 12.949 | 28.2433 | 28.244 | 6.1898 | 1.8409 | 4 | 4 | 19 | 6.3814 | 0.56304 | 3.593 |  | 0.038078
27 | 9.676 | 10.8347 | 10.8286 | 31.369 | 31.3693 | 6.3474 | 2.1054 | 6 | 6 | 144 | 6.4677 | 0.67922 | 4.393 |  | 0.038212
8 | 29.193 | 12.6316 | 12.6181 | 28.0235 | 28.0191 | 6.2395 | 1.939 | 3 | 3 | 34 | 6.3929 | 0.68436 | 4.375 |  | 0.040064
19 | 38.296 | 8.6671 | 8.6767 | 31.8629 | 31.8606 | 6.6357 | 0.65948 | 1 | 1 | 106 | 6.7613 | 0.12217 | 0.826 |  | 0.040627
30 | 2.108 | 16.7581 | 16.8344 | 28.8556 | 28.8305 | 5.7088 | 3.7725 | 7 | 7 | 166 | 5.8689 | 1.4081 | 8.264 |  | 0.04255
30 | 9.517 | 10.9612 | 10.9267 | 31.3747 | 31.3796 | 6.3299 | 2.3123 | 6 | 6 | 165 | 6.4511 | 0.76498 | 4.935 |  | 0.043492
17 | 29.987 | 9.2395 | 9.242 | 31.766 | 31.7683 | 6.5558 | 1.5495 | 3 | 3 | 91 | 6.6792 | 0.51772 | 3.458 |  | 0.044741
19 | 38.299 | 8.6649 | 8.6748 | 31.8652 | 31.8627 | 6.636 | 0.62897 | 2 | 2 | 107 | 6.7615 | 0.11314 | 0.765 |  | 0.04502
18 | 50.312 | 8.5572 | 8.5582 | 32.2177 | 32.2252 | 6.6368 | 1.1919 | 7 | 7 | 101 | 6.7715 | 0.36225 | 2.453 |  | 0.046543
7 | 1.733 | 14.2185 | 14.1395 | 24.8341 | 24.8789 | 6.1575 | 3.3709 | 7 | 7 | 31 | 6.3038 | 1.2267 | 7.733 |  | 0.047289
16 | 50.07 | 8.5545 | 8.5546 | 32.26 | 32.2593 | 6.6354 | 0.93478 | 4 | 4 | 84 | 6.7698 | 0.25126 | 1.701 | 6 | 0.048646
17 | 10.095 | 11.0506 | 10.9028 | 31.1759 | 31.2717 | 6.3258 | 2.2292 | 5 | 5 | 93 | 6.4484 | 0.82192 | 5.3 |  | 0.050031
26 | 9.654 | 10.8616 | 10.8156 | 31.3553 | 31.3724 | 6.3443 | 2.1163 | 5 | 5 | 137 | 6.4648 | 0.67195 | 4.344 |  | 0.050295
17 | 20.229 | 9.7485 | 9.7447 | 31.5779 | 31.5885 | 6.4905 | 1.7442 | 4 | 4 | 92 | 6.6101 | 0.61224 | 4.047 |  | 0.053663
21 | 44.07 | 8.7574 | 8.7646 | 31.8992 | 31.8968 | 6.6208 | 0.7549 | 1 | 1 | 108 | 6.7521 | 0.17861 | 1.206 |  | 0.055107
14 | 1.835 | 18.7663 | 18.6935 | 14.1095 | 14.9232 | 5.9969 | 3.2573 | 4 | 4 | 74 | 6.1841 | 1.2846 | 7.944 |  | 0.060526
16 | 74.653 | 8.4625 | 8.4628 | 32.3578 | 32.3572 | 6.6448 | 0.8317 | 3 | 3 | 83 | 6.7998 | 0.090297 | 0.614 |  | 0.066984
18 | 19.907 | 9.8209 | 9.8257 | 31.555 | 31.5793 | 6.4811 | 1.7162 | 9 | 9 | 103 | 6.6012 | 0.6161 | 4.067 |  | 0.069835
14 | 9.717 | 10.7921 | 10.837 | 31.2323 | 31.2241 | 6.3588 | 2.9663 | 3 | 3 | 73 | 6.4788 | 1.0249 | 6.64 |  | 0.071213
15 | 1.653 | 19.1398 | 18.662 | 19.0643 | 20.3953 | 5.7805 | 3.5901 | 6 | 6 | 80 | 5.9642 | 1.446 | 8.624 |  | 0.075521
24 | 1.762 | 17.0165 | 16.9175 | 27.4762 | 27.5497 | 5.7277 | 3.765 | 8 | 8 | 125 | 5.8905 | 1.3708 | 8.075 |  | 0.076512
25 | 9.542 | 11.0698 | 11.0712 | 31.2965 | 31.2988 | 6.3183 | 2.2793 | 6 | 6 | 131 | 6.4405 | 0.70662 | 4.551 | 6 | 0.087298
18 | 10.096 | 10.7003 | 10.7177 | 31.3365 | 31.3633 | 6.3671 | 1.9285 | 10 | 10 | 104 | 6.4867 | 0.7446 | 4.83 |  | 0.10494
29 | 9.851 | 11.1617 | 11.1768 | 31.3067 | 31.3024 | 6.3055 | 2.3678 | 6 | 6 | 158 | 6.4287 | 0.72628 | 4.669 |  | 0.1066
28 | 9.638 | 11.1039 | 11.0928 | 31.301 | 31.3021 | 6.3135 | 2.2662 | 6 | 6 | 151 | 6.4361 | 0.67821 | 4.365 |  | 0.10994
15 | 9.986 | 10.7521 | 10.782 | 31.2453 | 31.2428 | 6.3637 | 2.8577 | 5 | 5 | 79 | 6.4837 | 0.93327 | 6.051 |  | 0.11505
14 | 19.594 | 9.6411 | 9.6483 | 31.454 | 31.4534 | 6.511 | 2.1134 | 2 | 2 | 72 | 6.6293 | 0.59659 | 3.955 |  | 0.12439
3 | 1.752 | 14.2869 | 14.114 | 25.2603 | 25.1051 | 6.133 | 3.8326 | 6 | 6 | 10 | 6.2794 | 1.3116 | 8.236 |  | 0.16548
6 | 1.743 | 13.474 | 12.9461 | 25.1165 | 25.7671 | 6.2431 | 2.7804 | 9 | 9 | 24 | 6.3835 | 1.1993 | 7.656 |  | 0.18503
14 | 31.839 | 9.0566 | 9.1163 | 31.6711 | 31.6488 | 6.5865 | 0.94885 | 1 | 1 | 71 | 6.71 | 0.53517 | 3.591 |  | 0.32637
17 | 1.931 | 18.7803 | 18.2969 | 22.2588 | 22.7388 | 5.7111 | 2.831 | 6 | 6 | 94 | 5.8897 | 1.3705 | 8.072 |  | 0.33398
9 | 9.691 | 11.7389 | 11.745 | 27.4065 | 27.4084 | 6.3836 | 1.9445 | 5 | 5 | 42 | 6.5146 | 0.99515 | 6.483 |  | 0.34845
1 | 30.267 | 12.3546 | 12.3478 | 27.988 | 27.9777 | 6.2775 | 1.5598 | 4 | 4 | 4 | 6.4296 | 0.0069988 | 0.045 |  | 0.4705
18 | 2.071 | 15.9612 | 15.4796 | 27.9459 | 28.5623 | 5.8322 | 2.175 | 11 | 11 | 105 | 5.9881 | 1.3731 | 8.222 |  | 0.62499
28 | 74.695 | 8.0016 | 8.002 | 32.8238 | 32.8245 | 6.6938 | 1.0635 | 2 | 2 | 147 | 6.8447 | 0.19709 | 1.349 | 4 | 0.062153
28 | 49.968 | 8.7534 | 8.7563 | 32.1312 | 32.1297 | 6.6114 | 1.4142 | 3 | 3 | 148 | 6.7476 | 0.36043 | 2.432 | 4 | 0.053057
28 | 83.314 | 7.9704 | 7.9601 | 32.8694 | 32.883 | 6.6965 | 1.0483 | 1 | 1 | 146 | 6.8547 | 0.23984 | 1.644 | 4 | 0.012693
26 | 49.219 | 8.5869 | 8.6047 | 32.2619 | 32.2487 | 6.6305 | 1.32 | 2 | 2 | 134 | 6.7644 | 0.34918 | 2.362 | 3 | 0.02287
25 | 74.607 | 8.2088 | 8.2093 | 32.6063 | 32.6072 | 6.672 | 1.0723 | 2 | 2 | 127 | 6.8247 | 0.26023 | 1.776 | 3 | 0.0028612
29 | 29.587 | 9.4389 | 9.4359 | 31.7997 | 31.7977 | 6.5255 | 1.6676 | 4 | 4 | 156 | 6.6503 | 0.51968 | 3.456 | 3 | 0.0052258
23 | 154.771 | 8.2368 | 8.2376 | 32.7523 | 32.7518 | 6.6615 | 0.64548 | 1 | 1 | 109 | 6.8845 | 0.077566 | 0.534 | 3 | 0.0021837
30 | 29.786 | 9.4719 | 9.472 | 31.854 | 31.854 | 6.5185 | 1.7164 | 4 | 4 | 163 | 6.6437 | 0.54352 | 3.611 | 3 | 0.002827
26 | 20.357 | 9.8626 | 9.8636 | 31.6263 | 31.6267 | 6.4722 | 1.7727 | 4 | 4 | 136 | 6.593 | 0.56848 | 3.748 | 3 | 0.0026445
18 | 30.266 | 9.2051 | 9.2066 | 31.7569 | 31.7679 | 6.5612 | 1.515 | 8 | 8 | 102 | 6.6846 | 0.43129 | 2.883 | 36 | 0.026513

## ctd + bottle data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Event_Number | Station | Sample_Number | Oxygen:Dissolved | Temperature:Draw | Flag:Oxygen:Dissolved | Comments
1 | 49.473 | 13.5876 | 13.5701 | 29.4294 | 29.411 | 6.063 | 0.52935 | 1 | 1 | 1 | 1 | TI02 | 1 | 0 | 13.5 |  | OXY: Anoxic.
1 | 40.001 | 13.5879 | 13.5884 | 29.4328 | 29.4326 | 6.0628 | 0.51004 | 2 | 2 | 2 | 1 | TI02 | 2 | 0 | 13.5 |  | OXY: Anoxic.
1 | 34.893 | 13.5876 | 13.588 | 29.4328 | 29.4324 | 6.0629 | 0.5076 | 3 | 3 | 3 | 1 | TI02 | 3 | 0 | 13.5 |  | OXY: Anoxic.
1 | 30.267 | 12.3546 | 12.3478 | 27.988 | 27.9777 | 6.2775 | 1.5598 | 4 | 4 | 4 | 1 | TI02 | 4 | 0.045 | 13.3
3 | 68.614 | 13.1002 | 13.101 | 28.3228 | 28.3238 | 6.167 | 2.1209 | 1 | 1 | 5 | 3 | TI01 | 5 | 4.685 | 13
3 | 49.73 | 13.0573 | 13.0589 | 28.2528 | 28.2542 | 6.1752 | 2.1371 | 2 | 2 | 6 | 3 | TI01 | 6 | 4.701 | 13.1
3 | 29.818 | 12.7926 | 12.7905 | 28.0412 | 28.0401 | 6.2177 | 2.0623 | 3 | 3 | 7 | 3 | TI01 | 7 | 4.561 | 12.9
3 | 20.027 | 12.2496 | 12.2488 | 27.8147 | 27.8143 | 6.2983 | 1.9416 | 4 | 4 | 8 | 3 | TI01 | 8 | 4.243 | 12.7 | 6
3 | 9.469 | 11.6346 | 11.6352 | 27.32 | 27.32 | 6.4014 | 1.9665 | 5 | 5 | 9 | 3 | TI01 | 9 | 4.282 | 12.3
3 | 1.752 | 14.2869 | 14.114 | 25.2603 | 25.1051 | 6.133 | 3.8326 | 6 | 6 | 10 | 3 | TI01 | 10 | 8.236 | 15.3
4 | 47.367 | 13.5537 | 13.5543 | 29.4216 | 29.4221 | 6.0675 | 0.51555 | 1 | 1 | 11 | 4 | TH01 | 11 | 0 | 13.4 |  | OXY: Anoxic.
4 | 39.856 | 13.657 | 13.6576 | 29.41 | 29.4106 | 6.055 | 0.5101 | 2 | 2 | 12 | 4 | TH01 | 12 | 0 | 13.3 |  | OXY: Anoxic.
4 | 35.273 | 13.6138 | 13.6093 | 29.3418 | 29.3407 | 6.063 | 0.50661 | 3 | 3 | 13 | 4 | TH01 | 13 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
4 | 30.065 | 13.2306 | 13.2076 | 29.1619 | 29.1456 | 6.1181 | 0.5052 | 4 | 4 | 14 | 4 | TH01 | 14 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
4 | 25.22 | 12.1093 | 12.1082 | 27.7993 | 27.7989 | 6.3177 | 1.39 | 5 | 5 | 15 | 4 | TH01 | 15 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
6 | 120.04 | 13.9613 | 13.9616 | 29.4799 | 29.4808 | 6.0147 | 0.5144 | 1 | 1 | 16 | 6 | THAD1 | 16 | 0 | 13.9 | 6 | OXY: Anoxic.
6 | 120.092 | 13.9613 | 13.9617 | 29.4804 | 29.481 | 6.0147 | 0.51342 | 2 | 2 | 17 | 6 | THAD1 | 17 | 0.092 | 13.9 | 6
6 | 99.38 | 13.9546 | 13.9549 | 29.4363 | 29.437 | 6.0171 | 0.5084 | 3 | 3 | 18 | 6 | THAD1 | 18 | 0.035 | 13.9
6 | 74.381 | 12.9475 | 12.949 | 28.2433 | 28.244 | 6.1898 | 1.8409 | 4 | 4 | 19 | 6 | THAD1 | 19 | 3.593 | 13.9
6 | 51.288 | 12.8517 | 12.8522 | 28.167 | 28.1677 | 6.2051 | 1.914 | 5 | 5 | 20 | 6 | THAD1 | 20 | 4.088 | 12.9
6 | 31.132 | 12.4042 | 12.4053 | 27.9854 | 27.9862 | 6.271 | 1.8399 | 6 | 6 | 21 | 6 | THAD1 | 21 | 3.864 | 13.8
6 | 19.524 | 12.2711 | 12.2734 | 27.861 | 27.8622 | 6.2936 | 1.8957 | 7 | 7 | 22 | 6 | THAD1 | 22 | 4 | 12.2
6 | 9.311 | 11.58 | 11.5801 | 27.3371 | 27.3303 | 6.4082 | 1.735 | 8 | 8 | 23 | 6 | THAD1 | 23 | 3.693 | 11.6
6 | 1.743 | 13.474 | 12.9461 | 25.1165 | 25.7671 | 6.2431 | 2.7804 | 9 | 9 | 24 | 6 | THAD1 | 24 | 7.656 | 13.5
7 | 101.876 | 13.9255 | 13.9263 | 29.4249 | 29.4268 | 6.0211 | 0.52032 | 1 | 1 | 25 | 7 | THAD2 | 25 | 0 | 13.8 |  | OXY: Anoxic.
7 | 74.576 | 13.0725 | 13.0738 | 28.2767 | 28.2777 | 6.1723 | 2.061 | 2 | 2 | 26 | 7 | THAD2 | 26 | 4.351 | 13.7 | 6
7 | 49.343 | 13 | 13.001 | 28.2035 | 28.2044 | 6.1845 | 2.0929 | 3 | 3 | 27 | 7 | THAD2 | 27 | 4.574 | 12.9
7 | 29.645 | 12.4263 | 12.426 | 27.978 | 27.9786 | 6.2684 | 1.8597 | 4 | 4 | 28 | 7 | THAD2 | 28 | 3.986 | 12.5
7 | 19.078 | 12.1857 | 12.1801 | 27.8008 | 27.7986 | 6.3074 | 1.855 | 5 | 5 | 29 | 7 | THAD2 | 29 | 3.948 | 12.1
7 | 9.67 | 11.5781 | 11.5787 | 27.2885 | 27.2919 | 6.4104 | 1.8324 | 6 | 6 | 30 | 7 | THAD2 | 30 | 3.904 | 11.5
7 | 1.733 | 14.2185 | 14.1395 | 24.8341 | 24.8789 | 6.1575 | 3.3709 | 7 | 7 | 31 | 7 | THAD2 | 31 | 7.733 | 13.7
8 | 70.064 | 13.0896 | 13.0891 | 28.3009 | 28.3002 | 6.1692 | 2.1451 | 1 | 1 | 32 | 8 | THAD3 | 32 | -99 | 13.5 | 5 | OXY: Failed titration. Sample discarded. Object in flask
8 | 49.731 | 13.0502 | 13.0505 | 28.2538 | 28.2545 | 6.1761 | 2.1414 | 2 | 2 | 33 | 8 | THAD3 | 33 | 4.694 | 12.3
8 | 29.193 | 12.6316 | 12.6181 | 28.0235 | 28.0191 | 6.2395 | 1.939 | 3 | 3 | 34 | 8 | THAD3 | 34 | 4.375 | 13.9
8 | 19.457 | 12.1819 | 12.1872 | 27.8005 | 27.7995 | 6.3079 | 1.867 | 4 | 4 | 35 | 8 | THAD3 | 35 | 4.061 | 13.8
8 | 9.64 | 11.6645 | 11.668 | 27.337 | 27.3399 | 6.3966 | 2.0049 | 5 | 5 | 36 | 8 | THAD3 | 36 | 4.157 | 11.7
8 | 1.652 | 15.5244 | 15.2163 | 23.7485 | 24.1982 | 6.0356 | 3.6738 | 6 | 6 | 37 | 8 | THAD3 | 37 | -99 | 16.4 | 5 | OXY: First attempt, aborted due to bubble. Second attempt, no endpoint. Sample lost
9 | 65.122 | 13.1048 | 13.109 | 28.3303 | 28.3363 | 6.1661 | 2.1452 | 1 | 1 | 38 | 9 | TIAD2 | 38 | 4.698 | 13.1
9 | 49.473 | 13.0379 | 13.0396 | 28.244 | 28.2456 | 6.178 | 2.0877 | 2 | 2 | 39 | 9 | TIAD2 | 39 | 4.55 | 12.9
9 | 29.594 | 12.7574 | 12.7558 | 28.0446 | 28.0449 | 6.2222 | 2.0258 | 3 | 3 | 40 | 9 | TIAD2 | 40 | 4.431 | 12.7
9 | 19.692 | 12.2595 | 12.2592 | 27.8349 | 27.8355 | 6.2962 | 1.9044 | 4 | 4 | 41 | 9 | TIAD2 | 41 | 4.13 | 12.2
9 | 9.691 | 11.7389 | 11.745 | 27.4065 | 27.4084 | 6.3836 | 1.9445 | 5 | 5 | 42 | 9 | TIAD2 | 42 | 6.483 | 11.7
9 | 1.751 | 14.9184 | 14.8979 | 24.3814 | 24.35 | 6.0876 | 3.071 | 6 | 6 | 43 | 9 | TIAD2 | 43 | -99 | 13.4 | 5 | OXY: Missed endpoint. Stirrer turned off.
10 | 53.182 | 13.1304 | 13.1308 | 28.3687 | 28.3692 | 6.1613 | 2.1837 | 1 | 1 | 44 | 10 | TIAD3 | 44 | 4.834 | 13.3
10 | 30.112 | 13.002 | 13.0124 | 28.1415 | 28.1443 | 6.1866 | 2.17 | 2 | 2 | 45 | 10 | TIAD3 | 45 | 4.889 | 136.1
10 | 19.535 | 12.2858 | 12.2491 | 27.7585 | 27.7452 | 6.2957 | 1.9674 | 3 | 3 | 46 | 10 | TIAD3 | 46 | 4.36 | 12.4
10 | 9.79 | 12.6576 | 12.6592 | 27.6952 | 27.6878 | 6.2489 | 2.0876 | 4 | 4 | 47 | 10 | TIAD3 | 47 | 4.399 | 12.7
10 | 1.743 | 15.1137 | 15.1903 | 25.0036 | 24.8707 | 6.0398 | 2.7466 | 5 | 5 | 48 | 10 | TIAD3 | 48 | 6.02 | 13.4
11 | 73.687 | 12.9739 | 12.9743 | 29.1127 | 29.1135 | 6.1528 | 2.1821 | 1 | 1 | 49 | 11 | CL03 | 49 | 4.823 | 13.3
11 | 50.132 | 13.0121 | 13.0124 | 29.024 | 29.025 | 6.1513 | 2.209 | 2 | 2 | 50 | 11 | CL03 | 50 | 4.875 | 12.9
11 | 29.926 | 13.0733 | 13.0736 | 28.8054 | 28.8066 | 6.1519 | 2.2178 | 3 | 3 | 51 | 11 | CL03 | 51 | 4.907 | 12.9
11 | 19.955 | 13.0842 | 13.0846 | 28.7475 | 28.7483 | 6.1527 | 2.2155 | 4 | 4 | 52 | 11 | CL03 | 52 | 4.881 | 12.9
11 | 9.543 | 13.1583 | 13.1583 | 28.6013 | 28.6023 | 6.1488 | 2.2525 | 5 | 5 | 53 | 11 | CL03 | 53 | 4.965 | 13
11 | 1.881 | 14.1572 | 14.0797 | 28.2973 | 28.2958 | 6.0347 | 2.4411 | 6 | 6 | 54 | 11 | CL03 | 54 | 5.314 | 13.9 | 6
12 | 147.475 | 13.0117 | 13.012 | 29.1486 | 29.1493 | 6.1466 | 2.1923 | 1 | 1 | 55 | 12 | CL02 | 55 | 4.878 | 12.9
12 | 124.697 | 13.0046 | 13.0048 | 29.1437 | 29.1445 | 6.1477 | 2.1918 | 2 | 2 | 56 | 12 | CL02 | 56 | 4.883 | 12.9
12 | 99.812 | 13.0075 | 13.0078 | 29.1455 | 29.1463 | 6.1472 | 2.2021 | 3 | 3 | 57 | 12 | CL02 | 57 | 4.895 | 12.9
12 | 74.742 | 13.0231 | 13.0231 | 29.1119 | 29.1128 | 6.1465 | 2.2107 | 4 | 4 | 58 | 12 | CL02 | 58 | 4.89 | 12.9
12 | 49.441 | 13.0652 | 13.0652 | 29.0626 | 29.0636 | 6.143 | 2.2254 | 5 | 5 | 59 | 12 | CL02 | 59 | 4.93 | 12.9
12 | 30.006 | 13.13 | 13.1302 | 28.9579 | 28.9578 | 6.1388 | 2.2431 | 6 | 6 | 60 | 12 | CL02 | 60 | 4.96 | 12.9
12 | 19.67 | 13.1213 | 13.1218 | 28.8275 | 28.8288 | 6.1449 | 2.2393 | 7 | 7 | 61 | 12 | CL02 | 61 | 4.95 | 13.9
12 | 10.048 | 13.1144 | 13.1147 | 28.7868 | 28.7913 | 6.1473 | 2.237 | 8 | 8 | 62 | 12 | CL02 | 62 | 4.931 | 13.9
12 | 1.569 | 13.6436 | 13.6362 | 28.3835 | 28.3864 | 6.0955 | 2.3347 | 9 | 9 | 63 | 12 | CL02 | 63 | 5.148 | 13.4
13 | 88.205 | 12.676 | 12.6761 | 29.7323 | 29.7332 | 6.1673 | 2.1463 | 1 | 1 | 64 | 13 | CL01 | 64 | 4.757 | 12.7
13 | 74.898 | 12.7325 | 12.7321 | 29.6316 | 29.6333 | 6.1639 | 2.1764 | 2 | 2 | 65 | 13 | CL01 | 65 | 4.828 | 12.5
13 | 49.988 | 12.9708 | 12.9746 | 29.2613 | 29.257 | 6.1475 | 2.2226 | 3 | 3 | 66 | 13 | CL01 | 66 | 4.932 | 12.5
13 | 30.178 | 12.9694 | 12.9728 | 29.2623 | 29.2589 | 6.1476 | 2.2288 | 4 | 4 | 67 | 13 | CL01 | 67 | 4.936 | 12.5 | 6
13 | 19.277 | 13.1547 | 13.1574 | 28.955 | 28.9518 | 6.1357 | 2.2487 | 5 | 5 | 68 | 13 | CL01 | 68 | 4.944 | 12.5
13 | 9.647 | 13.1379 | 13.1379 | 28.9852 | 28.9869 | 6.1367 | 2.2537 | 6 | 6 | 69 | 13 | CL01 | 69 | 4.982 | 12.5
13 | 1.535 | 13.1422 | 13.1433 | 28.9798 | 28.9796 | 6.1364 | 2.2521 | 7 | 7 | 70 | 13 | CL01 | 70 | 4.986 | 13.7
14 | 31.839 | 9.0566 | 9.1163 | 31.6711 | 31.6488 | 6.5865 | 0.94885 | 1 | 1 | 71 | 14 | EF01 | 71 | 3.591 | 9.5
14 | 19.594 | 9.6411 | 9.6483 | 31.454 | 31.4534 | 6.511 | 2.1134 | 2 | 2 | 72 | 14 | EF01 | 72 | 3.955 | 10.7
14 | 9.717 | 10.7921 | 10.837 | 31.2323 | 31.2241 | 6.3588 | 2.9663 | 3 | 3 | 73 | 14 | EF01 | 73 | 6.64 | 10.4
14 | 1.835 | 18.7663 | 18.6935 | 14.1095 | 14.9232 | 5.9969 | 3.2573 | 4 | 4 | 74 | 14 | EF01 | 74 | 7.944 | 18.1
15 | 66.974 | 8.4878 | 8.4904 | 32.3286 | 32.3301 | 6.6423 | 0.90177 | 1 | 1 | 75 | 15 | EF02 | 75 | 1.129 | 8.4
15 | 49.478 | 8.5935 | 8.5945 | 32.2111 | 32.2114 | 6.6317 | 0.91137 | 2 | 2 | 76 | 15 | EF02 | 76 | 1.181 | 8.7 | 6
15 | 30.601 | 8.9604 | 8.9635 | 31.749 | 31.7489 | 6.5973 | 1.2685 | 3 | 3 | 77 | 15 | EF02 | 77 | 2.299 | 8.9
15 | 19.843 | 9.4694 | 9.4795 | 31.5234 | 31.5235 | 6.5328 | 1.8244 | 4 | 4 | 78 | 15 | EF02 | 78 | 3.835 | 9.3
15 | 9.986 | 10.7521 | 10.782 | 31.2453 | 31.2428 | 6.3637 | 2.8577 | 5 | 5 | 79 | 15 | EF02 | 79 | 6.051 | 10.2
15 | 1.653 | 19.1398 | 18.662 | 19.0643 | 20.3953 | 5.7805 | 3.5901 | 6 | 6 | 80 | 15 | EF02 | 80 | 8.624 | 17.2
16 | 113.973 | 8.1682 | 8.1689 | 32.5702 | 32.5703 | 6.6797 | 0.50991 | 1 | 1 | 81 | 16 | EF03 | 81 | 0 | 9.2 |  | OXY: Anoxic.
16 | 99.648 | 8.2149 | 8.2154 | 32.5594 | 32.559 | 6.6732 | 0.50664 | 2 | 2 | 82 | 16 | EF03 | 82 | 0 | 10.3 |  | OXY: Anoxic.
16 | 74.653 | 8.4625 | 8.4628 | 32.3578 | 32.3572 | 6.6448 | 0.8317 | 3 | 3 | 83 | 16 | EF03 | 83 | 0.614 | 8.7
16 | 50.07 | 8.5545 | 8.5546 | 32.26 | 32.2593 | 6.6354 | 0.93478 | 4 | 4 | 84 | 16 | EF03 | 84 | 1.701 | 8.7 | 6
16 | 30.006 | 9.0699 | 9.0702 | 31.7458 | 31.746 | 6.5814 | 1.5237 | 5 | 5 | 85 | 16 | EF03 | 85 | 2.941 | 9
16 | 19.927 | 9.6231 | 9.6472 | 31.5563 | 31.5455 | 6.5093 | 1.9494 | 6 | 6 | 86 | 16 | EF03 | 86 | 4.146 | 9.5
16 | 9.795 | 10.4286 | 10.4805 | 31.3544 | 31.3508 | 6.404 | 2.3438 | 7 | 7 | 87 | 16 | EF03 | 87 | 5.577 | 17.8
16 | 1.958 | 18.0887 | 18.0907 | 23.4182 | 23.9527 | 5.7479 | 3.7685 | 8 | 8 | 88 | 16 | EF03 | 88 | 8.606 |  |  | OXY: No draw T recorded
17 | 61.039 | 8.2675 | 8.2604 | 32.5152 | 32.5217 | 6.6672 | 1.1894 | 1 | 1 | 89 | 17 | EF04 | 89 | 2.256 | 12.1
17 | 50.217 | 8.578 | 8.5789 | 32.207 | 32.2078 | 6.6341 | 1.2965 | 2 | 2 | 90 | 17 | EF04 | 90 | 2.405 | 8.7 | 6
17 | 29.987 | 9.2395 | 9.242 | 31.766 | 31.7683 | 6.5558 | 1.5495 | 3 | 3 | 91 | 17 | EF04 | 91 | 3.458 | 9.2
17 | 20.229 | 9.7485 | 9.7447 | 31.5779 | 31.5885 | 6.4905 | 1.7442 | 4 | 4 | 92 | 17 | EF04 | 92 | 4.047 | 9.6
17 | 10.095 | 11.0506 | 10.9028 | 31.1759 | 31.2717 | 6.3258 | 2.2292 | 5 | 5 | 93 | 17 | EF04 | 93 | 5.3 | 10.4
17 | 1.931 | 18.7803 | 18.2969 | 22.2588 | 22.7388 | 5.7111 | 2.831 | 6 | 6 | 94 | 17 | EF04 | 94 | 8.072 | 17.1
18 | 197.637 | 8.2432 | 8.2638 | 32.8303 | 32.8164 | 6.6572 | 0.53263 | 1 | 1 | 95 | 18 | EF05 | 95 | 0.021 | 8.5
18 | 175.347 | 8.2486 | 8.2498 | 32.8194 | 32.8185 | 6.6568 | 0.56079 | 2 | 2 | 96 | 18 | EF05 | 96 | 0.084 | 8.4
18 | 149.99 | 8.2665 | 8.2683 | 32.7754 | 32.7749 | 6.6561 | 0.56417 | 3 | 3 | 97 | 18 | EF05 | 97 | 0.158 | 8.4
18 | 125.173 | 8.1747 | 8.1757 | 32.6424 | 32.6386 | 6.6756 | 1.0004 | 4 | 4 | 98 | 18 | EF05 | 98 | 1.596 | 8.4
18 | 100.146 | 8.172 | 8.173 | 32.6155 | 32.6143 | 6.6772 | 1.0205 | 5 | 5 | 99 | 18 | EF05 | 99 | 1.833 | 8.4
18 | 74.962 | 8.188 | 8.189 | 32.5906 | 32.5902 | 6.6758 | 1.038 | 6 | 6 | 100 | 18 | EF05 | 100 | 1.92 | 8.4
18 | 50.312 | 8.5572 | 8.5582 | 32.2177 | 32.2252 | 6.6368 | 1.1919 | 7 | 7 | 101 | 18 | EF05 | 101 | 2.453 | 8.2
18 | 30.266 | 9.2051 | 9.2066 | 31.7569 | 31.7679 | 6.5612 | 1.515 | 8 | 8 | 102 | 18 | EF05 | 102 | 2.883 | 9.4 | 36 | OXY: Replicate Outliers
18 | 19.907 | 9.8209 | 9.8257 | 31.555 | 31.5793 | 6.4811 | 1.7162 | 9 | 9 | 103 | 18 | EF05 | 103 | 4.067 | 9.7
18 | 10.096 | 10.7003 | 10.7177 | 31.3365 | 31.3633 | 6.3671 | 1.9285 | 10 | 10 | 104 | 18 | EF05 | 104 | 4.83 | 10.6
18 | 2.071 | 15.9612 | 15.4796 | 27.9459 | 28.5623 | 5.8322 | 2.175 | 11 | 11 | 105 | 18 | EF05 | 105 | 8.222 |  |  | OXY: No draw T recorded
19 | 38.296 | 8.6671 | 8.6767 | 31.8629 | 31.8606 | 6.6357 | 0.65948 | 1 | 1 | 106 | 19 | PI01 | 106 | 0.826 |  |  | OXY: No draw T recorded
19 | 38.299 | 8.6649 | 8.6748 | 31.8652 | 31.8627 | 6.636 | 0.62897 | 2 | 2 | 107 | 19 | PI01 | 107 | 0.765 |  |  | OXY: No draw T recorded
21 | 44.07 | 8.7574 | 8.7646 | 31.8992 | 31.8968 | 6.6208 | 0.7549 | 1 | 1 | 108 | 21 | PI03 | 108 | 1.206 |  |  | OXY: No draw T recorded
23 | 154.771 | 8.2368 | 8.2376 | 32.7523 | 32.7518 | 6.6615 | 0.64548 | 1 | 1 | 109 | 22 | EF06 | 109 | 0.534 | 8.3 | 3 | OXY: Bubble from H2SO4 dispenser
23 | 124.637 | 8.1816 | 8.1821 | 32.6387 | 32.6389 | 6.6747 | 1.0293 | 2 | 2 | 110 | 23 | EF06 | 110 | 1.649 | 8.3
23 | 99.446 | 8.1823 | 8.1829 | 32.6152 | 32.6154 | 6.6756 | 1.067 | 3 | 3 | 111 | 23 | EF06 | 111 | 1.799 | 8.5
23 | 74.612 | 8.2014 | 8.2019 | 32.5817 | 32.5818 | 6.6742 | 1.1034 | 4 | 4 | 112 | 23 | EF06 | 112 | 1.868 | 8.4
23 | 49.959 | 8.5811 | 8.5806 | 32.2019 | 32.2021 | 6.6339 | 1.299 | 5 | 5 | 113 | 23 | EF06 | 113 | 2.418 | 8.7
23 | 30.171 | 9.3752 | 9.3761 | 31.7709 | 31.7708 | 6.5359 | 1.6635 | 6 | 6 | 114 | 23 | EF06 | 114 | 3.424 | 9.3
23 | 19.965 | 9.8844 | 9.8902 | 31.5943 | 31.5926 | 6.4705 | 1.8512 | 7 | 7 | 115 | 23 | EF06 | 115 | 4.033 | 9.8
23 | 9.591 | 10.8141 | 10.7903 | 31.353 | 31.3563 | 6.3509 | 2.1329 | 8 | 8 | 116 | 23 | EF06 | 116 | 4.697 | 10.6 | 6
23 | 1.871 | 17.1391 | 17.1531 | 27.1652 | 27.1164 | 5.7247 | 3.7554 | 9 | 9 | 117 | 23 | EF06 | 117 | 8.341 | 15.5
24 | 117.886 | 8.1909 | 8.1912 | 32.6488 | 32.6491 | 6.6729 | 0.97053 | 1 | 1 | 118 | 24 | EF07 | 118 | 1.633 | 8.5
24 | 99.207 | 8.1873 | 8.1877 | 32.6151 | 32.6158 | 6.6749 | 1.0619 | 2 | 2 | 119 | 24 | EF07 | 119 | 1.76 | 8.7
24 | 74.906 | 8.2145 | 8.2149 | 32.591 | 32.5913 | 6.6719 | 1.0932 | 3 | 3 | 120 | 24 | EF07 | 120 | 1.848 | 8.6
24 | 49.842 | 8.5888 | 8.5913 | 32.2002 | 32.1989 | 6.6328 | 1.3132 | 4 | 4 | 121 | 24 | EF07 | 121 | 2.353 | 8.8
24 | 29.999 | 9.3638 | 9.363 | 31.7728 | 31.7731 | 6.5375 | 1.6618 | 5 | 5 | 122 | 24 | EF07 | 122 | 3.411 | 9.4
24 | 19.859 | 9.8116 | 9.8103 | 31.6347 | 31.6349 | 6.4791 | 1.7829 | 6 | 6 | 123 | 24 | EF07 | 123 | 3.845 | 9.9
24 | 9.744 | 10.8842 | 10.8809 | 31.3385 | 31.3404 | 6.3419 | 2.1754 | 7 | 7 | 124 | 24 | EF07 | 124 | 4.772 | 11.9 | 6
24 | 1.762 | 17.0165 | 16.9175 | 27.4762 | 27.5497 | 5.7277 | 3.765 | 8 | 8 | 125 | 24 | EF07 | 125 | 8.075 | 15.3
25 | 83.245 | 8.2102 | 8.2151 | 32.6115 | 32.6093 | 6.6716 | 1.0678 | 1 | 1 | 126 | 25 | EF08 | 126 | 1.776 | 8.5
25 | 74.607 | 8.2088 | 8.2093 | 32.6063 | 32.6072 | 6.672 | 1.0723 | 2 | 2 | 127 | 25 | EF08 | 127 | 1.776 | 9.1 | 3 | OXY: Bubble from H2SO4 dispenser
25 | 49.639 | 8.6057 | 8.6039 | 32.2095 | 32.2121 | 6.6299 | 1.3198 | 3 | 3 | 128 | 25 | EF08 | 128 | 2.307 | 8.6
25 | 29.42 | 9.3166 | 9.3136 | 31.8132 | 31.8143 | 6.5426 | 1.6283 | 4 | 4 | 129 | 25 | EF08 | 129 | 3.326 | 10
25 | 19.965 | 9.7968 | 9.7965 | 31.6384 | 31.6394 | 6.4811 | 1.7762 | 5 | 5 | 130 | 25 | EF08 | 130 | 3.699 | 10.7
25 | 9.542 | 11.0698 | 11.0712 | 31.2965 | 31.2988 | 6.3183 | 2.2793 | 6 | 6 | 131 | 25 | EF08 | 131 | 4.551 | 11 | 6
25 | 1.655 | 18.4717 | 18.6184 | 25.3271 | 25.0615 | 5.6407 | 3.5261 | 7 | 7 | 132 | 25 | EF08 | 132 | 7.738 | 18
26 | 59.789 | 8.2338 | 8.2346 | 32.5849 | 32.5862 | 6.6692 | 1.1225 | 1 | 1 | 133 | 26 | EF09 | 133 | 2.197 | 8.8
26 | 49.219 | 8.5869 | 8.6047 | 32.2619 | 32.2487 | 6.6305 | 1.32 | 2 | 2 | 134 | 26 | EF09 | 134 | 2.362 | 9.5 | 3 | OXY: Bubble from H2SO4 dispenser
26 | 30.126 | 9.3134 | 9.32 | 31.8262 | 31.8238 | 6.5426 | 1.6146 | 3 | 3 | 135 | 26 | EF09 | 135 | 3.279 | 9
26 | 20.357 | 9.8626 | 9.8636 | 31.6263 | 31.6267 | 6.4722 | 1.7727 | 4 | 4 | 136 | 26 | EF09 | 136 | 3.748 | 9.9 | 3 | OXY: Bubbles
26 | 9.654 | 10.8616 | 10.8156 | 31.3553 | 31.3724 | 6.3443 | 2.1163 | 5 | 5 | 137 | 26 | EF09 | 137 | 4.344 | 10.6
26 | 1.709 | 16.7685 | 16.0928 | 27.1323 | 27.6893 | 5.7781 | 3.3024 | 6 | 6 | 138 | 26 | EF09 | 138 | -99 | 15.6 | 5 | OXY: Bubbles. Sample lost but analyst did not give reason why
27 | 95.051 | 7.9687 | 7.9705 | 32.867 | 32.8678 | 6.6969 | 1.0006 | 1 | 1 | 139 | 27 | EF10 | 139 | 1.58 | 9.3
27 | 74.683 | 7.9927 | 7.9897 | 32.8333 | 32.8379 | 6.6947 | 1.0345 | 2 | 2 | 140 | 27 | EF10 | 140 | 1.644 | 8.7
27 | 49.897 | 8.77 | 8.7701 | 32.124 | 32.1245 | 6.6093 | 1.4051 | 3 | 3 | 141 | 27 | EF10 | 141 | 2.656 | 11.4
27 | 29.695 | 9.3773 | 9.378 | 31.8014 | 31.8014 | 6.5344 | 1.63 | 4 | 4 | 142 | 27 | EF10 | 142 | 3.292 |  |  | OXY: No draw T recorded
27 | 20.001 | 9.8262 | 9.8383 | 31.6408 | 31.6374 | 6.4768 | 1.7577 | 5 | 5 | 143 | 27 | EF10 | 143 | 3.665 |  |  | OXY: No draw T recorded
27 | 9.676 | 10.8347 | 10.8286 | 31.369 | 31.3693 | 6.3474 | 2.1054 | 6 | 6 | 144 | 27 | EF10 | 144 | 4.393 | 10.7
27 | 1.751 | 17.6719 | 17.4116 | 26.1086 | 26.3901 | 5.7022 | 3.4514 | 7 | 7 | 145 | 27 | EF10 | 145 | 7.866 | 17.5
28 | 83.314 | 7.9704 | 7.9601 | 32.8694 | 32.883 | 6.6965 | 1.0483 | 1 | 1 | 146 | 28 | EF11 | 146 | 1.644 | 8.7 | 4 | OXY: Titration ran amuck. Curve does not flatten. No endpoint
28 | 74.695 | 8.0016 | 8.002 | 32.8238 | 32.8245 | 6.6938 | 1.0635 | 2 | 2 | 147 | 28 | EF11 | 147 | 1.349 | 9.3 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint
28 | 49.968 | 8.7534 | 8.7563 | 32.1312 | 32.1297 | 6.6114 | 1.4142 | 3 | 3 | 148 | 28 | EF11 | 148 | 2.432 | 9.2 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint
28 | 29.879 | 9.3868 | 9.3907 | 31.7844 | 31.7825 | 6.5337 | 1.6333 | 4 | 4 | 149 | 28 | EF11 | 149 | 3.365 | 10.1
28 | 19.625 | 9.8848 | 9.8871 | 31.621 | 31.6207 | 6.4693 | 1.7613 | 5 | 5 | 150 | 28 | EF11 | 150 | 3.745 | 10.4
28 | 9.638 | 11.1039 | 11.0928 | 31.301 | 31.3021 | 6.3135 | 2.2662 | 6 | 6 | 151 | 28 | EF11 | 151 | 4.365 | 10.9
28 | 2.209 | 17.394 | 17.1697 | 27.228 | 27.358 | 5.6953 | 3.616 | 7 | 7 | 152 | 28 | EF11 | 152 | 7.944 | 13.5
29 | 85.664 | 7.9599 | 7.973 | 32.8891 | 32.8783 | 6.6973 | 1.1151 | 1 | 1 | 153 | 29 | EF12 | 153 | 1.913 | 8.3
29 | 74.598 | 8.0013 | 8.0022 | 32.8124 | 32.8129 | 6.6943 | 1.1214 | 2 | 2 | 154 | 29 | EF12 | 154 | 1.903 | 8.7
29 | 49.984 | 8.75 | 8.7506 | 32.1641 | 32.1641 | 6.6105 | 1.4434 | 3 | 3 | 155 | 29 | EF12 | 155 | 2.769 | 8.9
29 | 29.587 | 9.4389 | 9.4359 | 31.7997 | 31.7977 | 6.5255 | 1.6676 | 4 | 4 | 156 | 29 | EF12 | 156 | 3.456 | 11.7 | 3 | OXY: Bubble from H2SO4 dispenser
29 | 19.889 | 9.9241 | 9.9267 | 31.6206 | 31.6194 | 6.4638 | 1.7734 | 5 | 5 | 157 | 29 | EF12 | 157 | 3.916 | 10
29 | 9.851 | 11.1617 | 11.1768 | 31.3067 | 31.3024 | 6.3055 | 2.3678 | 6 | 6 | 158 | 29 | EF12 | 158 | 4.669 | 11.1
29 | 2.49 | 17.3668 | 17.5889 | 27.8874 | 27.7839 | 5.6745 | 3.5747 | 7 | 7 | 159 | 29 | EF12 | 159 | 7.939 | 16.5
30 | 96.816 | 7.8137 | 7.8177 | 33.0759 | 33.0748 | 6.7113 | 1.1568 | 1 | 1 | 160 | 30 | EF13 | 160 | 2.044 | 8.2
30 | 74.452 | 7.9712 | 7.9659 | 32.8616 | 32.8706 | 6.6967 | 1.1117 | 2 | 2 | 161 | 30 | EF13 | 161 | 1.926 | 8.7
30 | 49.978 | 8.7709 | 8.7631 | 32.1431 | 32.1481 | 6.6083 | 1.4547 | 3 | 3 | 162 | 30 | EF13 | 162 | 2.816 | 9.6
30 | 29.786 | 9.4719 | 9.472 | 31.854 | 31.854 | 6.5185 | 1.7164 | 4 | 4 | 163 | 30 | EF13 | 163 | 3.611 | 9.6 | 3 | OXY: Bubble from H2SO4 dispenser
30 | 19.732 | 10.0906 | 10.0757 | 31.5886 | 31.5948 | 6.4416 | 1.8712 | 5 | 5 | 164 | 30 | EF13 | 164 | 4.046 | 10.1
30 | 9.517 | 10.9612 | 10.9267 | 31.3747 | 31.3796 | 6.3299 | 2.3123 | 6 | 6 | 165 | 30 | EF13 | 165 | 4.935 | 13
30 | 2.108 | 16.7581 | 16.8344 | 28.8556 | 28.8305 | 5.7088 | 3.7725 | 7 | 7 | 166 | 30 | EF13 | 166 | 8.264 | 15.4

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers