## Notes
Analysis: | Kristen Cooke, WCVI onboard CCGS Vector, 13-AUG-11 to 21-AUG-11 using IOS Kit #4. Data processing by Kenny Scozzafava
2011-09-27 |  | KS
On August 16, 2011, the 1st set of standard titrations were analyzed (3 days into the cruise) and entered manually into standards history text file.
They would not load into the program properly as a result.
There is no confirmation wether or not system was working optimally since standards and blanks were only run once directly before sample analysis. Very difficult to assess
which single standard and blank titer from the three standardizations performed may have been best, given the great variability between them, to calculate all of the final oxygen concentration.
As a result, values were left as is with their respective standardization
Generally poor standard and blank precision throughout the cruise.
Thio Normality (a measure of system and chemical stability) fluctuated drastically throughout the cruise with the highest change (difference of 0.00128) occurring between August 18 to the 20th.
The daily maximum acceptable change in thio N is 0.0005.
There were 3 separate people involved in analysis. I sampler, 1 fixing samples, and 1 running the analysis
Many comments regarding bubbles in the reagents. Bottle top dispensers may have not been primed properly, may not have been tightened to bottle sufficiently, may have had leaky straw or
may have had leaky discharge valve. There was also evidence that the dispensers were switched between reagents (They are clearly labelled only to be used with their rerspective solution).
The reaction of the reagents inside works to gum up the dispensers and they will not dispense properly as a result.
Analyst also notes that the MnCl2 dispenser was sucking water into the tip when fixing oxygen samples on deck. This may be related to the problems
outlined above, since a dispenser that is working properly should not suck anything from the tip. Regardless, proper sampling requires the
dispensers to be filled in air and dispensed in the sample to avoid this problem.
Many inputting errors regarding sample numbers and associated depths and flask information. They are all corrected.
2011-12-28 |  | KS
Many Draw Temperatures were not recorded for duplicate samples. In those cases, I used the same T as recorded for the first of the two dups.
#13, 14, 15- Samples not drawn due to time constraints. Flagged as 5
#16- Anoxic duplicates. Not included in the precision analysis but still reported as an average.
Many samples were missing draw temperatures. I included a short note indicating that it was not recorded as opposed to a data entry error.
I can remove the comments if the spreadsheet is too cluttered.
#147-#149- Oxygen concentrations not trusted and flagged as 4. Analyst commented that the titrations ran amuck but with no other detail as to the cause.
Upon further investigation, it seems as though the program acted funny and split the titration as if it was aborted and restarted. Difficult to know what happened
#139-166- Final oxygen values were recalculated with a new balnk titer (0.0040 as opposed to 0.0047).
#102- Replicate outliers as determined by the Chauvenet criteria. Please check against ctd results. Reported as the average.
Duplicate precision is also quite poor even after the removal of one outlier.
There were not enough duplicate samples (13 out of 179 unique samples) to properly asses the precision of the data set.
Our goal is to collect at least 10% of the samples in duplicate and more if possible.
2012-03-08 |  | KS
Upon further review and due to the poor standard and blank precision, the great variability between standardizations, the lack of duplicate samples resulting in poor
sample precision (Sp= 0.027), all data are reported to the second decimal place only.
The quality of the ctd data is in question since they were averaged in a certain way as to maybe include anomolous spikes and other oddities in the data.
There were also high Dissolved oxygen gradients which also make it difficult to determine whether a certain bottle sample was in fact an outlier or not.
Cast 18 - sample 102 - It was not clear to Germaine what outlier, if any of the two, is better than the other. Left as is
Cast 23 - sample 109 - mild outlier in ctd comparison. Left flag as is for above reasons. Too much uncertainty.
Cast 25 - sample 127 - looks ok. Flag left as is for reasons above
Cast 26 - sample 134 - major outlier. Flag left as is for above reasons
Cast 26 - sample 136 - probably ok. Flag left as for above reasons.
Cast 28 - sample 146 - minor outlier. Because the titration itself was bad (there was no endpoint), flag is appropriate.
Cast 28 - sample 147 - major outlier. Flag left as is. Titration was bad
Cast 28 - sample 148 - major outlier. Flag left as is. Titration was bad
Cast 29 - sample 156 - looks ok. Flag left as is for above reasons.
Cast 30 - sample 163 - looks ok. Flag left as is for above reasons.

## QF 2011-63 OXY FINAL
HEADER: | Oxygen samples were analyzed at sea using an automated Winkler titration system
(Brinkmann Dosimat model 765 and a PC 950 Colorimeter controlled by AutoOxy software V3.4)
with modifications based on Carpenter (1965) and adhering to WOCE protocols (Culberson 1991). For details including a duplicate analysis, see document QF2011-63oxy.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Oxygen:Dissolved | Temperature:Draw | Flag:Oxygen:Dissolved | Comment
2011-63 | 1 | TI02 | 1 | 0 | 13.5 |  | OXY: Anoxic.
2011-63 | 1 | TI02 | 2 | 0 | 13.5 |  | OXY: Anoxic.
2011-63 | 1 | TI02 | 3 | 0 | 13.5 |  | OXY: Anoxic.
2011-63 | 1 | TI02 | 4 | 0.045475 | 13.3
2011-63 | 3 | TI01 | 5 | 4.6853 | 13
2011-63 | 3 | TI01 | 6 | 4.7007 | 13.1
2011-63 | 3 | TI01 | 7 | 4.5614 | 12.9
2011-63 | 3 | TI01 | 8 | 4.2428 | 12.7 | 6
2011-63 | 3 | TI01 | 9 | 4.2822 | 12.3
2011-63 | 3 | TI01 | 10 | 8.2362 | 15.3
2011-63 | 4 | TH01 | 11 | 0 | 13.4 |  | OXY: Anoxic.
2011-63 | 4 | TH01 | 12 | 0 | 13.3 |  | OXY: Anoxic.
2011-63 | 4 | TH01 | 13 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-63 | 4 | TH01 | 14 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-63 | 4 | TH01 | 15 | -99 |  | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-63 | 6 | THAD1 | 16 | 0 | 13.9 | 6 | OXY: Anoxic.
2011-63 | 6 | THAD1 | 17 | 0.091732 | 13.9 | 6
2011-63 | 6 | THAD1 | 18 | 0.035215 | 13.9
2011-63 | 6 | THAD1 | 19 | 3.5928 | 13.9
2011-63 | 6 | THAD1 | 20 | 4.0881 | 12.9
2011-63 | 6 | THAD1 | 21 | 3.864 | 13.8
2011-63 | 6 | THAD1 | 22 | 4.0002 | 12.2
2011-63 | 6 | THAD1 | 23 | 3.693 | 11.6
2011-63 | 6 | THAD1 | 24 | 7.6565 | 13.5
2011-63 | 7 | THAD2 | 25 | 0 | 13.8 |  | OXY: Anoxic.
2011-63 | 7 | THAD2 | 26 | 4.3508 | 13.7 | 6
2011-63 | 7 | THAD2 | 27 | 4.5741 | 12.9
2011-63 | 7 | THAD2 | 28 | 3.9858 | 12.5
2011-63 | 7 | THAD2 | 29 | 3.9485 | 12.1
2011-63 | 7 | THAD2 | 30 | 3.9041 | 11.5
2011-63 | 7 | THAD2 | 31 | 7.7326 | 13.7
2011-63 | 8 | THAD3 | 32 | -99 | 13.5 | 5 | OXY: Failed titration. Sample discarded. Object in flask
2011-63 | 8 | THAD3 | 33 | 4.6943 | 12.3
2011-63 | 8 | THAD3 | 34 | 4.3755 | 13.9
2011-63 | 8 | THAD3 | 35 | 4.0609 | 13.8
2011-63 | 8 | THAD3 | 36 | 4.1571 | 11.7
2011-63 | 8 | THAD3 | 37 | -99 | 16.4 | 5 | OXY: First attempt, aborted due to bubble. Second attempt, no endpoint. Sample lost
2011-63 | 9 | TIAD2 | 38 | 4.6983 | 13.1
2011-63 | 9 | TIAD2 | 39 | 4.5501 | 12.9
2011-63 | 9 | TIAD2 | 40 | 4.4309 | 12.7
2011-63 | 9 | TIAD2 | 41 | 4.1296 | 12.2
2011-63 | 9 | TIAD2 | 42 | 6.4832 | 11.7
2011-63 | 9 | TIAD2 | 43 | -99 | 13.4 | 5 | OXY: Missed endpoint. Stirrer turned off.
2011-63 | 10 | TIAD3 | 44 | 4.8343 | 13.3
2011-63 | 10 | TIAD3 | 45 | 4.8887 | 136.1
2011-63 | 10 | TIAD3 | 46 | 4.3605 | 12.4
2011-63 | 10 | TIAD3 | 47 | 4.399 | 12.7
2011-63 | 10 | TIAD3 | 48 | 6.0202 | 13.4
2011-63 | 11 | CL03 | 49 | 4.8232 | 13.3
2011-63 | 11 | CL03 | 50 | 4.8749 | 12.9
2011-63 | 11 | CL03 | 51 | 4.9072 | 12.9
2011-63 | 11 | CL03 | 52 | 4.8812 | 12.9
2011-63 | 11 | CL03 | 53 | 4.9648 | 13
2011-63 | 11 | CL03 | 54 | 5.3141 | 13.85 | 6
2011-63 | 12 | CL02 | 55 | 4.8781 | 12.9
2011-63 | 12 | CL02 | 56 | 4.8826 | 12.9
2011-63 | 12 | CL02 | 57 | 4.8947 | 12.9
2011-63 | 12 | CL02 | 58 | 4.8901 | 12.9
2011-63 | 12 | CL02 | 59 | 4.9296 | 12.9
2011-63 | 12 | CL02 | 60 | 4.9595 | 12.9
2011-63 | 12 | CL02 | 61 | 4.9498 | 13.9
2011-63 | 12 | CL02 | 62 | 4.9309 | 13.9
2011-63 | 12 | CL02 | 63 | 5.1476 | 13.4
2011-63 | 13 | CL01 | 64 | 4.7572 | 12.7
2011-63 | 13 | CL01 | 65 | 4.8281 | 12.5
2011-63 | 13 | CL01 | 66 | 4.9325 | 12.5
2011-63 | 13 | CL01 | 67 | 4.9362 | 12.5 | 6
2011-63 | 13 | CL01 | 68 | 4.944 | 12.5
2011-63 | 13 | CL01 | 69 | 4.9822 | 12.5
2011-63 | 13 | CL01 | 70 | 4.9864 | 13.7
2011-63 | 14 | EF01 | 71 | 3.591 | 9.5
2011-63 | 14 | EF01 | 72 | 3.9548 | 10.7
2011-63 | 14 | EF01 | 73 | 6.6396 | 10.4
2011-63 | 14 | EF01 | 74 | 7.944 | 18.1
2011-63 | 15 | EF02 | 75 | 1.1295 | 8.4
2011-63 | 15 | EF02 | 76 | 1.1813 | 8.7 | 6
2011-63 | 15 | EF02 | 77 | 2.2993 | 8.9
2011-63 | 15 | EF02 | 78 | 3.8354 | 9.3
2011-63 | 15 | EF02 | 79 | 6.0514 | 10.2
2011-63 | 15 | EF02 | 80 | 8.6244 | 17.2
2011-63 | 16 | EF03 | 81 | 0 | 9.2 |  | OXY: Anoxic.
2011-63 | 16 | EF03 | 82 | 0 | 10.3 |  | OXY: Anoxic.
2011-63 | 16 | EF03 | 83 | 0.61407 | 8.7
2011-63 | 16 | EF03 | 84 | 1.7011 | 8.7 | 6
2011-63 | 16 | EF03 | 85 | 2.9407 | 9
2011-63 | 16 | EF03 | 86 | 4.146 | 9.5
2011-63 | 16 | EF03 | 87 | 5.5771 | 17.8
2011-63 | 16 | EF03 | 88 | 8.606 |  |  | OXY: No draw T recorded
2011-63 | 17 | EF04 | 89 | 2.256 | 12.1
2011-63 | 17 | EF04 | 90 | 2.4048 | 8.7 | 6
2011-63 | 17 | EF04 | 91 | 3.458 | 9.2
2011-63 | 17 | EF04 | 92 | 4.0467 | 9.6
2011-63 | 17 | EF04 | 93 | 5.3003 | 10.4
2011-63 | 17 | EF04 | 94 | 8.0725 | 17.1
2011-63 | 18 | EF05 | 95 | 0.021294 | 8.5
2011-63 | 18 | EF05 | 96 | 0.084025 | 8.4
2011-63 | 18 | EF05 | 97 | 0.1577 | 8.4
2011-63 | 18 | EF05 | 98 | 1.5957 | 8.4
2011-63 | 18 | EF05 | 99 | 1.8334 | 8.4
2011-63 | 18 | EF05 | 100 | 1.92 | 8.4
2011-63 | 18 | EF05 | 101 | 2.4527 | 8.2
2011-63 | 18 | EF05 | 102 | 2.8831 | 9.4 | 36 | OXY: Replicate Outliers Not clear from SBE DO which replicate is better
2011-63 | 18 | EF05 | 103 | 4.0667 | 9.7
2011-63 | 18 | EF05 | 104 | 4.8296 | 10.6
2011-63 | 18 | EF05 | 105 | 8.2222 |  |  | OXY: No draw T recorded
2011-63 | 19 | PI01 | 106 | 0.82589 |  |  | OXY: No draw T recorded
2011-63 | 19 | PI01 | 107 | 0.76522 |  |  | OXY: No draw T recorded
2011-63 | 21 | PI03 | 108 | 1.2059 |  |  | OXY: No draw T recorded
2011-63 | 23 | EF06 | 109 | 0.53369 | 8.3 | 3 | OXY: Bubble from H2SO4 dispenser. Mild outlier in comparison to the ctd
2011-63 | 23 | EF06 | 110 | 1.6486 | 8.3
2011-63 | 23 | EF06 | 111 | 1.7989 | 8.5
2011-63 | 23 | EF06 | 112 | 1.8678 | 8.4
2011-63 | 23 | EF06 | 113 | 2.4177 | 8.7
2011-63 | 23 | EF06 | 114 | 3.4241 | 9.3
2011-63 | 23 | EF06 | 115 | 4.0329 | 9.8
2011-63 | 23 | EF06 | 116 | 4.6972 | 10.6 | 6
2011-63 | 23 | EF06 | 117 | 8.3408 | 15.5
2011-63 | 24 | EF07 | 118 | 1.6332 | 8.5
2011-63 | 24 | EF07 | 119 | 1.7596 | 8.7
2011-63 | 24 | EF07 | 120 | 1.8477 | 8.6
2011-63 | 24 | EF07 | 121 | 2.3533 | 8.8
2011-63 | 24 | EF07 | 122 | 3.4111 | 9.4
2011-63 | 24 | EF07 | 123 | 3.8449 | 9.9
2011-63 | 24 | EF07 | 124 | 4.7722 | 11.9 | 6
2011-63 | 24 | EF07 | 125 | 8.0753 | 15.3
2011-63 | 25 | EF08 | 126 | 1.7762 | 8.5
2011-63 | 25 | EF08 | 127 | 1.7761 | 9.1 | 3 | OXY: Bubble from H2SO4 dispenser. Looks ok in comparison to the ctd. Too much uncertainty to remove flag. See notes.
2011-63 | 25 | EF08 | 128 | 2.3067 | 8.6
2011-63 | 25 | EF08 | 129 | 3.3256 | 10
2011-63 | 25 | EF08 | 130 | 3.6991 | 10.7
2011-63 | 25 | EF08 | 131 | 4.5514 | 11 | 6
2011-63 | 25 | EF08 | 132 | 7.738 | 18
2011-63 | 26 | EF09 | 133 | 2.1966 | 8.8
2011-63 | 26 | EF09 | 134 | 2.3619 | 9.5 | 3 | OXY: Bubble from H2SO4 dispenser. Major outlier in comparison to the ctd. See notes.
2011-63 | 26 | EF09 | 135 | 3.2788 | 9
2011-63 | 26 | EF09 | 136 | 3.7478 | 9.9 | 3 | OXY: Bubbles. Looks ok in comparison to the ctd. Too much uncertainty to remove flag. See notes.
2011-63 | 26 | EF09 | 137 | 4.344 | 10.6
2011-63 | 26 | EF09 | 138 | -99 | 15.6 | 5 | OXY: Bubbles. Sample lost but analyst did not give reason why
2011-63 | 27 | EF10 | 139 | 1.5798 | 9.3
2011-63 | 27 | EF10 | 140 | 1.6438 | 8.7
2011-63 | 27 | EF10 | 141 | 2.6556 | 11.4
2011-63 | 27 | EF10 | 142 | 3.2917 |  |  | OXY: No draw T recorded
2011-63 | 27 | EF10 | 143 | 3.6647 |  |  | OXY: No draw T recorded
2011-63 | 27 | EF10 | 144 | 4.3932 | 10.7
2011-63 | 27 | EF10 | 145 | 7.8664 | 17.5
2011-63 | 28 | EF11 | 146 | 1.6436 | 8.7 | 4 | OXY: Titration ran amuck. Curve does not flatten. No endpoint. A minor outlier in comparison to the ctd. See notes.
2011-63 | 28 | EF11 | 147 | 1.3488 | 9.3 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint. Major outlier in comparison to the ctd. See notes
2011-63 | 28 | EF11 | 148 | 2.4318 | 9.2 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint. Major outlier in comparison to the ctd. See notes
2011-63 | 28 | EF11 | 149 | 3.3646 | 10.1
2011-63 | 28 | EF11 | 150 | 3.7454 | 10.4
2011-63 | 28 | EF11 | 151 | 4.3651 | 10.9
2011-63 | 28 | EF11 | 152 | 7.9442 | 13.5
2011-63 | 29 | EF12 | 153 | 1.9126 | 8.3
2011-63 | 29 | EF12 | 154 | 1.9029 | 8.7
2011-63 | 29 | EF12 | 155 | 2.7687 | 8.9
2011-63 | 29 | EF12 | 156 | 3.4563 | 11.7 | 3 | OXY: Bubbles. Looks ok in comparison to the ctd. Too much uncertainty to remove flag. See notes.
2011-63 | 29 | EF12 | 157 | 3.9158 | 10
2011-63 | 29 | EF12 | 158 | 4.6689 | 11.1
2011-63 | 29 | EF12 | 159 | 7.9388 | 16.5
2011-63 | 30 | EF13 | 160 | 2.0436 | 8.2
2011-63 | 30 | EF13 | 161 | 1.9255 | 8.7
2011-63 | 30 | EF13 | 162 | 2.8157 | 9.6
2011-63 | 30 | EF13 | 163 | 3.6114 | 9.6 | 3 | OXY: Bubbles. Looks ok in comparison to the ctd. Too much uncertainty to remove flag. See notes.
2011-63 | 30 | EF13 | 164 | 4.0461 | 10.1
2011-63 | 30 | EF13 | 165 | 4.9354 | 13
2011-63 | 30 | EF13 | 166 | 8.2638 | 15.4

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Sp for Dissolved Oxygen for the range 0.042 to 5.333 mL/L was 0.027, k = 12 with 1 outlier(s) removed
Precision Calculation for Duplicate Samples:
Precision was determined by analyzing replicate samples drawn from one Niskin
Pooled standard deviation of pairs of samples (Sp) was calucalated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs
Determination of Outliers
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded. The Chauvenet Statistic assumes a normal distribution
Duplicate Analysis for Samples Drawn From a Single Niskin bottle:
Event # | Sample# | Pressure | DO1 | DO2 | DOAvg | Draw T1 | Draw T2 | Draw TAvg | Difference | Diff2 | Chauvenet | Reject? | Comment | Critical values for Chauvenet statistic
3 | 8 | 20 | 4.2472 | 4.2384 | 4.2428 | 12.7 | 12.7 | 12.7 | -0.0087287 | 7.619e-05 | 0.30743 |  |  | n | Chauvenet
6 | 17 | 119 | 0.14179 | 0.041669 | 0.091732 | 13.9 | 13.9 | 13.9 | -0.10012 | 0.010025 | 0.63712 |  |  | 3 | 1.37
7 | 26 | 75 | 4.3409 | 4.3606 | 4.3508 | 13.7 | 13.7 | 13.7 | 0.019754 | 0.00039024 | 0.20469 |  |  | 4 | 1.53
11 | 54 | 1.8 | 5.3334 | 5.2948 | 5.3141 | 13.9 | 13.8 | 13.85 | -0.038572 | 0.0014878 | 0.41508 |  |  | 5 | 1.64
13 | 67 | 30 | 4.9378 | 4.9346 | 4.9362 | 12.5 | 12.5 | 12.5 | -0.0031921 | 1.019e-05 | 0.28746 |  |  | 6 | 1.73
15 | 76 | 50 | 1.1688 | 1.1937 | 1.1813 | 8.7 | 8.7 | 8.7 | 0.024858 | 0.00061793 | 0.18627 |  |  | 7 | 1.8
16 | 84 | 50 | 1.6967 | 1.7054 | 1.7011 | 8.7 | 8.7 | 8.7 | 0.008676 | 7.5273e-05 | 0.24465 |  |  | 8 | 1.87
17 | 90 | 50 | 2.4006 | 2.4091 | 2.4048 | 8.6 | 8.8 | 8.7 | 0.0085301 | 7.2763e-05 | 0.24517 |  |  | 9 | 1.91
18 | 102 | 30 | 2.4087 | 3.3575 | 2.8831 | 9.4 | 9.4 | 9.4 | 0.94871 | 0.90006 | 3.1463 | Y |  | 10 | 1.97
23 | 116 | 10 | 4.6989 | 4.6955 | 4.6972 | 10.6 | 10.6 | 10.6 | -0.0033773 | 1.1406e-05 | 0.28813 |  |  | 20 | 2.24
24 | 124 | 10 | 4.7672 | 4.7772 | 4.7722 | 11.9 | 11.9 | 11.9 | 0.0099861 | 9.9721e-05 | 0.23992 |  |  | 30 | 2.39
25 | 131 | 10 | 4.5257 | 4.5771 | 4.5514 | 11 | 11 | 11 | 0.051438 | 0.0026459 | 0.090393 |  |  | 50 | 2.57
[14 empty] | 80 | 2.74
[8 empty] | Mean Diff | 0.076497 |  |  |  |  | 100 | 2.81
[8 empty] | Stdev | 0.27722 |  |  |  |  | 150 | 2.93
[8 empty] | Count |  | 12 |  |  |  | 200 | 3.02
[8 empty] | Sum |  | 0.91557
[8 empty] | Sp |  | 0.19532
[8 empty] | MIN/MAX | 0.041669 | 5.3334
1 Outlier(s) Excluded
[8 empty] | Count |  | 11
[8 empty] | Sum |  | 0.015512
[8 empty] | Sp |  | 0.026554

## Flags
Institute of Ocean Sciences -Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## AutoOxy Raw Output
|  |  | Count | 179 | Min/Max | -0.032 | 8.624
File Name | LOC:EVENT_NUMBER | STN:STATION | Sample_number | Depth | Raw Oxygen:Dissolved | Thiosulfate_titer | Flask_Number | Flask_Volume | Blank | Standard | KIO3 Std volume ml | KIO3 Normality | THIO N | New Oxygen:Dissolved | Draw Temp | Flag | Comment
2011-630001.oxy | 1 | TI02 | 1 | 49 | 0 | 0 | 3040 | 132.56 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.5 |  | OXY: Anoxic. Titration aborted
2011-630001.oxy | 1 | TI02 | 2 | 40 | 0 | 0 | 2064 | 144.24 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.5 |  | OXY: Anoxic. Titration aborted
2011-630001.oxy | 1 | TI02 | 3 | 35 | 0 | 0 | 471 | 134.51 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.5 |  | OXY: Anoxic. Titration aborted
2011-630001.oxy | 1 | TI02 | 4 | 30 | 0.045 | 0.018 | 3160 | 144.9 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0.045475 | 13.3
2011-630001.oxy | 3 | TI01 | 5 | 68 | 4.685 | 0.856 | 3166 | 145.75 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.6853 | 13 |  | Curve noisy on flat portion but endpoint ok
2011-630001.oxy | 3 | TI01 | 6 | 49 | 4.701 | 0.795 | 3191 | 134.95 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.7007 | 13.1
2011-630001.oxy | 3 | TI01 | 7 | 30 | 4.561 | 0.819 | 2032 | 143.2 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.5614 | 12.9
2011-630001.oxy | 3 | TI01 | 8 | 20 | 4.247 | 0.694 | 1401 | 130.22 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.2472 | 12.7
2011-630001.oxy | 3 | TI01 | 8 | 20 | 4.238 | 0.714 | 2035 | 134.24 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.2384 | 12.7
2011-630001.oxy | 3 | TI01 | 9 | 9 | 4.282 | 0.777 | 3098 | 144.6 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.2822 | 12.3
2011-630001.oxy | 3 | TI01 | 10 | 1.8 | 8.236 | 1.456 | 3038 | 141.76 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 8.2362 | 15.3
2011-630001.oxy | 4 | TH01 | 11 | 47 | -0.005 | 0.009 | 1497 | 139.62 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.4 |  | OXY: Anoxic. Titration aborted
2011-630004.oxy | 4 | TH01 | 12 | 40 | 0 | 0 | 3104 | 137.0133 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.3 |  | OXY: Anoxic. Titration aborted
2011-630004.oxy | 4 | TH01 | 13 | 20 | [11 empty] | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-630004.oxy | 4 | TH01 | 14 | 10 | [11 empty] | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-630004.oxy | 4 | TH01 | 15 | 0 |  |  |  |  | 0.0077 | [6 empty] | 5 | OXY: Sample not drawn. Ran out of time. Only 2 minutes between casts
2011-630006.oxy | 6 | THAD1 | 16 | 119 | 0.018 | 0.013 | 1492 | 140 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.9 |  | OXY: Anoxic. Titration aborted
2011-630006.oxy | 6 | THAD1 | 16 | 119 | 0.018 | 0.013 | 3171 | 143.06 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.9 |  | OXY: Anoxic. Titration aborted
2011-630006.oxy | 6 | THAD1 | 17 | 119 | 0.142 | 0.035 | 3109 | 143.26 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0.14179 | 13.9
2011-630006.oxy | 6 | THAD1 | 17 | 119 | 0.042 | 0.017 | 1317 | 138.85 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0.041669 | 13.9
2011-630006.oxy | 6 | THAD1 | 18 | 100 | 0.035 | 0.016 | 3174 | 141.33 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0.035215 | 13.9
2011-630006.oxy | 6 | THAD1 | 19 | 74 | 3.593 | 0.596 | 3204 | 131.86 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.5928 | 13.9
2011-630006.oxy | 6 | THAD1 | 20 | 50 | 4.088 | 0.739 | 3142 | 143.97 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.0881 | 12.9
2011-630006.oxy | 6 | THAD1 | 21 | 30 | 3.864 | 0.683 | 3009 | 140.67 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.864 | 13.8
2011-630006.oxy | 6 | THAD1 | 22 | 20 | 4 | 0.681 | 3209 | 135.55 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.0002 | 12.2
2011-630006.oxy | 6 | THAD1 | 23 | 10 | 3.693 | 0.631 | 3184 | 135.88 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.693 | 11.6
2011-630006.oxy | 6 | THAD1 | 24 | 0 | 7.656 | 1.377 | 2084 | 144.13 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 7.6565 | 13.5
2011-630007.oxy | 7 | THAD2 | 25 | 102 | 0 | 0 | 1421 | 139.37 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 0 | 13.8 |  | OXY: Anoxic. Titration aborted
2011-630007.oxy | 7 | THAD2 | 26 | 75 | 4.341 | 0.762 | 708 | 139.92 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.3409 | 13.7
2011-630007.oxy | 7 | THAD2 | 26 | 75 | 4.366 | 0.756 | 1488 | 138.2 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.3606 | 13.7
2011-630007.oxy | 7 | THAD2 | 27 | 50 | 4.574 | 0.832 | 3181 | 145.07 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.5741 | 12.9
2011-630007.oxy | 7 | THAD2 | 28 | 30 | 3.986 | 0.681 | 1322 | 136.03 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.9858 | 12.5
2011-630007.oxy | 7 | THAD2 | 29 | 20 | 3.949 | 0.65 | 1481 | 131.05 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.9485 | 12.1
2011-630007.oxy | 7 | THAD2 | 30 | 10 | 3.904 | 0.695 | 3137 | 141.69 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 3.9041 | 11.5
2011-630007.oxy | 7 | THAD2 | 31 | 1.7 | 7.733 | 1.36 | 3055 | 140.98 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 7.7326 | 13.7
2011-630007.oxy | 8 | THAD3 | 32 | 50 |  |  | 3007 |  | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 |  | 13.5 | 5 | OXY: Failed titration. Sample discarded. Object in flask
2011-630007.oxy | 8 | THAD3 | 33 | 50 | 4.694 | 0.826 | 1400 | 140.39 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.6943 | 12.3
2011-630007.oxy | 8 | THAD3 | 34 | 50 | 4.376 | 0.788 | 2033 | 143.56 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.3755 | 13.9
2011-630007.oxy | 8 | THAD3 | 35 | 20 | 4.061 | 0.664 | 1311 | 130.22 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.0609 | 13.8
2011-630007.oxy | 8 | THAD3 | 36 | 10 | 4.157 | 0.709 | 1490 | 135.87 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.1571 | 11.7
2011-630007.oxy | 8 | THAD3 | 37 | 0 | 2.211 | 0.591 | 2069 | 143.17 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 |  | 16.4 | 5 | OXY: First attempt, aborted due to bubble. Second attempt, no endpoint. Sample lost
2011-630008.oxy | 9 | TIAD2 | 38 | 65 | 4.698 | 0.777 | 2065 | 131.97 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.6983 | 13.1
2011-630008.oxy | 9 | TIAD2 | 39 | 50 | 4.556 | 0.762 | 1489 | 133.58 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.5501 | 12.9
2011-630009.oxy | 9 | TIAD2 | 40 | 30 | 4.431 | 0.729 | 1464 | 131.19 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.4309 | 12.7
2011-630009.oxy | 9 | TIAD2 | 41 | 20 | 4.13 | 0.701 | 1387 | 135.22 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.1296 | 12.2
2011-630009.oxy | 9 | TIAD2 | 42 | 10 | 6.483 | 1.074 | 638 | 132.65 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 6.4832 | 11.7
2011-630009.oxy | 9 | TIAD2 | 43 | 0 |  |  | 1365 |  | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 |  | 13.4 | 5 | OXY: Missed endpoint. Stirrer turned off.
2011-630009.oxy | 10 | TIAD3 | 44 | 53 | 4.834 | 0.86 | 2068 | 141.98 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8343 | 13.3
2011-630009.oxy | 10 | TIAD3 | 45 | 30 | 4.895 | 0.838 | 1144 | 136.84 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8887 | 136.1 |  | Endpoint late by 0.001. Recalculated
2011-630009.oxy | 10 | TIAD3 | 46 | 20 | 4.361 | 0.746 | 631 | 136.38 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.3605 | 12.4
2011-630009.oxy | 10 | TIAD3 | 47 | 10 | 4.399 | 0.744 | 468 | 134.84 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.399 | 12.7
2011-630009.oxy | 10 | TIAD3 | 48 | 1.7 | 6.027 | 0.98 | 1475 | 130.27 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 6.0202 | 13.4 |  | Endpoint late by 0.001. Recalculated
2011-630009.oxy | 11 | CL03 | 49 | 73 | 4.823 | 0.786 | 745 | 130.09 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8232 | 13.3
2011-630009.oxy | 11 | CL03 | 50 | 50 | 4.875 | 0.838 | 1300 | 137.22 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8749 | 12.9
2011-630009.oxy | 11 | CL03 | 51 | 30 | 4.907 | 0.833 | 1395 | 135.52 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9072 | 12.9
2011-630009.oxy | 11 | CL03 | 52 | 20 | 4.881 | 0.833 | 1321 | 136.23 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8812 | 12.9
2011-630009.oxy | 11 | CL03 | 53 | 10 | 4.965 | 0.858 | 1406 | 137.98 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9648 | 13
2011-630009.oxy | 11 | CL03 | 54 | 1.8 | 5.333 | 0.952 | 1407 | 142.61 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 5.3334 | 13.9
2011-630009.oxy | 11 | CL03 | 54 | 1.8 | 5.295 | 0.878 | 1476 | 132.51 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 5.2948 | 13.8
2011-630009.oxy | 12 | CL02 | 55 | 146 | 4.878 | 0.831 | 1123 | 135.99 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8781 | 12.9
2011-630009.oxy | 12 | CL02 | 56 | 125 | 4.883 | 0.822 | 1330 | 134.4 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8826 | 12.9
2011-630009.oxy | 12 | CL02 | 57 | 100 | 4.895 | 0.859 | 1411 | 140.09 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8947 | 12.9
2011-630009.oxy | 12 | CL02 | 58 | 75 | 4.89 | 0.877 | 3151 | 143.15 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8901 | 12.9
2011-630009.oxy | 12 | CL02 | 59 | 50 | 4.929 | 0.898 | 1248 | 145.41 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9296 | 12.9
2011-630009.oxy | 12 | CL02 | 60 | 30 | 4.96 | 0.83 | 1318 | 133.63 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9595 | 12.9
2011-630012.oxy | 12 | CL02 | 61 | 20 | 4.95 | 0.905 | 2020 | 145.95 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9498 | 13.9
2011-630012.oxy | 12 | CL02 | 62 | 10 | 4.931 | 0.835 | 1373 | 135.2 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9309 | 13.9
2011-630012.oxy | 12 | CL02 | 63 | 1.6 | 5.148 | 0.869 | 1397 | 134.85 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 5.1476 | 13.4
2011-630012.oxy | 13 | CL01 | 64 | 88 | 4.757 | 0.852 | 1298 | 142.91 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.7572 | 12.7 |  | Corrected depth from 1.6 to 88 m
2011-630012.oxy | 13 | CL01 | 65 | 75 | 4.828 | 0.786 | 528 | 129.96 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.8281 | 12.5
2011-630012.oxy | 13 | CL01 | 66 | 50 | 4.933 | 0.807 | 1408 | 130.64 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9325 | 12.5
2011-630012.oxy | 13 | CL01 | 67 | 30 | 4.938 | 0.823 | 1307 | 133.08 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9378 | 12.5
2011-630012.oxy | 13 | CL01 | 67 | 30 | 4.935 | 0.827 | 3206 | 133.81 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9346 | 12.5
2011-630012.oxy | 13 | CL01 | 68 | 20 | 4.944 | 0.818 | 1392 | 132.11 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.944 | 12.5
2011-630012.oxy | 13 | CL01 | 69 | 10 | 4.988 | 0.836 | 3090 | 133.99 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9822 | 12.5 |  | Endpoint late by 0.001. Recalculated
2011-630012.oxy | 13 | CL01 | 70 | 1.6 | 4.986 | 0.912 | 1404 | 146.01 | 0.0077 | 0.7053 | 10.0127 | 0.0099064 | 0.14219 | 4.9864 | 13.7
2011-630012.oxy | 14 | EF01 | 71 | 32 | 3.591 | 0.607 | 3191 | 134.95 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.591 | 9.5
2011-630012.oxy | 14 | EF01 | 72 | 20 | 3.955 | 0.721 | 3166 | 145.75 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.9548 | 10.7
2011-630012.oxy | 14 | EF01 | 73 | 10 | 6.64 | 1.197 | 3160 | 144.9 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 6.6396 | 10.4
2011-630012.oxy | 14 | EF01 | 74 | 1.8 | 7.944 | 1.327 | 471 | 134.51 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 7.944 | 18.1
2011-630012.oxy | 15 | EF02 | 75 | 66 | 1.129 | 0.194 | 3040 | 132.56 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.1295 | 8.4
2011-630012.oxy | 15 | EF02 | 76 | 50 | 1.169 | 0.216 | 2032 | 143.2 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.1688 | 8.7
2011-630012.oxy | 15 | EF02 | 76 | 50 | 1.194 | 0.201 | 1401 | 130.22 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.1937 | 8.7
2011-630012.oxy | 15 | EF02 | 77 | 30 | 2.299 | 0.39 | 2035 | 134.24 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.2993 | 8.9
2011-630012.oxy | 15 | EF02 | 78 | 20 | 3.835 | 0.694 | 3098 | 144.6 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.8354 | 9.3
2011-630012.oxy | 15 | EF02 | 79 | 10 | 6.051 | 1.068 | 3038 | 141.76 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 6.0514 | 10.2
2011-630012.oxy | 15 | EF02 | 80 | 1.8 | 8.624 | 1.495 | 1497 | 139.62 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.6244 | 17.2
2011-630012.oxy | 16 | EF03 | 81 | 114 | 0 | 0 | 3174 | 141.33 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0 | 9.2 |  | OXY: Anoxic. Tirtration aborted
2011-630012.oxy | 16 | EF03 | 82 | 100 | -0.032 | 0.004 | 1317 | 138.85 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0 | 10.3 |  | OXY: Anoxic. Tirtration aborted
2011-630012.oxy | 16 | EF03 | 83 | 75 | 0.614 | 0.118 | 3109 | 143.26 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.61407 | 8.7
2011-630012.oxy | 16 | EF03 | 84 | 50 | 1.697 | 0.309 | 3171 | 143.06 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.6967 | 8.7
2011-630012.oxy | 16 | EF03 | 84 | 50 | 1.705 | 0.304 | 1492 | 140 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.7054 | 8.7
2011-630012.oxy | 16 | EF03 | 85 | 30 | 2.941 | 0.529 | 3104 | 143.16 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.9407 | 9
2011-630012.oxy | 16 | EF03 | 86 | 20 | 4.146 | 0.747 | 2084 | 144.13 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.146 | 9.5
2011-630012.oxy | 16 | EF03 | 87 | 10 | 5.577 | 0.944 | 3184 | 135.88 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 5.5771 | 17.8
2011-630012.oxy | 16 | EF03 | 88 | 1.9 | 8.606 | 1.448 | 3209 | 135.55 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.606 |  |  | OXY: No draw T recorded
2011-630012.oxy | 17 | EF04 | 89 | 60 | 2.256 | 0.401 | 3009 | 140.67 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.256 | 12.1
2011-630012.oxy | 17 | EF04 | 90 | 50 | 2.401 | 0.436 | 3142 | 143.97 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.4006 | 8.6
2011-630012.oxy | 17 | EF04 | 90 | 50 | 2.409 | 0.401 | 3204 | 131.86 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.4091 | 8.8
2011-630012.oxy | 17 | EF04 | 91 | 30 | 3.458 | 0.604 | 1421 | 139.37 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.458 | 9.2
2011-630012.oxy | 17 | EF04 | 92 | 20 | 4.047 | 0.708 | 708 | 139.92 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.0467 | 9.6
2011-630012.oxy | 17 | EF04 | 93 | 10 | 5.3 | 0.913 | 1488 | 138.2 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 5.3003 | 10.4
2011-630012.oxy | 17 | EF04 | 94 | 1.9 | 8.072 | 1.455 | 3181 | 145.07 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.0725 | 17.1
2011-630012.oxy | 18 | EF05 | 95 | 198 | 0.021 | 0.013 | 1322 | 136.03 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.021294 | 8.5 |  | H2S odour but plot looks good
2011-630012.oxy | 18 | EF05 | 96 | 175 | 0.078 | 0.023 | 1481 | 131.05 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.084025 | 8.4 |  | H2S odour but plot looks good
2011-630012.oxy | 18 | EF05 | 97 | 150 | 0.158 | 0.037 | 3137 | 141.69 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.1577 | 8.4
2011-630012.oxy | 18 | EF05 | 98 | 125 | 1.596 | 0.287 | 3055 | 140.98 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.5957 | 8.4
2011-630012.oxy | 18 | EF05 | 99 | 100 | 1.822 | 0.331 | 3007 | 142.13 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.8334 | 8.4 |  | Endpoint early by 0.002. recalculated
2011-630012.oxy | 18 | EF05 | 100 | 75 | 1.92 | 0.342 | 1400 | 140.39 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.92 | 8.4
2011-630012.oxy | 18 | EF05 | 101 | 50 | 2.453 | 0.444 | 2033 | 143.56 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.4527 | 8.2
2011-630012.oxy | 18 | EF05 | 102 | 30 | 2.409 | 0.396 | 1311 | 130.22 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.4087 | 9.4
2011-630012.oxy | 18 | EF05 | 102 | 30 | 3.358 | 0.572 | 1490 | 135.87 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.3575 | 9.4
2011-630012.oxy | 18 | EF05 | 103 | 20 | 4.067 | 0.728 | 2069 | 143.17 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.0667 | 9.7
2011-630012.oxy | 18 | EF05 | 104 | 10 | 4.83 | 0.796 | 690 | 132.12 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.8296 | 10.6
2011-630012.oxy | 18 | EF05 | 105 | 2.1 | 8.222 | 1.347 | 2065 | 131.97 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.2222 |  |  | OXY: No draw T recorded
2011-630019.oxy | 19 | PI01 | 106 | 38 | 0.826 | 0.152 | 708 | 139.92 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.82589 |  |  | OXY: No draw T recorded
2011-630019.oxy | 19 | PI01 | 107 | 38 | 0.765 | 0.141 | 1421 | 139.37 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.76522 |  |  | OXY: No draw T recorded
2011-630019.oxy | 21 | PI03 | 108 | 44 | 1.206 | 0.215 | 1488 | 138.2 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.2059 |  |  | OXY: No draw T recorded
2011-630019.oxy | 23 | EF06 | 109 | 153 | 0.534 | 0.105 | 3181 | 145.073 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 0.53369 | 8.3 | 3 | OXY: Bubble in H2SO4
2011-630019.oxy | 23 | EF06 | 110 | 125 | 1.649 | 0.286 | 1322 | 136.03 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.6486 | 8.3
2011-630019.oxy | 23 | EF06 | 111 | 100 | 1.799 | 0.3 | 1481 | 131.05 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.7989 | 8.5
2011-630019.oxy | 23 | EF06 | 112 | 75 | 1.868 | 0.336 | 3137 | 141.69 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.8678 | 8.4
2011-630019.oxy | 23 | EF06 | 113 | 50 | 2.418 | 0.43 | 3055 | 140.98 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.4177 | 8.7
2011-630019.oxy | 23 | EF06 | 114 | 30 | 3.424 | 0.61 | 3007 | 142.13 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.4241 | 9.3
2011-630019.oxy | 23 | EF06 | 115 | 20 | 4.033 | 0.708 | 1400 | 140.39 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.0329 | 9.8
2011-630019.oxy | 23 | EF06 | 116 | 10 | 4.699 | 0.842 | 2033 | 143.56 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.6989 | 10.6
2011-630019.oxy | 23 | EF06 | 116 | 10 | 4.696 | 0.763 | 1311 | 130.22 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.6955 | 10.6 |  | OXY: No draw T recorded
2011-630019.oxy | 23 | EF06 | 117 | 1.9 | 8.341 | 1.407 | 1490 | 135.87 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.3408 | 15.5
2011-630019.oxy | 24 | EF07 | 118 | 118 | 1.633 | 0.298 | 2069 | 143.17 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.6332 | 8.5 |  | Corrected depth info
2011-630019.oxy | 24 | EF07 | 119 | 100 | 1.76 | 0.296 | 690 | 132.116 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.7596 | 8.7
2011-630019.oxy | 24 | EF07 | 120 | 75 | 1.848 | 0.31 | 2065 | 131.97 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.8477 | 8.6 |  | Recalculated oxy. Previous sample's info was used in this points calculation
2011-630019.oxy | 24 | EF07 | 121 | 50 | 2.365 | 0.397 | 1489 | 133.58 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.3533 | 8.8 |  | Endpoint late by 0.002. Recalculated
2011-630019.oxy | 24 | EF07 | 122 | 30 | 3.411 | 0.561 | 1464 | 131.19 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.4111 | 9.4
2011-630019.oxy | 24 | EF07 | 123 | 20 | 3.845 | 0.653 | 1365 | 135.73 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.8449 | 9.9
2011-630019.oxy | 24 | EF07 | 124 | 10 | 4.767 | 0.789 | 638 | 132.65 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.7672 | 11.9
2011-630019.oxy | 24 | EF07 | 124 | 10 | 4.777 | 0.806 | 1387 | 135.22 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.7772 | 11.9 |  | Corrected depth info.
2011-630019.oxy | 24 | EF07 | 125 | 1.7 | 8.075 | 1.389 | 3207 | 138.49 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 8.0753 | 15.3
2011-630019.oxy | 25 | EF08 | 126 | 83 | 1.776 | 0.305 | 3191 | 134.95 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.7762 | 8.5 |  | Corrected cast info
2011-630019.oxy | 25 | EF08 | 127 | 75 | 1.776 | 0.329 | 3166 | 145.75 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 1.7761 | 9.1 | 3 | OXY: Bubble in H2SO4. Corrected sample, depth and flask info
2011-630019.oxy | 25 | EF08 | 128 | 50 | 2.307 | 0.422 | 3160 | 144.9 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.3067 | 8.6
2011-630019.oxy | 25 | EF08 | 129 | 30 | 3.326 | 0.561 | 471 | 134.51 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.3256 | 10
2011-630019.oxy | 25 | EF08 | 130 | 20 | 3.699 | 0.668 | 2064 | 144.24 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.6991 | 10.7
2011-630019.oxy | 25 | EF08 | 131 | 10 | 4.526 | 0.749 | 3040 | 132.56 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.5257 | 11
2011-630019.oxy | 25 | EF08 | 131 | 10 | 4.577 | 0.744 | 1401 | 130.22 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.5771 | 11 |  | OXY: Corrected Depth info.
2011-630019.oxy | 25 | EF08 | 132 | 1.6 | 7.738 | 1.377 | 2032 | 143.2 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 7.738 | 18
2011-630019.oxy | 26 | EF09 | 133 | 60 | 2.197 | 0.373 | 2035 | 134.24 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.1966 | 8.8
2011-630019.oxy | 26 | EF09 | 134 | 50 | 2.362 | 0.431 | 3098 | 144.6 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 2.3619 | 9.5 | 3 | OXY: Bubble in H2SO4
2011-630019.oxy | 26 | EF09 | 135 | 30 | 3.279 | 0.583 | 3038 | 141.76 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.2788 | 9 |  | Corrected sample, depth and flask info
2011-630019.oxy | 26 | EF09 | 136 | 20 | 3.748 | 0.655 | 1497 | 139.62 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 3.7478 | 9.9 | 3 | OXY: Bubbles
2011-630019.oxy | 26 | EF09 | 137 | 10 | 4.344 | 0.767 | 3174 | 141.33 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 | 4.344 | 10.6
2011-630019.oxy | 26 | EF09 | 138 | 1.7 | 0 | 0 | 1317 | 138.85 | 0.0073 | 0.7023 | 10.0127 | 0.0099064 | 0.14272 |  | 15.6 | 5 | OXY: Bubbles. Sample lost but analyst did not give reason why
2011-630027.oxy | 27 | EF10 | 139 | 95 | 1.577 | 0.288 | 3109 | 143.26 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.5798 | 9.3
2011-630027.oxy | 27 | EF10 | 140 | 75 | 1.641 | 0.299 | 3171 | 143.06 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.6438 | 8.7
2011-630027.oxy | 27 | EF10 | 141 | 50 | 2.654 | 0.469 | 1492 | 140 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 2.6556 | 11.4 |  | Corrected sample, depth and flask info
2011-630027.oxy | 27 | EF10 | 142 | 30 | 3.291 | 0.593 | 3104 | 143.16 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.2917 |  |  | OXY: No draw T recorded
2011-630027.oxy | 27 | EF10 | 143 | 20 | 3.664 | 0.664 | 2084 | 144.13 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.6647 |  |  | OXY: No draw T recorded
2011-630027.oxy | 27 | EF10 | 144 | 10 | 4.394 | 0.749 | 3184 | 135.88 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 4.3932 | 10.7
2011-630027.oxy | 27 | EF10 | 145 | 0 | 7.87 | 1.333 | 3209 | 135.55 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 7.8664 | 17.5
2011-630027.oxy | 28 | EF11 | 146 | 83 | 1.641 | 0.294 | 3009 | 140.67 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.6436 | 8.7 | 4 | OXY: Titration ran amuck. Curve does not flatten. No endpoint
2011-630027.oxy | 28 | EF11 | 147 | 75 | 1.346 | 0.248 | 3142 | 143.97 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.3488 | 9.3 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint
2011-630027.oxy | 28 | EF11 | 148 | 50 | 2.43 | 0.405 | 3204 | 131.86 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 2.4318 | 9.2 | 4 | OXY: Titration ran amuck. Bad curve. No endpoint
2011-630027.oxy | 28 | EF11 | 149 | 30 | 3.364 | 0.601 | 1144 | 141.98 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.3646 | 10.1 |  | Corrected depth info
2011-630027.oxy | 28 | EF11 | 150 | 20 | 3.745 | 0.644 | 1144 | 136.84 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.7454 | 10.4
2011-630027.oxy | 28 | EF11 | 151 | 10 | 4.365 | 0.747 | 631 | 136.38 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 4.3651 | 10.9
2011-630027.oxy | 28 | EF11 | 152 | 0 | 7.948 | 1.339 | 468 | 134.84 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 7.9442 | 13.5
2011-630027.oxy | 29 | EF12 | 153 | 85 | 1.91 | 0.316 | 1475 | 130.27 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.9126 | 8.3
2011-630027.oxy | 29 | EF12 | 154 | 75 | 1.9 | 0.314 | 745 | 130.09 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.9029 | 8.7
2011-630027.oxy | 29 | EF12 | 155 | 50 | 2.767 | 0.479 | 1300 | 137.22 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 2.7687 | 8.9
2011-630027.oxy | 29 | EF12 | 156 | 30 | 3.456 | 0.589 | 1395 | 135.52 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.4563 | 11.7 | 3 | OXY: Bubble in H2SO4
2011-630027.oxy | 29 | EF12 | 157 | 20 | 3.916 | 0.67 | 1321 | 136.23 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.9158 | 10
2011-630027.oxy | 29 | EF12 | 158 | 10 | 4.669 | 0.808 | 1406 | 137.98 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 4.6689 | 11.1
2011-630027.oxy | 29 | EF12 | 159 | 0 | 7.943 | 1.416 | 1407 | 142.61 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 7.9388 | 16.5
2011-630027.oxy | 30 | EF13 | 160 | 97 | 2.041 | 0.343 | 1476 | 132.51 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 2.0436 | 8.2
2011-630027.oxy | 30 | EF13 | 161 | 74 | 1.923 | 0.332 | 1123 | 135.99 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 1.9255 | 8.7
2011-630027.oxy | 30 | EF13 | 162 | 30 | 2.814 | 0.477 | 1330 | 134.4 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 2.8157 | 9.6
2011-630027.oxy | 30 | EF13 | 163 | 30 | 3.611 | 0.636 | 1411 | 140.09 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 3.6114 | 9.6 | 3 | OXY: Bubble in H2SO4
2011-630027.oxy | 30 | EF13 | 164 | 20 | 4.046 | 0.739 | 1248 | 145.41 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 4.0461 | 10.1
2011-630027.oxy | 30 | EF13 | 165 | 10 | 4.936 | 0.886 | 3151 | 143.15 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 4.9354 | 13
2011-630027.oxy | 30 | EF13 | 166 | 0 | 8.268 | 1.38 | 1318 | 133.63 | 0.004 | 0.7053 | 10.0127 | 0.0099064 | 0.14144 | 8.2638 | 15.4

## Thio N
Casts Analyzed | Stdization Date | Standard Titer | Spread | Blank | KIO3 Burette Volume | KIO3 N | Thio N | Diff | Comment
none | 2011-08-16 | 0.7035 | 0.001 | 0.0077 | 10.0127 | 0.0099064 | 0.14256 |  | Stds not used. Not recorded by program. Blanks not stable. Increasing trend in individual blank titers
#1-13 | 2011-08-16 | 0.7053 | 0.001 | 0.0077 | 10.0127 | 0.0099064 | 0.14219 | 0.00036783 | Blanks not stable. Increasing trend in individual blank titers
#15-26 | 2011-08-18 | 0.7023 | 0.002 | 0.0073 | 10.0127 | 0.0099064 | 0.14272 | -0.00053192 | Stds ok but should have analyzed one more to confirm stability and to know which one to reject. Titers are also quite a bit lower even though no reagent changes were made
none | 2011-08-20 | 0.7053 | 0.001 | 0.0047 | 10.0127 | 0.0099064 | 0.14158 | 0.0011408 | Stds ok. New batch H2S04. New blank but problems with dispensers and precision not good
#27-30 | 2011-12-28 | 0.7053 | 0.001 | 0.004 | 10.0127 | 0.0099064 | 0.14144 | 0.0012821 | Recalculated blank excluding one individual titer that was high. Used this value to recalculate final Oxy
| min | 0.7023 |  | 0.0047 |  |  | 0.14144
| max | 0.7053 |  | 0.0077 |  |  | 0.14272
| diff | 0.003 |  | 0.003 |  |  | 0.0012821
Inverse relationship between Std. Titer and thio Normality
Std titer most affected by T indirectly
Therefore Thio N and T have a direct relationship

## Titration Plots
Sample 37 Part 1
2 | 0.002 | 1.3 | 0.002
5 | 0.2 | 4.1 | 0.202
8 | 0.2 | 4.8 | 0.402
11 | 0.2 | 7 | 0.602
14 | 0.2 | 10.7 | 0.802
17 | 0.2 | 20.2 | 1.002
20 | 0.1 | 33.2 | 1.102
23 | 0.1 | 46.6 | 1.202
26 | 0.02 | 38.2 | 1.222
29 | 0.1 | 59.8 | 1.322
31 | 0.02 | 48.4 | 1.342
Sample 37 Part 2
2 | 0.002 | 53.1 | 0.002
4 | 0.02 | 57.2 | 0.022
6 | 0.02 | 58.7 | 0.042
8 | 0.02 | 63.1 | 0.062
10 | 0.02 | 68.7 | 0.082
12 | 0.02 | 72.3 | 0.102
14 | 0.008 | 74.9 | 0.11
16 | 0.008 | 77.7 | 0.118
18 | 0.008 | 80.8 | 0.126
20 | 0.004 | 83.6 | 0.13
22 | 0.004 | 83.5 | 0.134
24 | 0.004 | 84.5 | 0.138
26 | 0.004 | 84.6 | 0.142
28 | 0.004 | 86.4 | 0.146
30 | 0.004 | 87.9 | 0.15
32 | 0.004 | 88.6 | 0.154
34 | 0.004 | 90.5 | 0.158
36 | 0.002 | 91 | 0.16
38 | 0.002 | 91.6 | 0.162
40 | 0.002 | 92.2 | 0.164
42 | 0.002 | 92.6 | 0.166
44 | 0.002 | 93.3 | 0.168
46 | 0.002 | 93.9 | 0.17
48 | 0.002 | 94.5 | 0.172
50 | 0.002 | 94.9 | 0.174
52 | 0.002 | 95.6 | 0.176
54 | 0.001 | 96.4 | 0.177
56 | 0.001 | 96.5 | 0.178
58 | 0.001 | 97 | 0.179
60 | 0.001 | 97.1 | 0.18
62 | 0.001 | 97.4 | 0.181
64 | 0.001 | 97.6 | 0.182
66 | 0.001 | 98 | 0.183
68 | 0.001 | 98.3 | 0.184
70 | 0.001 | 98.6 | 0.185
72 | 0.001 | 98.9 | 0.186
74 | 0.001 | 99.1 | 0.187
76 | 0.001 | 99.3 | 0.188
78 | 0.001 | 99.5 | 0.189
80 | 0.001 | 99.2 | 0.19
82 | 0.001 | 99.5 | 0.191
84 | 0.001 | 99.5 | 0.192
86 | 0.001 | 99.6 | 0.193
88 | 0.001 | 99.6 | 0.194
90 | 0.001 | 99.6 | 0.195
92 | 0.001 | 99.6 | 0.196
94 | 0.001 | 99.6 | 0.197
Sample 147 Part 1
2 | 0.002 | 34.7 | 0.002
5 | 0.1 | 59.4 | 0.102
7 | 0.02 | 51 | 0.122
9 | 0.02 | 57.9 | 0.142
11 | 0.02 | 61.2 | 0.162
13 | 0.02 | 65.3 | 0.182
15 | 0.02 | 68.9 | 0.202
17 | 0.02 | 74.2 | 0.222
19 | 0.008 | 76.1 | 0.23
21 | 0.008 | 80.3 | 0.238
23 | 0.004 | 82.2 | 0.242
25 | 0.004 | 81.3 | 0.246
27 | 0.004 | 82.8 | 0.25
29 | 0.004 | 83.8 | 0.254
31 | 0.004 | 85.1 | 0.258
33 | 0.004 | 85.5 | 0.262
35 | 0.004 | 87 | 0.266
37 | 0.004 | 87.6 | 0.27
39 | 0.004 | 88.7 | 0.274
41 | 0.004 | 90.5 | 0.278
43 | 0.002 | 90.9 | 0.28
45 | 0.002 | 91.8 | 0.282
47 | 0.002 | 91.9 | 0.284
49 | 0.002 | 92.4 | 0.286
51 | 0.002 | 93.2 | 0.288
53 | 0.002 | 93.2 | 0.29
55 | 0.002 | 94.3 | 0.292
57 | 0.002 | 95.2 | 0.294
59 | 0.001 | 95.9 | 0.295
61 | 0.001 | 96.2 | 0.296
63 | 0.001 | 96.4 | 0.297
65 | 0.001 | 96.7 | 0.298
67 | 0.001 | 96.9 | 0.299
69 | 0.001 | 97.1 | 0.3
71 | 0.001 | 97.6 | 0.301
73 | 0.001 | 91.9 | 0.302
Sample 147 Part 2
2 | 0.002 | 34.9 | 0.002
5 | 0.1 | 54.3 | 0.102
7 | 0.02 | 44 | 0.122
10 | 0.05 | 61.4 | 0.172
12 | 0.02 | 60.3 | 0.192
14 | 0.02 | 68 | 0.212
16 | 0.02 | 73 | 0.232
18 | 0.008 | 74.9 | 0.24
20 | 0.008 | 78.1 | 0.248
22 | 0.008 | 78.6 | 0.256
24 | 0.008 | 78.8 | 0.264
26 | 0.008 | 78.9 | 0.272
28 | 0.008 | 78.3 | 0.28
30 | 0.008 | 78.2 | 0.288
32 | 0.008 | 78 | 0.296
34 | 0.008 | 77.9 | 0.304
36 | 0.008 | 77.9 | 0.312
Sample 148 Part 1
2 | 0.002 | 77.3 | 0.002
4 | 0.008 | 77.7 | 0.01
6 | 0.008 | 78.8 | 0.018
8 | 0.008 | 80.7 | 0.026
10 | 0.004 | 83.3 | 0.03
12 | 0.004 | 84.2 | 0.034
14 | 0.004 | 85.6 | 0.038
16 | 0.004 | 86.9 | 0.042
18 | 0.004 | 87.5 | 0.046
20 | 0.004 | 88.4 | 0.05
22 | 0.004 | 89.3 | 0.054
24 | 0.004 | 91 | 0.058
26 | 0.002 | 91.4 | 0.06
28 | 0.002 | 91.8 | 0.062
30 | 0.002 | 92.5 | 0.064
32 | 0.002 | 92.7 | 0.066
34 | 0.002 | 93.2 | 0.068
36 | 0.002 | 93.5 | 0.07
38 | 0.002 | 93.5 | 0.072
40 | 0.002 | 93.2 | 0.074
42 | 0.002 | 93.6 | 0.076
44 | 0.002 | 93.5 | 0.078
46 | 0.002 | 93.4 | 0.08
48 | 0.002 | 93.4 | 0.082
50 | 0.002 | 93.1 | 0.084
Sample 148 Part 2
3 | 0.002 | 23.3 | 0.002
6 | 0.1 | 37 | 0.102
9 | 0.1 | 54.5 | 0.202
11 | 0.02 | 46.3 | 0.222
14 | 0.02 | 59.4 | 0.242
16 | 0.02 | 57.2 | 0.262
18 | 0.02 | 63.2 | 0.282
20 | 0.02 | 67.4 | 0.302
22 | 0.02 | 71.4 | 0.322
24 | 0.008 | 75.1 | 0.33
26 | 0.008 | 77.4 | 0.338
28 | 0.008 | 79.4 | 0.346
30 | 0.008 | 82.2 | 0.354
32 | 0.004 | 83.4 | 0.358
34 | 0.004 | 85.5 | 0.362
36 | 0.004 | 87 | 0.366
38 | 0.004 | 87.2 | 0.37
40 | 0.004 | 89.3 | 0.374
42 | 0.004 | 89.9 | 0.378
44 | 0.004 | 91.7 | 0.382
46 | 0.002 | 92.5 | 0.384
48 | 0.002 | 93.3 | 0.386
50 | 0.002 | 93.7 | 0.388
52 | 0.002 | 94.2 | 0.39
54 | 0.002 | 95 | 0.392
56 | 0.002 | 95.5 | 0.394
58 | 0.001 | 96.2 | 0.395
60 | 0.001 | 96.8 | 0.396
62 | 0.001 | 97.2 | 0.397
64 | 0.001 | 97.4 | 0.398
66 | 0.001 | 97.8 | 0.399
68 | 0.001 | 98.1 | 0.4
70 | 0.001 | 98.4 | 0.401
72 | 0.001 | 98.7 | 0.402
74 | 0.001 | 99 | 0.403
76 | 0.001 | 99.1 | 0.404
78 | 0.001 | 99.2 | 0.405
80 | 0.001 | 99.3 | 0.406
82 | 0.001 | 99.3 | 0.407
84 | 0.001 | 99.4 | 0.408
86 | 0.001 | 99.3 | 0.409
88 | 0.001 | 99.3 | 0.41
90 | 0.001 | 99.2 | 0.411
92 | 0.001 | 99.3 | 0.412
94 | 0.001 | 99.3 | 0.413