## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. |

## PROCESSING NOTES

Cruise: 2011-76

Agency: OSD

Location: Strait of Georgia

Project: SOG Moorings

Party Chief: Macdonald R.W.

Platform: Vector

Date: 25 November 2011 – 28 November 2011

Processed by: Germaine Gatien

Date of Processing: 12 March 2012 – 19 March 2012

Number of original HEX files: 14

Number of CTD files:  3 (Only 3 sampled below 3m.)

Number of bottle casts: 12 (1 hex file does not belong to this cruise, 1 had no sampling)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0997), a Wet Labs Eco-AFL/FL Fluorometer (#2215), a pH sensor (#0692) and an altimeter.

The deck unit was an SBE 11+ (#0425). Seasave V7 was used.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The Rosette and CTD logs were generally in good order with a list of all equipment and clear notes about problems encountered. The event numbers of the first two casts were entered incorrectly on the rosette log sheets and in the headers.

The SBE data are considered of lower quality than normal because 24 scans were averaged in acquisition due to an improper setting in the configuration file. These averages may contain spikes that would normally be removed in processing and might be systematic in nature. Fine-tuning of alignment settings and CELLTM correction are also less sensitive with smoothed data. Unfortunately the same error was made on at least 4 other cruises using the same equipment, so we can not rely on settings determined from those data sets. Given that sampling was in an area that is not usually affected much by shed wakes and the CTD is believed to have worked well with few spikes, it is likely the errors caused by this are small.

CTD salinity was found to be very close to bottles for this cruise and the one that followed it, 2011-11.

Two bottles were meant to be fired at each level during casts #16 and #18; however, in each case Niskin #15 did not fire. Sample numbers were assigned as though water was taken from 2 different bottles, but all samples were actually taken from Niskin #16. Samples originally identified as #35 and #59 were reassigned as #36 and #60, respectively, at the end of processing. However, intermediate files and the analysts’ spreadsheets have sample numbers that are consistent with the numbers assigned to the samples at sea.

The tests for accuracy of Oxygen:Dissolved:SBE data are less reliable than usual, but the data are considered, very roughly, to be:

±0.2mL/L from 0 to 20db

±0.06mL/L from 20 to 150db

low by up to 0.05mL/L below 150db

Data from the SBE pH sensor were collected but are not to be archived at this time. Files were provided to the chief scientist with the pH data included.

Two changes have been made to processing methods for all cruises that occurred from January 2011 onwards:

- A new approach is being taken to the recalibration of the SBE Dissolved Oxygen data. The voltage channel is compared with bottles to find the slope and offset to enter in the configuration files. This method is the standard approach and is recommended by SeaBird.

- The transmissivity conversion has also been changed slightly so that it follows the method outlined in SeaBird Application Note 91. For more information on this see the document in folder: OSD_data_Archive\Cruise_Data\DOCUMENTS\Transmissivity

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names had a syntax error which was fixed.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained and included a list of instruments and science staff and clear indication that the TSG was not in use. The DO sensor is listed as #97 but is believed to be #997. No cruise report was available. There is confusion in the event numbers on the rosette sheets, but the correct choice is clear in the Daily Science log. One rosette bottle failed to fire on the only occasions it was meant to do so. This did not happen during 2011-11 which followed.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. Since the last factory calibration these sensors have been use on at 4 other cruises that have been processed, 2011-60, 2011-63, 2011-10 and 2011-11.

The configuration file did not change during the cruise. The calibration constants were checked for all instruments. There were a few errors in the con files:

- The transmissometer calibration was incorrect.

- The dissolved oxygen sensor serial number was changed from #0097 to #0997.

- The calibration for the Wetlabs Eco fluorometer could not be confirmed. The parameters entered are close to those on the characterization sheet from the factory, but a field test was done just before cruise 2011-60 in August and no record is available for that, but it looks reasonable

- The most serious error is that “Scans to average” was set to 24. This should be set to 1. There is no way to fix this and acquiring metre-averaged data only.  Single bad points will be included in the average since there is no opportunity to remove them first and fine-tuning of settings will be limited or impossible. (This problem also occurred during 2011-60, 2011-63 and 2011-10.)

A configuration file used at sea was corrected and saved as 2011-76-ctd.xmlcon.

There is a hex file for event #1 but it was not a CTD cast; the hex file could not be converted. The file contains data from November 23, 2012 with primary sensors 2374 and 3396 which were not used for 2011-76. This does not seem to belong to any other science cruise, so it may have been a test.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. Sample numbers were added to the ADDSAMP file based on rosette log records.

There were several occasions when bottles were fired but no sample number assigned. Those were removed from the ADDSAMP file.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2011-76-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling. A calculation was made of Ф using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees. Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

The fit using all data gave M = 0.4348 and B = -0.2110 with an R2 value of 0.6271. The initial estimates of M and B were used to calculate a difference between each point and the fit:
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, plots could be made varying the severity of the outlier removal. After the first 2 fits the M and B entries were updated. When removing a little more data has little effect on the fit, it is judged that a reasonable value has been found unless, in so doing, a whole class of points has been removed such as all high values, or all values from late in the cruise. For this cruise there were not many bottles and the scatter in the fit with all flagged bottles has a very low R2 value. As the # of bottles included is varied the changes are quite dramatic. The following table shows the results:

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4348 | -0.2110 | 0.4348 | -0.4853 |  |
| 23 | all unflagged bottles | 0.3850 | -0.0466 | 0.3850 | -0.1210 | 0.6271 |
| 19 | excl. outliers diff>0.1 | 0.4902 | -0.2363 | 0.4902 | -0.4820 | 0.9926 |
| 17 | excl. outliers diff>0.03 | 0.4860 | -0.2335 | 0.4860 | -0.4805 | 0.9977 |
| 15 | excl. outliers diff>0.015 | 0.4927 | -0.2432 | 0.4927 | -0.4936 | 0.9982 |
| 11 | excl. outliers diff>0.01 | 0.4905 | -0.2387 | 0.4905 | -0.4866 | 0.9992 |

The changes are more dramatic and removal of a few bottles has a large effect. The data never settle down even when more than half the bottles have been excluded. These data appear to be too noisy to achieve a believable fit. The Soc values are much higher than for the summer and September cruises on which this sensor was used. For 2011-11 in December the slope was also high at 0.4797, but for that cruise there were lots of bottles and the results settled down after a few outliers were removed. The reduced DO range may be affecting the fit. There could also be problems with the CTD data which cannot be assessed due to the averaging of scans in acquisition. One thing that distinguishes this cruise and some of the casts of 2011-11 is that there is a high DO gradient at 100db to 125db. The complexity of these profiles are challenging for the SBE sensor.

The values have not settled down when there are only 11 bottles left in the fit.

| Bottles Used | Cruise | m | b | Soc | Voffset | R2 | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 21 | 2011-60 | 0.4493 | -0.2214 | 0.4493 | -0.4928 | 0.9991 | few bottles |
| 93 | 2011-63 | 0.4474 | -0.2213 | 0.4474 | -0.4946 | 0.9991 | problems with bottles |
| 241 | 2011-10 | 0.4417 | -0.1979 | 0.4417 | -0.4480 | 0.9989 |  |
| 11 | 2011-76 | 0.4905 | -0.2387 | 0.4905 | -0.4866 | 0.9992 |  |
| 175 | 2011-11 | 0.4797 | -0.2324 | 0.4797 | -0.4845 | 0.9998 |  |

Using the results of cruise 2011-11 looks like the most appropriate choice since conditions were similar for the 2 cruises and there are so many bottles from 2011-11.

The configuration files were updated with the new parameters Soc and Voffset and saved with name 2011-76-ctd-new.xmlcon.

##### Hysteresis Study

This sensor has been recalibrated since the last hysteresis checks were done. There was no deep sampling during this cruise, so hysteresis is not going to be a problem and there is not enough deep data to enable reliable tests to be done to fine-tune H1, H3 and E.

##### BOTTLE FILE PREPARATION

Niskin #15 failed to close for casts #16 and 18, but sample #s were assigned as if it had closed. The samples that appear to be from Niskin #15 will be moved to Niskin #16 at the end of processing.

The ROS files were recreated with the new configuration parameters.

The files were put through CLEAN to create BOT files.

A track plot and header check turned up no problems.

There were no negative or off-scale fluorescence values.

Temperature and salinity were plotted for all BOT files. No outliers were noted which is not as meaningful as usual given the averaging on acquisition.

The addsamp.csv file prepared in the DO calibration step is already ordered on sample #s. It was converted to CST files to be used as a framework for the bottle files. It was also used to create SAM files with bottle #s and bottle positions. The SAM files were then bin-averaged.

Next, each of the analysis spreadsheets was examined to see what comments the analyst wanted included in the header file. These were used to create file 2011-76-bot-hdr.txt.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2011-76nuts*.xls which included a report on precision.  The file was simplified, reordered on sample numbers and the file was saved as 2011-76-nuts.csv. The file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet 2011-76oxy*.xls which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified; the file was then saved as 2011-76oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was done at IOS and saved in file 2011-76sal*.xls. The file was simplified, and saved as 2011-76sal.csv. There were no duplicate samples. The file was then converted to individual SAL files. The samples from FR1 are from event #4, not 5 – there is an error on the rosette sheet. Similarly, the samples from FR2 are from event #5. From event #6 onwards, all is well. The nutrient and DO samples were correctly identified.

The SAL, OXY and NUT files were merged with CST files in 3 steps. After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only, and SORT to get the files in bottle number order.

There was no sampling for cast #15, so the bottle file was deleted.

##### Compare

Salinity

The salinity samples were analyzed on January 9, 2012, 42 days after the last sample was collected.

Compare was run with pressure as reference channel. All outliers were from the top 3 or 4db and will not be flagged. No samples were flagged by the analyst.

When bottles from above 50db were excluded, the primary salinity was lower than the bottles by an average of 0.0011 and the secondary was high by an average of 0.0005.

For details see file 2011-76-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. The data come from only 2 casts. When plotted together there appears to be a lot of scatter, but the two casts are very different. For cast #16 the CTD is lower than the bottles except for 2 surface points that look like outliers; that cast has a high surface DO gradient and moderate gradient around 100m. For cast #18 the CTD data look low by about 0.02 to 0.04mL/L below 100m, is way too high at 100db and very close above 100m; that cast has a strong gradient at 100db and low gradient elsewhere. Both casts have a lot of structure that would certainly complicate the response and the errors are not large. For comparison, during 2011-11 with similar conditions, the differences when outliers were excluded were roughly ±0.05mL/L from ~20db to 75db and ±0.03mL/L below that.

A test was done to see if using a Soc/Voffset value found using the 2011-76 would have worked better. The values 0.4905/-.4866 was entered into a test configuration file, and new SAM files were derived. A test run of COMPARE shows worse results with the CTD high by ~0.03 to 0.06mL/L below 100m and high by 0.1mL/L near the surface. So using the 2011-11 DO calibration results looks like the better choice.

Only 1 bottle was flagged by the analyst and it is impossible to make a judgment about it based on COMPARE.

For details see file 2011-76-dox-comp1*.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and look reasonable; they show good agreement between bottles and CTD with the exception of a few surface bottles from cast #16.

##### Conversion of Full Files from Raw Data

All files were converted using 2011-76-ctd-new.con.

One cast was examined and all expected channels are present. The temperature and conductivity channels track reasonably well and upcast and downcasts are reasonably similar. The fluorescence, transmissivity, DO Voltage and pH look normal.

The altimetry has spikes at the bottom, but a reasonable estimate can be made.

##### WILDEDIT

Program WILDEDIT is usually run to remove spikes from the pressure, conductivity & temperature channels. Because the data are already averaged, this step will be skipped.

##### ALIGN DO

Tests to determine the optimal alignment of the DO voltage do not offer clear results. It is very hard to judge in these data because of the lack of detail, but a setting of +4.5s looks reasonable and was chosen for other 2011 cruises with the scans averaged on acquisition.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

On 4 other cruises using these sensors and for which the 24 scans were averaged on acquisition, tests to determine the best settings for CELLTM were inconclusive. The goal is to make upcasts look closer to downcasts on a T-S surface. There is little detail in the data to aid the judgment, but all choices looked better than no application, with the choice of (α = 0.0245, β=9.5) seeming best overall for both channels on those cruises and looks reasonable for this cruise.

CELLTM was run on all casts using setting (α = 0.0245, β=9.5) for both channels.

##### DERIVE

Program DERIVE was run twice:

- to calculate primary and secondary salinity and dissolved oxygen concentration. The Tau correction feature caused some bad data that prevented conversion to IOS headers. When the TAU correction was turned off, the conversion worked fine. This also happened during cruise 2011-11. Towards the end of processing 2011-11, it was found that there was some effect on higher gradient zones. For this cruise, since there were only 3 profiles to process and they both had high gradients and variability, the files were edited to remove some obvious bad data at the end of the upcasts. These bad data had normal pad values, so why they could not be converted is unknown, but removing them did allow conversion to IOS SHELL format.

- on cast #18 to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Cast #18 was plotted to check for agreement between the pairs of T and C sensors. The results were compared with 2 casts from the 2 cruises that bracketed 2011-76 and had the same T/C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-10-0047 | 290 | +0.0010 | +0.0001 | +0.0006 | High, Steady |
| 2011-10-0053 | 380 | +0.0009 | +0.00012 | +0.0003 | High, Steady |
| 2011-76-0018 | 250 | +0.0008 | +0.00015 | +0.001 | Moderate, F.Steady |
| 2011-11-0048 | 360 | +0.0007 | +0.00009 | +0.0003 | High, Steady |
| 2011-11-0053 | 360 | +0.00075 | +0.00008 | +0.0002 | High, Steady |

The differences are all slightly higher than for the other 2 cruises, though still small. The differences in salinity are in agreement with the results of COMPARE when the 2 channels differed by an average of 0.0002.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Only casts #15, #16 and 18 were converted since only those 3 went below the surface. Only bottle files are needed for casts 4-13.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Header Check was run and the minimum fluorescence value was -0.03mg/m^3. During the cruises before and after this one the minimum was -0.061mg/m^3. CALIBRATE will be used to add 0.061 so that there are no negative values. The bottle files had no negative values but the correction will be applied to those as well, for consistency. There were no off-scale fluorescence values.

Surface check was run and shows an average surface pressure for the cruise was 1.0db which is very low, but this was a deliberate choice for this cruise to sample surface conditions.

The cross-reference check was compared with the log book. One cast had the wrong station name, so that was corrected in the bottle file.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet. Plotting the data for the 3 files shows that the algorithm worked well.

For the bottle casts there are entries for only 6 casts. For the 3 that sampled to the bottom, the entries look appropriate. For the other 3 with a header entry, the one for cast #8 is not reasonable, so it was removed. The entries for casts #4 and 5 are not consistent with the log entry for bottom depth, though those that might be wrong. There are some spikes in the signal, so the header entries were removed.

The Water Depth headers all agree with the log book, though for the first 2 casts there was no header entry. There are entries in the log, but the altimetry suggests they may be inaccurate.

##### Shift

For 2011-76 there is little information on which to judge alignment. The shifts used for 2011-60, 2011-63, 2011-10 and 2011-11 were applied to these data.

Fluorescence

Fluorometers usually require alignment, either to remove the effects of pumping for SeaPoint sensors or to correct for slow response time in ECO sensors. Usually an advance of +48 records (2s) is found appropriate for the ECO sensor. Given the averaging of 24 records, 2s is equivalent to +2 records. Tests using that setting show the resulting offset in fluorescence sometimes seemed too high, sometimes too low, but overall, the results look reasonable. Low fluorescence values make the judgment even more difficult. SHIFT was applied with a +2 records setting.

Conductivity

The pre-averaging makes the alignment more difficult to assess. During 2011-60 tests were run using values from -0.1 records to +0.1 and none looked any better than the original. For 2011-63, 2011-10 and 2011-11 it was hard to find suitable features to aid this assessment. No shift was applied to these data.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked and no further alignment was found necessary. SHIFT was not run on DO.

pH

During 2011-63 a setting of +4 records was used for this data. As for fluorescence it is difficult to determine the best choice. SHIFT was applied with a +4 records setting.

##### 14. DELETE

The following DELETE parameters were used (adjustments were made to the usual settings to allow for the fact the data are already averaged):

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0

**Pressure filtered over 15 points**NOT APPLIED

Swells deleted. Warning message if pressure difference of 2.00

**Drop rates <   0.30m/s (calculated over 11 points) **NOT APPLIED.

**Drop rate applies in the range:  10db to 10db less than the maximum pressure

**NOT APPLIED

Sample interval:  taken from header

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  During 2011-60 in August there were only 5 bottles with a lot of scatter, but both CTD salinity channels appeared to be close to the bottles. There were no bottles from 2011-63 and there were many flagged bottles and outliers for 2011-10. From 2011-11 there was a better comparison and both CTD salinity channels were within 0.001 of the bottles. The secondary channels temperature and salinity were selected for archiving for all those cruises.

2. Dissolved Oxygen: The DO sensor was repaired and recalibrated in April 2011. During 2011-60 and 2011-63 there were too many problems with the comparison with bottles to justify a 2nd recalibration. For 2011-10 and 2011-11 the average offset was 0.003mL/L and 0.03mL/L; no 2nd recalibration was applied.

3. Pressure: The sensor was recalibrated in April 2011. Because all data were averaged on acquisition surface tests are not very useful, but no problem has been detected.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The SOGS data fell within the ranges, but the temperature and salinity were both low above 100m at SOGN. Similar results were seen during 2011-11 in the same area for salinity, though not temperature. It looks like there was strong vertical mixing. This looks real, not evidence of CTD calibration problems.

Repeat Casts –

There was one repeat cast. When plotted together the two SOGS casts show that there was active mixing. The changes along constant density lines show large changes in temperature but not in salinity.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

Both T/C sensor pairs produce salinity that is very close to the bottles. A decision was made to select the secondary T and S for archiving because in previous cruises they had fewer spikes than the primary channels.

CTDEDIT was used to edit the secondary salinity very lightly, remove surface and bottom records and a few records obviously corrupted by shed wakes. All casts needed some editing.

##### Initial Recalibration

File 2011-76-recal1.ccf was prepared to add 0.061mg/m^3 to the fluorescence to avoid negative values in the files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response-time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are compared to bottle data from the same pressure. For this cruise, the problem of data smoothed during acquisition means the comparison is unlikely to be sensitive enough to justify further recalibration, but is still worth doing to get some idea of the accuracy of the SBE DO data.

The downcast files are usually averaged in 0.5db bins, but in this case the full files were thinned to the levels at which bottles are usually fired. COMPARE was then run to compare the CTD and bottle data. When a few outliers are excluded the CTD DO is lower than the bottles by an average of 0.025mL/L, but the surface data are mostly close to the bottles. The fits are noisy, likely because of the complex DO gradients. A pressure-dependent correction might be appropriate, but the correction is overly dependent on the choice of outliers to be excluded, and the 2 casts give significantly different results. Given active mixing conditions and the fact that the comparison uses pressure rather than density and just 2 casts, no recalibration is justified.

##### Special Fluorometer Processing

No filtering is required for the ECO fluorescence.

##### BIN AVERAGE of CTD files

While the data were average over 24 scans, the descent rate is often less than 1m/s so Bin Average was still run on the COR1 files using the following settings:

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, and to add the following comments:

Data Processing Notes:

----------------------

The SBE data are considered of lower quality than normal because 24 scans

were averaged in acquisition. These averages may contain spikes that

would normally be removed in processing and might be systematic in nature.

Fine-tuning of alignment settings and CELLTM correction are also less

sensitive with smoothed data. Unfortunately the same error was made on

several other cruises using the same equipment, so we can not rely on

settings determined from those data sets. Given the area is in protected

waters and the CTD is believed to have worked well with few spikes, it is

likely the errors caused by this problem are small.

Fluorescence and Transmissivity data are nominal and unedited.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The tests for accuracy of Oxygen:Dissolved:SBE data are less reliable than

usual, but the data are considered, very roughly, to be:

±0.2mL/L from 0 to 20db

±0.06mL/L from 20 to 150db

low by up to 0.05mL/L below 150db

For details on the processing see processing report: 2011-76-proc.doc.

A second set of files was produced with pH:SBE channel included. These will not be archived.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

Profile and T-S plots were examined and no problems found.

The track plot looks ok.

Special files were prepared that included the pH channel and were given extension ctdph.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were between 86% and 90%. Those values look reasonable for the season.

##### 24. Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, SBE:pH, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

For casts #16 and 18, the samples from Niskin #15 were moved to Niskin #16, and the line for Niskin #15 was removed.

HEADER EDIT was run to change the chief scientist’s name to standard format, fix variable formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data. Those files were named *.CHE.

A note was added to the headers about Niskin #15 not firing.

A header check was run on the final files and no problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Data from the CHE files were exported to spreadsheet 2011-76-bottles-final.xls; no errors were found.

Standards check was run and the only warning was for the format for Temperature:Draw that was deliberately given a non-standard format.

Special files were prepared that included the pH channel; they were named with extension “cheph”.

##### 26. Producing final files

Comments were added to the CHE and CHEPH files for casts #16 and 18 explaining the reassignment of sample numbers for Niskin #15.

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (Notes from log book and sampling notes)

1. Hex file shows different sensors and data is before beginning of this cruise; source unknown, but likely a test.

4. Rosette sheet identifies this event (FR1) as #5. CTD log shows it is #4.

5. Rosette sheet identifies this event (FR2) as #6. CTD log shows it is #5.

6. Both CTD and Rosette log agree that FR4 was event #6.

4-13. Shallow casts – no CTD file to be prepared, just a CHE file.

15. No sampling because 2 bottles failed to fire – rerun, but CTD file was prepared.

16 & 18. Bottle #15 failed to fire.

CRUISE SUMMARY - CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 4Aug2011 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 23Apr2011 | Factory |  |  |
| Eco-AFL Fluorometer | 2215 | 4Aug2011 | IOS |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |
| pH | 692 | 29Dec0210 | Factory |  |  |
| Altimeter |  |  |  |  |  |