## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 6 January 2011 – 19 December 2011

Platform: Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 25 January 2012 – 31 January 2012

Number of original CTD casts: 64

Number of casts processed: 63 (1 file was corrupted)

INSTRUMENT SUMMARY

Several Sea Bird Model SBE 19 SEACAT CTDs were used. All were mounted with an altimeter.

SUMMARY OF QUALITY AND CONCERNS

Several CTDs were used and the calibrations are assumed to have been entered correctly as we do not have the means to check them.

There were 3 casts with sections of bad data in either all channels or just in salinity; in all cases the CTD involved was serial number 1513. This instrument performed properly for many casts, but it may need servicing.

Editing was applied to all casts. This involved removing a few near-surface records from many casts, bottom data from a few and editing of salinity for all. The salinity tends to “overreact” when temperature gradients are high, so the data were smoothed to produce stability in such situations. For 3 casts larger segments of data were removed. In all cases note was made in the header of what editing was done.

The soak period for the CTD was typically 1 -1.5 minutes long, but occasionally less than 1 minute. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. For most of these data the gradients are lower and the descent rates fairly steady; in those conditions errors in salinity are expected to be on the order of 0.005.

The position information was corrupted for file n11031701 so the following nominal position was entered:             LATITUDE               :  49  20.00000 N  ! (deg min)

LONGITUDE           : 124   5.00000 W  ! (deg min)

One file from 2011 was included in the data provided for 2010. It will be archived with the other 2011 data.

The times given in the SeaBird headers do not include a time zone, so it was assumed that they are in PST. All times were converted to UTC by adding 8 hours.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

63 hex files and corresponding configuration files were delivered in a zip file. One hex file had been delivered a year before with the 2010 files. A number of different CTDs were used.

- Conversion of Raw Data

- The raw data were converted by matching the configuration file to the input file name. Altimetry was converted. File n11072801.hex contained no data and was not converted. There was a CNV file in the zipped files, but it contains depth instead of pressure and no altimetry and the configuration file is also corrupted. DERIVE won’t work without pressure in the file. This file will not be processed.

There were 2 cases where there was an ACT file but no HEX file.

File n11032203.HEX does not have a corresponding ACT file; the start time in the file is just 13 minutes after n11032202, so the same position will be entered later.

- The soak period was less than a minute for some casts. For most casts there was a soak of between 1 and 2 minutes. A 2-minute soak is generally advised for the best salinity data, but in most cases the conductivity seems to have reached equilibrium before the downcasts started.

- No files were found with data clearly taken in air, so we cannot judge the pressure at which the CTD entered the water. There was one file which started with a slightly negative pressure, -0.3db, but the associated conductivity appears to be in water. Given that there are a variety of sensors, the pressure sensor might not have been fully equilibrated and these sensors are only considered ±1db, no recalibration will be applied to pressure.

- Temperature and conductivity generally look ok.

- The altimetry looks clean near the bottom of casts where the CTD got close to the bottom.

- WILDEDIT

No single-point spikes were noted, so this step was skipped.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 1 seconds to increase the pressure resolution.

- At this point it was discovered that file n11072801.hex had been copied over with some October data. There was a CNV file that must have been created before the corruption occurred. However, it has depth, not pressure and no altimetry.

- DERIVE

Program DERIVE was run to calculate salinity. The CON files had to be copied into folder FILTER to make this work.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

A spreadsheet of positions calculated from the *.act files was received from Peter Chandler together with a map in file NanooseCtdStations2011.xls. A column was added for Station Name and Nanoose was entered in that column for every file.

The spreadsheet did not contain the initial “n” found in the file names and was missing the final digit which was either a 1 or 2 since on some days there were 2 casts. Those items were added to the file names in the spreadsheet.

The format of latitude and longitude needed to be changed to DD MM.MM N and DDD MM.MM W. For latitude, this was done by concatenating INT(LAT), TRUNC((LAT - INT(LAT))*60),2) and N with spaces between each of those, entered using “ “. Longitude was adjusted in the same way. Do not add “! (deg min)” – that is done by the program. So for example:

=CONCATENATE(INT(C2),"  ",TRUNC(((C2)-INT(C2))*60,2)," N ")

where C2 is latitude.

The program MERGE CSV FILES TO HEADERS was then used to add that information to the headers.

8. Checking Headers

Header Check was run and no problems were found, though the nature of this data means we can only check that the headers are reasonable. Speed checks are impossible without event numbers and would not be useful even if there were. The CTD serial numbers are not in the headers. These could be added but given no information about calibration this does not seem to be worth the time it would take.

The track plot has one station that looks like it is on land - Station n11031701. The information in the ACT file appears to have been corrupted. The position for that file was changed to the following “nominal” position and an explanatory note was placed in the header:

LATITUDE               :  49  20.00000 N  ! (deg min)

LONGITUDE           : 124   5.00000 W  ! (deg min)

As noted earlier file n11032203 does not have a corresponding ACT file, but the start time in the file is just 13 minutes after n11032202, so the same position was entered in the later file.

Altimetry headers were exported to a spreadsheet. For most the CTD did not get within 15m of the bottom, but a sample of those with altimetry header entries were checked against plots and the algorithm worked well.

Surface check was run and the average surface pressure was 0.5db.There were a few negative values, but only 3 were lower than -0.2db. One cast had -0.7db with a corresponding conductivity of 0, so was likely out of the water. Two others had pressure ~-0.3db, so equivalent to zero given the precision of the instrument. One cast had high values at the surface and look like a spike that will be removed by DELETE and/or editing.

9. SHIFT

When this type of instrument was used during 2008-03, -18 and -38, the temperature channel was shifted by +2.25 records and for 2009-17 a setting of +1.75 looked best. For 2009 Nanoose data +1.9 looked best overall.

Since different CTDs are in use it is likely that no one setting suits all. Tests were run on casts from 3 different CTDs using settings from +1.75 to +2.75. A setting of +2.25 looking best overall, though all settings improved the data.

SHIFT was run on all casts using +2.25 records for Temperature. Salinity was recalculated.

11. DELETE

Then DELETE was run on the SHF files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.25m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings.

12. CTDEDIT

Plots were then opened in CTDEDIT. All casts required some editing. Surface records were removed from many files and bottom records were removed from a few casts where there was clearly some corruption. Salinity needed editing in all files to smooth unstable features that appear to be due to minor misalignments of T and C in the presence of sharp temperature gradients. A few casts required more severe editing due to bad salinity or bad temperature and salinity (casts 11082302, 11091402 and 1120701). All of the bad data came from a single CTD, serial number 1513, though most casts using that instrument did not have problems.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen and no problems were found.

Page plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – The only excursions from the local climatology were some slightly low salinity values near the bottom in November and around 30db in October and low temperature values around 40-50db in early June. These do not suggest calibration problems as the 3-standard deviation climatology is a little severe for casts this close to shore.

Repeat casts – There were some casts within a few hours of each other but they were not usually at the same place.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The time zone for the times in the headers is unknown. The times suggest either PST or PDT, since they are mostly between 8 and 10am, so it will be assumed it is PST for all. We generally find that computer times are not updated to mark the change from PST to PDT and vice versa. Program ADD TIME CHANNEL was used to add 8 hours to get UTC times. These times may well have an error of 1 hour.

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced. No errors were found.

Particulars

n11031701 – corrupted ACT file, used nominal position

n11032203 – no ACT file, used position from previous file starting 13 minutes earlier

n11022301 – ACT file but no HEX file

n11042002 – ACT file but no HEX file

n11072801 – no HEX file, only converted file without pressure

n11082302 – section of corrupted salinity removed

n11091402 – section of corrupted T and S; records removed

n11120701 - large sections of corrupted T and S data; records removed

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	6 January 2011	    End:	19 December 2011

Location: Nanoose Test Range 	Vessel:    Stikine

NOTE: Several SEACATs were used for this cruise and not all were identified. The calibration information used in conversion can be found in the headers of the CTD files.