## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | Updated channel names & formats in TOB & loop files. GG & SH |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 27 August 2020 | ### Added HPLC data. S.H. |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files and Loop. G.G. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2012-01

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: February 6, 2012 – February 21, 2012

Processed by: Germaine Gatien

Date of Processing: 3 April 2012 – 22 June 2012

Number of original HEX files: 33

Number of CTD files:  33

Number of bottle casts:	32

Number of original TSG files:   3

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0097), a Wet Labs Eco-AFL/FL Fluorometer (#2215), a Biospherical QSP-200L4S PAR sensor (#4615) and an altimeter (#1204).

A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

The data logging computer was #2 and Seasave version 7.21b was used.

The deck unit was a Seabird model 11, serial number 0471 #2.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A IOS rosette with 24 10L bottles was used.

The Oxygen Kit was model 865+665, Kit Number 1.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, plus details about the rosette, TSG and oxygen sampling kit. The Chief Scientist provided sampling notes with a good description of problems relevant to processing.

The only significant problem not mentioned in the log was that the configuration file changed partway through the cruise and later reverted to the original one. This occurred because of a computer crash and duplicate folders that each appeared to contain the correct configuration file; the one selected had the wrong parameters for 2 sensors and was missing some header information, though no problem has been found to be caused by those missing lines.

There were underway loops logs, oxygen analysis logs and rosette logs which were all in good order. Notes from the Chief Scientist were very helpful in dealing with problems encountered at sea.

Over the past 2 years severe equipment problems and delays in salinity analyses led to a large backlog in processing. We are now largely “caught up” so that processing can start closer to the time of a cruise. A problem with configuration files was quickly solved because staff remembered what happened.

Salinity samples were analyzed within 10 days of the end of the cruise, and shows primary salinity calibration errors to be <0.001. The secondary salinity read lower and showed some pressure-dependence which can be indicative of conductivity cell damage. However, the temperature also shows more pressure-dependence than usual, so the explanation may be complex.

Cast # 44 had bad downcast data below 116db, affecting sensors on the primary pump including primary temperature, primary conductivity and dissolved oxygen data. While the secondary channels were not affected, the questions about secondary T and C calibration make choosing the secondary channels unwise. Records below 116db were removed from the CTD file. The problem cleared up at the bottom so the CHE file is fine.

The WetLabs ECO fluorometer was mounted on the CTD for this cruise. There are doubts about these data as the sensors have not been adequately inter-calibrated with the SeaPoint fluorometers or extracted chlorophyll, and there have been changes to sampling rates and ranges from one cruise to another. It would be wise to use both types of fluorometer or only the SeaPoint until more testing is done, and such testing must include sampling a wide range of chlorophyll values. Initial comparisons with extracted chlorophyll indicated that the fluorometer was reading too low. After studying the results of cruise 2012-14 it was concluded that a range change thought to have occurred just before that cruise must have actually occurred before 2012-01. So the ECO fluorometer data were recalibrated by multiplying by 25/6. That change produced fluorescence values that are much too high for very low values of extracted chlorophyll. As extracted CHL values rise, the ratio of the fluorometer readings to extracted CHL samples decreases, reaching ~1 when extracted CHL values approach the cruise maximum of ~1ug/L.

An improvement in the ECO fluorometer data is that there is more detail in the traces due to a change in the sampling rate of the instrument.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.3mL/L from 0 to 300db

±0.1mL/L from 300db to 800db

±0.05mL/L below 800db

The TSG position data became stuck on one occasion; ship positions and times were available from another source, so the faulty data were replaced in the TSG file.

There were two unusual features about the TSG data. Firstly, there was a very sudden shift in the TSG data at 12:15UTC on Feb. 9th with temperature rising by about 0.7C° and salinity dropping by ~1.9 within 7 minutes and two thirds of that change occurred in a 1-minute window. Fluorescence rose during that time. The change occurs around 48.47°N 124.56°W. There is no indication that this shift was instrumental. Secondly, the ship made 3 passes through the area eastward of P2 and westward of Victoria within a day and a half, offering some possibilities for study of short-term change in the TSG data. In Juan de Fuca 2 of the passes are towards the northern side and 1 closer to the southern side. Closer to P2 the 3 tracks are fairly close.

There are patches with extremely noisy salinity data; these are likely related to very rough seas.  Salinity is smoother during stops for CTD casts. The differences between the intake temperature and lab temperature are also extremely noisy. An estimate was made that the TSG salinity is low by 0.03 based on comparisons with loop samples and CTD salinity, excluding many outliers.

The TSG fluorescence data are raw with volts as units. The TSG fluorometer was cleaned prior to this cruise and a pre-cruise calibration was done to obtain a blank voltage. However, using that value leads to extensive negative values no matter what scale factor is used. In the absence of a reliable calibration, only the nominal results are available. The ratio of TSG raw fluorescence to extracted CHL rosette samples varies from ~0.6 for low CHL values (~0.3ug/L) to ~0.2 for higher CHL (CHL>1ug/L); the mean ratio is 0.38. The comparison with loop samples is similar with a mean ratio of 0.33.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as sampling notes summarizing problems and points of interest with reference to processing.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated shortly before this cruise.

The calibration constants were checked for all instruments. The only error found was in the serial number of the SBE DO sensor. That was corrected and the file was saved as 2012-01-ctd.xmlcon.

An initial conversion to check that the configuration did not change showed that it actually did change for casts #6 – 19. There is no note of this in the log but the science personnel remember that the computer rebooted and there was confusion about what configuration file to use since there were two folders with con files that looked correct. There were 4 differences in the con file:

- The PAR parameters are set so all values will be 0

- The Transmissometer model and serial number were missing and the parameters were wrong

- The following lines are missing from the headers, but not those for casts #6-20:

* SBE 11 Firmware Version 4.6

* GPIB address = 1

* underwater unit scan rate = 24 Hz

* Paroscientific Digiquartz pressure sensor with high resolution counter

* number of scans to advance conductivity = 1.750

* two 12 bit A/D voltages per word start with word # 5

* number of words from the deck unit = 12

- The Number of Bytes Per Scan is 41 whereas it is 37 for casts with the initial con file.

A new configuration file, 2012-01-ctd2, was prepared with the headers as used for the odd casts and corrections to the PAR and Transmissivity entries. Both PAR and transmissivity look ok when conversion was done using the new file.

There may be problems with the data for these casts, one being that conductivity may need to be advanced by 1.75 records to reduce salinity spiking. This will be checked alter. There may be other effects due to the missing lines, so this will have to be kept in mind whenever checks are done.

The PAR sensor was not always mounted. Based on notes from the chief scientist cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. (Since bottles were fired out of order, the file was 1st ordered on bottle position, sample #s added and it was then reordered on bottle number.) Sample numbers were added to the ADDSAMP file based on rosette log records.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2012-01-DO-cal.csv. The titrated DO values were added to that file and lines removed for which there was no DO sampling.   A calculations was made of  Ф  using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

When all values are included the R2 value was 0.9981. Removing values flagged “3” or “4” produced better results. Most of the hypoxic samples from Cast #1 had been flagged, but the one that remains does not look out of line. It is difficult and tedious to pick our outliers on the plots and find and eliminate them from the fit. A simpler approach was to use the M and B values from the factory calibration to determine the difference from the fit for each sample, as follows: 
  	Difference = M*Voltage – B – DO/Phi

Watching how the plots change, the process stops when obvious outliers have been removed as judged by visual inspection and the R2 value, being careful to stop before the DO range is significantly reduced. (The M and B from the first fit with all bottles was substituted for the factory values in the calculation of differences for subsequent fits.)

When the data were sorted on that difference, fits were done excluding flagged DO samples. Then the M and B were updated to the value from that comparison and differences were calculated again. The changes were fairly small as more data were removed, reflecting the fact that there were few major outliers. Since this sensor was last recalibrated the only tests were done on CTD data that was of lower quality than usual, so looking at previous results is not terribly useful. The results are shown in the following table:

| Summary of Soc Voffset including the original values in the factory calibration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Factory setting | 0.4348 | -0.2110 | 0.4348 | -0.4853 |  |
| 148 | All data | 0.4822 | -0.2390 | 0.4822 | -0.4956 | 0.9981 |
| 142 | all data except 3 and 4 flags | 0.4811 | -0.2365 | 0.4811 | -0.4916 | 0.9992 |
| 138 | excl. outliers diff>0.02 | 0.4810 | -0.2354 | 0.4810 | -0.4894 | 0.9998 |
| 130 | excl. outliers diff>0.01 | 0.4807 | -0.2340 | 0.4807 | -0.4868 | 0.9999 |
| 118 | excl. outliers diff>0.005 | 0.4800 | -0.2320 | 0.4800 | -0.4833 | 0.9999 |
| 105 | excl. outliers diff>0.001 | 0.4800 | -0.2313 | 0.4800 | -0.4819 | 1.0000 |

To test for hysteresis a plot was made of the differences from the fits versus titrated DO and some data stood out from the others; an examination of the data shows they are mostly from below 2500db indicating hysteresis. An experiment was run to see if varying E in the equation of Phi would reduce this. Reducing E improved the fits, with a value of E=0.03 producing results that were flat. This seems a large change from the factory setting of E=0.036, but the results do look encouraging. The equivalent table to that given above is:

| Summary of Soc Voffset including the original values in the factory calibration when E=0.030 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4348 | -0.2110 | 0.4348 | -0.4853 |  |
| 148 | all | 0.4809 | -0.2345 | 0.4809 | -0.4876 | 0.9982 |
| 141 | all data except 3 and 4 flags | 0.4799 | -0.2321 | 0.4799 | -0.4836 | 0.9992 |
| 138 | excl. outliers diff>0.02 | 0.4798 | -0.2310 | 0.4798 | -0.4815 | 0.9999 |
| 136 | excl. outliers diff>0.01 | 0.4801 | -0.2312 | 0.4801 | -0.4816 | 0.9999 |
| 128 | excl. outliers diff>0.005 | 0.4797 | -0.2299 | 0.4797 | -0.4793 | 0.9999 |
| 118 | excl. outliers diff>0.001 | 0.4801 | -0.2302 | 0.4801 | -0.4795 | 1.0000 |

The 2nd set of values will be used, picking the one with differences >0.01 excluded (plus flagged values). The study with E=0.036 was saved as 2012-01-DO-study-e36.xls and the 2nd study was saved as 2012-01-DO-study.xls.

The only major outliers are 2 surface bottles and a 100m bottle from cast #1. The latter is likely near the sudden rise in DO above the hypoxic layer. Of the bottles already flagged, sample #199 is a major outlier and samples #444 and 484 somewhat out of line. The others are not particularly out of line. No changes to bottle flags look necessary, but they can be checked later in COMPARE.

(For more details see files 2012-01-do-cal-study.xls and 2012-01-do-cal-study-e36.xls.)

The configuration files were updated with the new values for parameters Soc. Voffset and E and saved with names 2012-01-ctd-new.con.

##### Hysteresis Study

Hysteresis tests were run in the previous section working on just the value of E. Since tests on other cruises have shown that other relevant parameters (H1 and H3) do not seem to change much, this is considered sufficient. The value E=3.0 will be used and H1 and H3 will be left with nominal values.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. The only data that looked odd were a few small spikes in secondary salinity for casts #19 and 38.CTDEDIT was used to clean those files and the edited files were copied to BOT.

(NOTE: These steps were repeated later for casts 1, 2, 4, 5, 23, 28, 30 and 49 because the PAR sensor channel had not been converted.)

A preliminary header check turned up no problems and the maximum fluorescence value is ~1.7 so there is no off-scale fluorescence, however, the minimum values are slightly negative, as was found for other recent cruises.

The addsamp.csv file prepared in the DO calibration step was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-01-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2012-01chl.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2012-01chl.csv which was then converted to individual CHL files. Loop data were not included, but were added to file 2012-01-loops.csv.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-01oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-01oxy.csv. That file was converted into individual *.OXY files. Loop data were not included, but were added to file 2012-01-loops.csv.

SALINITY

Salinity analysis was provided in spreadsheet QF2012-01SAL.xls. The file was simplified and saved as 2012-01sal.csv. That file was converted to individual SAL files. Loop data were not included, but were added to file 2012-01-loops.csv. The salinity data were analyzed within 10 days of the end of the cruise.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-01nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and saved as 2012-01-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There was only 1 severe outlier and that is from a surface sample and is associated with very high standard deviation in the CTD salinity channels, so no quality flag is warranted.

Samples that had already been previously flagged were examined. In consultation with the analyst, the following actions were taken:

- Sample #86 – cast #10 – the average looks ok in comparison to the CTD salinity though both readings seem slightly out of line. No change to flags.

- Sample #200 – cast #25 – looks ok in comparison with CTD salinity and duplicates agree; flag was dropped by analyst.

- Sample #231 – cast #27 – This may be an outlier in the comparison with CTD salinity, but the CTD data are fairly noisy, so this is not clear. No change to flag.

- Sample #528 – cast #65 - This looks ok in comparison to CTD salinity, but flag left.

Comments about the COMPARE observations were added to header comments where appropriate.

When only 1 outlier is excluded the CTD primary salinity is found to be low by an average of 0.0011 with a standard deviation of 0.0025 and the secondary is low by an average of 0.0024 and standard deviation of 0.0019. When data above 400db are excluded as well the average differences are -0.0004 and -0.0028 with standard deviations of 0.0006 and 0.0010. The primary looks flat with pressure but the secondary shows much more pressure dependence than usual with values low by about 0.001 at 500db but by 0.005 at 4000db. The CTD technician was informed that sensor #2128 likely needs service.

The 22 bottles fired at 2000m during cast #36 were studied next. The standard deviations were 0.0004 for both salinity channels; the primary was low by an average of 0.0004 and the secondary by 0.0030. This is excellent repeatability and shows signs of both good sampling protocol and analysis. The salinity data were analyzed very quickly.

For full details for the COMPARE run see file 2012-01-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected the fit of differences against DO concentration is quite flat. All differences are within ±0.2mL/L when 3 major outliers are excluded and most are within ±0.04mL/L. Excluding the 3 major outliers the average difference shows the sensor to be high by 0.001 with a standard deviation of 0.033 and excluding all outliers >0.04 the average indicates the sensor is low by an average of 0.002 with a standard deviation of 0.015. These are good results.

The following samples were investigated because they were outliers or had been f lagged by the analyst:

- Sample 6, cast #1 – Just above hypoxic layer, likely slow response in DO sensor, not bottle problem. No flag added.

- Sample 11, cast #1 – Near-surface in high gradient, likely poor vertical match or slow DO sensor response accounts for difference. No flag added.

- Sample 103, cast #10 – flagged “3” but looks ok in comparison of titrated samples and CTD dissolved oxygen.with CTD DO. Flag not changed.

- Sample 199, cast #25 –severe outlier in comparison - flagged “5” and pad value entered

- Sample 200, cast #25 – originally flagged “3” but looks ok in comparison; flag changed to “2”

- Sample 215, cast #25 – originally flagged “3” but looks ok in comparison; flag changed to “2”.

- Sample 444, cast #50 – flagged “4”, minor outlier in comparison. Flag not changed.

- Sample 458, cast #50 – originally flagged “4”, seems ok in COMPARE and in profile. Flag changed to “2”.

- Sample 484, cast #60 – flagged “3”, minor outlier in comparison

The changes were made in consultation with the analyst.

As a final test for hysteresis a plot was made of differences against DO concentration and points from below 2000db were excluded so the deep bottles show up in red. The plot shows no obvious difference above and below 2000db.

For more details see 2012-01-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

Fluorescence

COMPARE was run using the Wet Labs ECO CTD Fluorescence and the Extracted Chlorophyll from bottles. The ECO fluorescence is too high for CHL<0.1 ug/, but otherwise reading too low. It looks like the ECO/CHL ratio gradually falls as CHL goes up so that ECO readings are ~0.3ug/L when CHL is ~1ug/L. CHL never got above 1.1ug/L. We need to learn more about this sensor.

NOTE: At the end of processing it was realized that the range was probably wrong in the configuration file. This error was noted in the paper work for 2012-14, but not for 2012-01. The assumption that the range change was made before the later cruise appears to be wrong. It was probably just discovered during 2012-14. The scale factor should have been 25, not 6. This would bring the fluorescence values into better agreement with chlorophyll for CHL~1, though the values are too high when CHL is low. This is generally noted for both types of fluorometers. This will be corrected later using a second run of CALIBRATE.

The unwanted lines were removed from the MRG files for cast #1 and 60. CLEAN was rerun.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. Errors were found in cast #4 – 1 missing salinity value – and cast #17 nutrients because the event # was wrong in the spreadsheet. The latter was fixed and the nutrient spreadsheet reconverted. The MERGE steps were repeated. The chief scientist was alerted as to the missing salinity sample. There is a stored sample that may be run; that can be added later.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-01-ctd-new.con and 2012-01-ctd2-new.con, as appropriate.

A few casts were examined and all expected channels are present. There is more detail in the fluorescence channel than seen in 2011 cruises. This is believed to be due to a change made to the sampling rate for the instrument.

The descent rate is extremely noisy for many casts with many complete reversals of direction during the descent and obvious shed wake corruption.

The two temperature channels are fairly close during the downcasts though both channels have occasional spikes. During the upcasts traces differ more and there are odd excursions that are often seen in upcasts. This is likely something to do with how the CTD is mounted. The conductivity channels are similar to temperature.

Altimetry looks useful when the CTD got near the bottom, dissolved oxygen, fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 4 casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is so noisy on the upcast, but for 3 casts a 3.5s advance appears to align downcast features best. The other cast was best with a 2s advance, but it was a shallow cast with an odd temperature profile.

ALIGNCTD was used to advance the DO Voltage by 3.5s relative to the pressure.

##### ALIGN Conductivity

Some lines were missing from the headers of the 8 casts that were acquired using the wrong configuration file. Those included information about the alignment of the conductivity sensors, so tests were done on casts #6-19 to see if the alignment had been applied in acquisition or not. ALIGNCTD was used to advance by +0.073s the primary conductivity relative to pressure. The results look much worse after the ALIGNCTD step than when salinity is derived without that step. Fine-tuning of conductivity alignment will be done later, so there will be a second chance to test this. Conductivity will  not be aligned at this stage.

##### CELLTM

The upcast data are extremely noisy so the usual tests for CELLTM settings were a little hard to interpret; however, all settings improved the data and the best  choice was (α = 0.0245, β=9.5) for the primary and (α = 0.03, β=9) for the secondary. Other cruises that have used this equipment since they were last used had a problem in that data had been averaged on acquisition; the parameters chosen for both channels for those cruises was also (α = 0.0245, β=9.5).

CELLTM was run using (α = 0.0245, β=9.5) for the primary and (α = 0.03, β=9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### STRIP

Because some files were too large to convert to IOS SHELL format, program STRIP was used to remove the Pump Status and Descent-Rate channels. Even that was not enough to allow conversion of cast #50, so the secondary temperature, secondary conductivity and oxygen voltage channels were also removed for just that file.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-01-0017 | 800 1400 | +0.0007 +0.0010 | -0.00007 -0.00010 | -0.0015 -0.0024 | F.High, X Noisy |
| 2012-01-0034 | 800 1400 2500 3500 | +0.0006 +0.0007 +0.0015 +0.0016 | -0.00007 -0.00010 -0.00014 -0.00014 | -0.0014 -0.0021 -0.0032 -0.0039 | High, X Noisy |
| 2012-01-0050 | 800 1400 2500 3500 4200 | +0.0005 +0.0007 +0.0014 +0.0020 +0.0021 | -0.00010 -0.00011 -0.00014 -0.00016 -0.00015 | -0.0016 -0.0022 -0.0034 -0.0041 -0.0044 | High, XX Noisy |

The differences in conductivity are higher than usual. Temperature, conductivity and salinity differences show some pressure-dependence. Some of this may be due to slight misalignment in high gradients leading to a higher or lower difference than seen in deep water where the differences are not sensitive to small alignment differences. The pressure-dependence does seem less obvious at great depth. However, the salinity comparison discussed in section 11 does indicate serious pressure-dependence in salinity. This could be the result of conductivity cell damage, but it is possible there is some problem with the temperature sensor as well. There is no obvious time dependence.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some negative values in fluorescence, as usual. The maximum fluorescence value was 29.9ug/L which is at or very close to the maximum (~30 for range 6); however, this value occurs in a surface spike, so there are unlikely to be off-scale values in the final data. There are no negative pressure values. Speeds look reasonable.

Surface check was run and shows an average surface pressure for the cruise was 4.1db which looks a little deep for the Tully, but the conditions were very rough so the CTD was probably not brought as close to the surface as usual. Cast #38 has a section of data at the end of the upcast where the CTD rose to about 0.8db a few times, then sank again; the pumps were running. When the pressure is lowest, first one, then the other salinity channels goes to very low values, suggesting that the sensor was bobbing around very close to the surface, perhaps dipping out of the water for a very brief time. Most salinity values at 0.8db are clearly “in water”, so this does not suggest an error in pressure.

The cross-reference check was compared with the log book and the only problems found were:

- The format for the station name in 4 casts; P 1, P 4 and P 6 were changed to P1, P4 and P6, respectively.

- The date for P14 in the log book is Feb. 18th but in the file it is Feb 19th. The file date makes more sense and is consistent with the TSG data, so no change was made.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet. Most casts did not get within 15m of the bottom so there are no header entries. There were only 2 problems in the altimetry headers:

- Cast #46 did not get close to bottom; spikes in altimetry were misinterpreted. The altimetry statement was removed from the header of the CLN file.

- Cast #17 got near the bottom, but no bottles were fired close to bottom, so the altimetry statement was removed from the SAMAVG file.

Water depths were checked against log entries and some discrepancies were found and corrected:

- Cast #14 had 1481m in the header but that is too shallow for P6 and the log gives 2541m, so that figure was put in the headers of the CLN and SAMAVG files.

- Casts #19, 25 and 27 at P12 had a variety of entries in the headers; they matched original log entries that had been crossed out and substituted with 3280m. The CLN and SAMAVG files were changed to 3280m.

- Cast #40 was changed to match the log entry in CLN and SAMAVG files.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence trace look like that of the temperature trace. This task was complicated by:

- Noisy upcast temperature data so the vertical offset is confused.

- Many stops for bottles on upcasts so that fluorescence gets opportunities to “catch up”

- Extremely noisy descent and ascent rates resulting in very confused traces.

- Very low fluorescence values so there is little variability to aid estimate of vertical offset.

Comparisons do indicate that applying either +24 or +48 records improves the data. The optimal choice is likely +48 but in places that looks like it might be too high. The value used on other recent uses of the ECO sensor was +48 records, but there have been some changes to the range and sampling rates of the instrument. However, in the absence of a clear  result from the tests, +48 records will be applied. The fluorescence is so low that the difference is unlikely to be significant.

SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records.

Conductivity

Tests were run on the primary conductivity channel only since there are problems with the secondary salinity calibration. The results were plotted for 3 casts using a variety of shifts and the best results were with an advance of -0.3 records.

SHIFT was run on all casts using that setting.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in late March 2011 and were used for 2011-60, 2011-63, 2011-10, 2011-11 and 2011-76. In all cases the data were averaged on acquisition. The secondary sensors were selected for archiving in each case due to more noise in the primary. The comparisons for 1 cruise were based on only 5 bottles and another had bottles that had been stored for 4 months before analysis, so the comparison was noisy and not trusted. The best 2 comparisons showed the primary to be low by 0.0011 and 0.0005 and the secondary to be low by 0.0005 and 0.0003.

2. Dissolved Oxygen

The DO sensor was recalibrated in Dec.2010. It was used for 2011-60, 2011-63, 2011-10, 2011-11 and 2011-76. As mentioned above the CTD data were averaged over 24 scans on acquisition. Most bottle comparisons were noisy and/or had few bottles. The best results were from 2011-10.

3. Pressure

The sensor was recalibrated in April 2011 and was used for was used for 2011-60, 2011-63, 2011-10, 2011-11 and 2011-76. No further offset was applied to any of those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions were in salinity at the bottom of the halocline for stations P18 and P20 when salinity was a little high. At P21 it was close to the top of the range. These excursions look real, not indicative of instrument malfunction.

Repeat Casts –

There were repeat casts at P4, P12, P20 and P26. When plotted together in T-S space two P26 casts taken 11 hours apart had temperature differences of <0.01C° and salinity differences of <0.001 at about 1200db along lines of constant density. For P20 there is only 1 cast below 1500db and above that level there is active mixing, so the changes are large. For P12 at 1700db and 6 hours apart the differences were <0.005C° and <0.0005 in salinity. For P4 at ~1000db and 4 hours apart the differences were ~0.004C° and ~0.0005. This is good repeatability, especially given the extraordinarily rough conditions.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The bottle comparison shows that the secondary salinity is pressure-dependent, so the primary channels were chosen for archiving.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. Most of the casts had extremely noisy descent rates so that despite a high average rate, there are many reversals corrupting data over many metres. Heavy editing was required for most casts.

For cast #44, the primary temperature, salinity and dissolved oxygen were bad below 386db and 503db, so those records were removed. While the secondary could be used, the dissolved oxygen will have to be removed and there are some doubts about the secondary salinity data. The bottle file is fine, since the values returned to normal values at the bottom of the cast. It is assumed there was a blockage that cleared – the pumps were on throughout. The bad section started with a spike.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

The dissolved oxygen look fine in COMPARE and the primary salinity was within 0.001 of bottles, so does not required recalibration.

The fluorescence channel will be adjusted by adding 0.022ug/L to avoid negative numbers.

CALIBRATE was run using file 2012-01-recal1.ccf to apply the offset to the Fluorescence channel.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration there is a lot of scatter with values close to zero for low DO, a little low for mid-range DO and a little high for high DO. The differences versus pressure are noisy near the surface with outliers in both directions and a slight bias towards the SBE DO being high. Below 2000db the SBE DO looks very slightly low, by ~0.03mL/l. The errors are small and given the nature of the comparison are as close as we can expect. The plot against file pair number shows no hint of time-dependence; the fist cast looks out of line, but this is normal for Saanich Inlet where the change from 0 at 125db to 6mL/L at the surface is always challenging for the DO sensor. No further recalibration will be applied..

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

CLEAN was run to remove the SeaBird Headers and comments from the secondary file.

SORT was run to rearrange data by increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

CALIBRATE was run using file 2012-01-recal2.ccf to multiply the fluorescence by 4.1667.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence, Transmissivity and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.3mL/L from 0 to 300db

±0.1mL/L from 300db to 800db

±0.05mL/L below 800db

For details on the processing see processing report: 2012-01-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 96% to 102% with most very close to 100%.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was on casts with no PAR sensor to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

The MRGCOR1 files were put through CALIBRATE using file 2012-01-recal2.ccf to multiply the fluorescence by 4.1667.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run on the final files and no problems were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up errors in station name format; those were fixed.

The track plot was produced on screen and no further errors were found.

##### 25. Thermosalinograph Data

Data were provided in 3 hex files. The 1st two names were missing a zero so were changed to standard format.

The loop said to be from P1 was actually taken while sailing and should be called Loop 1.

Loop data were combined in file 2012-01-loops.csv. These include 7 salinity and extracted chlorophyll samples, 6 nutrient samples and 1 dissolved oxygen sample. Time and date were added to the file to enable addition of the TSG data later.

None of the loop samples coincide with rosette casts.

a.) Checking calibrations

The calibrations were checked and the only problem concerns the fluorometry. The configuration file contains a Vblank value of 0.186. While this was said to be from a calibration on 18 Sept 2001, it is known that there was a check just before this cruise. The value from the 2001 check was 0.068 and that value was used during 2012-14 which followed this cruise. A few test conversions were done and the value of 0.186 led to many large negative values. The shape of the data looks normal, so the negative values are not due to an instrument malfunction. A conversion using the 2001 parameters produces values that are too high, with no values <2ug/L. A few checks were made against extracted CHL loop samples to see if there is a simple error in the Vblank entry, but none looks plausible. There will be an opportunity to check the fluorescence against the CTD values later, though there are problems with the CTD fluorometer so this may not be very useful. It was decided to use Vblank=0 and scale factor=1 for the conversion. Later, we can attempt to recalibrate if we have more information.

After that change the CON file was saved as 2012-01-tsg.con. The remote temperature calibration parameters could not be checked. They are not listed in the configuration file or header records. This is because it is connected serially with the new TSG system and details can only be checked through SEATERM when it is connected to a computer directly. This is awkward and we have no way of being sure it was done correctly. This should be checked at the beginning of every cruise to be sure no-one has changed it. If we did find an error at the end, there does not appear to be any easy way to correct it.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

A check was made of fluorescence and the minimum values measured were 0.105, 0.118 and 0.121V for the 3 files. This shows that the Vblank must be lower than 0.105.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. There is a lot of noise in the salinity at times. This could be due to the rough conditions. There was one sudden drop in salinity accompanied by increases in temperature and a slight drop in the flow rate. This occurred near the mouth of Juan de Fuca Strait and is not associated with a change in direction. The fluorescence data look ok in shape, but the values are bad, with large negative values. The changes may well be real given the location, so the data should be left as is. Positions did not update for about 1.5 hours at the end of file #2. A frequent pattern seen in salinity is a gradual drop in values followed by a sharp return to values similar to those seen before the drop. The most notable bad section was during the run from P16 to P24 and it appears that while stopped the salinity was much quieter than during those runs. Could a build-up of bubbles in the system account for the lower salinity until there is a sudden clearing of bubbles? Changes of ship direction may have an effect.

Ship position data for every 10s during the period for which they were missing in the TSG file were obtained in file Feb2 2012 RMC.xls. These data were thinned to every 3rd point to match the TSG data and the position formats were changed. Opening the ATC file in Ultraedit allowed a simple cut and paste operation to substitute these new values for the missing data.

Having a cruise where the ship went over the same area three times in a short time span enables some study of variability. The overlap occurs eastward of P2, an area of high variability and the outward and southward tracks through Juan de Fuca Strait are on opposite sides of the Strait, but there are some areas where the tracks are close. Another interesting feature is a sudden change towards lower density water around 48.47°N 124.57°W in Juan de Fuca Strait. Temperature rose by ~0.7° and salinity fell by ~1.9 in 7 minutes and about 2/3rds of that occurred within 1 minute. Fluorescence increased.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2012-01-ctd-tsg-comp.xls.

All files overlapped for at least 1 CTD casts so they were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 25 matches. TSG values were also found for times of underway loop sampling and added to file 2011-01-tsg-loop-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.0001° and in longitude <0.00024°. The median differences were 0.00000° and 0.00001°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. The differences between the two temperatures were mostly between 0.15 and 0.35Cº, but the traces are extremely noisy. Early in the cruise the differences are close to 0.25 Cº, but it drops westward of P21 and mostly stays low until the ship was eastward of P20. This is likely due to water temperatures being lower offshore while the ship temperature probably does not vary much, so heating would increase.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. During stops the intake and lab temperatures differed by an average of 0.029Cº and a median of 0.028Cº

When all data were included the TSG intake temperature was higher than the CTD by average of 0.01Cº; the median difference was the same and the standard deviation was 0.004Cº.

The TSG salinity is very noisy at times. The TSG is lower than the CTD salinity by an average of 0.104, a median of 0.081 and a standard deviation of 0.106. The comparison is very noisy. When the only obvious outlier was excluded the TSG is low by an average of 0.089 and a median of 0.081; the standard deviation is 0.07. Excluding cases with standard deviation in the TSG salinity > 0.003 and differences between TSG and CTD >0.1, the average difference is 0.035, the median is 0.032 and the standard deviation of 0.03.

The ratio of TSG raw fluorescence to CTD fluorescence ranges from 0.5 to 1.3 and a median of 0.7 and standard deviation of 0.24. The highest ratios occurred for the lowest fluorescence values where fluorometers tend not to work so well. (See 2012-01-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons  A spreadsheet was prepared by combining all loop sample data with TSG salinity and TSG fluorescence (median values over 2 minutes). There were no rosette casts at the same time as loop sampling.  (See 2012-01-tsg-loop-comp.xls.)

The TSG salinity was lower than the loop salinity by an average of 0.063 and a median of 0.047; the standard deviation was 0.06. If 2 outliers are excluded the TSG salinity is low by an average of 0.032 and a median of 0.023 with a standard deviation of 0.02.

The TSG fluorescence is raw, so we do not expect a good comparison. When all data are included the TSG values are an average of 31% of the extracted CHL values, a mean of 33% and a standard deviation of 11%. Excluding 1 outlier reduces the ratio slightly to an average of 28%, a median of 27% and a standard deviation of 9%.  If the TSG is roughly 30% of the Extracted CHL and the CTD fluorescence is about 50% of the extracted chlorophyll, then the TSG is roughly 60% of the CTD fluorescence values. The fluorometers tend not to work well at low values, pt for CHL~1 the relationship is that TSG is about 20% and CTD 33% of extracted chlorophyll, so again the TSG is roughly 60% of the CTD fluorescence. But this is raw TSG fluorescence, so can we learn anything about how to calibrate it? If we assume that the raw TSG and loop CHL are equal, then by assuming a scale factor we can derive the Vblank value. In order to obtain a Vblank close to 0.168, we require a scale factor of ~29. Until further information is received about this sensor it seems wisest to leave the raw values in the file, with a header note explaining that the data are nominal.

- 5m rosette samples – To do a few more checks against CHL, the 5m rosette values were extracted and combined with the TSG fluorescence and salinity. The times were matched but we only have start times for the casts, so there will be significant errors due to the difference in time. (See file 2012-01-Rosette-tsg.xls.)

The TSG raw fluorescence has a median value of ~0.38 times the rosette bottle values based on 16 bottles. The ratio is ~0.6 when CHL is about 0.3ug/L and ~0.20 when CHL is about 1ug/L. The median value is higher than in the loop bottle comparison, but that is likely because there are fewer high CHL values in the rosette comparison. Given the many sources of error the results of the two comparisons are reasonably consistent.

The TSG salinity is lower than the bottles by 0.050 when all bottles are included, but when differences >0.1 are excluded the TSG salinity is low by 0.038. It is lower than the CTD salinity during the bottle stop by 0.031.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and there is no record of any other uses since that time.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was fairly steady.

3. The temperature in the loop increases by about 0.25Cº inshore of P20 and about 0.30Cº to 0.35Cº offshore of P20; the difference is likely due to colder offshore waters. The values are reasonable for winter.

4. The TSG intake temperature may be high by 0.01Cº, but the extreme conditions during CTD casts and general noisiness of the TSG data does not support recalibration of temperature.

5. The TSG Salinity is very noisy at times. The comparison with loop samples shows it to be low by a median of 0.023 when a few outliers are excluded. The comparison with CTD casts shows the salinity to be low by a median of 0.08 when all bottles are included but the standard deviation is 0.1. When many outliers are excluded it was found to be low by a median of 0.035 with a standard deviation of 0.03. An addition of 0.02 will be applied to salinity.

6. The fluorescence data have been converted with Vblank = 0 and Scale Factor=1, so the units are volts and the data are nominal. The pre-cruise measurement of Vblank as 0.168 appears to be wrong. Many attempts were made to fine a combination of Vblank and Scale Factor that would lead to reasonable fluorescence values, but none was found. It is recommended that this sensor be serviced at the factory.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. Single-point spikes in salinity that are not associated with temperature spikes were cleaned in all files, but most were in file #2.

A few records were removed from the end of file #3 because the flow was turned off.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

File 2012-01-tsg-recal1.ccf was prepared to adjust salinity by adding 0.02. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag. Position:New was removed since there were very few 0 values, so it is not very useful.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Primary and Temperature:Lab. The reorder is to ensure that programs pick the proxy intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

File 2012-01-che-surface.csv with data from CTD rosette bottles from 5db was combined with file 2012-01-tsg-loop-comp.xls with loop samples; the combined file was called 2012-01-surface-loop.csv.

That spreadsheet was simplified, date calculated in DD/MM/YYYY format. A 6-line header was added and the file was saved as 2011-01-loop-6linehdr.csv. It was converted to IOS format, put through CLEAN and HEADEDIT to get start and stop times and positions, and to add general comments and specific comments for flagged values. The final file was renamed 2011-01-surface.loop. A track plot looks reasonable and a plot of salinity versus date looks right.

# Particulars

# 1. Niskins 12-24 are not needed in final bottle files – only closed for testing.

# 2. Loop mentioned on rosette log actually occurred underway – should be called Loop 1.

# 4. Long time at surface, pump 2 slow to pick up.

8. Cast stopped on way down – had to reterminate wire. Event is crossed out in log book but there are data for the downcast to 125m, so this should be processed.

# 38. Niskins closed in pressure order, NOT as written on log sheet. Refer to Hydro file for details.

# 60. Niskin 23 closed by mistake; do not include in bottle file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1396DR | 26Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 23Apr11 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2215 | 7feb2012 |  |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2012-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 26Mar11 | Factory |  |  |
| Conductivity | 2487 | 26Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 1Feb12 | IOS |  |  |
| Temperature:Secondary | 0603 | 03Mar11 | “ |  |  |